## DATA
Site: | [9 empty] | Parameter | Corrected | Rejected |  | Linear Drift Correction |  | pH4 = | 4.25 | Turbidity Standard | Turbidity Reading
LTD Type/# |  | Dataset Comments: | [5 empty] | Comments: Should Reflect Qualifiers (if used) |  |  | Qualifier | Qualifier |  | Parameter: | pH su | pH7 = | 7.2 | 0 | 0.5
Lat/Long: | [9 empty] | Temp, °C | CT | RT |  | Start Date+Time: | 2021-08-06 12:30 | pH10 = | 10.1 | 100
Parameter |  | T (C°) | SC (µS/cm) | DO Sat (%) | DO Conc (mg/L) | pH (su) | Turb (NTU) | * | * | SC µS/cm | CSC | RSC |  | Stop Date+Time: | 2021-08-17 13:15 | mmHg: | 650 | 1000 | 1070
Assessability | [7 empty] | Codes: Both= either FS or NS, Non= NS only, Neither=Not Assessable |  | DO, % | C% | R% |  | Interval: | 1:00 | Count: | 265.75
Average | [7 empty] | Note: Summary statistics do not include Rejected Data |  | DO, mg/L | CDO | RDO |  | Correction Amount: | -0.74 | Slope: | -0.0027846
Maximum | [7 empty] | Rejected: Based on instrument vericfication beyond Maximum Allowable Limits, exposure/burial, or malfunction |  | pH | CPH | RPH
Minimum | [9 empty] | Turb, NTU | CY | RY
DATE | DATE+TIME | Temp | SpCond | DO Sat | DO Conc | pH | Turbidity | Qualifier | Comments | [7 empty] | Drift Correction Instructions
[11 empty] | MACROS | [5 empty] | 1. Copy and paste data into the data fields to the left.
2. Select parameter from drop down list (Cell S2).
3. Enter time series start/stop times and interval.
4. Enter correction amount (upwards positive, downwards negative), except pH and Turbidity.
5. For pH and Turbidity correction, enter the standard reading values.
6. For DO correction, enter local barometric pressure to nearest 10 mmHg at deployment.
7. Press Correct! key. Corrected data will populate column K.
[11 empty] | Temp Stats Results (C°) |  | If you cannot find the saved file, check the My Documents folder on your PC
4T3
6T3
MWAT
Max 24hr ΔDO (mg/L)
Result
Turbidity Min (NTU)
3 Day
4 Day
5 Day
6 Day
7 Day
14 Day
30 Day
Import Thermograph
QC LTD Data
Run 4T3/6T3/MWAT
Save/Create Upload
Import Sonde File
Import DOlog.csv File
Import Previous UL
Correct!
Calculate Delta DO
Run Turbidity Results
Note on S6: Kristopher Barrios: The correction amount entered here varies by parameter: DO%: subtract the percent saturation verification value from 100%. Values above 100% are negative, below 100% positive. DO mg/L calculated also. DO mg/L (only for Hobo manual corrections): enter the percent difference of the QA check applied to the last valid value of the time series. DO Sat calculated also. Ca = 100*(DO(100% sat) – DO(instrument))/DO(instrument) Temp: enter the percent difference of the QA check applied to the last valid value of the time series. Ca = 100*(T(water bath) – T(instrument))/T(instrument) SpCond: enter the percent difference of the QA check applied to the last valid value of the time series. Ca = 100*(SC(standard) – SC(instrument))/SC(instrument) pH: use results table to right Turb: use results table to right Note: If a significant amount of time has passed between end of file and verification time, only a portion of the total drift occurred during deployment. Use the following equation to create factor for adjusting total drift: (end time - start time)/(verification time - start time)

## CALCS
MAXIMUM
Temperature | Qualifier
<>*RT*
|  | Statistics | T | SC | DO | DO | pH | Turb
|  | Summary | (C°) | (uS/cm) | % | (mg/l) | su | (NTU)
Conductivity | Qualifier | Average
<>*RSC*
D.O. % | Qualifier
<>*R%*
D.O. Conc. | Qualifier
<>*RDO*
pH | Qualifier
<>*RPH*
Turbidity | Qualifier
<>*RY*
MINIMUM
Temperature | Qualifier
<>*RT*
Conductivity | Qualifier
<>*RSC*
D.O. % | Qualifier
<>*R%*
D.O. Conc. | Qualifier
<>*RDO*
pH | Qualifier
<>*RPH*
Turbidity | Qualifier
<>*RY*

## Statistical Outliers
Statistical Outliers
Parameter | High Outlier | Low Outlier | High Count | Low Count
| Quartile3 + 3*IQR | Quartile1 - 3*IQR
Temp | #NUM! | #NUM! | 0 | 0
SpCond | #NUM! | #NUM! | 0 | 0
DO Sat | #NUM! | #NUM! | 0 | 0
DO Conc | #NUM! | #NUM! | 0 | 0
pH | #NUM! | #NUM! | 0 | 0
Turbidity | #NUM! | #NUM! | 0 | 0

## DO-CALCS
DATE | TIME | TEMP | SpCond | ODOSat | ODO Conc | pH | Turbidity | Qualifier | Comments | 24hr_DO_Delta

## MAX
DATE | TIME | TEMP | Qualifier

## 4T3_6T3
DATE | TIME | TEMP | Qualifier | MIN(4h) | MIN(6h)

## MWAT
DATE | TIME | TEMP | Qualifier | AVG(168h)

## R-C_DATA
DATE | TIME | TEMP | SpCond | ODOSat | ODO Conc | pH | Turbidity | Qualifier | Comments

## Tracking
Action | Date+Time | User_Name
Data Import
Data QA and Qualified
4T3 and 6T3 Calculated
Upload File Created
File Path Loaded:

## Lists
Parameter | Interval
Temperature C | 0:15
SpCond uS | 1:00
DOsat %
DO mg/L
pH su
Turbidity NTU

## DO-Temp_Table
DOTABLES Result
Oxygen Solubility Table
Solubility of oxygen in fresh water at various temperatures and pressures
[Solubility shown in milligrams per liter. Values based on published equations by Benson and Krause (1980 and 1984).
°C, degrees Celsius; mm Hg, millimeters of mercury]
Temp. | Barometric Pressure (mm Hg)
(°C) | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730 | 740 | 750 | 760 | 770 | 780 | 790 | 800
Series | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42
0 | 7.66 | 7.85 | 8.04 | 8.24 | 8.43 | 8.62 | 8.82 | 9.01 | 9.2 | 9.4 | 9.59 | 9.79 | 9.98 | 10.17 | 10.37 | 10.56 | 10.75 | 10.95 | 11.14 | 11.33 | 11.53 | 11.72 | 11.91 | 12.11 | 12.3 | 12.49 | 12.69 | 12.88 | 13.07 | 13.27 | 13.46 | 13.65 | 13.85 | 14.04 | 14.23 | 14.43 | 14.62 | 14.81 | 15.01 | 15.2 | 15.39
0.01 | 7.65 | 7.85 | 8.04 | 8.23 | 8.43 | 8.62 | 8.82 | 9.01 | 9.2 | 9.4 | 9.59 | 9.78 | 9.98 | 10.17 | 10.36 | 10.56 | 10.75 | 10.94 | 11.14 | 11.33 | 11.52 | 11.72 | 11.91 | 12.1 | 12.3 | 12.49 | 12.68 | 12.88 | 13.07 | 13.26 | 13.46 | 13.65 | 13.84 | 14.04 | 14.23 | 14.42 | 14.62 | 14.81 | 15 | 15.2 | 15.39
0.02 | 7.65 | 7.85 | 8.04 | 8.23 | 8.43 | 8.62 | 8.81 | 9.01 | 9.2 | 9.39 | 9.59 | 9.78 | 9.97 | 10.17 | 10.36 | 10.55 | 10.75 | 10.94 | 11.13 | 11.33 | 11.52 | 11.71 | 11.91 | 12.1 | 12.29 | 12.49 | 12.68 | 12.87 | 13.07 | 13.26 | 13.45 | 13.65 | 13.84 | 14.03 | 14.23 | 14.42 | 14.61 | 14.81 | 15 | 15.19 | 15.39
0.03 | 7.65 | 7.84 | 8.04 | 8.23 | 8.42 | 8.62 | 8.81 | 9 | 9.2 | 9.39 | 9.58 | 9.78 | 9.97 | 10.16 | 10.36 | 10.55 | 10.74 | 10.94 | 11.13 | 11.32 | 11.52 | 11.71 | 11.9 | 12.1 | 12.29 | 12.48 | 12.68 | 12.87 | 13.06 | 13.26 | 13.45 | 13.64 | 13.84 | 14.03 | 14.22 | 14.42 | 14.61 | 14.8 | 14.99 | 15.19 | 15.38
0.04 | 7.65 | 7.84 | 8.03 | 8.23 | 8.42 | 8.61 | 8.81 | 9 | 9.19 | 9.39 | 9.58 | 9.77 | 9.97 | 10.16 | 10.35 | 10.55 | 10.74 | 10.93 | 11.13 | 11.32 | 11.51 | 11.71 | 11.9 | 12.09 | 12.29 | 12.48 | 12.67 | 12.87 | 13.06 | 13.25 | 13.45 | 13.64 | 13.83 | 14.02 | 14.22 | 14.41 | 14.6 | 14.8 | 14.99 | 15.18 | 15.38
0.05 | 7.65 | 7.84 | 8.03 | 8.23 | 8.42 | 8.61 | 8.81 | 9 | 9.19 | 9.38 | 9.58 | 9.77 | 9.96 | 10.16 | 10.35 | 10.54 | 10.74 | 10.93 | 11.12 | 11.32 | 11.51 | 11.7 | 11.9 | 12.09 | 12.28 | 12.48 | 12.67 | 12.86 | 13.06 | 13.25 | 13.44 | 13.63 | 13.83 | 14.02 | 14.21 | 14.41 | 14.6 | 14.79 | 14.99 | 15.18 | 15.37
0.06 | 7.64 | 7.84 | 8.03 | 8.22 | 8.42 | 8.61 | 8.8 | 9 | 9.19 | 9.38 | 9.58 | 9.77 | 9.96 | 10.15 | 10.35 | 10.54 | 10.73 | 10.93 | 11.12 | 11.31 | 11.51 | 11.7 | 11.89 | 12.09 | 12.28 | 12.47 | 12.67 | 12.86 | 13.05 | 13.24 | 13.44 | 13.63 | 13.82 | 14.02 | 14.21 | 14.4 | 14.6 | 14.79 | 14.98 | 15.18 | 15.37
0.07 | 7.64 | 7.83 | 8.03 | 8.22 | 8.41 | 8.61 | 8.8 | 8.99 | 9.19 | 9.38 | 9.57 | 9.77 | 9.96 | 10.15 | 10.35 | 10.54 | 10.73 | 10.92 | 11.12 | 11.31 | 11.5 | 11.7 | 11.89 | 12.08 | 12.28 | 12.47 | 12.66 | 12.85 | 13.05 | 13.24 | 13.43 | 13.63 | 13.82 | 14.01 | 14.21 | 14.4 | 14.59 | 14.78 | 14.98 | 15.17 | 15.36
0.08 | 7.64 | 7.83 | 8.03 | 8.22 | 8.41 | 8.6 | 8.8 | 8.99 | 9.18 | 9.38 | 9.57 | 9.76 | 9.96 | 10.15 | 10.34 | 10.54 | 10.73 | 10.92 | 11.11 | 11.31 | 11.5 | 11.69 | 11.89 | 12.08 | 12.27 | 12.47 | 12.66 | 12.85 | 13.04 | 13.24 | 13.43 | 13.62 | 13.82 | 14.01 | 14.2 | 14.39 | 14.59 | 14.78 | 14.97 | 15.17 | 15.36
0.09 | 7.64 | 7.83 | 8.02 | 8.22 | 8.41 | 8.6 | 8.8 | 8.99 | 9.18 | 9.37 | 9.57 | 9.76 | 9.95 | 10.15 | 10.34 | 10.53 | 10.73 | 10.92 | 11.11 | 11.3 | 11.5 | 11.69 | 11.88 | 12.08 | 12.27 | 12.46 | 12.65 | 12.85 | 13.04 | 13.23 | 13.43 | 13.62 | 13.81 | 14.01 | 14.2 | 14.39 | 14.58 | 14.78 | 14.97 | 15.16 | 15.36
0.1 | 7.63 | 7.83 | 8.02 | 8.21 | 8.41 | 8.6 | 8.79 | 8.99 | 9.18 | 9.37 | 9.56 | 9.76 | 9.95 | 10.14 | 10.34 | 10.53 | 10.72 | 10.92 | 11.11 | 11.3 | 11.49 | 11.69 | 11.88 | 12.07 | 12.27 | 12.46 | 12.65 | 12.84 | 13.04 | 13.23 | 13.42 | 13.62 | 13.81 | 14 | 14.19 | 14.39 | 14.58 | 14.77 | 14.97 | 15.16 | 15.35
0.11 | 7.63 | 7.83 | 8.02 | 8.21 | 8.4 | 8.6 | 8.79 | 8.98 | 9.18 | 9.37 | 9.56 | 9.75 | 9.95 | 10.14 | 10.33 | 10.53 | 10.72 | 10.91 | 11.1 | 11.3 | 11.49 | 11.68 | 11.88 | 12.07 | 12.26 | 12.45 | 12.65 | 12.84 | 13.03 | 13.23 | 13.42 | 13.61 | 13.8 | 14 | 14.19 | 14.38 | 14.58 | 14.77 | 14.96 | 15.15 | 15.35
0.12 | 7.63 | 7.82 | 8.02 | 8.21 | 8.4 | 8.59 | 8.79 | 8.98 | 9.17 | 9.37 | 9.56 | 9.75 | 9.94 | 10.14 | 10.33 | 10.52 | 10.72 | 10.91 | 11.1 | 11.29 | 11.49 | 11.68 | 11.87 | 12.07 | 12.26 | 12.45 | 12.64 | 12.84 | 13.03 | 13.22 | 13.41 | 13.61 | 13.8 | 13.99 | 14.19 | 14.38 | 14.57 | 14.76 | 14.96 | 15.15 | 15.34
0.13 | 7.63 | 7.82 | 8.01 | 8.21 | 8.4 | 8.59 | 8.79 | 8.98 | 9.17 | 9.36 | 9.56 | 9.75 | 9.94 | 10.13 | 10.33 | 10.52 | 10.71 | 10.91 | 11.1 | 11.29 | 11.48 | 11.68 | 11.87 | 12.06 | 12.25 | 12.45 | 12.64 | 12.83 | 13.03 | 13.22 | 13.41 | 13.6 | 13.8 | 13.99 | 14.18 | 14.37 | 14.57 | 14.76 | 14.95 | 15.15 | 15.34
0.14 | 7.63 | 7.82 | 8.01 | 8.2 | 8.4 | 8.59 | 8.78 | 8.98 | 9.17 | 9.36 | 9.55 | 9.75 | 9.94 | 10.13 | 10.32 | 10.52 | 10.71 | 10.9 | 11.1 | 11.29 | 11.48 | 11.67 | 11.87 | 12.06 | 12.25 | 12.44 | 12.64 | 12.83 | 13.02 | 13.21 | 13.41 | 13.6 | 13.79 | 13.99 | 14.18 | 14.37 | 14.56 | 14.76 | 14.95 | 15.14 | 15.33
0.15 | 7.62 | 7.82 | 8.01 | 8.2 | 8.39 | 8.59 | 8.78 | 8.97 | 9.17 | 9.36 | 9.55 | 9.74 | 9.94 | 10.13 | 10.32 | 10.51 | 10.71 | 10.9 | 11.09 | 11.28 | 11.48 | 11.67 | 11.86 | 12.06 | 12.25 | 12.44 | 12.63 | 12.83 | 13.02 | 13.21 | 13.4 | 13.6 | 13.79 | 13.98 | 14.17 | 14.37 | 14.56 | 14.75 | 14.94 | 15.14 | 15.33
0.16 | 7.62 | 7.81 | 8.01 | 8.2 | 8.39 | 8.59 | 8.78 | 8.97 | 9.16 | 9.36 | 9.55 | 9.74 | 9.93 | 10.13 | 10.32 | 10.51 | 10.7 | 10.9 | 11.09 | 11.28 | 11.47 | 11.67 | 11.86 | 12.05 | 12.24 | 12.44 | 12.63 | 12.82 | 13.01 | 13.21 | 13.4 | 13.59 | 13.78 | 13.98 | 14.17 | 14.36 | 14.55 | 14.75 | 14.94 | 15.13 | 15.32
0.17 | 7.62 | 7.81 | 8 | 8.2 | 8.39 | 8.58 | 8.78 | 8.97 | 9.16 | 9.35 | 9.55 | 9.74 | 9.93 | 10.12 | 10.32 | 10.51 | 10.7 | 10.89 | 11.09 | 11.28 | 11.47 | 11.66 | 11.86 | 12.05 | 12.24 | 12.43 | 12.63 | 12.82 | 13.01 | 13.2 | 13.4 | 13.59 | 13.78 | 13.97 | 14.17 | 14.36 | 14.55 | 14.74 | 14.94 | 15.13 | 15.32
0.18 | 7.62 | 7.81 | 8 | 8.2 | 8.39 | 8.58 | 8.77 | 8.97 | 9.16 | 9.35 | 9.54 | 9.74 | 9.93 | 10.12 | 10.31 | 10.51 | 10.7 | 10.89 | 11.08 | 11.28 | 11.47 | 11.66 | 11.85 | 12.05 | 12.24 | 12.43 | 12.62 | 12.81 | 13.01 | 13.2 | 13.39 | 13.58 | 13.78 | 13.97 | 14.16 | 14.35 | 14.55 | 14.74 | 14.93 | 15.12 | 15.32
0.19 | 7.62 | 7.81 | 8 | 8.19 | 8.39 | 8.58 | 8.77 | 8.96 | 9.16 | 9.35 | 9.54 | 9.73 | 9.92 | 10.12 | 10.31 | 10.5 | 10.69 | 10.89 | 11.08 | 11.27 | 11.46 | 11.66 | 11.85 | 12.04 | 12.23 | 12.43 | 12.62 | 12.81 | 13 | 13.2 | 13.39 | 13.58 | 13.77 | 13.97 | 14.16 | 14.35 | 14.54 | 14.73 | 14.93 | 15.12 | 15.31
0.2 | 7.61 | 7.81 | 8 | 8.19 | 8.38 | 8.58 | 8.77 | 8.96 | 9.15 | 9.34 | 9.54 | 9.73 | 9.92 | 10.11 | 10.31 | 10.5 | 10.69 | 10.88 | 11.08 | 11.27 | 11.46 | 11.65 | 11.85 | 12.04 | 12.23 | 12.42 | 12.62 | 12.81 | 13 | 13.19 | 13.38 | 13.58 | 13.77 | 13.96 | 14.15 | 14.35 | 14.54 | 14.73 | 14.92 | 15.12 | 15.31
0.21 | 7.61 | 7.8 | 8 | 8.19 | 8.38 | 8.57 | 8.77 | 8.96 | 9.15 | 9.34 | 9.53 | 9.73 | 9.92 | 10.11 | 10.3 | 10.5 | 10.69 | 10.88 | 11.07 | 11.27 | 11.46 | 11.65 | 11.84 | 12.03 | 12.23 | 12.42 | 12.61 | 12.8 | 13 | 13.19 | 13.38 | 13.57 | 13.77 | 13.96 | 14.15 | 14.34 | 14.53 | 14.73 | 14.92 | 15.11 | 15.3
0.22 | 7.61 | 7.8 | 7.99 | 8.19 | 8.38 | 8.57 | 8.76 | 8.95 | 9.15 | 9.34 | 9.53 | 9.72 | 9.92 | 10.11 | 10.3 | 10.49 | 10.69 | 10.88 | 11.07 | 11.26 | 11.45 | 11.65 | 11.84 | 12.03 | 12.22 | 12.42 | 12.61 | 12.8 | 12.99 | 13.18 | 13.38 | 13.57 | 13.76 | 13.95 | 14.15 | 14.34 | 14.53 | 14.72 | 14.91 | 15.11 | 15.3
0.23 | 7.61 | 7.8 | 7.99 | 8.18 | 8.38 | 8.57 | 8.76 | 8.95 | 9.14 | 9.34 | 9.53 | 9.72 | 9.91 | 10.11 | 10.3 | 10.49 | 10.68 | 10.87 | 11.07 | 11.26 | 11.45 | 11.64 | 11.84 | 12.03 | 12.22 | 12.41 | 12.6 | 12.8 | 12.99 | 13.18 | 13.37 | 13.57 | 13.76 | 13.95 | 14.14 | 14.33 | 14.53 | 14.72 | 14.91 | 15.1 | 15.29
0.24 | 7.6 | 7.8 | 7.99 | 8.18 | 8.37 | 8.57 | 8.76 | 8.95 | 9.14 | 9.33 | 9.53 | 9.72 | 9.91 | 10.1 | 10.3 | 10.49 | 10.68 | 10.87 | 11.06 | 11.26 | 11.45 | 11.64 | 11.83 | 12.02 | 12.22 | 12.41 | 12.6 | 12.79 | 12.99 | 13.18 | 13.37 | 13.56 | 13.75 | 13.95 | 14.14 | 14.33 | 14.52 | 14.71 | 14.91 | 15.1 | 15.29
0.25 | 7.6 | 7.79 | 7.99 | 8.18 | 8.37 | 8.56 | 8.76 | 8.95 | 9.14 | 9.33 | 9.52 | 9.72 | 9.91 | 10.1 | 10.29 | 10.48 | 10.68 | 10.87 | 11.06 | 11.25 | 11.45 | 11.64 | 11.83 | 12.02 | 12.21 | 12.41 | 12.6 | 12.79 | 12.98 | 13.17 | 13.37 | 13.56 | 13.75 | 13.94 | 14.13 | 14.33 | 14.52 | 14.71 | 14.9 | 15.09 | 15.29
0.26 | 7.6 | 7.79 | 7.98 | 8.18 | 8.37 | 8.56 | 8.75 | 8.94 | 9.14 | 9.33 | 9.52 | 9.71 | 9.91 | 10.1 | 10.29 | 10.48 | 10.67 | 10.87 | 11.06 | 11.25 | 11.44 | 11.63 | 11.83 | 12.02 | 12.21 | 12.4 | 12.59 | 12.79 | 12.98 | 13.17 | 13.36 | 13.55 | 13.75 | 13.94 | 14.13 | 14.32 | 14.51 | 14.71 | 14.9 | 15.09 | 15.28
0.27 | 7.6 | 7.79 | 7.98 | 8.17 | 8.37 | 8.56 | 8.75 | 8.94 | 9.13 | 9.33 | 9.52 | 9.71 | 9.9 | 10.09 | 10.29 | 10.48 | 10.67 | 10.86 | 11.05 | 11.25 | 11.44 | 11.63 | 11.82 | 12.01 | 12.21 | 12.4 | 12.59 | 12.78 | 12.97 | 13.17 | 13.36 | 13.55 | 13.74 | 13.93 | 14.13 | 14.32 | 14.51 | 14.7 | 14.89 | 15.09 | 15.28
0.28 | 7.6 | 7.79 | 7.98 | 8.17 | 8.36 | 8.56 | 8.75 | 8.94 | 9.13 | 9.32 | 9.52 | 9.71 | 9.9 | 10.09 | 10.28 | 10.48 | 10.67 | 10.86 | 11.05 | 11.24 | 11.44 | 11.63 | 11.82 | 12.01 | 12.2 | 12.39 | 12.59 | 12.78 | 12.97 | 13.16 | 13.35 | 13.55 | 13.74 | 13.93 | 14.12 | 14.31 | 14.51 | 14.7 | 14.89 | 15.08 | 15.27
0.29 | 7.59 | 7.79 | 7.98 | 8.17 | 8.36 | 8.55 | 8.75 | 8.94 | 9.13 | 9.32 | 9.51 | 9.7 | 9.9 | 10.09 | 10.28 | 10.47 | 10.66 | 10.86 | 11.05 | 11.24 | 11.43 | 11.62 | 11.82 | 12.01 | 12.2 | 12.39 | 12.58 | 12.78 | 12.97 | 13.16 | 13.35 | 13.54 | 13.73 | 13.93 | 14.12 | 14.31 | 14.5 | 14.69 | 14.89 | 15.08 | 15.27
0.3 | 7.59 | 7.78 | 7.98 | 8.17 | 8.36 | 8.55 | 8.74 | 8.93 | 9.13 | 9.32 | 9.51 | 9.7 | 9.89 | 10.09 | 10.28 | 10.47 | 10.66 | 10.85 | 11.05 | 11.24 | 11.43 | 11.62 | 11.81 | 12 | 12.2 | 12.39 | 12.58 | 12.77 | 12.96 | 13.16 | 13.35 | 13.54 | 13.73 | 13.92 | 14.11 | 14.31 | 14.5 | 14.69 | 14.88 | 15.07 | 15.26
0.31 | 7.59 | 7.78 | 7.97 | 8.16 | 8.36 | 8.55 | 8.74 | 8.93 | 9.12 | 9.32 | 9.51 | 9.7 | 9.89 | 10.08 | 10.27 | 10.47 | 10.66 | 10.85 | 11.04 | 11.23 | 11.43 | 11.62 | 11.81 | 12 | 12.19 | 12.38 | 12.58 | 12.77 | 12.96 | 13.15 | 13.34 | 13.53 | 13.73 | 13.92 | 14.11 | 14.3 | 14.49 | 14.69 | 14.88 | 15.07 | 15.26
0.32 | 7.59 | 7.78 | 7.97 | 8.16 | 8.35 | 8.55 | 8.74 | 8.93 | 9.12 | 9.31 | 9.5 | 9.7 | 9.89 | 10.08 | 10.27 | 10.46 | 10.66 | 10.85 | 11.04 | 11.23 | 11.42 | 11.61 | 11.81 | 12 | 12.19 | 12.38 | 12.57 | 12.76 | 12.96 | 13.15 | 13.34 | 13.53 | 13.72 | 13.91 | 14.11 | 14.3 | 14.49 | 14.68 | 14.87 | 15.06 | 15.26
0.33 | 7.58 | 7.78 | 7.97 | 8.16 | 8.35 | 8.54 | 8.74 | 8.93 | 9.12 | 9.31 | 9.5 | 9.69 | 9.89 | 10.08 | 10.27 | 10.46 | 10.65 | 10.84 | 11.04 | 11.23 | 11.42 | 11.61 | 11.8 | 11.99 | 12.19 | 12.38 | 12.57 | 12.76 | 12.95 | 13.14 | 13.34 | 13.53 | 13.72 | 13.91 | 14.1 | 14.29 | 14.49 | 14.68 | 14.87 | 15.06 | 15.25
0.34 | 7.58 | 7.77 | 7.97 | 8.16 | 8.35 | 8.54 | 8.73 | 8.92 | 9.12 | 9.31 | 9.5 | 9.69 | 9.88 | 10.07 | 10.27 | 10.46 | 10.65 | 10.84 | 11.03 | 11.22 | 11.42 | 11.61 | 11.8 | 11.99 | 12.18 | 12.37 | 12.57 | 12.76 | 12.95 | 13.14 | 13.33 | 13.52 | 13.72 | 13.91 | 14.1 | 14.29 | 14.48 | 14.67 | 14.86 | 15.06 | 15.25
0.35 | 7.58 | 7.77 | 7.96 | 8.16 | 8.35 | 8.54 | 8.73 | 8.92 | 9.11 | 9.31 | 9.5 | 9.69 | 9.88 | 10.07 | 10.26 | 10.45 | 10.65 | 10.84 | 11.03 | 11.22 | 11.41 | 11.6 | 11.8 | 11.99 | 12.18 | 12.37 | 12.56 | 12.75 | 12.95 | 13.14 | 13.33 | 13.52 | 13.71 | 13.9 | 14.09 | 14.29 | 14.48 | 14.67 | 14.86 | 15.05 | 15.24
0.36 | 7.58 | 7.77 | 7.96 | 8.15 | 8.34 | 8.54 | 8.73 | 8.92 | 9.11 | 9.3 | 9.49 | 9.69 | 9.88 | 10.07 | 10.26 | 10.45 | 10.64 | 10.83 | 11.03 | 11.22 | 11.41 | 11.6 | 11.79 | 11.98 | 12.18 | 12.37 | 12.56 | 12.75 | 12.94 | 13.13 | 13.32 | 13.52 | 13.71 | 13.9 | 14.09 | 14.28 | 14.47 | 14.66 | 14.86 | 15.05 | 15.24
0.37 | 7.58 | 7.77 | 7.96 | 8.15 | 8.34 | 8.53 | 8.73 | 8.92 | 9.11 | 9.3 | 9.49 | 9.68 | 9.87 | 10.07 | 10.26 | 10.45 | 10.64 | 10.83 | 11.02 | 11.21 | 11.41 | 11.6 | 11.79 | 11.98 | 12.17 | 12.36 | 12.55 | 12.75 | 12.94 | 13.13 | 13.32 | 13.51 | 13.7 | 13.89 | 14.09 | 14.28 | 14.47 | 14.66 | 14.85 | 15.04 | 15.23
0.38 | 7.57 | 7.77 | 7.96 | 8.15 | 8.34 | 8.53 | 8.72 | 8.91 | 9.11 | 9.3 | 9.49 | 9.68 | 9.87 | 10.06 | 10.25 | 10.45 | 10.64 | 10.83 | 11.02 | 11.21 | 11.4 | 11.59 | 11.79 | 11.98 | 12.17 | 12.36 | 12.55 | 12.74 | 12.93 | 13.13 | 13.32 | 13.51 | 13.7 | 13.89 | 14.08 | 14.27 | 14.46 | 14.66 | 14.85 | 15.04 | 15.23
0.39 | 7.57 | 7.76 | 7.95 | 8.15 | 8.34 | 8.53 | 8.72 | 8.91 | 9.1 | 9.29 | 9.49 | 9.68 | 9.87 | 10.06 | 10.25 | 10.44 | 10.63 | 10.83 | 11.02 | 11.21 | 11.4 | 11.59 | 11.78 | 11.97 | 12.17 | 12.36 | 12.55 | 12.74 | 12.93 | 13.12 | 13.31 | 13.5 | 13.7 | 13.89 | 14.08 | 14.27 | 14.46 | 14.65 | 14.84 | 15.03 | 15.23
0.4 | 7.57 | 7.76 | 7.95 | 8.14 | 8.34 | 8.53 | 8.72 | 8.91 | 9.1 | 9.29 | 9.48 | 9.67 | 9.87 | 10.06 | 10.25 | 10.44 | 10.63 | 10.82 | 11.01 | 11.21 | 11.4 | 11.59 | 11.78 | 11.97 | 12.16 | 12.35 | 12.54 | 12.74 | 12.93 | 13.12 | 13.31 | 13.5 | 13.69 | 13.88 | 14.07 | 14.27 | 14.46 | 14.65 | 14.84 | 15.03 | 15.22
0.41 | 7.57 | 7.76 | 7.95 | 8.14 | 8.33 | 8.52 | 8.72 | 8.91 | 9.1 | 9.29 | 9.48 | 9.67 | 9.86 | 10.05 | 10.25 | 10.44 | 10.63 | 10.82 | 11.01 | 11.2 | 11.39 | 11.58 | 11.78 | 11.97 | 12.16 | 12.35 | 12.54 | 12.73 | 12.92 | 13.11 | 13.31 | 13.5 | 13.69 | 13.88 | 14.07 | 14.26 | 14.45 | 14.64 | 14.84 | 15.03 | 15.22
0.42 | 7.57 | 7.76 | 7.95 | 8.14 | 8.33 | 8.52 | 8.71 | 8.9 | 9.1 | 9.29 | 9.48 | 9.67 | 9.86 | 10.05 | 10.24 | 10.43 | 10.63 | 10.82 | 11.01 | 11.2 | 11.39 | 11.58 | 11.77 | 11.96 | 12.15 | 12.35 | 12.54 | 12.73 | 12.92 | 13.11 | 13.3 | 13.49 | 13.68 | 13.88 | 14.07 | 14.26 | 14.45 | 14.64 | 14.83 | 15.02 | 15.21
0.43 | 7.56 | 7.75 | 7.95 | 8.14 | 8.33 | 8.52 | 8.71 | 8.9 | 9.09 | 9.28 | 9.48 | 9.67 | 9.86 | 10.05 | 10.24 | 10.43 | 10.62 | 10.81 | 11 | 11.2 | 11.39 | 11.58 | 11.77 | 11.96 | 12.15 | 12.34 | 12.53 | 12.72 | 12.92 | 13.11 | 13.3 | 13.49 | 13.68 | 13.87 | 14.06 | 14.25 | 14.44 | 14.64 | 14.83 | 15.02 | 15.21
0.44 | 7.56 | 7.75 | 7.94 | 8.13 | 8.33 | 8.52 | 8.71 | 8.9 | 9.09 | 9.28 | 9.47 | 9.66 | 9.85 | 10.05 | 10.24 | 10.43 | 10.62 | 10.81 | 11 | 11.19 | 11.38 | 11.57 | 11.77 | 11.96 | 12.15 | 12.34 | 12.53 | 12.72 | 12.91 | 13.1 | 13.29 | 13.49 | 13.68 | 13.87 | 14.06 | 14.25 | 14.44 | 14.63 | 14.82 | 15.01 | 15.2
0.45 | 7.56 | 7.75 | 7.94 | 8.13 | 8.32 | 8.51 | 8.71 | 8.9 | 9.09 | 9.28 | 9.47 | 9.66 | 9.85 | 10.04 | 10.23 | 10.43 | 10.62 | 10.81 | 11 | 11.19 | 11.38 | 11.57 | 11.76 | 11.95 | 12.14 | 12.34 | 12.53 | 12.72 | 12.91 | 13.1 | 13.29 | 13.48 | 13.67 | 13.86 | 14.05 | 14.25 | 14.44 | 14.63 | 14.82 | 15.01 | 15.2
0.46 | 7.56 | 7.75 | 7.94 | 8.13 | 8.32 | 8.51 | 8.7 | 8.89 | 9.09 | 9.28 | 9.47 | 9.66 | 9.85 | 10.04 | 10.23 | 10.42 | 10.61 | 10.8 | 11 | 11.19 | 11.38 | 11.57 | 11.76 | 11.95 | 12.14 | 12.33 | 12.52 | 12.71 | 12.91 | 13.1 | 13.29 | 13.48 | 13.67 | 13.86 | 14.05 | 14.24 | 14.43 | 14.62 | 14.81 | 15.01 | 15.2
0.47 | 7.55 | 7.75 | 7.94 | 8.13 | 8.32 | 8.51 | 8.7 | 8.89 | 9.08 | 9.27 | 9.46 | 9.66 | 9.85 | 10.04 | 10.23 | 10.42 | 10.61 | 10.8 | 10.99 | 11.18 | 11.37 | 11.57 | 11.76 | 11.95 | 12.14 | 12.33 | 12.52 | 12.71 | 12.9 | 13.09 | 13.28 | 13.47 | 13.66 | 13.86 | 14.05 | 14.24 | 14.43 | 14.62 | 14.81 | 15 | 15.19
0.48 | 7.55 | 7.74 | 7.93 | 8.13 | 8.32 | 8.51 | 8.7 | 8.89 | 9.08 | 9.27 | 9.46 | 9.65 | 9.84 | 10.03 | 10.23 | 10.42 | 10.61 | 10.8 | 10.99 | 11.18 | 11.37 | 11.56 | 11.75 | 11.94 | 12.13 | 12.33 | 12.52 | 12.71 | 12.9 | 13.09 | 13.28 | 13.47 | 13.66 | 13.85 | 14.04 | 14.23 | 14.42 | 14.62 | 14.81 | 15 | 15.19
0.49 | 7.55 | 7.74 | 7.93 | 8.12 | 8.31 | 8.5 | 8.7 | 8.89 | 9.08 | 9.27 | 9.46 | 9.65 | 9.84 | 10.03 | 10.22 | 10.41 | 10.6 | 10.8 | 10.99 | 11.18 | 11.37 | 11.56 | 11.75 | 11.94 | 12.13 | 12.32 | 12.51 | 12.7 | 12.89 | 13.08 | 13.28 | 13.47 | 13.66 | 13.85 | 14.04 | 14.23 | 14.42 | 14.61 | 14.8 | 14.99 | 15.18
0.5 | 7.55 | 7.74 | 7.93 | 8.12 | 8.31 | 8.5 | 8.69 | 8.88 | 9.07 | 9.27 | 9.46 | 9.65 | 9.84 | 10.03 | 10.22 | 10.41 | 10.6 | 10.79 | 10.98 | 11.17 | 11.36 | 11.56 | 11.75 | 11.94 | 12.13 | 12.32 | 12.51 | 12.7 | 12.89 | 13.08 | 13.27 | 13.46 | 13.65 | 13.84 | 14.03 | 14.23 | 14.42 | 14.61 | 14.8 | 14.99 | 15.18
0.51 | 7.55 | 7.74 | 7.93 | 8.12 | 8.31 | 8.5 | 8.69 | 8.88 | 9.07 | 9.26 | 9.45 | 9.64 | 9.84 | 10.03 | 10.22 | 10.41 | 10.6 | 10.79 | 10.98 | 11.17 | 11.36 | 11.55 | 11.74 | 11.93 | 12.12 | 12.31 | 12.51 | 12.7 | 12.89 | 13.08 | 13.27 | 13.46 | 13.65 | 13.84 | 14.03 | 14.22 | 14.41 | 14.6 | 14.79 | 14.98 | 15.17
0.52 | 7.54 | 7.73 | 7.93 | 8.12 | 8.31 | 8.5 | 8.69 | 8.88 | 9.07 | 9.26 | 9.45 | 9.64 | 9.83 | 10.02 | 10.21 | 10.4 | 10.6 | 10.79 | 10.98 | 11.17 | 11.36 | 11.55 | 11.74 | 11.93 | 12.12 | 12.31 | 12.5 | 12.69 | 12.88 | 13.07 | 13.26 | 13.46 | 13.65 | 13.84 | 14.03 | 14.22 | 14.41 | 14.6 | 14.79 | 14.98 | 15.17
0.53 | 7.54 | 7.73 | 7.92 | 8.11 | 8.3 | 8.5 | 8.69 | 8.88 | 9.07 | 9.26 | 9.45 | 9.64 | 9.83 | 10.02 | 10.21 | 10.4 | 10.59 | 10.78 | 10.97 | 11.16 | 11.35 | 11.55 | 11.74 | 11.93 | 12.12 | 12.31 | 12.5 | 12.69 | 12.88 | 13.07 | 13.26 | 13.45 | 13.64 | 13.83 | 14.02 | 14.21 | 14.4 | 14.59 | 14.79 | 14.98 | 15.17
0.54 | 7.54 | 7.73 | 7.92 | 8.11 | 8.3 | 8.49 | 8.68 | 8.87 | 9.06 | 9.26 | 9.45 | 9.64 | 9.83 | 10.02 | 10.21 | 10.4 | 10.59 | 10.78 | 10.97 | 11.16 | 11.35 | 11.54 | 11.73 | 11.92 | 12.11 | 12.3 | 12.5 | 12.69 | 12.88 | 13.07 | 13.26 | 13.45 | 13.64 | 13.83 | 14.02 | 14.21 | 14.4 | 14.59 | 14.78 | 14.97 | 15.16
0.55 | 7.54 | 7.73 | 7.92 | 8.11 | 8.3 | 8.49 | 8.68 | 8.87 | 9.06 | 9.25 | 9.44 | 9.63 | 9.82 | 10.01 | 10.21 | 10.4 | 10.59 | 10.78 | 10.97 | 11.16 | 11.35 | 11.54 | 11.73 | 11.92 | 12.11 | 12.3 | 12.49 | 12.68 | 12.87 | 13.06 | 13.25 | 13.44 | 13.63 | 13.82 | 14.02 | 14.21 | 14.4 | 14.59 | 14.78 | 14.97 | 15.16
0.56 | 7.54 | 7.73 | 7.92 | 8.11 | 8.3 | 8.49 | 8.68 | 8.87 | 9.06 | 9.25 | 9.44 | 9.63 | 9.82 | 10.01 | 10.2 | 10.39 | 10.58 | 10.77 | 10.96 | 11.15 | 11.35 | 11.54 | 11.73 | 11.92 | 12.11 | 12.3 | 12.49 | 12.68 | 12.87 | 13.06 | 13.25 | 13.44 | 13.63 | 13.82 | 14.01 | 14.2 | 14.39 | 14.58 | 14.77 | 14.96 | 15.15
0.57 | 7.53 | 7.72 | 7.91 | 8.1 | 8.29 | 8.49 | 8.68 | 8.87 | 9.06 | 9.25 | 9.44 | 9.63 | 9.82 | 10.01 | 10.2 | 10.39 | 10.58 | 10.77 | 10.96 | 11.15 | 11.34 | 11.53 | 11.72 | 11.91 | 12.1 | 12.29 | 12.48 | 12.67 | 12.87 | 13.06 | 13.25 | 13.44 | 13.63 | 13.82 | 14.01 | 14.2 | 14.39 | 14.58 | 14.77 | 14.96 | 15.15
0.58 | 7.53 | 7.72 | 7.91 | 8.1 | 8.29 | 8.48 | 8.67 | 8.86 | 9.05 | 9.24 | 9.44 | 9.63 | 9.82 | 10.01 | 10.2 | 10.39 | 10.58 | 10.77 | 10.96 | 11.15 | 11.34 | 11.53 | 11.72 | 11.91 | 12.1 | 12.29 | 12.48 | 12.67 | 12.86 | 13.05 | 13.24 | 13.43 | 13.62 | 13.81 | 14 | 14.19 | 14.38 | 14.57 | 14.76 | 14.95 | 15.15
0.59 | 7.53 | 7.72 | 7.91 | 8.1 | 8.29 | 8.48 | 8.67 | 8.86 | 9.05 | 9.24 | 9.43 | 9.62 | 9.81 | 10 | 10.19 | 10.38 | 10.57 | 10.76 | 10.96 | 11.15 | 11.34 | 11.53 | 11.72 | 11.91 | 12.1 | 12.29 | 12.48 | 12.67 | 12.86 | 13.05 | 13.24 | 13.43 | 13.62 | 13.81 | 14 | 14.19 | 14.38 | 14.57 | 14.76 | 14.95 | 15.14
0.6 | 7.53 | 7.72 | 7.91 | 8.1 | 8.29 | 8.48 | 8.67 | 8.86 | 9.05 | 9.24 | 9.43 | 9.62 | 9.81 | 10 | 10.19 | 10.38 | 10.57 | 10.76 | 10.95 | 11.14 | 11.33 | 11.52 | 11.71 | 11.9 | 12.09 | 12.28 | 12.47 | 12.66 | 12.85 | 13.04 | 13.23 | 13.43 | 13.62 | 13.81 | 14 | 14.19 | 14.38 | 14.57 | 14.76 | 14.95 | 15.14
0.61 | 7.52 | 7.71 | 7.91 | 8.1 | 8.29 | 8.48 | 8.67 | 8.86 | 9.05 | 9.24 | 9.43 | 9.62 | 9.81 | 10 | 10.19 | 10.38 | 10.57 | 10.76 | 10.95 | 11.14 | 11.33 | 11.52 | 11.71 | 11.9 | 12.09 | 12.28 | 12.47 | 12.66 | 12.85 | 13.04 | 13.23 | 13.42 | 13.61 | 13.8 | 13.99 | 14.18 | 14.37 | 14.56 | 14.75 | 14.94 | 15.13
0.62 | 7.52 | 7.71 | 7.9 | 8.09 | 8.28 | 8.47 | 8.66 | 8.85 | 9.04 | 9.23 | 9.42 | 9.61 | 9.8 | 10 | 10.19 | 10.38 | 10.57 | 10.76 | 10.95 | 11.14 | 11.33 | 11.52 | 11.71 | 11.9 | 12.09 | 12.28 | 12.47 | 12.66 | 12.85 | 13.04 | 13.23 | 13.42 | 13.61 | 13.8 | 13.99 | 14.18 | 14.37 | 14.56 | 14.75 | 14.94 | 15.13
0.63 | 7.52 | 7.71 | 7.9 | 8.09 | 8.28 | 8.47 | 8.66 | 8.85 | 9.04 | 9.23 | 9.42 | 9.61 | 9.8 | 9.99 | 10.18 | 10.37 | 10.56 | 10.75 | 10.94 | 11.13 | 11.32 | 11.51 | 11.7 | 11.89 | 12.08 | 12.27 | 12.46 | 12.65 | 12.84 | 13.03 | 13.22 | 13.41 | 13.6 | 13.79 | 13.98 | 14.17 | 14.36 | 14.55 | 14.74 | 14.93 | 15.12
0.64 | 7.52 | 7.71 | 7.9 | 8.09 | 8.28 | 8.47 | 8.66 | 8.85 | 9.04 | 9.23 | 9.42 | 9.61 | 9.8 | 9.99 | 10.18 | 10.37 | 10.56 | 10.75 | 10.94 | 11.13 | 11.32 | 11.51 | 11.7 | 11.89 | 12.08 | 12.27 | 12.46 | 12.65 | 12.84 | 13.03 | 13.22 | 13.41 | 13.6 | 13.79 | 13.98 | 14.17 | 14.36 | 14.55 | 14.74 | 14.93 | 15.12
0.65 | 7.52 | 7.71 | 7.9 | 8.09 | 8.28 | 8.47 | 8.66 | 8.85 | 9.04 | 9.23 | 9.42 | 9.61 | 9.8 | 9.99 | 10.18 | 10.37 | 10.56 | 10.75 | 10.94 | 11.13 | 11.32 | 11.51 | 11.7 | 11.89 | 12.08 | 12.27 | 12.46 | 12.65 | 12.84 | 13.03 | 13.22 | 13.41 | 13.6 | 13.79 | 13.98 | 14.17 | 14.36 | 14.55 | 14.74 | 14.93 | 15.12
0.66 | 7.51 | 7.7 | 7.89 | 8.08 | 8.27 | 8.46 | 8.65 | 8.84 | 9.03 | 9.22 | 9.41 | 9.6 | 9.79 | 9.98 | 10.17 | 10.36 | 10.55 | 10.74 | 10.93 | 11.12 | 11.31 | 11.5 | 11.69 | 11.88 | 12.07 | 12.26 | 12.45 | 12.64 | 12.83 | 13.02 | 13.21 | 13.4 | 13.59 | 13.78 | 13.97 | 14.16 | 14.35 | 14.54 | 14.73 | 14.92 | 15.11
0.67 | 7.51 | 7.7 | 7.89 | 8.08 | 8.27 | 8.46 | 8.65 | 8.84 | 9.03 | 9.22 | 9.41 | 9.6 | 9.79 | 9.98 | 10.17 | 10.36 | 10.55 | 10.74 | 10.93 | 11.12 | 11.31 | 11.5 | 11.69 | 11.88 | 12.07 | 12.26 | 12.45 | 12.64 | 12.83 | 13.02 | 13.21 | 13.4 | 13.59 | 13.78 | 13.97 | 14.16 | 14.35 | 14.54 | 14.73 | 14.92 | 15.11
0.68 | 7.51 | 7.7 | 7.89 | 8.08 | 8.27 | 8.46 | 8.65 | 8.84 | 9.03 | 9.22 | 9.41 | 9.6 | 9.79 | 9.98 | 10.17 | 10.36 | 10.55 | 10.74 | 10.93 | 11.12 | 11.31 | 11.5 | 11.69 | 11.88 | 12.07 | 12.26 | 12.45 | 12.64 | 12.83 | 13.02 | 13.21 | 13.39 | 13.58 | 13.77 | 13.96 | 14.15 | 14.34 | 14.53 | 14.72 | 14.91 | 15.1
0.69 | 7.51 | 7.7 | 7.89 | 8.08 | 8.27 | 8.46 | 8.65 | 8.84 | 9.03 | 9.22 | 9.41 | 9.6 | 9.79 | 9.98 | 10.17 | 10.35 | 10.54 | 10.73 | 10.92 | 11.11 | 11.3 | 11.49 | 11.68 | 11.87 | 12.06 | 12.25 | 12.44 | 12.63 | 12.82 | 13.01 | 13.2 | 13.39 | 13.58 | 13.77 | 13.96 | 14.15 | 14.34 | 14.53 | 14.72 | 14.91 | 15.1
0.7 | 7.51 | 7.7 | 7.88 | 8.07 | 8.26 | 8.45 | 8.64 | 8.83 | 9.02 | 9.21 | 9.4 | 9.59 | 9.78 | 9.97 | 10.16 | 10.35 | 10.54 | 10.73 | 10.92 | 11.11 | 11.3 | 11.49 | 11.68 | 11.87 | 12.06 | 12.25 | 12.44 | 12.63 | 12.82 | 13.01 | 13.2 | 13.39 | 13.58 | 13.77 | 13.96 | 14.15 | 14.34 | 14.53 | 14.72 | 14.9 | 15.09
0.71 | 7.5 | 7.69 | 7.88 | 8.07 | 8.26 | 8.45 | 8.64 | 8.83 | 9.02 | 9.21 | 9.4 | 9.59 | 9.78 | 9.97 | 10.16 | 10.35 | 10.54 | 10.73 | 10.92 | 11.11 | 11.3 | 11.49 | 11.68 | 11.87 | 12.06 | 12.25 | 12.44 | 12.63 | 12.81 | 13 | 13.19 | 13.38 | 13.57 | 13.76 | 13.95 | 14.14 | 14.33 | 14.52 | 14.71 | 14.9 | 15.09
0.72 | 7.5 | 7.69 | 7.88 | 8.07 | 8.26 | 8.45 | 8.64 | 8.83 | 9.02 | 9.21 | 9.4 | 9.59 | 9.78 | 9.97 | 10.16 | 10.35 | 10.54 | 10.73 | 10.92 | 11.1 | 11.29 | 11.48 | 11.67 | 11.86 | 12.05 | 12.24 | 12.43 | 12.62 | 12.81 | 13 | 13.19 | 13.38 | 13.57 | 13.76 | 13.95 | 14.14 | 14.33 | 14.52 | 14.71 | 14.9 | 15.09
0.73 | 7.5 | 7.69 | 7.88 | 8.07 | 8.26 | 8.45 | 8.64 | 8.83 | 9.02 | 9.21 | 9.4 | 9.58 | 9.77 | 9.96 | 10.15 | 10.34 | 10.53 | 10.72 | 10.91 | 11.1 | 11.29 | 11.48 | 11.67 | 11.86 | 12.05 | 12.24 | 12.43 | 12.62 | 12.81 | 13 | 13.19 | 13.38 | 13.57 | 13.76 | 13.94 | 14.13 | 14.32 | 14.51 | 14.7 | 14.89 | 15.08
0.74 | 7.5 | 7.69 | 7.88 | 8.07 | 8.26 | 8.44 | 8.63 | 8.82 | 9.01 | 9.2 | 9.39 | 9.58 | 9.77 | 9.96 | 10.15 | 10.34 | 10.53 | 10.72 | 10.91 | 11.1 | 11.29 | 11.48 | 11.67 | 11.86 | 12.05 | 12.24 | 12.43 | 12.61 | 12.8 | 12.99 | 13.18 | 13.37 | 13.56 | 13.75 | 13.94 | 14.13 | 14.32 | 14.51 | 14.7 | 14.89 | 15.08
0.75 | 7.49 | 7.68 | 7.87 | 8.06 | 8.25 | 8.44 | 8.63 | 8.82 | 9.01 | 9.2 | 9.39 | 9.58 | 9.77 | 9.96 | 10.15 | 10.34 | 10.53 | 10.72 | 10.91 | 11.1 | 11.28 | 11.47 | 11.66 | 11.85 | 12.04 | 12.23 | 12.42 | 12.61 | 12.8 | 12.99 | 13.18 | 13.37 | 13.56 | 13.75 | 13.94 | 14.13 | 14.32 | 14.51 | 14.69 | 14.88 | 15.07
0.76 | 7.49 | 7.68 | 7.87 | 8.06 | 8.25 | 8.44 | 8.63 | 8.82 | 9.01 | 9.2 | 9.39 | 9.58 | 9.77 | 9.96 | 10.15 | 10.33 | 10.52 | 10.71 | 10.9 | 11.09 | 11.28 | 11.47 | 11.66 | 11.85 | 12.04 | 12.23 | 12.42 | 12.61 | 12.8 | 12.99 | 13.18 | 13.37 | 13.55 | 13.74 | 13.93 | 14.12 | 14.31 | 14.5 | 14.69 | 14.88 | 15.07
0.77 | 7.49 | 7.68 | 7.87 | 8.06 | 8.25 | 8.44 | 8.63 | 8.82 | 9.01 | 9.2 | 9.38 | 9.57 | 9.76 | 9.95 | 10.14 | 10.33 | 10.52 | 10.71 | 10.9 | 11.09 | 11.28 | 11.47 | 11.66 | 11.85 | 12.04 | 12.23 | 12.41 | 12.6 | 12.79 | 12.98 | 13.17 | 13.36 | 13.55 | 13.74 | 13.93 | 14.12 | 14.31 | 14.5 | 14.69 | 14.88 | 15.06
0.78 | 7.49 | 7.68 | 7.87 | 8.06 | 8.25 | 8.44 | 8.62 | 8.81 | 9 | 9.19 | 9.38 | 9.57 | 9.76 | 9.95 | 10.14 | 10.33 | 10.52 | 10.71 | 10.9 | 11.09 | 11.28 | 11.46 | 11.65 | 11.84 | 12.03 | 12.22 | 12.41 | 12.6 | 12.79 | 12.98 | 13.17 | 13.36 | 13.55 | 13.74 | 13.93 | 14.11 | 14.3 | 14.49 | 14.68 | 14.87 | 15.06
0.79 | 7.49 | 7.68 | 7.86 | 8.05 | 8.24 | 8.43 | 8.62 | 8.81 | 9 | 9.19 | 9.38 | 9.57 | 9.76 | 9.95 | 10.14 | 10.33 | 10.52 | 10.7 | 10.89 | 11.08 | 11.27 | 11.46 | 11.65 | 11.84 | 12.03 | 12.22 | 12.41 | 12.6 | 12.79 | 12.98 | 13.16 | 13.35 | 13.54 | 13.73 | 13.92 | 14.11 | 14.3 | 14.49 | 14.68 | 14.87 | 15.06
0.8 | 7.48 | 7.67 | 7.86 | 8.05 | 8.24 | 8.43 | 8.62 | 8.81 | 9 | 9.19 | 9.38 | 9.57 | 9.76 | 9.94 | 10.13 | 10.32 | 10.51 | 10.7 | 10.89 | 11.08 | 11.27 | 11.46 | 11.65 | 11.84 | 12.03 | 12.22 | 12.4 | 12.59 | 12.78 | 12.97 | 13.16 | 13.35 | 13.54 | 13.73 | 13.92 | 14.11 | 14.3 | 14.48 | 14.67 | 14.86 | 15.05
0.81 | 7.48 | 7.67 | 7.86 | 8.05 | 8.24 | 8.43 | 8.62 | 8.81 | 9 | 9.18 | 9.37 | 9.56 | 9.75 | 9.94 | 10.13 | 10.32 | 10.51 | 10.7 | 10.89 | 11.08 | 11.27 | 11.46 | 11.64 | 11.83 | 12.02 | 12.21 | 12.4 | 12.59 | 12.78 | 12.97 | 13.16 | 13.35 | 13.54 | 13.72 | 13.91 | 14.1 | 14.29 | 14.48 | 14.67 | 14.86 | 15.05
0.82 | 7.48 | 7.67 | 7.86 | 8.05 | 8.24 | 8.43 | 8.61 | 8.8 | 8.99 | 9.18 | 9.37 | 9.56 | 9.75 | 9.94 | 10.13 | 10.32 | 10.51 | 10.7 | 10.88 | 11.07 | 11.26 | 11.45 | 11.64 | 11.83 | 12.02 | 12.21 | 12.4 | 12.59 | 12.78 | 12.96 | 13.15 | 13.34 | 13.53 | 13.72 | 13.91 | 14.1 | 14.29 | 14.48 | 14.67 | 14.85 | 15.04
0.83 | 7.48 | 7.67 | 7.86 | 8.05 | 8.23 | 8.42 | 8.61 | 8.8 | 8.99 | 9.18 | 9.37 | 9.56 | 9.75 | 9.94 | 10.13 | 10.31 | 10.5 | 10.69 | 10.88 | 11.07 | 11.26 | 11.45 | 11.64 | 11.83 | 12.02 | 12.2 | 12.39 | 12.58 | 12.77 | 12.96 | 13.15 | 13.34 | 13.53 | 13.72 | 13.91 | 14.09 | 14.28 | 14.47 | 14.66 | 14.85 | 15.04
0.84 | 7.48 | 7.66 | 7.85 | 8.04 | 8.23 | 8.42 | 8.61 | 8.8 | 8.99 | 9.18 | 9.37 | 9.56 | 9.74 | 9.93 | 10.12 | 10.31 | 10.5 | 10.69 | 10.88 | 11.07 | 11.26 | 11.45 | 11.63 | 11.82 | 12.01 | 12.2 | 12.39 | 12.58 | 12.77 | 12.96 | 13.15 | 13.34 | 13.52 | 13.71 | 13.9 | 14.09 | 14.28 | 14.47 | 14.66 | 14.85 | 15.04
0.85 | 7.47 | 7.66 | 7.85 | 8.04 | 8.23 | 8.42 | 8.61 | 8.8 | 8.99 | 9.17 | 9.36 | 9.55 | 9.74 | 9.93 | 10.12 | 10.31 | 10.5 | 10.69 | 10.88 | 11.06 | 11.25 | 11.44 | 11.63 | 11.82 | 12.01 | 12.2 | 12.39 | 12.58 | 12.76 | 12.95 | 13.14 | 13.33 | 13.52 | 13.71 | 13.9 | 14.09 | 14.28 | 14.46 | 14.65 | 14.84 | 15.03
0.86 | 7.47 | 7.66 | 7.85 | 8.04 | 8.23 | 8.42 | 8.61 | 8.79 | 8.98 | 9.17 | 9.36 | 9.55 | 9.74 | 9.93 | 10.12 | 10.31 | 10.49 | 10.68 | 10.87 | 11.06 | 11.25 | 11.44 | 11.63 | 11.82 | 12.01 | 12.19 | 12.38 | 12.57 | 12.76 | 12.95 | 13.14 | 13.33 | 13.52 | 13.71 | 13.89 | 14.08 | 14.27 | 14.46 | 14.65 | 14.84 | 15.03
0.87 | 7.47 | 7.66 | 7.85 | 8.04 | 8.22 | 8.41 | 8.6 | 8.79 | 8.98 | 9.17 | 9.36 | 9.55 | 9.74 | 9.93 | 10.11 | 10.3 | 10.49 | 10.68 | 10.87 | 11.06 | 11.25 | 11.44 | 11.62 | 11.81 | 12 | 12.19 | 12.38 | 12.57 | 12.76 | 12.95 | 13.14 | 13.32 | 13.51 | 13.7 | 13.89 | 14.08 | 14.27 | 14.46 | 14.65 | 14.83 | 15.02
0.88 | 7.47 | 7.66 | 7.84 | 8.03 | 8.22 | 8.41 | 8.6 | 8.79 | 8.98 | 9.17 | 9.36 | 9.54 | 9.73 | 9.92 | 10.11 | 10.3 | 10.49 | 10.68 | 10.87 | 11.06 | 11.24 | 11.43 | 11.62 | 11.81 | 12 | 12.19 | 12.38 | 12.57 | 12.75 | 12.94 | 13.13 | 13.32 | 13.51 | 13.7 | 13.89 | 14.08 | 14.26 | 14.45 | 14.64 | 14.83 | 15.02
0.89 | 7.46 | 7.65 | 7.84 | 8.03 | 8.22 | 8.41 | 8.6 | 8.79 | 8.98 | 9.16 | 9.35 | 9.54 | 9.73 | 9.92 | 10.11 | 10.3 | 10.49 | 10.67 | 10.86 | 11.05 | 11.24 | 11.43 | 11.62 | 11.81 | 12 | 12.18 | 12.37 | 12.56 | 12.75 | 12.94 | 13.13 | 13.32 | 13.51 | 13.69 | 13.88 | 14.07 | 14.26 | 14.45 | 14.64 | 14.83 | 15.01
0.9 | 7.46 | 7.65 | 7.84 | 8.03 | 8.22 | 8.41 | 8.6 | 8.78 | 8.97 | 9.16 | 9.35 | 9.54 | 9.73 | 9.92 | 10.11 | 10.29 | 10.48 | 10.67 | 10.86 | 11.05 | 11.24 | 11.43 | 11.61 | 11.8 | 11.99 | 12.18 | 12.37 | 12.56 | 12.75 | 12.94 | 13.12 | 13.31 | 13.5 | 13.69 | 13.88 | 14.07 | 14.26 | 14.44 | 14.63 | 14.82 | 15.01
0.91 | 7.46 | 7.65 | 7.84 | 8.03 | 8.22 | 8.4 | 8.59 | 8.78 | 8.97 | 9.16 | 9.35 | 9.54 | 9.73 | 9.91 | 10.1 | 10.29 | 10.48 | 10.67 | 10.86 | 11.05 | 11.23 | 11.42 | 11.61 | 11.8 | 11.99 | 12.18 | 12.37 | 12.55 | 12.74 | 12.93 | 13.12 | 13.31 | 13.5 | 13.69 | 13.87 | 14.06 | 14.25 | 14.44 | 14.63 | 14.82 | 15.01
0.92 | 7.46 | 7.65 | 7.84 | 8.02 | 8.21 | 8.4 | 8.59 | 8.78 | 8.97 | 9.16 | 9.35 | 9.53 | 9.72 | 9.91 | 10.1 | 10.29 | 10.48 | 10.67 | 10.85 | 11.04 | 11.23 | 11.42 | 11.61 | 11.8 | 11.99 | 12.17 | 12.36 | 12.55 | 12.74 | 12.93 | 13.12 | 13.31 | 13.49 | 13.68 | 13.87 | 14.06 | 14.25 | 14.44 | 14.63 | 14.81 | 15
0.93 | 7.46 | 7.65 | 7.83 | 8.02 | 8.21 | 8.4 | 8.59 | 8.78 | 8.97 | 9.15 | 9.34 | 9.53 | 9.72 | 9.91 | 10.1 | 10.29 | 10.47 | 10.66 | 10.85 | 11.04 | 11.23 | 11.42 | 11.61 | 11.79 | 11.98 | 12.17 | 12.36 | 12.55 | 12.74 | 12.92 | 13.11 | 13.3 | 13.49 | 13.68 | 13.87 | 14.06 | 14.24 | 14.43 | 14.62 | 14.81 | 15
0.94 | 7.45 | 7.64 | 7.83 | 8.02 | 8.21 | 8.4 | 8.59 | 8.77 | 8.96 | 9.15 | 9.34 | 9.53 | 9.72 | 9.91 | 10.09 | 10.28 | 10.47 | 10.66 | 10.85 | 11.04 | 11.23 | 11.41 | 11.6 | 11.79 | 11.98 | 12.17 | 12.36 | 12.54 | 12.73 | 12.92 | 13.11 | 13.3 | 13.49 | 13.67 | 13.86 | 14.05 | 14.24 | 14.43 | 14.62 | 14.81 | 14.99
0.95 | 7.45 | 7.64 | 7.83 | 8.02 | 8.21 | 8.39 | 8.58 | 8.77 | 8.96 | 9.15 | 9.34 | 9.53 | 9.71 | 9.9 | 10.09 | 10.28 | 10.47 | 10.66 | 10.85 | 11.03 | 11.22 | 11.41 | 11.6 | 11.79 | 11.98 | 12.16 | 12.35 | 12.54 | 12.73 | 12.92 | 13.11 | 13.29 | 13.48 | 13.67 | 13.86 | 14.05 | 14.24 | 14.42 | 14.61 | 14.8 | 14.99
0.96 | 7.45 | 7.64 | 7.83 | 8.02 | 8.2 | 8.39 | 8.58 | 8.77 | 8.96 | 9.15 | 9.33 | 9.52 | 9.71 | 9.9 | 10.09 | 10.28 | 10.47 | 10.65 | 10.84 | 11.03 | 11.22 | 11.41 | 11.6 | 11.78 | 11.97 | 12.16 | 12.35 | 12.54 | 12.73 | 12.91 | 13.1 | 13.29 | 13.48 | 13.67 | 13.86 | 14.04 | 14.23 | 14.42 | 14.61 | 14.8 | 14.99
0.97 | 7.45 | 7.64 | 7.82 | 8.01 | 8.2 | 8.39 | 8.58 | 8.77 | 8.96 | 9.14 | 9.33 | 9.52 | 9.71 | 9.9 | 10.09 | 10.27 | 10.46 | 10.65 | 10.84 | 11.03 | 11.22 | 11.4 | 11.59 | 11.78 | 11.97 | 12.16 | 12.35 | 12.53 | 12.72 | 12.91 | 13.1 | 13.29 | 13.48 | 13.66 | 13.85 | 14.04 | 14.23 | 14.42 | 14.6 | 14.79 | 14.98
0.98 | 7.45 | 7.63 | 7.82 | 8.01 | 8.2 | 8.39 | 8.58 | 8.76 | 8.95 | 9.14 | 9.33 | 9.52 | 9.71 | 9.89 | 10.08 | 10.27 | 10.46 | 10.65 | 10.84 | 11.02 | 11.21 | 11.4 | 11.59 | 11.78 | 11.97 | 12.15 | 12.34 | 12.53 | 12.72 | 12.91 | 13.1 | 13.28 | 13.47 | 13.66 | 13.85 | 14.04 | 14.22 | 14.41 | 14.6 | 14.79 | 14.98
0.99 | 7.44 | 7.63 | 7.82 | 8.01 | 8.2 | 8.39 | 8.57 | 8.76 | 8.95 | 9.14 | 9.33 | 9.52 | 9.7 | 9.89 | 10.08 | 10.27 | 10.46 | 10.64 | 10.83 | 11.02 | 11.21 | 11.4 | 11.59 | 11.77 | 11.96 | 12.15 | 12.34 | 12.53 | 12.72 | 12.9 | 13.09 | 13.28 | 13.47 | 13.66 | 13.84 | 14.03 | 14.22 | 14.41 | 14.6 | 14.78 | 14.97
1 | 7.44 | 7.63 | 7.82 | 8.01 | 8.19 | 8.38 | 8.57 | 8.76 | 8.95 | 9.14 | 9.32 | 9.51 | 9.7 | 9.89 | 10.08 | 10.27 | 10.45 | 10.64 | 10.83 | 11.02 | 11.21 | 11.39 | 11.58 | 11.77 | 11.96 | 12.15 | 12.34 | 12.52 | 12.71 | 12.9 | 13.09 | 13.28 | 13.46 | 13.65 | 13.84 | 14.03 | 14.22 | 14.4 | 14.59 | 14.78 | 14.97
1.01 | 7.44 | 7.63 | 7.82 | 8 | 8.19 | 8.38 | 8.57 | 8.76 | 8.95 | 9.13 | 9.32 | 9.51 | 9.7 | 9.89 | 10.07 | 10.26 | 10.45 | 10.64 | 10.83 | 11.01 | 11.2 | 11.39 | 11.58 | 11.77 | 11.96 | 12.14 | 12.33 | 12.52 | 12.71 | 12.9 | 13.08 | 13.27 | 13.46 | 13.65 | 13.84 | 14.02 | 14.21 | 14.4 | 14.59 | 14.78 | 14.96
1.02 | 7.44 | 7.63 | 7.81 | 8 | 8.19 | 8.38 | 8.57 | 8.75 | 8.94 | 9.13 | 9.32 | 9.51 | 9.7 | 9.88 | 10.07 | 10.26 | 10.45 | 10.64 | 10.82 | 11.01 | 11.2 | 11.39 | 11.58 | 11.76 | 11.95 | 12.14 | 12.33 | 12.52 | 12.7 | 12.89 | 13.08 | 13.27 | 13.46 | 13.64 | 13.83 | 14.02 | 14.21 | 14.4 | 14.58 | 14.77 | 14.96
1.03 | 7.44 | 7.62 | 7.81 | 8 | 8.19 | 8.38 | 8.56 | 8.75 | 8.94 | 9.13 | 9.32 | 9.5 | 9.69 | 9.88 | 10.07 | 10.26 | 10.44 | 10.63 | 10.82 | 11.01 | 11.2 | 11.38 | 11.57 | 11.76 | 11.95 | 12.14 | 12.32 | 12.51 | 12.7 | 12.89 | 13.08 | 13.26 | 13.45 | 13.64 | 13.83 | 14.02 | 14.2 | 14.39 | 14.58 | 14.77 | 14.96
1.04 | 7.43 | 7.62 | 7.81 | 8 | 8.19 | 8.37 | 8.56 | 8.75 | 8.94 | 9.13 | 9.31 | 9.5 | 9.69 | 9.88 | 10.07 | 10.25 | 10.44 | 10.63 | 10.82 | 11.01 | 11.19 | 11.38 | 11.57 | 11.76 | 11.95 | 12.13 | 12.32 | 12.51 | 12.7 | 12.89 | 13.07 | 13.26 | 13.45 | 13.64 | 13.82 | 14.01 | 14.2 | 14.39 | 14.58 | 14.76 | 14.95
1.05 | 7.43 | 7.62 | 7.81 | 8 | 8.18 | 8.37 | 8.56 | 8.75 | 8.94 | 9.12 | 9.31 | 9.5 | 9.69 | 9.88 | 10.06 | 10.25 | 10.44 | 10.63 | 10.81 | 11 | 11.19 | 11.38 | 11.57 | 11.75 | 11.94 | 12.13 | 12.32 | 12.51 | 12.69 | 12.88 | 13.07 | 13.26 | 13.45 | 13.63 | 13.82 | 14.01 | 14.2 | 14.38 | 14.57 | 14.76 | 14.95
1.06 | 7.43 | 7.62 | 7.81 | 7.99 | 8.18 | 8.37 | 8.56 | 8.74 | 8.93 | 9.12 | 9.31 | 9.5 | 9.68 | 9.87 | 10.06 | 10.25 | 10.44 | 10.62 | 10.81 | 11 | 11.19 | 11.38 | 11.56 | 11.75 | 11.94 | 12.13 | 12.31 | 12.5 | 12.69 | 12.88 | 13.07 | 13.25 | 13.44 | 13.63 | 13.82 | 14 | 14.19 | 14.38 | 14.57 | 14.76 | 14.94
1.07 | 7.43 | 7.62 | 7.8 | 7.99 | 8.18 | 8.37 | 8.55 | 8.74 | 8.93 | 9.12 | 9.31 | 9.49 | 9.68 | 9.87 | 10.06 | 10.25 | 10.43 | 10.62 | 10.81 | 11 | 11.18 | 11.37 | 11.56 | 11.75 | 11.94 | 12.12 | 12.31 | 12.5 | 12.69 | 12.87 | 13.06 | 13.25 | 13.44 | 13.63 | 13.81 | 14 | 14.19 | 14.38 | 14.56 | 14.75 | 14.94
1.08 | 7.43 | 7.61 | 7.8 | 7.99 | 8.18 | 8.36 | 8.55 | 8.74 | 8.93 | 9.12 | 9.3 | 9.49 | 9.68 | 9.87 | 10.05 | 10.24 | 10.43 | 10.62 | 10.81 | 10.99 | 11.18 | 11.37 | 11.56 | 11.74 | 11.93 | 12.12 | 12.31 | 12.5 | 12.68 | 12.87 | 13.06 | 13.25 | 13.43 | 13.62 | 13.81 | 14 | 14.18 | 14.37 | 14.56 | 14.75 | 14.94
1.09 | 7.42 | 7.61 | 7.8 | 7.99 | 8.17 | 8.36 | 8.55 | 8.74 | 8.93 | 9.11 | 9.3 | 9.49 | 9.68 | 9.86 | 10.05 | 10.24 | 10.43 | 10.61 | 10.8 | 10.99 | 11.18 | 11.37 | 11.55 | 11.74 | 11.93 | 12.12 | 12.3 | 12.49 | 12.68 | 12.87 | 13.06 | 13.24 | 13.43 | 13.62 | 13.81 | 13.99 | 14.18 | 14.37 | 14.56 | 14.74 | 14.93
1.1 | 7.42 | 7.61 | 7.8 | 7.98 | 8.17 | 8.36 | 8.55 | 8.74 | 8.92 | 9.11 | 9.3 | 9.49 | 9.67 | 9.86 | 10.05 | 10.24 | 10.42 | 10.61 | 10.8 | 10.99 | 11.18 | 11.36 | 11.55 | 11.74 | 11.93 | 12.11 | 12.3 | 12.49 | 12.68 | 12.86 | 13.05 | 13.24 | 13.43 | 13.61 | 13.8 | 13.99 | 14.18 | 14.36 | 14.55 | 14.74 | 14.93
1.11 | 7.42 | 7.61 | 7.79 | 7.98 | 8.17 | 8.36 | 8.54 | 8.73 | 8.92 | 9.11 | 9.3 | 9.48 | 9.67 | 9.86 | 10.05 | 10.23 | 10.42 | 10.61 | 10.8 | 10.98 | 11.17 | 11.36 | 11.55 | 11.73 | 11.92 | 12.11 | 12.3 | 12.49 | 12.67 | 12.86 | 13.05 | 13.24 | 13.42 | 13.61 | 13.8 | 13.99 | 14.17 | 14.36 | 14.55 | 14.74 | 14.92
1.12 | 7.42 | 7.6 | 7.79 | 7.98 | 8.17 | 8.35 | 8.54 | 8.73 | 8.92 | 9.11 | 9.29 | 9.48 | 9.67 | 9.86 | 10.04 | 10.23 | 10.42 | 10.61 | 10.79 | 10.98 | 11.17 | 11.36 | 11.54 | 11.73 | 11.92 | 12.11 | 12.29 | 12.48 | 12.67 | 12.86 | 13.04 | 13.23 | 13.42 | 13.61 | 13.79 | 13.98 | 14.17 | 14.36 | 14.54 | 14.73 | 14.92
1.13 | 7.41 | 7.6 | 7.79 | 7.98 | 8.16 | 8.35 | 8.54 | 8.73 | 8.92 | 9.1 | 9.29 | 9.48 | 9.67 | 9.85 | 10.04 | 10.23 | 10.42 | 10.6 | 10.79 | 10.98 | 11.17 | 11.35 | 11.54 | 11.73 | 11.92 | 12.1 | 12.29 | 12.48 | 12.67 | 12.85 | 13.04 | 13.23 | 13.42 | 13.6 | 13.79 | 13.98 | 14.17 | 14.35 | 14.54 | 14.73 | 14.91
1.14 | 7.41 | 7.6 | 7.79 | 7.98 | 8.16 | 8.35 | 8.54 | 8.73 | 8.91 | 9.1 | 9.29 | 9.48 | 9.66 | 9.85 | 10.04 | 10.23 | 10.41 | 10.6 | 10.79 | 10.98 | 11.16 | 11.35 | 11.54 | 11.72 | 11.91 | 12.1 | 12.29 | 12.47 | 12.66 | 12.85 | 13.04 | 13.22 | 13.41 | 13.6 | 13.79 | 13.97 | 14.16 | 14.35 | 14.54 | 14.72 | 14.91
1.15 | 7.41 | 7.6 | 7.79 | 7.97 | 8.16 | 8.35 | 8.54 | 8.72 | 8.91 | 9.1 | 9.29 | 9.47 | 9.66 | 9.85 | 10.03 | 10.22 | 10.41 | 10.6 | 10.78 | 10.97 | 11.16 | 11.35 | 11.53 | 11.72 | 11.91 | 12.1 | 12.28 | 12.47 | 12.66 | 12.85 | 13.03 | 13.22 | 13.41 | 13.6 | 13.78 | 13.97 | 14.16 | 14.34 | 14.53 | 14.72 | 14.91
1.16 | 7.41 | 7.6 | 7.78 | 7.97 | 8.16 | 8.35 | 8.53 | 8.72 | 8.91 | 9.1 | 9.28 | 9.47 | 9.66 | 9.84 | 10.03 | 10.22 | 10.41 | 10.59 | 10.78 | 10.97 | 11.16 | 11.34 | 11.53 | 11.72 | 11.91 | 12.09 | 12.28 | 12.47 | 12.66 | 12.84 | 13.03 | 13.22 | 13.4 | 13.59 | 13.78 | 13.97 | 14.15 | 14.34 | 14.53 | 14.72 | 14.9
1.17 | 7.41 | 7.59 | 7.78 | 7.97 | 8.16 | 8.34 | 8.53 | 8.72 | 8.91 | 9.09 | 9.28 | 9.47 | 9.65 | 9.84 | 10.03 | 10.22 | 10.4 | 10.59 | 10.78 | 10.97 | 11.15 | 11.34 | 11.53 | 11.72 | 11.9 | 12.09 | 12.28 | 12.46 | 12.65 | 12.84 | 13.03 | 13.21 | 13.4 | 13.59 | 13.78 | 13.96 | 14.15 | 14.34 | 14.52 | 14.71 | 14.9
1.18 | 7.4 | 7.59 | 7.78 | 7.97 | 8.15 | 8.34 | 8.53 | 8.72 | 8.9 | 9.09 | 9.28 | 9.46 | 9.65 | 9.84 | 10.03 | 10.21 | 10.4 | 10.59 | 10.78 | 10.96 | 11.15 | 11.34 | 11.52 | 11.71 | 11.9 | 12.09 | 12.27 | 12.46 | 12.65 | 12.84 | 13.02 | 13.21 | 13.4 | 13.58 | 13.77 | 13.96 | 14.15 | 14.33 | 14.52 | 14.71 | 14.89
1.19 | 7.4 | 7.59 | 7.78 | 7.96 | 8.15 | 8.34 | 8.53 | 8.71 | 8.9 | 9.09 | 9.27 | 9.46 | 9.65 | 9.84 | 10.02 | 10.21 | 10.4 | 10.59 | 10.77 | 10.96 | 11.15 | 11.33 | 11.52 | 11.71 | 11.9 | 12.08 | 12.27 | 12.46 | 12.64 | 12.83 | 13.02 | 13.21 | 13.39 | 13.58 | 13.77 | 13.95 | 14.14 | 14.33 | 14.52 | 14.7 | 14.89
1.2 | 7.4 | 7.59 | 7.77 | 7.96 | 8.15 | 8.34 | 8.52 | 8.71 | 8.9 | 9.08 | 9.27 | 9.46 | 9.65 | 9.83 | 10.02 | 10.21 | 10.4 | 10.58 | 10.77 | 10.96 | 11.14 | 11.33 | 11.52 | 11.71 | 11.89 | 12.08 | 12.27 | 12.45 | 12.64 | 12.83 | 13.02 | 13.2 | 13.39 | 13.58 | 13.76 | 13.95 | 14.14 | 14.32 | 14.51 | 14.7 | 14.89
1.21 | 7.4 | 7.59 | 7.77 | 7.96 | 8.15 | 8.33 | 8.52 | 8.71 | 8.9 | 9.08 | 9.27 | 9.46 | 9.64 | 9.83 | 10.02 | 10.21 | 10.39 | 10.58 | 10.77 | 10.95 | 11.14 | 11.33 | 11.52 | 11.7 | 11.89 | 12.08 | 12.26 | 12.45 | 12.64 | 12.82 | 13.01 | 13.2 | 13.39 | 13.57 | 13.76 | 13.95 | 14.13 | 14.32 | 14.51 | 14.69 | 14.88
1.22 | 7.4 | 7.58 | 7.77 | 7.96 | 8.14 | 8.33 | 8.52 | 8.71 | 8.89 | 9.08 | 9.27 | 9.45 | 9.64 | 9.83 | 10.02 | 10.2 | 10.39 | 10.58 | 10.76 | 10.95 | 11.14 | 11.32 | 11.51 | 11.7 | 11.89 | 12.07 | 12.26 | 12.45 | 12.63 | 12.82 | 13.01 | 13.19 | 13.38 | 13.57 | 13.76 | 13.94 | 14.13 | 14.32 | 14.5 | 14.69 | 14.88
1.23 | 7.39 | 7.58 | 7.77 | 7.95 | 8.14 | 8.33 | 8.52 | 8.7 | 8.89 | 9.08 | 9.26 | 9.45 | 9.64 | 9.83 | 10.01 | 10.2 | 10.39 | 10.57 | 10.76 | 10.95 | 11.13 | 11.32 | 11.51 | 11.7 | 11.88 | 12.07 | 12.26 | 12.44 | 12.63 | 12.82 | 13 | 13.19 | 13.38 | 13.57 | 13.75 | 13.94 | 14.13 | 14.31 | 14.5 | 14.69 | 14.87
1.24 | 7.39 | 7.58 | 7.77 | 7.95 | 8.14 | 8.33 | 8.51 | 8.7 | 8.89 | 9.07 | 9.26 | 9.45 | 9.64 | 9.82 | 10.01 | 10.2 | 10.38 | 10.57 | 10.76 | 10.94 | 11.13 | 11.32 | 11.51 | 11.69 | 11.88 | 12.07 | 12.25 | 12.44 | 12.63 | 12.81 | 13 | 13.19 | 13.37 | 13.56 | 13.75 | 13.94 | 14.12 | 14.31 | 14.5 | 14.68 | 14.87
1.25 | 7.39 | 7.58 | 7.76 | 7.95 | 8.14 | 8.32 | 8.51 | 8.7 | 8.89 | 9.07 | 9.26 | 9.45 | 9.63 | 9.82 | 10.01 | 10.19 | 10.38 | 10.57 | 10.75 | 10.94 | 11.13 | 11.32 | 11.5 | 11.69 | 11.88 | 12.06 | 12.25 | 12.44 | 12.62 | 12.81 | 13 | 13.18 | 13.37 | 13.56 | 13.74 | 13.93 | 14.12 | 14.31 | 14.49 | 14.68 | 14.87
1.26 | 7.39 | 7.57 | 7.76 | 7.95 | 8.14 | 8.32 | 8.51 | 8.7 | 8.88 | 9.07 | 9.26 | 9.44 | 9.63 | 9.82 | 10 | 10.19 | 10.38 | 10.56 | 10.75 | 10.94 | 11.13 | 11.31 | 11.5 | 11.69 | 11.87 | 12.06 | 12.25 | 12.43 | 12.62 | 12.81 | 12.99 | 13.18 | 13.37 | 13.55 | 13.74 | 13.93 | 14.11 | 14.3 | 14.49 | 14.67 | 14.86
1.27 | 7.39 | 7.57 | 7.76 | 7.95 | 8.13 | 8.32 | 8.51 | 8.69 | 8.88 | 9.07 | 9.25 | 9.44 | 9.63 | 9.81 | 10 | 10.19 | 10.38 | 10.56 | 10.75 | 10.94 | 11.12 | 11.31 | 11.5 | 11.68 | 11.87 | 12.06 | 12.24 | 12.43 | 12.62 | 12.8 | 12.99 | 13.18 | 13.36 | 13.55 | 13.74 | 13.92 | 14.11 | 14.3 | 14.48 | 14.67 | 14.86
1.28 | 7.38 | 7.57 | 7.76 | 7.94 | 8.13 | 8.32 | 8.5 | 8.69 | 8.88 | 9.06 | 9.25 | 9.44 | 9.63 | 9.81 | 10 | 10.19 | 10.37 | 10.56 | 10.75 | 10.93 | 11.12 | 11.31 | 11.49 | 11.68 | 11.87 | 12.05 | 12.24 | 12.43 | 12.61 | 12.8 | 12.99 | 13.17 | 13.36 | 13.55 | 13.73 | 13.92 | 14.11 | 14.29 | 14.48 | 14.67 | 14.85
1.29 | 7.38 | 7.57 | 7.75 | 7.94 | 8.13 | 8.32 | 8.5 | 8.69 | 8.88 | 9.06 | 9.25 | 9.44 | 9.62 | 9.81 | 10 | 10.18 | 10.37 | 10.56 | 10.74 | 10.93 | 11.12 | 11.3 | 11.49 | 11.68 | 11.86 | 12.05 | 12.24 | 12.42 | 12.61 | 12.8 | 12.98 | 13.17 | 13.36 | 13.54 | 13.73 | 13.92 | 14.1 | 14.29 | 14.48 | 14.66 | 14.85
1.3 | 7.38 | 7.57 | 7.75 | 7.94 | 8.13 | 8.31 | 8.5 | 8.69 | 8.87 | 9.06 | 9.25 | 9.43 | 9.62 | 9.81 | 9.99 | 10.18 | 10.37 | 10.55 | 10.74 | 10.93 | 11.11 | 11.3 | 11.49 | 11.67 | 11.86 | 12.05 | 12.23 | 12.42 | 12.61 | 12.79 | 12.98 | 13.17 | 13.35 | 13.54 | 13.73 | 13.91 | 14.1 | 14.29 | 14.47 | 14.66 | 14.84
1.31 | 7.38 | 7.56 | 7.75 | 7.94 | 8.12 | 8.31 | 8.5 | 8.68 | 8.87 | 9.06 | 9.24 | 9.43 | 9.62 | 9.8 | 9.99 | 10.18 | 10.36 | 10.55 | 10.74 | 10.92 | 11.11 | 11.3 | 11.48 | 11.67 | 11.86 | 12.04 | 12.23 | 12.42 | 12.6 | 12.79 | 12.98 | 13.16 | 13.35 | 13.54 | 13.72 | 13.91 | 14.09 | 14.28 | 14.47 | 14.65 | 14.84
1.32 | 7.38 | 7.56 | 7.75 | 7.93 | 8.12 | 8.31 | 8.49 | 8.68 | 8.87 | 9.05 | 9.24 | 9.43 | 9.61 | 9.8 | 9.99 | 10.17 | 10.36 | 10.55 | 10.73 | 10.92 | 11.11 | 11.29 | 11.48 | 11.67 | 11.85 | 12.04 | 12.23 | 12.41 | 12.6 | 12.79 | 12.97 | 13.16 | 13.34 | 13.53 | 13.72 | 13.9 | 14.09 | 14.28 | 14.46 | 14.65 | 14.84
1.33 | 7.37 | 7.56 | 7.75 | 7.93 | 8.12 | 8.31 | 8.49 | 8.68 | 8.87 | 9.05 | 9.24 | 9.43 | 9.61 | 9.8 | 9.98 | 10.17 | 10.36 | 10.54 | 10.73 | 10.92 | 11.1 | 11.29 | 11.48 | 11.66 | 11.85 | 12.04 | 12.22 | 12.41 | 12.6 | 12.78 | 12.97 | 13.15 | 13.34 | 13.53 | 13.71 | 13.9 | 14.09 | 14.27 | 14.46 | 14.65 | 14.83
1.34 | 7.37 | 7.56 | 7.74 | 7.93 | 8.12 | 8.3 | 8.49 | 8.68 | 8.86 | 9.05 | 9.24 | 9.42 | 9.61 | 9.8 | 9.98 | 10.17 | 10.35 | 10.54 | 10.73 | 10.91 | 11.1 | 11.29 | 11.47 | 11.66 | 11.85 | 12.03 | 12.22 | 12.41 | 12.59 | 12.78 | 12.96 | 13.15 | 13.34 | 13.52 | 13.71 | 13.9 | 14.08 | 14.27 | 14.46 | 14.64 | 14.83
1.35 | 7.37 | 7.56 | 7.74 | 7.93 | 8.11 | 8.3 | 8.49 | 8.67 | 8.86 | 9.05 | 9.23 | 9.42 | 9.61 | 9.79 | 9.98 | 10.17 | 10.35 | 10.54 | 10.72 | 10.91 | 11.1 | 11.28 | 11.47 | 11.66 | 11.84 | 12.03 | 12.22 | 12.4 | 12.59 | 12.77 | 12.96 | 13.15 | 13.33 | 13.52 | 13.71 | 13.89 | 14.08 | 14.27 | 14.45 | 14.64 | 14.82
1.36 | 7.37 | 7.55 | 7.74 | 7.93 | 8.11 | 8.3 | 8.49 | 8.67 | 8.86 | 9.04 | 9.23 | 9.42 | 9.6 | 9.79 | 9.98 | 10.16 | 10.35 | 10.54 | 10.72 | 10.91 | 11.09 | 11.28 | 11.47 | 11.65 | 11.84 | 12.03 | 12.21 | 12.4 | 12.59 | 12.77 | 12.96 | 13.14 | 13.33 | 13.52 | 13.7 | 13.89 | 14.08 | 14.26 | 14.45 | 14.63 | 14.82
1.37 | 7.36 | 7.55 | 7.74 | 7.92 | 8.11 | 8.3 | 8.48 | 8.67 | 8.86 | 9.04 | 9.23 | 9.41 | 9.6 | 9.79 | 9.97 | 10.16 | 10.35 | 10.53 | 10.72 | 10.91 | 11.09 | 11.28 | 11.46 | 11.65 | 11.84 | 12.02 | 12.21 | 12.4 | 12.58 | 12.77 | 12.95 | 13.14 | 13.33 | 13.51 | 13.7 | 13.89 | 14.07 | 14.26 | 14.44 | 14.63 | 14.82
1.38 | 7.36 | 7.55 | 7.74 | 7.92 | 8.11 | 8.29 | 8.48 | 8.67 | 8.85 | 9.04 | 9.23 | 9.41 | 9.6 | 9.78 | 9.97 | 10.16 | 10.34 | 10.53 | 10.72 | 10.9 | 11.09 | 11.27 | 11.46 | 11.65 | 11.83 | 12.02 | 12.21 | 12.39 | 12.58 | 12.76 | 12.95 | 13.14 | 13.32 | 13.51 | 13.7 | 13.88 | 14.07 | 14.25 | 14.44 | 14.63 | 14.81
1.39 | 7.36 | 7.55 | 7.73 | 7.92 | 8.11 | 8.29 | 8.48 | 8.66 | 8.85 | 9.04 | 9.22 | 9.41 | 9.6 | 9.78 | 9.97 | 10.15 | 10.34 | 10.53 | 10.71 | 10.9 | 11.09 | 11.27 | 11.46 | 11.64 | 11.83 | 12.02 | 12.2 | 12.39 | 12.57 | 12.76 | 12.95 | 13.13 | 13.32 | 13.51 | 13.69 | 13.88 | 14.06 | 14.25 | 14.44 | 14.62 | 14.81
1.4 | 7.36 | 7.54 | 7.73 | 7.92 | 8.1 | 8.29 | 8.48 | 8.66 | 8.85 | 9.03 | 9.22 | 9.41 | 9.59 | 9.78 | 9.97 | 10.15 | 10.34 | 10.52 | 10.71 | 10.9 | 11.08 | 11.27 | 11.45 | 11.64 | 11.83 | 12.01 | 12.2 | 12.39 | 12.57 | 12.76 | 12.94 | 13.13 | 13.32 | 13.5 | 13.69 | 13.87 | 14.06 | 14.25 | 14.43 | 14.62 | 14.8
1.41 | 7.36 | 7.54 | 7.73 | 7.91 | 8.1 | 8.29 | 8.47 | 8.66 | 8.85 | 9.03 | 9.22 | 9.4 | 9.59 | 9.78 | 9.96 | 10.15 | 10.33 | 10.52 | 10.71 | 10.89 | 11.08 | 11.27 | 11.45 | 11.64 | 11.82 | 12.01 | 12.2 | 12.38 | 12.57 | 12.75 | 12.94 | 13.13 | 13.31 | 13.5 | 13.68 | 13.87 | 14.06 | 14.24 | 14.43 | 14.61 | 14.8
1.42 | 7.35 | 7.54 | 7.73 | 7.91 | 8.1 | 8.28 | 8.47 | 8.66 | 8.84 | 9.03 | 9.22 | 9.4 | 9.59 | 9.77 | 9.96 | 10.15 | 10.33 | 10.52 | 10.7 | 10.89 | 11.08 | 11.26 | 11.45 | 11.63 | 11.82 | 12.01 | 12.19 | 12.38 | 12.56 | 12.75 | 12.94 | 13.12 | 13.31 | 13.49 | 13.68 | 13.87 | 14.05 | 14.24 | 14.42 | 14.61 | 14.8
1.43 | 7.35 | 7.54 | 7.72 | 7.91 | 8.1 | 8.28 | 8.47 | 8.65 | 8.84 | 9.03 | 9.21 | 9.4 | 9.58 | 9.77 | 9.96 | 10.14 | 10.33 | 10.52 | 10.7 | 10.89 | 11.07 | 11.26 | 11.44 | 11.63 | 11.82 | 12 | 12.19 | 12.37 | 12.56 | 12.75 | 12.93 | 13.12 | 13.3 | 13.49 | 13.68 | 13.86 | 14.05 | 14.23 | 14.42 | 14.61 | 14.79
1.44 | 7.35 | 7.54 | 7.72 | 7.91 | 8.09 | 8.28 | 8.47 | 8.65 | 8.84 | 9.02 | 9.21 | 9.4 | 9.58 | 9.77 | 9.95 | 10.14 | 10.33 | 10.51 | 10.7 | 10.88 | 11.07 | 11.26 | 11.44 | 11.63 | 11.81 | 12 | 12.19 | 12.37 | 12.56 | 12.74 | 12.93 | 13.11 | 13.3 | 13.49 | 13.67 | 13.86 | 14.04 | 14.23 | 14.42 | 14.6 | 14.79
1.45 | 7.35 | 7.53 | 7.72 | 7.91 | 8.09 | 8.28 | 8.46 | 8.65 | 8.84 | 9.02 | 9.21 | 9.39 | 9.58 | 9.77 | 9.95 | 10.14 | 10.32 | 10.51 | 10.7 | 10.88 | 11.07 | 11.25 | 11.44 | 11.62 | 11.81 | 12 | 12.18 | 12.37 | 12.55 | 12.74 | 12.93 | 13.11 | 13.3 | 13.48 | 13.67 | 13.85 | 14.04 | 14.23 | 14.41 | 14.6 | 14.78
1.46 | 7.35 | 7.53 | 7.72 | 7.9 | 8.09 | 8.28 | 8.46 | 8.65 | 8.83 | 9.02 | 9.21 | 9.39 | 9.58 | 9.76 | 9.95 | 10.13 | 10.32 | 10.51 | 10.69 | 10.88 | 11.06 | 11.25 | 11.44 | 11.62 | 11.81 | 11.99 | 12.18 | 12.36 | 12.55 | 12.74 | 12.92 | 13.11 | 13.29 | 13.48 | 13.67 | 13.85 | 14.04 | 14.22 | 14.41 | 14.59 | 14.78
1.47 | 7.34 | 7.53 | 7.72 | 7.9 | 8.09 | 8.27 | 8.46 | 8.65 | 8.83 | 9.02 | 9.2 | 9.39 | 9.57 | 9.76 | 9.95 | 10.13 | 10.32 | 10.5 | 10.69 | 10.87 | 11.06 | 11.25 | 11.43 | 11.62 | 11.8 | 11.99 | 12.18 | 12.36 | 12.55 | 12.73 | 12.92 | 13.1 | 13.29 | 13.48 | 13.66 | 13.85 | 14.03 | 14.22 | 14.4 | 14.59 | 14.78
1.48 | 7.34 | 7.53 | 7.71 | 7.9 | 8.09 | 8.27 | 8.46 | 8.64 | 8.83 | 9.01 | 9.2 | 9.39 | 9.57 | 9.76 | 9.94 | 10.13 | 10.31 | 10.5 | 10.69 | 10.87 | 11.06 | 11.24 | 11.43 | 11.61 | 11.8 | 11.99 | 12.17 | 12.36 | 12.54 | 12.73 | 12.91 | 13.1 | 13.29 | 13.47 | 13.66 | 13.84 | 14.03 | 14.21 | 14.4 | 14.59 | 14.77
1.49 | 7.34 | 7.53 | 7.71 | 7.9 | 8.08 | 8.27 | 8.45 | 8.64 | 8.83 | 9.01 | 9.2 | 9.38 | 9.57 | 9.75 | 9.94 | 10.13 | 10.31 | 10.5 | 10.68 | 10.87 | 11.05 | 11.24 | 11.43 | 11.61 | 11.8 | 11.98 | 12.17 | 12.35 | 12.54 | 12.73 | 12.91 | 13.1 | 13.28 | 13.47 | 13.65 | 13.84 | 14.02 | 14.21 | 14.4 | 14.58 | 14.77
1.5 | 7.34 | 7.52 | 7.71 | 7.9 | 8.08 | 8.27 | 8.45 | 8.64 | 8.82 | 9.01 | 9.19 | 9.38 | 9.57 | 9.75 | 9.94 | 10.12 | 10.31 | 10.49 | 10.68 | 10.87 | 11.05 | 11.24 | 11.42 | 11.61 | 11.79 | 11.98 | 12.17 | 12.35 | 12.54 | 12.72 | 12.91 | 13.09 | 13.28 | 13.46 | 13.65 | 13.84 | 14.02 | 14.21 | 14.39 | 14.58 | 14.76
1.51 | 7.34 | 7.52 | 7.71 | 7.89 | 8.08 | 8.26 | 8.45 | 8.64 | 8.82 | 9.01 | 9.19 | 9.38 | 9.56 | 9.75 | 9.93 | 10.12 | 10.31 | 10.49 | 10.68 | 10.86 | 11.05 | 11.23 | 11.42 | 11.61 | 11.79 | 11.98 | 12.16 | 12.35 | 12.53 | 12.72 | 12.9 | 13.09 | 13.28 | 13.46 | 13.65 | 13.83 | 14.02 | 14.2 | 14.39 | 14.57 | 14.76
1.52 | 7.33 | 7.52 | 7.71 | 7.89 | 8.08 | 8.26 | 8.45 | 8.63 | 8.82 | 9 | 9.19 | 9.38 | 9.56 | 9.75 | 9.93 | 10.12 | 10.3 | 10.49 | 10.67 | 10.86 | 11.05 | 11.23 | 11.42 | 11.6 | 11.79 | 11.97 | 12.16 | 12.34 | 12.53 | 12.72 | 12.9 | 13.09 | 13.27 | 13.46 | 13.64 | 13.83 | 14.01 | 14.2 | 14.38 | 14.57 | 14.76
1.53 | 7.33 | 7.52 | 7.7 | 7.89 | 8.07 | 8.26 | 8.45 | 8.63 | 8.82 | 9 | 9.19 | 9.37 | 9.56 | 9.74 | 9.93 | 10.11 | 10.3 | 10.49 | 10.67 | 10.86 | 11.04 | 11.23 | 11.41 | 11.6 | 11.78 | 11.97 | 12.16 | 12.34 | 12.53 | 12.71 | 12.9 | 13.08 | 13.27 | 13.45 | 13.64 | 13.82 | 14.01 | 14.19 | 14.38 | 14.57 | 14.75
1.54 | 7.33 | 7.52 | 7.7 | 7.89 | 8.07 | 8.26 | 8.44 | 8.63 | 8.81 | 9 | 9.18 | 9.37 | 9.56 | 9.74 | 9.93 | 10.11 | 10.3 | 10.48 | 10.67 | 10.85 | 11.04 | 11.22 | 11.41 | 11.6 | 11.78 | 11.97 | 12.15 | 12.34 | 12.52 | 12.71 | 12.89 | 13.08 | 13.26 | 13.45 | 13.63 | 13.82 | 14.01 | 14.19 | 14.38 | 14.56 | 14.75
1.55 | 7.33 | 7.51 | 7.7 | 7.88 | 8.07 | 8.25 | 8.44 | 8.63 | 8.81 | 9 | 9.18 | 9.37 | 9.55 | 9.74 | 9.92 | 10.11 | 10.29 | 10.48 | 10.67 | 10.85 | 11.04 | 11.22 | 11.41 | 11.59 | 11.78 | 11.96 | 12.15 | 12.33 | 12.52 | 12.7 | 12.89 | 13.08 | 13.26 | 13.45 | 13.63 | 13.82 | 14 | 14.19 | 14.37 | 14.56 | 14.74
1.56 | 7.33 | 7.51 | 7.7 | 7.88 | 8.07 | 8.25 | 8.44 | 8.62 | 8.81 | 8.99 | 9.18 | 9.37 | 9.55 | 9.74 | 9.92 | 10.11 | 10.29 | 10.48 | 10.66 | 10.85 | 11.03 | 11.22 | 11.4 | 11.59 | 11.77 | 11.96 | 12.15 | 12.33 | 12.52 | 12.7 | 12.89 | 13.07 | 13.26 | 13.44 | 13.63 | 13.81 | 14 | 14.18 | 14.37 | 14.55 | 14.74
1.57 | 7.32 | 7.51 | 7.69 | 7.88 | 8.07 | 8.25 | 8.44 | 8.62 | 8.81 | 8.99 | 9.18 | 9.36 | 9.55 | 9.73 | 9.92 | 10.1 | 10.29 | 10.47 | 10.66 | 10.84 | 11.03 | 11.22 | 11.4 | 11.59 | 11.77 | 11.96 | 12.14 | 12.33 | 12.51 | 12.7 | 12.88 | 13.07 | 13.25 | 13.44 | 13.62 | 13.81 | 13.99 | 14.18 | 14.36 | 14.55 | 14.73
1.58 | 7.32 | 7.51 | 7.69 | 7.88 | 8.06 | 8.25 | 8.43 | 8.62 | 8.8 | 8.99 | 9.17 | 9.36 | 9.55 | 9.73 | 9.92 | 10.1 | 10.29 | 10.47 | 10.66 | 10.84 | 11.03 | 11.21 | 11.4 | 11.58 | 11.77 | 11.95 | 12.14 | 12.32 | 12.51 | 12.69 | 12.88 | 13.06 | 13.25 | 13.43 | 13.62 | 13.81 | 13.99 | 14.18 | 14.36 | 14.55 | 14.73
1.59 | 7.32 | 7.5 | 7.69 | 7.88 | 8.06 | 8.25 | 8.43 | 8.62 | 8.8 | 8.99 | 9.17 | 9.36 | 9.54 | 9.73 | 9.91 | 10.1 | 10.28 | 10.47 | 10.65 | 10.84 | 11.02 | 11.21 | 11.39 | 11.58 | 11.76 | 11.95 | 12.14 | 12.32 | 12.51 | 12.69 | 12.88 | 13.06 | 13.25 | 13.43 | 13.62 | 13.8 | 13.99 | 14.17 | 14.36 | 14.54 | 14.73
1.6 | 7.32 | 7.5 | 7.69 | 7.87 | 8.06 | 8.24 | 8.43 | 8.61 | 8.8 | 8.98 | 9.17 | 9.35 | 9.54 | 9.72 | 9.91 | 10.1 | 10.28 | 10.47 | 10.65 | 10.84 | 11.02 | 11.21 | 11.39 | 11.58 | 11.76 | 11.95 | 12.13 | 12.32 | 12.5 | 12.69 | 12.87 | 13.06 | 13.24 | 13.43 | 13.61 | 13.8 | 13.98 | 14.17 | 14.35 | 14.54 | 14.72
1.61 | 7.32 | 7.5 | 7.69 | 7.87 | 8.06 | 8.24 | 8.43 | 8.61 | 8.8 | 8.98 | 9.17 | 9.35 | 9.54 | 9.72 | 9.91 | 10.09 | 10.28 | 10.46 | 10.65 | 10.83 | 11.02 | 11.2 | 11.39 | 11.57 | 11.76 | 11.94 | 12.13 | 12.31 | 12.5 | 12.68 | 12.87 | 13.05 | 13.24 | 13.42 | 13.61 | 13.79 | 13.98 | 14.16 | 14.35 | 14.53 | 14.72
1.62 | 7.31 | 7.5 | 7.68 | 7.87 | 8.05 | 8.24 | 8.42 | 8.61 | 8.79 | 8.98 | 9.16 | 9.35 | 9.53 | 9.72 | 9.9 | 10.09 | 10.27 | 10.46 | 10.64 | 10.83 | 11.01 | 11.2 | 11.38 | 11.57 | 11.76 | 11.94 | 12.13 | 12.31 | 12.5 | 12.68 | 12.87 | 13.05 | 13.23 | 13.42 | 13.6 | 13.79 | 13.97 | 14.16 | 14.34 | 14.53 | 14.71
1.63 | 7.31 | 7.5 | 7.68 | 7.87 | 8.05 | 8.24 | 8.42 | 8.61 | 8.79 | 8.98 | 9.16 | 9.35 | 9.53 | 9.72 | 9.9 | 10.09 | 10.27 | 10.46 | 10.64 | 10.83 | 11.01 | 11.2 | 11.38 | 11.57 | 11.75 | 11.94 | 12.12 | 12.31 | 12.49 | 12.68 | 12.86 | 13.05 | 13.23 | 13.42 | 13.6 | 13.79 | 13.97 | 14.16 | 14.34 | 14.53 | 14.71
1.64 | 7.31 | 7.49 | 7.68 | 7.86 | 8.05 | 8.23 | 8.42 | 8.6 | 8.79 | 8.97 | 9.16 | 9.34 | 9.53 | 9.71 | 9.9 | 10.08 | 10.27 | 10.45 | 10.64 | 10.82 | 11.01 | 11.19 | 11.38 | 11.56 | 11.75 | 11.93 | 12.12 | 12.3 | 12.49 | 12.67 | 12.86 | 13.04 | 13.23 | 13.41 | 13.6 | 13.78 | 13.97 | 14.15 | 14.34 | 14.52 | 14.71
1.65 | 7.31 | 7.49 | 7.68 | 7.86 | 8.05 | 8.23 | 8.42 | 8.6 | 8.79 | 8.97 | 9.16 | 9.34 | 9.53 | 9.71 | 9.9 | 10.08 | 10.27 | 10.45 | 10.64 | 10.82 | 11.01 | 11.19 | 11.38 | 11.56 | 11.75 | 11.93 | 12.12 | 12.3 | 12.48 | 12.67 | 12.85 | 13.04 | 13.22 | 13.41 | 13.59 | 13.78 | 13.96 | 14.15 | 14.33 | 14.52 | 14.7
1.66 | 7.31 | 7.49 | 7.68 | 7.86 | 8.04 | 8.23 | 8.41 | 8.6 | 8.78 | 8.97 | 9.15 | 9.34 | 9.52 | 9.71 | 9.89 | 10.08 | 10.26 | 10.45 | 10.63 | 10.82 | 11 | 11.19 | 11.37 | 11.56 | 11.74 | 11.93 | 12.11 | 12.3 | 12.48 | 12.67 | 12.85 | 13.04 | 13.22 | 13.41 | 13.59 | 13.77 | 13.96 | 14.14 | 14.33 | 14.51 | 14.7
1.67 | 7.3 | 7.49 | 7.67 | 7.86 | 8.04 | 8.23 | 8.41 | 8.6 | 8.78 | 8.97 | 9.15 | 9.34 | 9.52 | 9.71 | 9.89 | 10.08 | 10.26 | 10.45 | 10.63 | 10.81 | 11 | 11.18 | 11.37 | 11.55 | 11.74 | 11.92 | 12.11 | 12.29 | 12.48 | 12.66 | 12.85 | 13.03 | 13.22 | 13.4 | 13.59 | 13.77 | 13.96 | 14.14 | 14.33 | 14.51 | 14.69
1.68 | 7.3 | 7.49 | 7.67 | 7.86 | 8.04 | 8.23 | 8.41 | 8.59 | 8.78 | 8.96 | 9.15 | 9.33 | 9.52 | 9.7 | 9.89 | 10.07 | 10.26 | 10.44 | 10.63 | 10.81 | 11 | 11.18 | 11.37 | 11.55 | 11.74 | 11.92 | 12.11 | 12.29 | 12.47 | 12.66 | 12.84 | 13.03 | 13.21 | 13.4 | 13.58 | 13.77 | 13.95 | 14.14 | 14.32 | 14.51 | 14.69
1.69 | 7.3 | 7.48 | 7.67 | 7.85 | 8.04 | 8.22 | 8.41 | 8.59 | 8.78 | 8.96 | 9.15 | 9.33 | 9.52 | 9.7 | 9.89 | 10.07 | 10.25 | 10.44 | 10.62 | 10.81 | 10.99 | 11.18 | 11.36 | 11.55 | 11.73 | 11.92 | 12.1 | 12.29 | 12.47 | 12.66 | 12.84 | 13.02 | 13.21 | 13.39 | 13.58 | 13.76 | 13.95 | 14.13 | 14.32 | 14.5 | 14.69
1.7 | 7.3 | 7.48 | 7.67 | 7.85 | 8.04 | 8.22 | 8.41 | 8.59 | 8.77 | 8.96 | 9.14 | 9.33 | 9.51 | 9.7 | 9.88 | 10.07 | 10.25 | 10.44 | 10.62 | 10.81 | 10.99 | 11.18 | 11.36 | 11.54 | 11.73 | 11.91 | 12.1 | 12.28 | 12.47 | 12.65 | 12.84 | 13.02 | 13.21 | 13.39 | 13.58 | 13.76 | 13.94 | 14.13 | 14.31 | 14.5 | 14.68
1.71 | 7.29 | 7.48 | 7.66 | 7.85 | 8.03 | 8.22 | 8.4 | 8.59 | 8.77 | 8.96 | 9.14 | 9.33 | 9.51 | 9.7 | 9.88 | 10.06 | 10.25 | 10.43 | 10.62 | 10.8 | 10.99 | 11.17 | 11.36 | 11.54 | 11.73 | 11.91 | 12.1 | 12.28 | 12.46 | 12.65 | 12.83 | 13.02 | 13.2 | 13.39 | 13.57 | 13.76 | 13.94 | 14.12 | 14.31 | 14.49 | 14.68
1.72 | 7.29 | 7.48 | 7.66 | 7.85 | 8.03 | 8.22 | 8.4 | 8.59 | 8.77 | 8.95 | 9.14 | 9.32 | 9.51 | 9.69 | 9.88 | 10.06 | 10.25 | 10.43 | 10.62 | 10.8 | 10.98 | 11.17 | 11.35 | 11.54 | 11.72 | 11.91 | 12.09 | 12.28 | 12.46 | 12.65 | 12.83 | 13.01 | 13.2 | 13.38 | 13.57 | 13.75 | 13.94 | 14.12 | 14.31 | 14.49 | 14.67
1.73 | 7.29 | 7.48 | 7.66 | 7.84 | 8.03 | 8.21 | 8.4 | 8.58 | 8.77 | 8.95 | 9.14 | 9.32 | 9.51 | 9.69 | 9.87 | 10.06 | 10.24 | 10.43 | 10.61 | 10.8 | 10.98 | 11.17 | 11.35 | 11.53 | 11.72 | 11.9 | 12.09 | 12.27 | 12.46 | 12.64 | 12.83 | 13.01 | 13.2 | 13.38 | 13.56 | 13.75 | 13.93 | 14.12 | 14.3 | 14.49 | 14.67
1.74 | 7.29 | 7.47 | 7.66 | 7.84 | 8.03 | 8.21 | 8.4 | 8.58 | 8.76 | 8.95 | 9.13 | 9.32 | 9.5 | 9.69 | 9.87 | 10.06 | 10.24 | 10.43 | 10.61 | 10.79 | 10.98 | 11.16 | 11.35 | 11.53 | 11.72 | 11.9 | 12.09 | 12.27 | 12.45 | 12.64 | 12.82 | 13.01 | 13.19 | 13.38 | 13.56 | 13.74 | 13.93 | 14.11 | 14.3 | 14.48 | 14.67
1.75 | 7.29 | 7.47 | 7.66 | 7.84 | 8.02 | 8.21 | 8.39 | 8.58 | 8.76 | 8.95 | 9.13 | 9.32 | 9.5 | 9.68 | 9.87 | 10.05 | 10.24 | 10.42 | 10.61 | 10.79 | 10.98 | 11.16 | 11.34 | 11.53 | 11.71 | 11.9 | 12.08 | 12.27 | 12.45 | 12.63 | 12.82 | 13 | 13.19 | 13.37 | 13.56 | 13.74 | 13.93 | 14.11 | 14.29 | 14.48 | 14.66
1.76 | 7.28 | 7.47 | 7.65 | 7.84 | 8.02 | 8.21 | 8.39 | 8.58 | 8.76 | 8.94 | 9.13 | 9.31 | 9.5 | 9.68 | 9.87 | 10.05 | 10.24 | 10.42 | 10.6 | 10.79 | 10.97 | 11.16 | 11.34 | 11.53 | 11.71 | 11.89 | 12.08 | 12.26 | 12.45 | 12.63 | 12.82 | 13 | 13.18 | 13.37 | 13.55 | 13.74 | 13.92 | 14.11 | 14.29 | 14.47 | 14.66
1.77 | 7.28 | 7.47 | 7.65 | 7.84 | 8.02 | 8.2 | 8.39 | 8.57 | 8.76 | 8.94 | 9.13 | 9.31 | 9.5 | 9.68 | 9.86 | 10.05 | 10.23 | 10.42 | 10.6 | 10.79 | 10.97 | 11.15 | 11.34 | 11.52 | 11.71 | 11.89 | 12.08 | 12.26 | 12.44 | 12.63 | 12.81 | 13 | 13.18 | 13.36 | 13.55 | 13.73 | 13.92 | 14.1 | 14.29 | 14.47 | 14.65
1.78 | 7.28 | 7.47 | 7.65 | 7.83 | 8.02 | 8.2 | 8.39 | 8.57 | 8.76 | 8.94 | 9.12 | 9.31 | 9.49 | 9.68 | 9.86 | 10.05 | 10.23 | 10.41 | 10.6 | 10.78 | 10.97 | 11.15 | 11.33 | 11.52 | 11.7 | 11.89 | 12.07 | 12.26 | 12.44 | 12.62 | 12.81 | 12.99 | 13.18 | 13.36 | 13.55 | 13.73 | 13.91 | 14.1 | 14.28 | 14.47 | 14.65
1.79 | 7.28 | 7.46 | 7.65 | 7.83 | 8.02 | 8.2 | 8.38 | 8.57 | 8.75 | 8.94 | 9.12 | 9.31 | 9.49 | 9.67 | 9.86 | 10.04 | 10.23 | 10.41 | 10.6 | 10.78 | 10.96 | 11.15 | 11.33 | 11.52 | 11.7 | 11.88 | 12.07 | 12.25 | 12.44 | 12.62 | 12.81 | 12.99 | 13.17 | 13.36 | 13.54 | 13.73 | 13.91 | 14.09 | 14.28 | 14.46 | 14.65
1.8 | 7.28 | 7.46 | 7.65 | 7.83 | 8.01 | 8.2 | 8.38 | 8.57 | 8.75 | 8.93 | 9.12 | 9.3 | 9.49 | 9.67 | 9.86 | 10.04 | 10.22 | 10.41 | 10.59 | 10.78 | 10.96 | 11.14 | 11.33 | 11.51 | 11.7 | 11.88 | 12.07 | 12.25 | 12.43 | 12.62 | 12.8 | 12.99 | 13.17 | 13.35 | 13.54 | 13.72 | 13.91 | 14.09 | 14.27 | 14.46 | 14.64
1.81 | 7.27 | 7.46 | 7.64 | 7.83 | 8.01 | 8.2 | 8.38 | 8.56 | 8.75 | 8.93 | 9.12 | 9.3 | 9.48 | 9.67 | 9.85 | 10.04 | 10.22 | 10.41 | 10.59 | 10.77 | 10.96 | 11.14 | 11.33 | 11.51 | 11.69 | 11.88 | 12.06 | 12.25 | 12.43 | 12.61 | 12.8 | 12.98 | 13.17 | 13.35 | 13.53 | 13.72 | 13.9 | 14.09 | 14.27 | 14.45 | 14.64
1.82 | 7.27 | 7.46 | 7.64 | 7.83 | 8.01 | 8.19 | 8.38 | 8.56 | 8.75 | 8.93 | 9.11 | 9.3 | 9.48 | 9.67 | 9.85 | 10.03 | 10.22 | 10.4 | 10.59 | 10.77 | 10.95 | 11.14 | 11.32 | 11.51 | 11.69 | 11.87 | 12.06 | 12.24 | 12.43 | 12.61 | 12.79 | 12.98 | 13.16 | 13.35 | 13.53 | 13.71 | 13.9 | 14.08 | 14.27 | 14.45 | 14.63
1.83 | 7.27 | 7.45 | 7.64 | 7.82 | 8.01 | 8.19 | 8.38 | 8.56 | 8.74 | 8.93 | 9.11 | 9.3 | 9.48 | 9.66 | 9.85 | 10.03 | 10.22 | 10.4 | 10.58 | 10.77 | 10.95 | 11.14 | 11.32 | 11.5 | 11.69 | 11.87 | 12.06 | 12.24 | 12.42 | 12.61 | 12.79 | 12.97 | 13.16 | 13.34 | 13.53 | 13.71 | 13.89 | 14.08 | 14.26 | 14.45 | 14.63
1.84 | 7.27 | 7.45 | 7.64 | 7.82 | 8 | 8.19 | 8.37 | 8.56 | 8.74 | 8.92 | 9.11 | 9.29 | 9.48 | 9.66 | 9.84 | 10.03 | 10.21 | 10.4 | 10.58 | 10.76 | 10.95 | 11.13 | 11.32 | 11.5 | 11.68 | 11.87 | 12.05 | 12.24 | 12.42 | 12.6 | 12.79 | 12.97 | 13.16 | 13.34 | 13.52 | 13.71 | 13.89 | 14.07 | 14.26 | 14.44 | 14.63
1.85 | 7.27 | 7.45 | 7.63 | 7.82 | 8 | 8.19 | 8.37 | 8.55 | 8.74 | 8.92 | 9.11 | 9.29 | 9.47 | 9.66 | 9.84 | 10.03 | 10.21 | 10.39 | 10.58 | 10.76 | 10.95 | 11.13 | 11.31 | 11.5 | 11.68 | 11.86 | 12.05 | 12.23 | 12.42 | 12.6 | 12.78 | 12.97 | 13.15 | 13.34 | 13.52 | 13.7 | 13.89 | 14.07 | 14.25 | 14.44 | 14.62
1.86 | 7.26 | 7.45 | 7.63 | 7.82 | 8 | 8.18 | 8.37 | 8.55 | 8.74 | 8.92 | 9.1 | 9.29 | 9.47 | 9.66 | 9.84 | 10.02 | 10.21 | 10.39 | 10.57 | 10.76 | 10.94 | 11.13 | 11.31 | 11.49 | 11.68 | 11.86 | 12.05 | 12.23 | 12.41 | 12.6 | 12.78 | 12.96 | 13.15 | 13.33 | 13.52 | 13.7 | 13.88 | 14.07 | 14.25 | 14.43 | 14.62
1.87 | 7.26 | 7.45 | 7.63 | 7.81 | 8 | 8.18 | 8.37 | 8.55 | 8.73 | 8.92 | 9.1 | 9.28 | 9.47 | 9.65 | 9.84 | 10.02 | 10.2 | 10.39 | 10.57 | 10.76 | 10.94 | 11.12 | 11.31 | 11.49 | 11.67 | 11.86 | 12.04 | 12.23 | 12.41 | 12.59 | 12.78 | 12.96 | 13.14 | 13.33 | 13.51 | 13.7 | 13.88 | 14.06 | 14.25 | 14.43 | 14.61
1.88 | 7.26 | 7.44 | 7.63 | 7.81 | 8 | 8.18 | 8.36 | 8.55 | 8.73 | 8.91 | 9.1 | 9.28 | 9.47 | 9.65 | 9.83 | 10.02 | 10.2 | 10.39 | 10.57 | 10.75 | 10.94 | 11.12 | 11.3 | 11.49 | 11.67 | 11.85 | 12.04 | 12.22 | 12.41 | 12.59 | 12.77 | 12.96 | 13.14 | 13.32 | 13.51 | 13.69 | 13.88 | 14.06 | 14.24 | 14.43 | 14.61
1.89 | 7.26 | 7.44 | 7.63 | 7.81 | 7.99 | 8.18 | 8.36 | 8.54 | 8.73 | 8.91 | 9.1 | 9.28 | 9.46 | 9.65 | 9.83 | 10.01 | 10.2 | 10.38 | 10.57 | 10.75 | 10.93 | 11.12 | 11.3 | 11.48 | 11.67 | 11.85 | 12.04 | 12.22 | 12.4 | 12.59 | 12.77 | 12.95 | 13.14 | 13.32 | 13.5 | 13.69 | 13.87 | 14.06 | 14.24 | 14.42 | 14.61
1.9 | 7.26 | 7.44 | 7.62 | 7.81 | 7.99 | 8.18 | 8.36 | 8.54 | 8.73 | 8.91 | 9.09 | 9.28 | 9.46 | 9.64 | 9.83 | 10.01 | 10.2 | 10.38 | 10.56 | 10.75 | 10.93 | 11.11 | 11.3 | 11.48 | 11.66 | 11.85 | 12.03 | 12.22 | 12.4 | 12.58 | 12.77 | 12.95 | 13.13 | 13.32 | 13.5 | 13.68 | 13.87 | 14.05 | 14.24 | 14.42 | 14.6
1.91 | 7.25 | 7.44 | 7.62 | 7.81 | 7.99 | 8.17 | 8.36 | 8.54 | 8.72 | 8.91 | 9.09 | 9.27 | 9.46 | 9.64 | 9.83 | 10.01 | 10.19 | 10.38 | 10.56 | 10.74 | 10.93 | 11.11 | 11.29 | 11.48 | 11.66 | 11.85 | 12.03 | 12.21 | 12.4 | 12.58 | 12.76 | 12.95 | 13.13 | 13.31 | 13.5 | 13.68 | 13.86 | 14.05 | 14.23 | 14.41 | 14.6
1.92 | 7.25 | 7.44 | 7.62 | 7.8 | 7.99 | 8.17 | 8.35 | 8.54 | 8.72 | 8.91 | 9.09 | 9.27 | 9.46 | 9.64 | 9.82 | 10.01 | 10.19 | 10.37 | 10.56 | 10.74 | 10.92 | 11.11 | 11.29 | 11.47 | 11.66 | 11.84 | 12.03 | 12.21 | 12.39 | 12.58 | 12.76 | 12.94 | 13.13 | 13.31 | 13.49 | 13.68 | 13.86 | 14.04 | 14.23 | 14.41 | 14.59
1.93 | 7.25 | 7.43 | 7.62 | 7.8 | 7.98 | 8.17 | 8.35 | 8.54 | 8.72 | 8.9 | 9.09 | 9.27 | 9.45 | 9.64 | 9.82 | 10 | 10.19 | 10.37 | 10.55 | 10.74 | 10.92 | 11.1 | 11.29 | 11.47 | 11.66 | 11.84 | 12.02 | 12.21 | 12.39 | 12.57 | 12.76 | 12.94 | 13.12 | 13.31 | 13.49 | 13.67 | 13.86 | 14.04 | 14.22 | 14.41 | 14.59
1.94 | 7.25 | 7.43 | 7.62 | 7.8 | 7.98 | 8.17 | 8.35 | 8.53 | 8.72 | 8.9 | 9.08 | 9.27 | 9.45 | 9.63 | 9.82 | 10 | 10.18 | 10.37 | 10.55 | 10.73 | 10.92 | 11.1 | 11.29 | 11.47 | 11.65 | 11.84 | 12.02 | 12.2 | 12.39 | 12.57 | 12.75 | 12.94 | 13.12 | 13.3 | 13.49 | 13.67 | 13.85 | 14.04 | 14.22 | 14.4 | 14.59
1.95 | 7.25 | 7.43 | 7.61 | 7.8 | 7.98 | 8.16 | 8.35 | 8.53 | 8.71 | 8.9 | 9.08 | 9.26 | 9.45 | 9.63 | 9.81 | 10 | 10.18 | 10.37 | 10.55 | 10.73 | 10.92 | 11.1 | 11.28 | 11.47 | 11.65 | 11.83 | 12.02 | 12.2 | 12.38 | 12.57 | 12.75 | 12.93 | 13.12 | 13.3 | 13.48 | 13.67 | 13.85 | 14.03 | 14.22 | 14.4 | 14.58
1.96 | 7.24 | 7.43 | 7.61 | 7.79 | 7.98 | 8.16 | 8.34 | 8.53 | 8.71 | 8.9 | 9.08 | 9.26 | 9.45 | 9.63 | 9.81 | 10 | 10.18 | 10.36 | 10.55 | 10.73 | 10.91 | 11.1 | 11.28 | 11.46 | 11.65 | 11.83 | 12.01 | 12.2 | 12.38 | 12.56 | 12.75 | 12.93 | 13.11 | 13.3 | 13.48 | 13.66 | 13.85 | 14.03 | 14.21 | 14.4 | 14.58
1.97 | 7.24 | 7.43 | 7.61 | 7.79 | 7.98 | 8.16 | 8.34 | 8.53 | 8.71 | 8.89 | 9.08 | 9.26 | 9.44 | 9.63 | 9.81 | 9.99 | 10.18 | 10.36 | 10.54 | 10.73 | 10.91 | 11.09 | 11.28 | 11.46 | 11.64 | 11.83 | 12.01 | 12.19 | 12.38 | 12.56 | 12.74 | 12.93 | 13.11 | 13.29 | 13.48 | 13.66 | 13.84 | 14.02 | 14.21 | 14.39 | 14.57
1.98 | 7.24 | 7.42 | 7.61 | 7.79 | 7.97 | 8.16 | 8.34 | 8.52 | 8.71 | 8.89 | 9.07 | 9.26 | 9.44 | 9.62 | 9.81 | 9.99 | 10.17 | 10.36 | 10.54 | 10.72 | 10.91 | 11.09 | 11.27 | 11.46 | 11.64 | 11.82 | 12.01 | 12.19 | 12.37 | 12.56 | 12.74 | 12.92 | 13.1 | 13.29 | 13.47 | 13.65 | 13.84 | 14.02 | 14.2 | 14.39 | 14.57
1.99 | 7.24 | 7.42 | 7.6 | 7.79 | 7.97 | 8.15 | 8.34 | 8.52 | 8.7 | 8.89 | 9.07 | 9.25 | 9.44 | 9.62 | 9.8 | 9.99 | 10.17 | 10.35 | 10.54 | 10.72 | 10.9 | 11.09 | 11.27 | 11.45 | 11.64 | 11.82 | 12 | 12.19 | 12.37 | 12.55 | 12.74 | 12.92 | 13.1 | 13.28 | 13.47 | 13.65 | 13.83 | 14.02 | 14.2 | 14.38 | 14.57
2 | 7.24 | 7.42 | 7.6 | 7.79 | 7.97 | 8.15 | 8.34 | 8.52 | 8.7 | 8.89 | 9.07 | 9.25 | 9.43 | 9.62 | 9.8 | 9.98 | 10.17 | 10.35 | 10.53 | 10.72 | 10.9 | 11.08 | 11.27 | 11.45 | 11.63 | 11.82 | 12 | 12.18 | 12.37 | 12.55 | 12.73 | 12.91 | 13.1 | 13.28 | 13.46 | 13.65 | 13.83 | 14.01 | 14.2 | 14.38 | 14.56
2.01 | 7.23 | 7.42 | 7.6 | 7.78 | 7.97 | 8.15 | 8.33 | 8.52 | 8.7 | 8.88 | 9.07 | 9.25 | 9.43 | 9.62 | 9.8 | 9.98 | 10.16 | 10.35 | 10.53 | 10.71 | 10.9 | 11.08 | 11.26 | 11.45 | 11.63 | 11.81 | 12 | 12.18 | 12.36 | 12.55 | 12.73 | 12.91 | 13.09 | 13.28 | 13.46 | 13.64 | 13.83 | 14.01 | 14.19 | 14.38 | 14.56
2.02 | 7.23 | 7.42 | 7.6 | 7.78 | 7.96 | 8.15 | 8.33 | 8.51 | 8.7 | 8.88 | 9.06 | 9.25 | 9.43 | 9.61 | 9.8 | 9.98 | 10.16 | 10.35 | 10.53 | 10.71 | 10.89 | 11.08 | 11.26 | 11.44 | 11.63 | 11.81 | 11.99 | 12.18 | 12.36 | 12.54 | 12.72 | 12.91 | 13.09 | 13.27 | 13.46 | 13.64 | 13.82 | 14.01 | 14.19 | 14.37 | 14.55
2.03 | 7.23 | 7.41 | 7.6 | 7.78 | 7.96 | 8.15 | 8.33 | 8.51 | 8.69 | 8.88 | 9.06 | 9.24 | 9.43 | 9.61 | 9.79 | 9.98 | 10.16 | 10.34 | 10.53 | 10.71 | 10.89 | 11.07 | 11.26 | 11.44 | 11.62 | 11.81 | 11.99 | 12.17 | 12.36 | 12.54 | 12.72 | 12.9 | 13.09 | 13.27 | 13.45 | 13.64 | 13.82 | 14 | 14.18 | 14.37 | 14.55
2.04 | 7.23 | 7.41 | 7.59 | 7.78 | 7.96 | 8.14 | 8.33 | 8.51 | 8.69 | 8.88 | 9.06 | 9.24 | 9.42 | 9.61 | 9.79 | 9.97 | 10.16 | 10.34 | 10.52 | 10.71 | 10.89 | 11.07 | 11.25 | 11.44 | 11.62 | 11.8 | 11.99 | 12.17 | 12.35 | 12.53 | 12.72 | 12.9 | 13.08 | 13.27 | 13.45 | 13.63 | 13.82 | 14 | 14.18 | 14.36 | 14.55
2.05 | 7.23 | 7.41 | 7.59 | 7.78 | 7.96 | 8.14 | 8.32 | 8.51 | 8.69 | 8.87 | 9.06 | 9.24 | 9.42 | 9.6 | 9.79 | 9.97 | 10.15 | 10.34 | 10.52 | 10.7 | 10.89 | 11.07 | 11.25 | 11.43 | 11.62 | 11.8 | 11.98 | 12.17 | 12.35 | 12.53 | 12.71 | 12.9 | 13.08 | 13.26 | 13.45 | 13.63 | 13.81 | 13.99 | 14.18 | 14.36 | 14.54
2.06 | 7.22 | 7.41 | 7.59 | 7.77 | 7.96 | 8.14 | 8.32 | 8.5 | 8.69 | 8.87 | 9.05 | 9.24 | 9.42 | 9.6 | 9.79 | 9.97 | 10.15 | 10.33 | 10.52 | 10.7 | 10.88 | 11.07 | 11.25 | 11.43 | 11.61 | 11.8 | 11.98 | 12.16 | 12.35 | 12.53 | 12.71 | 12.89 | 13.08 | 13.26 | 13.44 | 13.62 | 13.81 | 13.99 | 14.17 | 14.36 | 14.54
2.07 | 7.22 | 7.41 | 7.59 | 7.77 | 7.95 | 8.14 | 8.32 | 8.5 | 8.69 | 8.87 | 9.05 | 9.23 | 9.42 | 9.6 | 9.78 | 9.97 | 10.15 | 10.33 | 10.51 | 10.7 | 10.88 | 11.06 | 11.25 | 11.43 | 11.61 | 11.79 | 11.98 | 12.16 | 12.34 | 12.52 | 12.71 | 12.89 | 13.07 | 13.26 | 13.44 | 13.62 | 13.8 | 13.99 | 14.17 | 14.35 | 14.53
2.08 | 7.22 | 7.4 | 7.59 | 7.77 | 7.95 | 8.13 | 8.32 | 8.5 | 8.68 | 8.87 | 9.05 | 9.23 | 9.41 | 9.6 | 9.78 | 9.96 | 10.15 | 10.33 | 10.51 | 10.69 | 10.88 | 11.06 | 11.24 | 11.42 | 11.61 | 11.79 | 11.97 | 12.16 | 12.34 | 12.52 | 12.7 | 12.89 | 13.07 | 13.25 | 13.43 | 13.62 | 13.8 | 13.98 | 14.17 | 14.35 | 14.53
2.09 | 7.22 | 7.4 | 7.58 | 7.77 | 7.95 | 8.13 | 8.32 | 8.5 | 8.68 | 8.86 | 9.05 | 9.23 | 9.41 | 9.59 | 9.78 | 9.96 | 10.14 | 10.33 | 10.51 | 10.69 | 10.87 | 11.06 | 11.24 | 11.42 | 11.6 | 11.79 | 11.97 | 12.15 | 12.34 | 12.52 | 12.7 | 12.88 | 13.07 | 13.25 | 13.43 | 13.61 | 13.8 | 13.98 | 14.16 | 14.34 | 14.53
2.1 | 7.22 | 7.4 | 7.58 | 7.76 | 7.95 | 8.13 | 8.31 | 8.5 | 8.68 | 8.86 | 9.04 | 9.23 | 9.41 | 9.59 | 9.77 | 9.96 | 10.14 | 10.32 | 10.51 | 10.69 | 10.87 | 11.05 | 11.24 | 11.42 | 11.6 | 11.78 | 11.97 | 12.15 | 12.33 | 12.51 | 12.7 | 12.88 | 13.06 | 13.24 | 13.43 | 13.61 | 13.79 | 13.98 | 14.16 | 14.34 | 14.52
2.11 | 7.21 | 7.4 | 7.58 | 7.76 | 7.95 | 8.13 | 8.31 | 8.49 | 8.68 | 8.86 | 9.04 | 9.22 | 9.41 | 9.59 | 9.77 | 9.95 | 10.14 | 10.32 | 10.5 | 10.68 | 10.87 | 11.05 | 11.23 | 11.42 | 11.6 | 11.78 | 11.96 | 12.15 | 12.33 | 12.51 | 12.69 | 12.88 | 13.06 | 13.24 | 13.42 | 13.61 | 13.79 | 13.97 | 14.15 | 14.34 | 14.52
2.12 | 7.21 | 7.39 | 7.58 | 7.76 | 7.94 | 8.13 | 8.31 | 8.49 | 8.67 | 8.86 | 9.04 | 9.22 | 9.4 | 9.59 | 9.77 | 9.95 | 10.13 | 10.32 | 10.5 | 10.68 | 10.86 | 11.05 | 11.23 | 11.41 | 11.59 | 11.78 | 11.96 | 12.14 | 12.32 | 12.51 | 12.69 | 12.87 | 13.05 | 13.24 | 13.42 | 13.6 | 13.79 | 13.97 | 14.15 | 14.33 | 14.51
2.13 | 7.21 | 7.39 | 7.58 | 7.76 | 7.94 | 8.12 | 8.31 | 8.49 | 8.67 | 8.85 | 9.04 | 9.22 | 9.4 | 9.58 | 9.77 | 9.95 | 10.13 | 10.31 | 10.5 | 10.68 | 10.86 | 11.04 | 11.23 | 11.41 | 11.59 | 11.77 | 11.96 | 12.14 | 12.32 | 12.5 | 12.69 | 12.87 | 13.05 | 13.23 | 13.42 | 13.6 | 13.78 | 13.96 | 14.15 | 14.33 | 14.51
2.14 | 7.21 | 7.39 | 7.57 | 7.76 | 7.94 | 8.12 | 8.3 | 8.49 | 8.67 | 8.85 | 9.03 | 9.22 | 9.4 | 9.58 | 9.76 | 9.95 | 10.13 | 10.31 | 10.49 | 10.68 | 10.86 | 11.04 | 11.22 | 11.41 | 11.59 | 11.77 | 11.95 | 12.14 | 12.32 | 12.5 | 12.68 | 12.87 | 13.05 | 13.23 | 13.41 | 13.6 | 13.78 | 13.96 | 14.14 | 14.32 | 14.51
2.15 | 7.21 | 7.39 | 7.57 | 7.75 | 7.94 | 8.12 | 8.3 | 8.48 | 8.67 | 8.85 | 9.03 | 9.21 | 9.4 | 9.58 | 9.76 | 9.94 | 10.13 | 10.31 | 10.49 | 10.67 | 10.86 | 11.04 | 11.22 | 11.4 | 11.59 | 11.77 | 11.95 | 12.13 | 12.31 | 12.5 | 12.68 | 12.86 | 13.04 | 13.23 | 13.41 | 13.59 | 13.77 | 13.96 | 14.14 | 14.32 | 14.5
2.16 | 7.2 | 7.39 | 7.57 | 7.75 | 7.93 | 8.12 | 8.3 | 8.48 | 8.66 | 8.85 | 9.03 | 9.21 | 9.39 | 9.58 | 9.76 | 9.94 | 10.12 | 10.31 | 10.49 | 10.67 | 10.85 | 11.04 | 11.22 | 11.4 | 11.58 | 11.76 | 11.95 | 12.13 | 12.31 | 12.49 | 12.68 | 12.86 | 13.04 | 13.22 | 13.41 | 13.59 | 13.77 | 13.95 | 14.13 | 14.32 | 14.5
2.17 | 7.2 | 7.38 | 7.57 | 7.75 | 7.93 | 8.11 | 8.3 | 8.48 | 8.66 | 8.84 | 9.03 | 9.21 | 9.39 | 9.57 | 9.76 | 9.94 | 10.12 | 10.3 | 10.49 | 10.67 | 10.85 | 11.03 | 11.21 | 11.4 | 11.58 | 11.76 | 11.94 | 12.13 | 12.31 | 12.49 | 12.67 | 12.85 | 13.04 | 13.22 | 13.4 | 13.58 | 13.77 | 13.95 | 14.13 | 14.31 | 14.5
2.18 | 7.2 | 7.38 | 7.57 | 7.75 | 7.93 | 8.11 | 8.29 | 8.48 | 8.66 | 8.84 | 9.02 | 9.21 | 9.39 | 9.57 | 9.75 | 9.94 | 10.12 | 10.3 | 10.48 | 10.66 | 10.85 | 11.03 | 11.21 | 11.39 | 11.58 | 11.76 | 11.94 | 12.12 | 12.3 | 12.49 | 12.67 | 12.85 | 13.03 | 13.22 | 13.4 | 13.58 | 13.76 | 13.94 | 14.13 | 14.31 | 14.49
2.19 | 7.2 | 7.38 | 7.56 | 7.75 | 7.93 | 8.11 | 8.29 | 8.47 | 8.66 | 8.84 | 9.02 | 9.2 | 9.39 | 9.57 | 9.75 | 9.93 | 10.11 | 10.3 | 10.48 | 10.66 | 10.84 | 11.03 | 11.21 | 11.39 | 11.57 | 11.75 | 11.94 | 12.12 | 12.3 | 12.48 | 12.67 | 12.85 | 13.03 | 13.21 | 13.39 | 13.58 | 13.76 | 13.94 | 14.12 | 14.31 | 14.49
2.2 | 7.2 | 7.38 | 7.56 | 7.74 | 7.93 | 8.11 | 8.29 | 8.47 | 8.65 | 8.84 | 9.02 | 9.2 | 9.38 | 9.57 | 9.75 | 9.93 | 10.11 | 10.29 | 10.48 | 10.66 | 10.84 | 11.02 | 11.21 | 11.39 | 11.57 | 11.75 | 11.93 | 12.12 | 12.3 | 12.48 | 12.66 | 12.84 | 13.03 | 13.21 | 13.39 | 13.57 | 13.76 | 13.94 | 14.12 | 14.3 | 14.48
2.21 | 7.19 | 7.38 | 7.56 | 7.74 | 7.92 | 8.11 | 8.29 | 8.47 | 8.65 | 8.83 | 9.02 | 9.2 | 9.38 | 9.56 | 9.74 | 9.93 | 10.11 | 10.29 | 10.47 | 10.66 | 10.84 | 11.02 | 11.2 | 11.38 | 11.57 | 11.75 | 11.93 | 12.11 | 12.29 | 12.48 | 12.66 | 12.84 | 13.02 | 13.21 | 13.39 | 13.57 | 13.75 | 13.93 | 14.12 | 14.3 | 14.48
2.22 | 7.19 | 7.37 | 7.56 | 7.74 | 7.92 | 8.1 | 8.29 | 8.47 | 8.65 | 8.83 | 9.01 | 9.2 | 9.38 | 9.56 | 9.74 | 9.92 | 10.11 | 10.29 | 10.47 | 10.65 | 10.83 | 11.02 | 11.2 | 11.38 | 11.56 | 11.75 | 11.93 | 12.11 | 12.29 | 12.47 | 12.66 | 12.84 | 13.02 | 13.2 | 13.38 | 13.57 | 13.75 | 13.93 | 14.11 | 14.29 | 14.48
2.23 | 7.19 | 7.37 | 7.55 | 7.74 | 7.92 | 8.1 | 8.28 | 8.47 | 8.65 | 8.83 | 9.01 | 9.19 | 9.38 | 9.56 | 9.74 | 9.92 | 10.1 | 10.29 | 10.47 | 10.65 | 10.83 | 11.01 | 11.2 | 11.38 | 11.56 | 11.74 | 11.92 | 12.11 | 12.29 | 12.47 | 12.65 | 12.83 | 13.02 | 13.2 | 13.38 | 13.56 | 13.74 | 13.93 | 14.11 | 14.29 | 14.47
2.24 | 7.19 | 7.37 | 7.55 | 7.73 | 7.92 | 8.1 | 8.28 | 8.46 | 8.64 | 8.83 | 9.01 | 9.19 | 9.37 | 9.55 | 9.74 | 9.92 | 10.1 | 10.28 | 10.46 | 10.65 | 10.83 | 11.01 | 11.19 | 11.37 | 11.56 | 11.74 | 11.92 | 12.1 | 12.28 | 12.47 | 12.65 | 12.83 | 13.01 | 13.19 | 13.38 | 13.56 | 13.74 | 13.92 | 14.1 | 14.29 | 14.47
2.25 | 7.19 | 7.37 | 7.55 | 7.73 | 7.91 | 8.1 | 8.28 | 8.46 | 8.64 | 8.82 | 9.01 | 9.19 | 9.37 | 9.55 | 9.73 | 9.92 | 10.1 | 10.28 | 10.46 | 10.64 | 10.83 | 11.01 | 11.19 | 11.37 | 11.55 | 11.74 | 11.92 | 12.1 | 12.28 | 12.46 | 12.65 | 12.83 | 13.01 | 13.19 | 13.37 | 13.55 | 13.74 | 13.92 | 14.1 | 14.28 | 14.46
2.26 | 7.18 | 7.37 | 7.55 | 7.73 | 7.91 | 8.09 | 8.28 | 8.46 | 8.64 | 8.82 | 9 | 9.19 | 9.37 | 9.55 | 9.73 | 9.91 | 10.1 | 10.28 | 10.46 | 10.64 | 10.82 | 11 | 11.19 | 11.37 | 11.55 | 11.73 | 11.91 | 12.1 | 12.28 | 12.46 | 12.64 | 12.82 | 13.01 | 13.19 | 13.37 | 13.55 | 13.73 | 13.91 | 14.1 | 14.28 | 14.46
2.27 | 7.18 | 7.36 | 7.55 | 7.73 | 7.91 | 8.09 | 8.27 | 8.46 | 8.64 | 8.82 | 9 | 9.18 | 9.37 | 9.55 | 9.73 | 9.91 | 10.09 | 10.27 | 10.46 | 10.64 | 10.82 | 11 | 11.18 | 11.37 | 11.55 | 11.73 | 11.91 | 12.09 | 12.27 | 12.46 | 12.64 | 12.82 | 13 | 13.18 | 13.37 | 13.55 | 13.73 | 13.91 | 14.09 | 14.27 | 14.46
2.28 | 7.18 | 7.36 | 7.54 | 7.73 | 7.91 | 8.09 | 8.27 | 8.45 | 8.64 | 8.82 | 9 | 9.18 | 9.36 | 9.54 | 9.73 | 9.91 | 10.09 | 10.27 | 10.45 | 10.64 | 10.82 | 11 | 11.18 | 11.36 | 11.54 | 11.73 | 11.91 | 12.09 | 12.27 | 12.45 | 12.63 | 12.82 | 13 | 13.18 | 13.36 | 13.54 | 13.73 | 13.91 | 14.09 | 14.27 | 14.45
2.29 | 7.18 | 7.36 | 7.54 | 7.72 | 7.91 | 8.09 | 8.27 | 8.45 | 8.63 | 8.81 | 9 | 9.18 | 9.36 | 9.54 | 9.72 | 9.91 | 10.09 | 10.27 | 10.45 | 10.63 | 10.81 | 11 | 11.18 | 11.36 | 11.54 | 11.72 | 11.9 | 12.09 | 12.27 | 12.45 | 12.63 | 12.81 | 12.99 | 13.18 | 13.36 | 13.54 | 13.72 | 13.9 | 14.08 | 14.27 | 14.45
2.3 | 7.18 | 7.36 | 7.54 | 7.72 | 7.9 | 8.09 | 8.27 | 8.45 | 8.63 | 8.81 | 8.99 | 9.18 | 9.36 | 9.54 | 9.72 | 9.9 | 10.08 | 10.27 | 10.45 | 10.63 | 10.81 | 10.99 | 11.17 | 11.36 | 11.54 | 11.72 | 11.9 | 12.08 | 12.26 | 12.45 | 12.63 | 12.81 | 12.99 | 13.17 | 13.35 | 13.54 | 13.72 | 13.9 | 14.08 | 14.26 | 14.44
2.31 | 7.17 | 7.36 | 7.54 | 7.72 | 7.9 | 8.08 | 8.26 | 8.45 | 8.63 | 8.81 | 8.99 | 9.17 | 9.35 | 9.54 | 9.72 | 9.9 | 10.08 | 10.26 | 10.44 | 10.63 | 10.81 | 10.99 | 11.17 | 11.35 | 11.53 | 11.72 | 11.9 | 12.08 | 12.26 | 12.44 | 12.62 | 12.81 | 12.99 | 13.17 | 13.35 | 13.53 | 13.71 | 13.9 | 14.08 | 14.26 | 14.44
2.32 | 7.17 | 7.35 | 7.54 | 7.72 | 7.9 | 8.08 | 8.26 | 8.44 | 8.63 | 8.81 | 8.99 | 9.17 | 9.35 | 9.53 | 9.72 | 9.9 | 10.08 | 10.26 | 10.44 | 10.62 | 10.81 | 10.99 | 11.17 | 11.35 | 11.53 | 11.71 | 11.89 | 12.08 | 12.26 | 12.44 | 12.62 | 12.8 | 12.98 | 13.17 | 13.35 | 13.53 | 13.71 | 13.89 | 14.07 | 14.25 | 14.44
2.33 | 7.17 | 7.35 | 7.53 | 7.72 | 7.9 | 8.08 | 8.26 | 8.44 | 8.62 | 8.81 | 8.99 | 9.17 | 9.35 | 9.53 | 9.71 | 9.89 | 10.08 | 10.26 | 10.44 | 10.62 | 10.8 | 10.98 | 11.17 | 11.35 | 11.53 | 11.71 | 11.89 | 12.07 | 12.25 | 12.44 | 12.62 | 12.8 | 12.98 | 13.16 | 13.34 | 13.53 | 13.71 | 13.89 | 14.07 | 14.25 | 14.43
2.34 | 7.17 | 7.35 | 7.53 | 7.71 | 7.89 | 8.08 | 8.26 | 8.44 | 8.62 | 8.8 | 8.98 | 9.17 | 9.35 | 9.53 | 9.71 | 9.89 | 10.07 | 10.25 | 10.44 | 10.62 | 10.8 | 10.98 | 11.16 | 11.34 | 11.53 | 11.71 | 11.89 | 12.07 | 12.25 | 12.43 | 12.61 | 12.8 | 12.98 | 13.16 | 13.34 | 13.52 | 13.7 | 13.88 | 14.07 | 14.25 | 14.43
2.35 | 7.17 | 7.35 | 7.53 | 7.71 | 7.89 | 8.07 | 8.26 | 8.44 | 8.62 | 8.8 | 8.98 | 9.16 | 9.34 | 9.53 | 9.71 | 9.89 | 10.07 | 10.25 | 10.43 | 10.62 | 10.8 | 10.98 | 11.16 | 11.34 | 11.52 | 11.7 | 11.89 | 12.07 | 12.25 | 12.43 | 12.61 | 12.79 | 12.97 | 13.15 | 13.34 | 13.52 | 13.7 | 13.88 | 14.06 | 14.24 | 14.42
2.36 | 7.16 | 7.35 | 7.53 | 7.71 | 7.89 | 8.07 | 8.25 | 8.43 | 8.62 | 8.8 | 8.98 | 9.16 | 9.34 | 9.52 | 9.71 | 9.89 | 10.07 | 10.25 | 10.43 | 10.61 | 10.79 | 10.97 | 11.16 | 11.34 | 11.52 | 11.7 | 11.88 | 12.06 | 12.24 | 12.43 | 12.61 | 12.79 | 12.97 | 13.15 | 13.33 | 13.51 | 13.7 | 13.88 | 14.06 | 14.24 | 14.42
2.37 | 7.16 | 7.34 | 7.53 | 7.71 | 7.89 | 8.07 | 8.25 | 8.43 | 8.61 | 8.8 | 8.98 | 9.16 | 9.34 | 9.52 | 9.7 | 9.88 | 10.07 | 10.25 | 10.43 | 10.61 | 10.79 | 10.97 | 11.15 | 11.33 | 11.52 | 11.7 | 11.88 | 12.06 | 12.24 | 12.42 | 12.6 | 12.79 | 12.97 | 13.15 | 13.33 | 13.51 | 13.69 | 13.87 | 14.05 | 14.24 | 14.42
2.38 | 7.16 | 7.34 | 7.52 | 7.7 | 7.89 | 8.07 | 8.25 | 8.43 | 8.61 | 8.79 | 8.97 | 9.16 | 9.34 | 9.52 | 9.7 | 9.88 | 10.06 | 10.24 | 10.43 | 10.61 | 10.79 | 10.97 | 11.15 | 11.33 | 11.51 | 11.69 | 11.88 | 12.06 | 12.24 | 12.42 | 12.6 | 12.78 | 12.96 | 13.14 | 13.33 | 13.51 | 13.69 | 13.87 | 14.05 | 14.23 | 14.41
2.39 | 7.16 | 7.34 | 7.52 | 7.7 | 7.88 | 8.07 | 8.25 | 8.43 | 8.61 | 8.79 | 8.97 | 9.15 | 9.33 | 9.52 | 9.7 | 9.88 | 10.06 | 10.24 | 10.42 | 10.6 | 10.78 | 10.97 | 11.15 | 11.33 | 11.51 | 11.69 | 11.87 | 12.05 | 12.23 | 12.42 | 12.6 | 12.78 | 12.96 | 13.14 | 13.32 | 13.5 | 13.68 | 13.87 | 14.05 | 14.23 | 14.41
2.4 | 7.16 | 7.34 | 7.52 | 7.7 | 7.88 | 8.06 | 8.24 | 8.43 | 8.61 | 8.79 | 8.97 | 9.15 | 9.33 | 9.51 | 9.69 | 9.88 | 10.06 | 10.24 | 10.42 | 10.6 | 10.78 | 10.96 | 11.14 | 11.33 | 11.51 | 11.69 | 11.87 | 12.05 | 12.23 | 12.41 | 12.59 | 12.77 | 12.96 | 13.14 | 13.32 | 13.5 | 13.68 | 13.86 | 14.04 | 14.22 | 14.41
2.41 | 7.15 | 7.34 | 7.52 | 7.7 | 7.88 | 8.06 | 8.24 | 8.42 | 8.6 | 8.79 | 8.97 | 9.15 | 9.33 | 9.51 | 9.69 | 9.87 | 10.05 | 10.24 | 10.42 | 10.6 | 10.78 | 10.96 | 11.14 | 11.32 | 11.5 | 11.68 | 11.87 | 12.05 | 12.23 | 12.41 | 12.59 | 12.77 | 12.95 | 13.13 | 13.31 | 13.5 | 13.68 | 13.86 | 14.04 | 14.22 | 14.4
2.42 | 7.15 | 7.33 | 7.52 | 7.7 | 7.88 | 8.06 | 8.24 | 8.42 | 8.6 | 8.78 | 8.96 | 9.15 | 9.33 | 9.51 | 9.69 | 9.87 | 10.05 | 10.23 | 10.41 | 10.59 | 10.78 | 10.96 | 11.14 | 11.32 | 11.5 | 11.68 | 11.86 | 12.04 | 12.22 | 12.41 | 12.59 | 12.77 | 12.95 | 13.13 | 13.31 | 13.49 | 13.67 | 13.85 | 14.04 | 14.22 | 14.4
2.43 | 7.15 | 7.33 | 7.51 | 7.69 | 7.88 | 8.06 | 8.24 | 8.42 | 8.6 | 8.78 | 8.96 | 9.14 | 9.32 | 9.51 | 9.69 | 9.87 | 10.05 | 10.23 | 10.41 | 10.59 | 10.77 | 10.95 | 11.14 | 11.32 | 11.5 | 11.68 | 11.86 | 12.04 | 12.22 | 12.4 | 12.58 | 12.76 | 12.95 | 13.13 | 13.31 | 13.49 | 13.67 | 13.85 | 14.03 | 14.21 | 14.39
2.44 | 7.15 | 7.33 | 7.51 | 7.69 | 7.87 | 8.05 | 8.24 | 8.42 | 8.6 | 8.78 | 8.96 | 9.14 | 9.32 | 9.5 | 9.68 | 9.86 | 10.05 | 10.23 | 10.41 | 10.59 | 10.77 | 10.95 | 11.13 | 11.31 | 11.49 | 11.68 | 11.86 | 12.04 | 12.22 | 12.4 | 12.58 | 12.76 | 12.94 | 13.12 | 13.3 | 13.48 | 13.67 | 13.85 | 14.03 | 14.21 | 14.39
2.45 | 7.15 | 7.33 | 7.51 | 7.69 | 7.87 | 8.05 | 8.23 | 8.41 | 8.6 | 8.78 | 8.96 | 9.14 | 9.32 | 9.5 | 9.68 | 9.86 | 10.04 | 10.22 | 10.41 | 10.59 | 10.77 | 10.95 | 11.13 | 11.31 | 11.49 | 11.67 | 11.85 | 12.03 | 12.21 | 12.4 | 12.58 | 12.76 | 12.94 | 13.12 | 13.3 | 13.48 | 13.66 | 13.84 | 14.02 | 14.2 | 14.39
2.46 | 7.14 | 7.33 | 7.51 | 7.69 | 7.87 | 8.05 | 8.23 | 8.41 | 8.59 | 8.77 | 8.95 | 9.14 | 9.32 | 9.5 | 9.68 | 9.86 | 10.04 | 10.22 | 10.4 | 10.58 | 10.76 | 10.95 | 11.13 | 11.31 | 11.49 | 11.67 | 11.85 | 12.03 | 12.21 | 12.39 | 12.57 | 12.75 | 12.93 | 13.12 | 13.3 | 13.48 | 13.66 | 13.84 | 14.02 | 14.2 | 14.38
2.47 | 7.14 | 7.32 | 7.5 | 7.69 | 7.87 | 8.05 | 8.23 | 8.41 | 8.59 | 8.77 | 8.95 | 9.13 | 9.31 | 9.5 | 9.68 | 9.86 | 10.04 | 10.22 | 10.4 | 10.58 | 10.76 | 10.94 | 11.12 | 11.3 | 11.48 | 11.67 | 11.85 | 12.03 | 12.21 | 12.39 | 12.57 | 12.75 | 12.93 | 13.11 | 13.29 | 13.47 | 13.65 | 13.84 | 14.02 | 14.2 | 14.38
2.48 | 7.14 | 7.32 | 7.5 | 7.68 | 7.86 | 8.05 | 8.23 | 8.41 | 8.59 | 8.77 | 8.95 | 9.13 | 9.31 | 9.49 | 9.67 | 9.85 | 10.04 | 10.22 | 10.4 | 10.58 | 10.76 | 10.94 | 11.12 | 11.3 | 11.48 | 11.66 | 11.84 | 12.02 | 12.2 | 12.39 | 12.57 | 12.75 | 12.93 | 13.11 | 13.29 | 13.47 | 13.65 | 13.83 | 14.01 | 14.19 | 14.37
2.49 | 7.14 | 7.32 | 7.5 | 7.68 | 7.86 | 8.04 | 8.22 | 8.4 | 8.59 | 8.77 | 8.95 | 9.13 | 9.31 | 9.49 | 9.67 | 9.85 | 10.03 | 10.21 | 10.39 | 10.57 | 10.76 | 10.94 | 11.12 | 11.3 | 11.48 | 11.66 | 11.84 | 12.02 | 12.2 | 12.38 | 12.56 | 12.74 | 12.92 | 13.11 | 13.29 | 13.47 | 13.65 | 13.83 | 14.01 | 14.19 | 14.37
2.5 | 7.14 | 7.32 | 7.5 | 7.68 | 7.86 | 8.04 | 8.22 | 8.4 | 8.58 | 8.76 | 8.94 | 9.13 | 9.31 | 9.49 | 9.67 | 9.85 | 10.03 | 10.21 | 10.39 | 10.57 | 10.75 | 10.93 | 11.11 | 11.29 | 11.48 | 11.66 | 11.84 | 12.02 | 12.2 | 12.38 | 12.56 | 12.74 | 12.92 | 13.1 | 13.28 | 13.46 | 13.64 | 13.82 | 14 | 14.19 | 14.37
2.51 | 7.14 | 7.32 | 7.5 | 7.68 | 7.86 | 8.04 | 8.22 | 8.4 | 8.58 | 8.76 | 8.94 | 9.12 | 9.3 | 9.48 | 9.67 | 9.85 | 10.03 | 10.21 | 10.39 | 10.57 | 10.75 | 10.93 | 11.11 | 11.29 | 11.47 | 11.65 | 11.83 | 12.01 | 12.19 | 12.38 | 12.56 | 12.74 | 12.92 | 13.1 | 13.28 | 13.46 | 13.64 | 13.82 | 14 | 14.18 | 14.36
2.52 | 7.13 | 7.31 | 7.49 | 7.68 | 7.86 | 8.04 | 8.22 | 8.4 | 8.58 | 8.76 | 8.94 | 9.12 | 9.3 | 9.48 | 9.66 | 9.84 | 10.02 | 10.2 | 10.39 | 10.57 | 10.75 | 10.93 | 11.11 | 11.29 | 11.47 | 11.65 | 11.83 | 12.01 | 12.19 | 12.37 | 12.55 | 12.73 | 12.91 | 13.09 | 13.28 | 13.46 | 13.64 | 13.82 | 14 | 14.18 | 14.36
2.53 | 7.13 | 7.31 | 7.49 | 7.67 | 7.85 | 8.03 | 8.22 | 8.4 | 8.58 | 8.76 | 8.94 | 9.12 | 9.3 | 9.48 | 9.66 | 9.84 | 10.02 | 10.2 | 10.38 | 10.56 | 10.74 | 10.92 | 11.1 | 11.29 | 11.47 | 11.65 | 11.83 | 12.01 | 12.19 | 12.37 | 12.55 | 12.73 | 12.91 | 13.09 | 13.27 | 13.45 | 13.63 | 13.81 | 13.99 | 14.17 | 14.35
2.54 | 7.13 | 7.31 | 7.49 | 7.67 | 7.85 | 8.03 | 8.21 | 8.39 | 8.57 | 8.75 | 8.94 | 9.12 | 9.3 | 9.48 | 9.66 | 9.84 | 10.02 | 10.2 | 10.38 | 10.56 | 10.74 | 10.92 | 11.1 | 11.28 | 11.46 | 11.64 | 11.82 | 12 | 12.18 | 12.37 | 12.55 | 12.73 | 12.91 | 13.09 | 13.27 | 13.45 | 13.63 | 13.81 | 13.99 | 14.17 | 14.35
2.55 | 7.13 | 7.31 | 7.49 | 7.67 | 7.85 | 8.03 | 8.21 | 8.39 | 8.57 | 8.75 | 8.93 | 9.11 | 9.29 | 9.47 | 9.65 | 9.84 | 10.02 | 10.2 | 10.38 | 10.56 | 10.74 | 10.92 | 11.1 | 11.28 | 11.46 | 11.64 | 11.82 | 12 | 12.18 | 12.36 | 12.54 | 12.72 | 12.9 | 13.08 | 13.26 | 13.44 | 13.63 | 13.81 | 13.99 | 14.17 | 14.35
2.56 | 7.13 | 7.31 | 7.49 | 7.67 | 7.85 | 8.03 | 8.21 | 8.39 | 8.57 | 8.75 | 8.93 | 9.11 | 9.29 | 9.47 | 9.65 | 9.83 | 10.01 | 10.19 | 10.37 | 10.55 | 10.73 | 10.92 | 11.1 | 11.28 | 11.46 | 11.64 | 11.82 | 12 | 12.18 | 12.36 | 12.54 | 12.72 | 12.9 | 13.08 | 13.26 | 13.44 | 13.62 | 13.8 | 13.98 | 14.16 | 14.34
2.57 | 7.12 | 7.3 | 7.48 | 7.66 | 7.85 | 8.03 | 8.21 | 8.39 | 8.57 | 8.75 | 8.93 | 9.11 | 9.29 | 9.47 | 9.65 | 9.83 | 10.01 | 10.19 | 10.37 | 10.55 | 10.73 | 10.91 | 11.09 | 11.27 | 11.45 | 11.63 | 11.81 | 11.99 | 12.18 | 12.36 | 12.54 | 12.72 | 12.9 | 13.08 | 13.26 | 13.44 | 13.62 | 13.8 | 13.98 | 14.16 | 14.34
2.58 | 7.12 | 7.3 | 7.48 | 7.66 | 7.84 | 8.02 | 8.2 | 8.38 | 8.56 | 8.75 | 8.93 | 9.11 | 9.29 | 9.47 | 9.65 | 9.83 | 10.01 | 10.19 | 10.37 | 10.55 | 10.73 | 10.91 | 11.09 | 11.27 | 11.45 | 11.63 | 11.81 | 11.99 | 12.17 | 12.35 | 12.53 | 12.71 | 12.89 | 13.07 | 13.25 | 13.43 | 13.61 | 13.79 | 13.97 | 14.15 | 14.34
2.59 | 7.12 | 7.3 | 7.48 | 7.66 | 7.84 | 8.02 | 8.2 | 8.38 | 8.56 | 8.74 | 8.92 | 9.1 | 9.28 | 9.46 | 9.64 | 9.82 | 10.01 | 10.19 | 10.37 | 10.55 | 10.73 | 10.91 | 11.09 | 11.27 | 11.45 | 11.63 | 11.81 | 11.99 | 12.17 | 12.35 | 12.53 | 12.71 | 12.89 | 13.07 | 13.25 | 13.43 | 13.61 | 13.79 | 13.97 | 14.15 | 14.33
2.6 | 7.12 | 7.3 | 7.48 | 7.66 | 7.84 | 8.02 | 8.2 | 8.38 | 8.56 | 8.74 | 8.92 | 9.1 | 9.28 | 9.46 | 9.64 | 9.82 | 10 | 10.18 | 10.36 | 10.54 | 10.72 | 10.9 | 11.08 | 11.26 | 11.44 | 11.62 | 11.8 | 11.98 | 12.17 | 12.35 | 12.53 | 12.71 | 12.89 | 13.07 | 13.25 | 13.43 | 13.61 | 13.79 | 13.97 | 14.15 | 14.33
2.61 | 7.12 | 7.3 | 7.48 | 7.66 | 7.84 | 8.02 | 8.2 | 8.38 | 8.56 | 8.74 | 8.92 | 9.1 | 9.28 | 9.46 | 9.64 | 9.82 | 10 | 10.18 | 10.36 | 10.54 | 10.72 | 10.9 | 11.08 | 11.26 | 11.44 | 11.62 | 11.8 | 11.98 | 12.16 | 12.34 | 12.52 | 12.7 | 12.88 | 13.06 | 13.24 | 13.42 | 13.6 | 13.78 | 13.96 | 14.14 | 14.32
2.62 | 7.11 | 7.29 | 7.47 | 7.65 | 7.83 | 8.01 | 8.19 | 8.38 | 8.56 | 8.74 | 8.92 | 9.1 | 9.28 | 9.46 | 9.64 | 9.82 | 10 | 10.18 | 10.36 | 10.54 | 10.72 | 10.9 | 11.08 | 11.26 | 11.44 | 11.62 | 11.8 | 11.98 | 12.16 | 12.34 | 12.52 | 12.7 | 12.88 | 13.06 | 13.24 | 13.42 | 13.6 | 13.78 | 13.96 | 14.14 | 14.32
2.63 | 7.11 | 7.29 | 7.47 | 7.65 | 7.83 | 8.01 | 8.19 | 8.37 | 8.55 | 8.73 | 8.91 | 9.09 | 9.27 | 9.45 | 9.63 | 9.81 | 9.99 | 10.17 | 10.35 | 10.53 | 10.71 | 10.89 | 11.07 | 11.25 | 11.43 | 11.62 | 11.8 | 11.98 | 12.16 | 12.34 | 12.52 | 12.7 | 12.88 | 13.06 | 13.24 | 13.42 | 13.6 | 13.78 | 13.96 | 14.14 | 14.32
2.64 | 7.11 | 7.29 | 7.47 | 7.65 | 7.83 | 8.01 | 8.19 | 8.37 | 8.55 | 8.73 | 8.91 | 9.09 | 9.27 | 9.45 | 9.63 | 9.81 | 9.99 | 10.17 | 10.35 | 10.53 | 10.71 | 10.89 | 11.07 | 11.25 | 11.43 | 11.61 | 11.79 | 11.97 | 12.15 | 12.33 | 12.51 | 12.69 | 12.87 | 13.05 | 13.23 | 13.41 | 13.59 | 13.77 | 13.95 | 14.13 | 14.31
2.65 | 7.11 | 7.29 | 7.47 | 7.65 | 7.83 | 8.01 | 8.19 | 8.37 | 8.55 | 8.73 | 8.91 | 9.09 | 9.27 | 9.45 | 9.63 | 9.81 | 9.99 | 10.17 | 10.35 | 10.53 | 10.71 | 10.89 | 11.07 | 11.25 | 11.43 | 11.61 | 11.79 | 11.97 | 12.15 | 12.33 | 12.51 | 12.69 | 12.87 | 13.05 | 13.23 | 13.41 | 13.59 | 13.77 | 13.95 | 14.13 | 14.31
2.66 | 7.11 | 7.29 | 7.47 | 7.65 | 7.83 | 8.01 | 8.19 | 8.37 | 8.55 | 8.73 | 8.91 | 9.09 | 9.27 | 9.45 | 9.63 | 9.81 | 9.99 | 10.17 | 10.35 | 10.53 | 10.71 | 10.89 | 11.07 | 11.25 | 11.43 | 11.61 | 11.79 | 11.97 | 12.15 | 12.33 | 12.51 | 12.69 | 12.87 | 13.05 | 13.22 | 13.4 | 13.58 | 13.76 | 13.94 | 14.12 | 14.3
2.67 | 7.1 | 7.28 | 7.46 | 7.64 | 7.82 | 8 | 8.18 | 8.36 | 8.54 | 8.72 | 8.9 | 9.08 | 9.26 | 9.44 | 9.62 | 9.8 | 9.98 | 10.16 | 10.34 | 10.52 | 10.7 | 10.88 | 11.06 | 11.24 | 11.42 | 11.6 | 11.78 | 11.96 | 12.14 | 12.32 | 12.5 | 12.68 | 12.86 | 13.04 | 13.22 | 13.4 | 13.58 | 13.76 | 13.94 | 14.12 | 14.3
2.68 | 7.1 | 7.28 | 7.46 | 7.64 | 7.82 | 8 | 8.18 | 8.36 | 8.54 | 8.72 | 8.9 | 9.08 | 9.26 | 9.44 | 9.62 | 9.8 | 9.98 | 10.16 | 10.34 | 10.52 | 10.7 | 10.88 | 11.06 | 11.24 | 11.42 | 11.6 | 11.78 | 11.96 | 12.14 | 12.32 | 12.5 | 12.68 | 12.86 | 13.04 | 13.22 | 13.4 | 13.58 | 13.76 | 13.94 | 14.12 | 14.3
2.69 | 7.1 | 7.28 | 7.46 | 7.64 | 7.82 | 8 | 8.18 | 8.36 | 8.54 | 8.72 | 8.9 | 9.08 | 9.26 | 9.44 | 9.62 | 9.8 | 9.98 | 10.16 | 10.34 | 10.52 | 10.7 | 10.88 | 11.06 | 11.24 | 11.42 | 11.6 | 11.78 | 11.96 | 12.14 | 12.32 | 12.5 | 12.67 | 12.85 | 13.03 | 13.21 | 13.39 | 13.57 | 13.75 | 13.93 | 14.11 | 14.29
2.7 | 7.1 | 7.28 | 7.46 | 7.64 | 7.82 | 8 | 8.18 | 8.36 | 8.54 | 8.72 | 8.9 | 9.08 | 9.26 | 9.44 | 9.62 | 9.8 | 9.98 | 10.15 | 10.33 | 10.51 | 10.69 | 10.87 | 11.05 | 11.23 | 11.41 | 11.59 | 11.77 | 11.95 | 12.13 | 12.31 | 12.49 | 12.67 | 12.85 | 13.03 | 13.21 | 13.39 | 13.57 | 13.75 | 13.93 | 14.11 | 14.29
2.71 | 7.1 | 7.28 | 7.46 | 7.64 | 7.82 | 7.99 | 8.17 | 8.35 | 8.53 | 8.71 | 8.89 | 9.07 | 9.25 | 9.43 | 9.61 | 9.79 | 9.97 | 10.15 | 10.33 | 10.51 | 10.69 | 10.87 | 11.05 | 11.23 | 11.41 | 11.59 | 11.77 | 11.95 | 12.13 | 12.31 | 12.49 | 12.67 | 12.85 | 13.03 | 13.21 | 13.39 | 13.57 | 13.75 | 13.93 | 14.11 | 14.29
2.72 | 7.09 | 7.27 | 7.45 | 7.63 | 7.81 | 7.99 | 8.17 | 8.35 | 8.53 | 8.71 | 8.89 | 9.07 | 9.25 | 9.43 | 9.61 | 9.79 | 9.97 | 10.15 | 10.33 | 10.51 | 10.69 | 10.87 | 11.05 | 11.23 | 11.41 | 11.59 | 11.77 | 11.95 | 12.13 | 12.31 | 12.49 | 12.66 | 12.84 | 13.02 | 13.2 | 13.38 | 13.56 | 13.74 | 13.92 | 14.1 | 14.28
2.73 | 7.09 | 7.27 | 7.45 | 7.63 | 7.81 | 7.99 | 8.17 | 8.35 | 8.53 | 8.71 | 8.89 | 9.07 | 9.25 | 9.43 | 9.61 | 9.79 | 9.97 | 10.15 | 10.33 | 10.51 | 10.69 | 10.87 | 11.04 | 11.22 | 11.4 | 11.58 | 11.76 | 11.94 | 12.12 | 12.3 | 12.48 | 12.66 | 12.84 | 13.02 | 13.2 | 13.38 | 13.56 | 13.74 | 13.92 | 14.1 | 14.28
2.74 | 7.09 | 7.27 | 7.45 | 7.63 | 7.81 | 7.99 | 8.17 | 8.35 | 8.53 | 8.71 | 8.89 | 9.07 | 9.25 | 9.43 | 9.61 | 9.78 | 9.96 | 10.14 | 10.32 | 10.5 | 10.68 | 10.86 | 11.04 | 11.22 | 11.4 | 11.58 | 11.76 | 11.94 | 12.12 | 12.3 | 12.48 | 12.66 | 12.84 | 13.02 | 13.2 | 13.38 | 13.56 | 13.73 | 13.91 | 14.09 | 14.27
2.75 | 7.09 | 7.27 | 7.45 | 7.63 | 7.81 | 7.99 | 8.17 | 8.35 | 8.53 | 8.7 | 8.88 | 9.06 | 9.24 | 9.42 | 9.6 | 9.78 | 9.96 | 10.14 | 10.32 | 10.5 | 10.68 | 10.86 | 11.04 | 11.22 | 11.4 | 11.58 | 11.76 | 11.94 | 12.12 | 12.3 | 12.47 | 12.65 | 12.83 | 13.01 | 13.19 | 13.37 | 13.55 | 13.73 | 13.91 | 14.09 | 14.27
2.76 | 7.09 | 7.27 | 7.45 | 7.62 | 7.8 | 7.98 | 8.16 | 8.34 | 8.52 | 8.7 | 8.88 | 9.06 | 9.24 | 9.42 | 9.6 | 9.78 | 9.96 | 10.14 | 10.32 | 10.5 | 10.68 | 10.86 | 11.04 | 11.22 | 11.39 | 11.57 | 11.75 | 11.93 | 12.11 | 12.29 | 12.47 | 12.65 | 12.83 | 13.01 | 13.19 | 13.37 | 13.55 | 13.73 | 13.91 | 14.09 | 14.27
2.77 | 7.08 | 7.26 | 7.44 | 7.62 | 7.8 | 7.98 | 8.16 | 8.34 | 8.52 | 8.7 | 8.88 | 9.06 | 9.24 | 9.42 | 9.6 | 9.78 | 9.96 | 10.14 | 10.32 | 10.49 | 10.67 | 10.85 | 11.03 | 11.21 | 11.39 | 11.57 | 11.75 | 11.93 | 12.11 | 12.29 | 12.47 | 12.65 | 12.83 | 13.01 | 13.19 | 13.37 | 13.54 | 13.72 | 13.9 | 14.08 | 14.26
2.78 | 7.08 | 7.26 | 7.44 | 7.62 | 7.8 | 7.98 | 8.16 | 8.34 | 8.52 | 8.7 | 8.88 | 9.06 | 9.24 | 9.42 | 9.59 | 9.77 | 9.95 | 10.13 | 10.31 | 10.49 | 10.67 | 10.85 | 11.03 | 11.21 | 11.39 | 11.57 | 11.75 | 11.93 | 12.11 | 12.29 | 12.46 | 12.64 | 12.82 | 13 | 13.18 | 13.36 | 13.54 | 13.72 | 13.9 | 14.08 | 14.26
2.79 | 7.08 | 7.26 | 7.44 | 7.62 | 7.8 | 7.98 | 8.16 | 8.34 | 8.52 | 8.7 | 8.87 | 9.05 | 9.23 | 9.41 | 9.59 | 9.77 | 9.95 | 10.13 | 10.31 | 10.49 | 10.67 | 10.85 | 11.03 | 11.21 | 11.39 | 11.56 | 11.74 | 11.92 | 12.1 | 12.28 | 12.46 | 12.64 | 12.82 | 13 | 13.18 | 13.36 | 13.54 | 13.72 | 13.9 | 14.08 | 14.25
2.8 | 7.08 | 7.26 | 7.44 | 7.62 | 7.8 | 7.98 | 8.15 | 8.33 | 8.51 | 8.69 | 8.87 | 9.05 | 9.23 | 9.41 | 9.59 | 9.77 | 9.95 | 10.13 | 10.31 | 10.49 | 10.67 | 10.84 | 11.02 | 11.2 | 11.38 | 11.56 | 11.74 | 11.92 | 12.1 | 12.28 | 12.46 | 12.64 | 12.82 | 13 | 13.18 | 13.35 | 13.53 | 13.71 | 13.89 | 14.07 | 14.25
2.81 | 7.08 | 7.26 | 7.44 | 7.61 | 7.79 | 7.97 | 8.15 | 8.33 | 8.51 | 8.69 | 8.87 | 9.05 | 9.23 | 9.41 | 9.59 | 9.77 | 9.95 | 10.12 | 10.3 | 10.48 | 10.66 | 10.84 | 11.02 | 11.2 | 11.38 | 11.56 | 11.74 | 11.92 | 12.1 | 12.28 | 12.45 | 12.63 | 12.81 | 12.99 | 13.17 | 13.35 | 13.53 | 13.71 | 13.89 | 14.07 | 14.25
2.82 | 7.07 | 7.25 | 7.43 | 7.61 | 7.79 | 7.97 | 8.15 | 8.33 | 8.51 | 8.69 | 8.87 | 9.05 | 9.23 | 9.41 | 9.58 | 9.76 | 9.94 | 10.12 | 10.3 | 10.48 | 10.66 | 10.84 | 11.02 | 11.2 | 11.38 | 11.56 | 11.73 | 11.91 | 12.09 | 12.27 | 12.45 | 12.63 | 12.81 | 12.99 | 13.17 | 13.35 | 13.53 | 13.71 | 13.88 | 14.06 | 14.24
2.83 | 7.07 | 7.25 | 7.43 | 7.61 | 7.79 | 7.97 | 8.15 | 8.33 | 8.51 | 8.69 | 8.86 | 9.04 | 9.22 | 9.4 | 9.58 | 9.76 | 9.94 | 10.12 | 10.3 | 10.48 | 10.66 | 10.84 | 11.02 | 11.19 | 11.37 | 11.55 | 11.73 | 11.91 | 12.09 | 12.27 | 12.45 | 12.63 | 12.81 | 12.99 | 13.16 | 13.34 | 13.52 | 13.7 | 13.88 | 14.06 | 14.24
2.84 | 7.07 | 7.25 | 7.43 | 7.61 | 7.79 | 7.97 | 8.15 | 8.33 | 8.5 | 8.68 | 8.86 | 9.04 | 9.22 | 9.4 | 9.58 | 9.76 | 9.94 | 10.12 | 10.3 | 10.47 | 10.65 | 10.83 | 11.01 | 11.19 | 11.37 | 11.55 | 11.73 | 11.91 | 12.09 | 12.27 | 12.44 | 12.62 | 12.8 | 12.98 | 13.16 | 13.34 | 13.52 | 13.7 | 13.88 | 14.06 | 14.24
2.85 | 7.07 | 7.25 | 7.43 | 7.61 | 7.79 | 7.96 | 8.14 | 8.32 | 8.5 | 8.68 | 8.86 | 9.04 | 9.22 | 9.4 | 9.58 | 9.76 | 9.93 | 10.11 | 10.29 | 10.47 | 10.65 | 10.83 | 11.01 | 11.19 | 11.37 | 11.55 | 11.73 | 11.9 | 12.08 | 12.26 | 12.44 | 12.62 | 12.8 | 12.98 | 13.16 | 13.34 | 13.52 | 13.69 | 13.87 | 14.05 | 14.23
2.86 | 7.07 | 7.25 | 7.43 | 7.6 | 7.78 | 7.96 | 8.14 | 8.32 | 8.5 | 8.68 | 8.86 | 9.04 | 9.22 | 9.39 | 9.57 | 9.75 | 9.93 | 10.11 | 10.29 | 10.47 | 10.65 | 10.83 | 11.01 | 11.19 | 11.36 | 11.54 | 11.72 | 11.9 | 12.08 | 12.26 | 12.44 | 12.62 | 12.8 | 12.97 | 13.15 | 13.33 | 13.51 | 13.69 | 13.87 | 14.05 | 14.23
2.87 | 7.06 | 7.24 | 7.42 | 7.6 | 7.78 | 7.96 | 8.14 | 8.32 | 8.5 | 8.68 | 8.86 | 9.03 | 9.21 | 9.39 | 9.57 | 9.75 | 9.93 | 10.11 | 10.29 | 10.47 | 10.65 | 10.82 | 11 | 11.18 | 11.36 | 11.54 | 11.72 | 11.9 | 12.08 | 12.26 | 12.43 | 12.61 | 12.79 | 12.97 | 13.15 | 13.33 | 13.51 | 13.69 | 13.87 | 14.04 | 14.22
2.88 | 7.06 | 7.24 | 7.42 | 7.6 | 7.78 | 7.96 | 8.14 | 8.32 | 8.49 | 8.67 | 8.85 | 9.03 | 9.21 | 9.39 | 9.57 | 9.75 | 9.93 | 10.11 | 10.28 | 10.46 | 10.64 | 10.82 | 11 | 11.18 | 11.36 | 11.54 | 11.72 | 11.89 | 12.07 | 12.25 | 12.43 | 12.61 | 12.79 | 12.97 | 13.15 | 13.33 | 13.5 | 13.68 | 13.86 | 14.04 | 14.22
2.89 | 7.06 | 7.24 | 7.42 | 7.6 | 7.78 | 7.96 | 8.13 | 8.31 | 8.49 | 8.67 | 8.85 | 9.03 | 9.21 | 9.39 | 9.57 | 9.75 | 9.92 | 10.1 | 10.28 | 10.46 | 10.64 | 10.82 | 11 | 11.18 | 11.35 | 11.53 | 11.71 | 11.89 | 12.07 | 12.25 | 12.43 | 12.61 | 12.79 | 12.96 | 13.14 | 13.32 | 13.5 | 13.68 | 13.86 | 14.04 | 14.22
2.9 | 7.06 | 7.24 | 7.42 | 7.6 | 7.77 | 7.95 | 8.13 | 8.31 | 8.49 | 8.67 | 8.85 | 9.03 | 9.21 | 9.38 | 9.56 | 9.74 | 9.92 | 10.1 | 10.28 | 10.46 | 10.64 | 10.82 | 10.99 | 11.17 | 11.35 | 11.53 | 11.71 | 11.89 | 12.07 | 12.25 | 12.42 | 12.6 | 12.78 | 12.96 | 13.14 | 13.32 | 13.5 | 13.68 | 13.85 | 14.03 | 14.21
2.91 | 7.06 | 7.24 | 7.41 | 7.59 | 7.77 | 7.95 | 8.13 | 8.31 | 8.49 | 8.67 | 8.85 | 9.02 | 9.2 | 9.38 | 9.56 | 9.74 | 9.92 | 10.1 | 10.28 | 10.45 | 10.63 | 10.81 | 10.99 | 11.17 | 11.35 | 11.53 | 11.71 | 11.89 | 12.06 | 12.24 | 12.42 | 12.6 | 12.78 | 12.96 | 13.14 | 13.31 | 13.49 | 13.67 | 13.85 | 14.03 | 14.21
2.92 | 7.06 | 7.23 | 7.41 | 7.59 | 7.77 | 7.95 | 8.13 | 8.31 | 8.49 | 8.66 | 8.84 | 9.02 | 9.2 | 9.38 | 9.56 | 9.74 | 9.92 | 10.09 | 10.27 | 10.45 | 10.63 | 10.81 | 10.99 | 11.17 | 11.35 | 11.52 | 11.7 | 11.88 | 12.06 | 12.24 | 12.42 | 12.6 | 12.78 | 12.95 | 13.13 | 13.31 | 13.49 | 13.67 | 13.85 | 14.03 | 14.2
2.93 | 7.05 | 7.23 | 7.41 | 7.59 | 7.77 | 7.95 | 8.13 | 8.3 | 8.48 | 8.66 | 8.84 | 9.02 | 9.2 | 9.38 | 9.56 | 9.73 | 9.91 | 10.09 | 10.27 | 10.45 | 10.63 | 10.81 | 10.99 | 11.16 | 11.34 | 11.52 | 11.7 | 11.88 | 12.06 | 12.24 | 12.41 | 12.59 | 12.77 | 12.95 | 13.13 | 13.31 | 13.49 | 13.67 | 13.84 | 14.02 | 14.2
2.94 | 7.05 | 7.23 | 7.41 | 7.59 | 7.77 | 7.95 | 8.12 | 8.3 | 8.48 | 8.66 | 8.84 | 9.02 | 9.2 | 9.37 | 9.55 | 9.73 | 9.91 | 10.09 | 10.27 | 10.45 | 10.63 | 10.8 | 10.98 | 11.16 | 11.34 | 11.52 | 11.7 | 11.88 | 12.05 | 12.23 | 12.41 | 12.59 | 12.77 | 12.95 | 13.13 | 13.3 | 13.48 | 13.66 | 13.84 | 14.02 | 14.2
2.95 | 7.05 | 7.23 | 7.41 | 7.59 | 7.76 | 7.94 | 8.12 | 8.3 | 8.48 | 8.66 | 8.84 | 9.01 | 9.19 | 9.37 | 9.55 | 9.73 | 9.91 | 10.09 | 10.27 | 10.44 | 10.62 | 10.8 | 10.98 | 11.16 | 11.34 | 11.52 | 11.69 | 11.87 | 12.05 | 12.23 | 12.41 | 12.59 | 12.77 | 12.94 | 13.12 | 13.3 | 13.48 | 13.66 | 13.84 | 14.01 | 14.19
2.96 | 7.05 | 7.23 | 7.4 | 7.58 | 7.76 | 7.94 | 8.12 | 8.3 | 8.48 | 8.66 | 8.83 | 9.01 | 9.19 | 9.37 | 9.55 | 9.73 | 9.91 | 10.08 | 10.26 | 10.44 | 10.62 | 10.8 | 10.98 | 11.16 | 11.33 | 11.51 | 11.69 | 11.87 | 12.05 | 12.23 | 12.4 | 12.58 | 12.76 | 12.94 | 13.12 | 13.3 | 13.48 | 13.65 | 13.83 | 14.01 | 14.19
2.97 | 7.05 | 7.22 | 7.4 | 7.58 | 7.76 | 7.94 | 8.12 | 8.3 | 8.47 | 8.65 | 8.83 | 9.01 | 9.19 | 9.37 | 9.55 | 9.72 | 9.9 | 10.08 | 10.26 | 10.44 | 10.62 | 10.8 | 10.97 | 11.15 | 11.33 | 11.51 | 11.69 | 11.87 | 12.04 | 12.22 | 12.4 | 12.58 | 12.76 | 12.94 | 13.12 | 13.29 | 13.47 | 13.65 | 13.83 | 14.01 | 14.19
2.98 | 7.04 | 7.22 | 7.4 | 7.58 | 7.76 | 7.94 | 8.11 | 8.29 | 8.47 | 8.65 | 8.83 | 9.01 | 9.19 | 9.36 | 9.54 | 9.72 | 9.9 | 10.08 | 10.26 | 10.44 | 10.61 | 10.79 | 10.97 | 11.15 | 11.33 | 11.51 | 11.68 | 11.86 | 12.04 | 12.22 | 12.4 | 12.58 | 12.75 | 12.93 | 13.11 | 13.29 | 13.47 | 13.65 | 13.83 | 14 | 14.18
2.99 | 7.04 | 7.22 | 7.4 | 7.58 | 7.76 | 7.93 | 8.11 | 8.29 | 8.47 | 8.65 | 8.83 | 9 | 9.18 | 9.36 | 9.54 | 9.72 | 9.9 | 10.08 | 10.25 | 10.43 | 10.61 | 10.79 | 10.97 | 11.15 | 11.32 | 11.5 | 11.68 | 11.86 | 12.04 | 12.22 | 12.39 | 12.57 | 12.75 | 12.93 | 13.11 | 13.29 | 13.46 | 13.64 | 13.82 | 14 | 14.18
3 | 7.04 | 7.22 | 7.4 | 7.58 | 7.75 | 7.93 | 8.11 | 8.29 | 8.47 | 8.65 | 8.82 | 9 | 9.18 | 9.36 | 9.54 | 9.72 | 9.89 | 10.07 | 10.25 | 10.43 | 10.61 | 10.79 | 10.96 | 11.14 | 11.32 | 11.5 | 11.68 | 11.86 | 12.03 | 12.21 | 12.39 | 12.57 | 12.75 | 12.93 | 13.1 | 13.28 | 13.46 | 13.64 | 13.82 | 14 | 14.17
3.01 | 7.04 | 7.22 | 7.39 | 7.57 | 7.75 | 7.93 | 8.11 | 8.29 | 8.47 | 8.64 | 8.82 | 9 | 9.18 | 9.36 | 9.54 | 9.71 | 9.89 | 10.07 | 10.25 | 10.43 | 10.61 | 10.78 | 10.96 | 11.14 | 11.32 | 11.5 | 11.67 | 11.85 | 12.03 | 12.21 | 12.39 | 12.57 | 12.74 | 12.92 | 13.1 | 13.28 | 13.46 | 13.64 | 13.81 | 13.99 | 14.17
3.02 | 7.04 | 7.21 | 7.39 | 7.57 | 7.75 | 7.93 | 8.11 | 8.28 | 8.46 | 8.64 | 8.82 | 9 | 9.18 | 9.35 | 9.53 | 9.71 | 9.89 | 10.07 | 10.25 | 10.42 | 10.6 | 10.78 | 10.96 | 11.14 | 11.32 | 11.49 | 11.67 | 11.85 | 12.03 | 12.21 | 12.38 | 12.56 | 12.74 | 12.92 | 13.1 | 13.28 | 13.45 | 13.63 | 13.81 | 13.99 | 14.17
3.03 | 7.03 | 7.21 | 7.39 | 7.57 | 7.75 | 7.93 | 8.1 | 8.28 | 8.46 | 8.64 | 8.82 | 9 | 9.17 | 9.35 | 9.53 | 9.71 | 9.89 | 10.06 | 10.24 | 10.42 | 10.6 | 10.78 | 10.96 | 11.13 | 11.31 | 11.49 | 11.67 | 11.85 | 12.02 | 12.2 | 12.38 | 12.56 | 12.74 | 12.92 | 13.09 | 13.27 | 13.45 | 13.63 | 13.81 | 13.98 | 14.16
3.04 | 7.03 | 7.21 | 7.39 | 7.57 | 7.75 | 7.92 | 8.1 | 8.28 | 8.46 | 8.64 | 8.81 | 8.99 | 9.17 | 9.35 | 9.53 | 9.71 | 9.88 | 10.06 | 10.24 | 10.42 | 10.6 | 10.77 | 10.95 | 11.13 | 11.31 | 11.49 | 11.67 | 11.84 | 12.02 | 12.2 | 12.38 | 12.56 | 12.73 | 12.91 | 13.09 | 13.27 | 13.45 | 13.62 | 13.8 | 13.98 | 14.16
3.05 | 7.03 | 7.21 | 7.39 | 7.56 | 7.74 | 7.92 | 8.1 | 8.28 | 8.46 | 8.63 | 8.81 | 8.99 | 9.17 | 9.35 | 9.52 | 9.7 | 9.88 | 10.06 | 10.24 | 10.42 | 10.59 | 10.77 | 10.95 | 11.13 | 11.31 | 11.48 | 11.66 | 11.84 | 12.02 | 12.2 | 12.37 | 12.55 | 12.73 | 12.91 | 13.09 | 13.27 | 13.44 | 13.62 | 13.8 | 13.98 | 14.16
3.06 | 7.03 | 7.21 | 7.38 | 7.56 | 7.74 | 7.92 | 8.1 | 8.28 | 8.45 | 8.63 | 8.81 | 8.99 | 9.17 | 9.34 | 9.52 | 9.7 | 9.88 | 10.06 | 10.23 | 10.41 | 10.59 | 10.77 | 10.95 | 11.13 | 11.3 | 11.48 | 11.66 | 11.84 | 12.02 | 12.19 | 12.37 | 12.55 | 12.73 | 12.91 | 13.08 | 13.26 | 13.44 | 13.62 | 13.8 | 13.97 | 14.15
3.07 | 7.03 | 7.2 | 7.38 | 7.56 | 7.74 | 7.92 | 8.1 | 8.27 | 8.45 | 8.63 | 8.81 | 8.99 | 9.16 | 9.34 | 9.52 | 9.7 | 9.88 | 10.05 | 10.23 | 10.41 | 10.59 | 10.77 | 10.94 | 11.12 | 11.3 | 11.48 | 11.66 | 11.83 | 12.01 | 12.19 | 12.37 | 12.55 | 12.72 | 12.9 | 13.08 | 13.26 | 13.44 | 13.61 | 13.79 | 13.97 | 14.15
3.08 | 7.02 | 7.2 | 7.38 | 7.56 | 7.74 | 7.91 | 8.09 | 8.27 | 8.45 | 8.63 | 8.81 | 8.98 | 9.16 | 9.34 | 9.52 | 9.7 | 9.87 | 10.05 | 10.23 | 10.41 | 10.59 | 10.76 | 10.94 | 11.12 | 11.3 | 11.47 | 11.65 | 11.83 | 12.01 | 12.19 | 12.36 | 12.54 | 12.72 | 12.9 | 13.08 | 13.25 | 13.43 | 13.61 | 13.79 | 13.97 | 14.14
3.09 | 7.02 | 7.2 | 7.38 | 7.56 | 7.73 | 7.91 | 8.09 | 8.27 | 8.45 | 8.62 | 8.8 | 8.98 | 9.16 | 9.34 | 9.51 | 9.69 | 9.87 | 10.05 | 10.23 | 10.4 | 10.58 | 10.76 | 10.94 | 11.12 | 11.29 | 11.47 | 11.65 | 11.83 | 12.01 | 12.18 | 12.36 | 12.54 | 12.72 | 12.9 | 13.07 | 13.25 | 13.43 | 13.61 | 13.78 | 13.96 | 14.14
3.1 | 7.02 | 7.2 | 7.38 | 7.55 | 7.73 | 7.91 | 8.09 | 8.27 | 8.44 | 8.62 | 8.8 | 8.98 | 9.16 | 9.33 | 9.51 | 9.69 | 9.87 | 10.05 | 10.22 | 10.4 | 10.58 | 10.76 | 10.94 | 11.11 | 11.29 | 11.47 | 11.65 | 11.82 | 12 | 12.18 | 12.36 | 12.54 | 12.71 | 12.89 | 13.07 | 13.25 | 13.43 | 13.6 | 13.78 | 13.96 | 14.14
3.11 | 7.02 | 7.2 | 7.37 | 7.55 | 7.73 | 7.91 | 8.09 | 8.26 | 8.44 | 8.62 | 8.8 | 8.98 | 9.15 | 9.33 | 9.51 | 9.69 | 9.87 | 10.04 | 10.22 | 10.4 | 10.58 | 10.75 | 10.93 | 11.11 | 11.29 | 11.47 | 11.64 | 11.82 | 12 | 12.18 | 12.35 | 12.53 | 12.71 | 12.89 | 13.07 | 13.24 | 13.42 | 13.6 | 13.78 | 13.95 | 14.13
3.12 | 7.02 | 7.19 | 7.37 | 7.55 | 7.73 | 7.91 | 8.08 | 8.26 | 8.44 | 8.62 | 8.8 | 8.97 | 9.15 | 9.33 | 9.51 | 9.68 | 9.86 | 10.04 | 10.22 | 10.4 | 10.57 | 10.75 | 10.93 | 11.11 | 11.28 | 11.46 | 11.64 | 11.82 | 12 | 12.17 | 12.35 | 12.53 | 12.71 | 12.88 | 13.06 | 13.24 | 13.42 | 13.6 | 13.77 | 13.95 | 14.13
3.13 | 7.01 | 7.19 | 7.37 | 7.55 | 7.73 | 7.9 | 8.08 | 8.26 | 8.44 | 8.62 | 8.79 | 8.97 | 9.15 | 9.33 | 9.5 | 9.68 | 9.86 | 10.04 | 10.22 | 10.39 | 10.57 | 10.75 | 10.93 | 11.1 | 11.28 | 11.46 | 11.64 | 11.82 | 11.99 | 12.17 | 12.35 | 12.53 | 12.7 | 12.88 | 13.06 | 13.24 | 13.41 | 13.59 | 13.77 | 13.95 | 14.13
3.14 | 7.01 | 7.19 | 7.37 | 7.55 | 7.72 | 7.9 | 8.08 | 8.26 | 8.44 | 8.61 | 8.79 | 8.97 | 9.15 | 9.32 | 9.5 | 9.68 | 9.86 | 10.03 | 10.21 | 10.39 | 10.57 | 10.75 | 10.92 | 11.1 | 11.28 | 11.46 | 11.63 | 11.81 | 11.99 | 12.17 | 12.34 | 12.52 | 12.7 | 12.88 | 13.06 | 13.23 | 13.41 | 13.59 | 13.77 | 13.94 | 14.12
3.15 | 7.01 | 7.19 | 7.37 | 7.54 | 7.72 | 7.9 | 8.08 | 8.26 | 8.43 | 8.61 | 8.79 | 8.97 | 9.14 | 9.32 | 9.5 | 9.68 | 9.85 | 10.03 | 10.21 | 10.39 | 10.57 | 10.74 | 10.92 | 11.1 | 11.28 | 11.45 | 11.63 | 11.81 | 11.99 | 12.16 | 12.34 | 12.52 | 12.7 | 12.87 | 13.05 | 13.23 | 13.41 | 13.58 | 13.76 | 13.94 | 14.12
3.16 | 7.01 | 7.19 | 7.36 | 7.54 | 7.72 | 7.9 | 8.08 | 8.25 | 8.43 | 8.61 | 8.79 | 8.96 | 9.14 | 9.32 | 9.5 | 9.67 | 9.85 | 10.03 | 10.21 | 10.38 | 10.56 | 10.74 | 10.92 | 11.1 | 11.27 | 11.45 | 11.63 | 11.81 | 11.98 | 12.16 | 12.34 | 12.52 | 12.69 | 12.87 | 13.05 | 13.23 | 13.4 | 13.58 | 13.76 | 13.94 | 14.11
3.17 | 7.01 | 7.18 | 7.36 | 7.54 | 7.72 | 7.9 | 8.07 | 8.25 | 8.43 | 8.61 | 8.78 | 8.96 | 9.14 | 9.32 | 9.49 | 9.67 | 9.85 | 10.03 | 10.2 | 10.38 | 10.56 | 10.74 | 10.91 | 11.09 | 11.27 | 11.45 | 11.62 | 11.8 | 11.98 | 12.16 | 12.33 | 12.51 | 12.69 | 12.87 | 13.05 | 13.22 | 13.4 | 13.58 | 13.76 | 13.93 | 14.11
3.18 | 7.01 | 7.18 | 7.36 | 7.54 | 7.72 | 7.89 | 8.07 | 8.25 | 8.43 | 8.6 | 8.78 | 8.96 | 9.14 | 9.31 | 9.49 | 9.67 | 9.85 | 10.02 | 10.2 | 10.38 | 10.56 | 10.73 | 10.91 | 11.09 | 11.27 | 11.44 | 11.62 | 11.8 | 11.98 | 12.15 | 12.33 | 12.51 | 12.69 | 12.86 | 13.04 | 13.22 | 13.4 | 13.57 | 13.75 | 13.93 | 14.11
3.19 | 7 | 7.18 | 7.36 | 7.54 | 7.71 | 7.89 | 8.07 | 8.25 | 8.42 | 8.6 | 8.78 | 8.96 | 9.13 | 9.31 | 9.49 | 9.67 | 9.84 | 10.02 | 10.2 | 10.38 | 10.55 | 10.73 | 10.91 | 11.09 | 11.26 | 11.44 | 11.62 | 11.8 | 11.97 | 12.15 | 12.33 | 12.51 | 12.68 | 12.86 | 13.04 | 13.22 | 13.39 | 13.57 | 13.75 | 13.93 | 14.1
3.2 | 7 | 7.18 | 7.36 | 7.53 | 7.71 | 7.89 | 8.07 | 8.24 | 8.42 | 8.6 | 8.78 | 8.95 | 9.13 | 9.31 | 9.49 | 9.66 | 9.84 | 10.02 | 10.2 | 10.37 | 10.55 | 10.73 | 10.91 | 11.08 | 11.26 | 11.44 | 11.62 | 11.79 | 11.97 | 12.15 | 12.33 | 12.5 | 12.68 | 12.86 | 13.03 | 13.21 | 13.39 | 13.57 | 13.74 | 13.92 | 14.1
3.21 | 7 | 7.18 | 7.35 | 7.53 | 7.71 | 7.89 | 8.06 | 8.24 | 8.42 | 8.6 | 8.77 | 8.95 | 9.13 | 9.31 | 9.48 | 9.66 | 9.84 | 10.02 | 10.19 | 10.37 | 10.55 | 10.73 | 10.9 | 11.08 | 11.26 | 11.44 | 11.61 | 11.79 | 11.97 | 12.14 | 12.32 | 12.5 | 12.68 | 12.85 | 13.03 | 13.21 | 13.39 | 13.56 | 13.74 | 13.92 | 14.09
3.22 | 7 | 7.18 | 7.35 | 7.53 | 7.71 | 7.88 | 8.06 | 8.24 | 8.42 | 8.59 | 8.77 | 8.95 | 9.13 | 9.3 | 9.48 | 9.66 | 9.84 | 10.01 | 10.19 | 10.37 | 10.55 | 10.72 | 10.9 | 11.08 | 11.25 | 11.43 | 11.61 | 11.79 | 11.96 | 12.14 | 12.32 | 12.5 | 12.67 | 12.85 | 13.03 | 13.2 | 13.38 | 13.56 | 13.74 | 13.91 | 14.09
3.23 | 7 | 7.17 | 7.35 | 7.53 | 7.71 | 7.88 | 8.06 | 8.24 | 8.41 | 8.59 | 8.77 | 8.95 | 9.12 | 9.3 | 9.48 | 9.66 | 9.83 | 10.01 | 10.19 | 10.37 | 10.54 | 10.72 | 10.9 | 11.07 | 11.25 | 11.43 | 11.61 | 11.78 | 11.96 | 12.14 | 12.32 | 12.49 | 12.67 | 12.85 | 13.02 | 13.2 | 13.38 | 13.56 | 13.73 | 13.91 | 14.09
3.24 | 6.99 | 7.17 | 7.35 | 7.53 | 7.7 | 7.88 | 8.06 | 8.24 | 8.41 | 8.59 | 8.77 | 8.94 | 9.12 | 9.3 | 9.48 | 9.65 | 9.83 | 10.01 | 10.19 | 10.36 | 10.54 | 10.72 | 10.89 | 11.07 | 11.25 | 11.43 | 11.6 | 11.78 | 11.96 | 12.13 | 12.31 | 12.49 | 12.67 | 12.84 | 13.02 | 13.2 | 13.38 | 13.55 | 13.73 | 13.91 | 14.08
3.25 | 6.99 | 7.17 | 7.35 | 7.52 | 7.7 | 7.88 | 8.06 | 8.23 | 8.41 | 8.59 | 8.76 | 8.94 | 9.12 | 9.3 | 9.47 | 9.65 | 9.83 | 10.01 | 10.18 | 10.36 | 10.54 | 10.71 | 10.89 | 11.07 | 11.25 | 11.42 | 11.6 | 11.78 | 11.95 | 12.13 | 12.31 | 12.49 | 12.66 | 12.84 | 13.02 | 13.19 | 13.37 | 13.55 | 13.73 | 13.9 | 14.08
3.26 | 6.99 | 7.17 | 7.34 | 7.52 | 7.7 | 7.88 | 8.05 | 8.23 | 8.41 | 8.59 | 8.76 | 8.94 | 9.12 | 9.29 | 9.47 | 9.65 | 9.83 | 10 | 10.18 | 10.36 | 10.53 | 10.71 | 10.89 | 11.07 | 11.24 | 11.42 | 11.6 | 11.77 | 11.95 | 12.13 | 12.31 | 12.48 | 12.66 | 12.84 | 13.01 | 13.19 | 13.37 | 13.54 | 13.72 | 13.9 | 14.08
3.27 | 6.99 | 7.17 | 7.34 | 7.52 | 7.7 | 7.87 | 8.05 | 8.23 | 8.41 | 8.58 | 8.76 | 8.94 | 9.11 | 9.29 | 9.47 | 9.65 | 9.82 | 10 | 10.18 | 10.35 | 10.53 | 10.71 | 10.89 | 11.06 | 11.24 | 11.42 | 11.59 | 11.77 | 11.95 | 12.12 | 12.3 | 12.48 | 12.66 | 12.83 | 13.01 | 13.19 | 13.36 | 13.54 | 13.72 | 13.9 | 14.07
3.28 | 6.99 | 7.16 | 7.34 | 7.52 | 7.69 | 7.87 | 8.05 | 8.23 | 8.4 | 8.58 | 8.76 | 8.93 | 9.11 | 9.29 | 9.47 | 9.64 | 9.82 | 10 | 10.17 | 10.35 | 10.53 | 10.71 | 10.88 | 11.06 | 11.24 | 11.41 | 11.59 | 11.77 | 11.94 | 12.12 | 12.3 | 12.48 | 12.65 | 12.83 | 13.01 | 13.18 | 13.36 | 13.54 | 13.71 | 13.89 | 14.07
3.29 | 6.98 | 7.16 | 7.34 | 7.52 | 7.69 | 7.87 | 8.05 | 8.22 | 8.4 | 8.58 | 8.76 | 8.93 | 9.11 | 9.29 | 9.46 | 9.64 | 9.82 | 9.99 | 10.17 | 10.35 | 10.53 | 10.7 | 10.88 | 11.06 | 11.23 | 11.41 | 11.59 | 11.76 | 11.94 | 12.12 | 12.3 | 12.47 | 12.65 | 12.83 | 13 | 13.18 | 13.36 | 13.53 | 13.71 | 13.89 | 14.06
3.3 | 6.98 | 7.16 | 7.34 | 7.51 | 7.69 | 7.87 | 8.04 | 8.22 | 8.4 | 8.58 | 8.75 | 8.93 | 9.11 | 9.28 | 9.46 | 9.64 | 9.81 | 9.99 | 10.17 | 10.35 | 10.52 | 10.7 | 10.88 | 11.05 | 11.23 | 11.41 | 11.58 | 11.76 | 11.94 | 12.12 | 12.29 | 12.47 | 12.65 | 12.82 | 13 | 13.18 | 13.35 | 13.53 | 13.71 | 13.88 | 14.06
3.31 | 6.98 | 7.16 | 7.33 | 7.51 | 7.69 | 7.87 | 8.04 | 8.22 | 8.4 | 8.57 | 8.75 | 8.93 | 9.1 | 9.28 | 9.46 | 9.64 | 9.81 | 9.99 | 10.17 | 10.34 | 10.52 | 10.7 | 10.87 | 11.05 | 11.23 | 11.4 | 11.58 | 11.76 | 11.94 | 12.11 | 12.29 | 12.47 | 12.64 | 12.82 | 13 | 13.17 | 13.35 | 13.53 | 13.7 | 13.88 | 14.06
3.32 | 6.98 | 7.16 | 7.33 | 7.51 | 7.69 | 7.86 | 8.04 | 8.22 | 8.39 | 8.57 | 8.75 | 8.93 | 9.1 | 9.28 | 9.46 | 9.63 | 9.81 | 9.99 | 10.16 | 10.34 | 10.52 | 10.69 | 10.87 | 11.05 | 11.22 | 11.4 | 11.58 | 11.76 | 11.93 | 12.11 | 12.29 | 12.46 | 12.64 | 12.82 | 12.99 | 13.17 | 13.35 | 13.52 | 13.7 | 13.88 | 14.05
3.33 | 6.98 | 7.15 | 7.33 | 7.51 | 7.68 | 7.86 | 8.04 | 8.22 | 8.39 | 8.57 | 8.75 | 8.92 | 9.1 | 9.28 | 9.45 | 9.63 | 9.81 | 9.98 | 10.16 | 10.34 | 10.51 | 10.69 | 10.87 | 11.04 | 11.22 | 11.4 | 11.58 | 11.75 | 11.93 | 12.11 | 12.28 | 12.46 | 12.64 | 12.81 | 12.99 | 13.17 | 13.34 | 13.52 | 13.7 | 13.87 | 14.05
3.34 | 6.98 | 7.15 | 7.33 | 7.51 | 7.68 | 7.86 | 8.04 | 8.21 | 8.39 | 8.57 | 8.74 | 8.92 | 9.1 | 9.27 | 9.45 | 9.63 | 9.8 | 9.98 | 10.16 | 10.33 | 10.51 | 10.69 | 10.87 | 11.04 | 11.22 | 11.4 | 11.57 | 11.75 | 11.93 | 12.1 | 12.28 | 12.46 | 12.63 | 12.81 | 12.99 | 13.16 | 13.34 | 13.52 | 13.69 | 13.87 | 14.05
3.35 | 6.97 | 7.15 | 7.33 | 7.5 | 7.68 | 7.86 | 8.03 | 8.21 | 8.39 | 8.56 | 8.74 | 8.92 | 9.09 | 9.27 | 9.45 | 9.62 | 9.8 | 9.98 | 10.16 | 10.33 | 10.51 | 10.69 | 10.86 | 11.04 | 11.22 | 11.39 | 11.57 | 11.75 | 11.92 | 12.1 | 12.28 | 12.45 | 12.63 | 12.81 | 12.98 | 13.16 | 13.34 | 13.51 | 13.69 | 13.87 | 14.04
3.36 | 6.97 | 7.15 | 7.32 | 7.5 | 7.68 | 7.86 | 8.03 | 8.21 | 8.39 | 8.56 | 8.74 | 8.92 | 9.09 | 9.27 | 9.45 | 9.62 | 9.8 | 9.98 | 10.15 | 10.33 | 10.51 | 10.68 | 10.86 | 11.04 | 11.21 | 11.39 | 11.57 | 11.74 | 11.92 | 12.1 | 12.27 | 12.45 | 12.63 | 12.8 | 12.98 | 13.16 | 13.33 | 13.51 | 13.69 | 13.86 | 14.04
3.37 | 6.97 | 7.15 | 7.32 | 7.5 | 7.68 | 7.85 | 8.03 | 8.21 | 8.38 | 8.56 | 8.74 | 8.91 | 9.09 | 9.27 | 9.44 | 9.62 | 9.8 | 9.97 | 10.15 | 10.33 | 10.5 | 10.68 | 10.86 | 11.03 | 11.21 | 11.39 | 11.56 | 11.74 | 11.92 | 12.09 | 12.27 | 12.45 | 12.62 | 12.8 | 12.98 | 13.15 | 13.33 | 13.51 | 13.68 | 13.86 | 14.04
3.38 | 6.97 | 7.14 | 7.32 | 7.5 | 7.67 | 7.85 | 8.03 | 8.2 | 8.38 | 8.56 | 8.73 | 8.91 | 9.09 | 9.26 | 9.44 | 9.62 | 9.79 | 9.97 | 10.15 | 10.32 | 10.5 | 10.68 | 10.85 | 11.03 | 11.21 | 11.38 | 11.56 | 11.74 | 11.91 | 12.09 | 12.27 | 12.44 | 12.62 | 12.8 | 12.97 | 13.15 | 13.33 | 13.5 | 13.68 | 13.85 | 14.03
3.39 | 6.97 | 7.14 | 7.32 | 7.5 | 7.67 | 7.85 | 8.03 | 8.2 | 8.38 | 8.56 | 8.73 | 8.91 | 9.08 | 9.26 | 9.44 | 9.61 | 9.79 | 9.97 | 10.14 | 10.32 | 10.5 | 10.67 | 10.85 | 11.03 | 11.2 | 11.38 | 11.56 | 11.73 | 11.91 | 12.09 | 12.26 | 12.44 | 12.62 | 12.79 | 12.97 | 13.15 | 13.32 | 13.5 | 13.67 | 13.85 | 14.03
3.4 | 6.96 | 7.14 | 7.32 | 7.49 | 7.67 | 7.85 | 8.02 | 8.2 | 8.38 | 8.55 | 8.73 | 8.91 | 9.08 | 9.26 | 9.44 | 9.61 | 9.79 | 9.97 | 10.14 | 10.32 | 10.49 | 10.67 | 10.85 | 11.02 | 11.2 | 11.38 | 11.55 | 11.73 | 11.91 | 12.08 | 12.26 | 12.44 | 12.61 | 12.79 | 12.97 | 13.14 | 13.32 | 13.49 | 13.67 | 13.85 | 14.02
3.41 | 6.96 | 7.14 | 7.31 | 7.49 | 7.67 | 7.84 | 8.02 | 8.2 | 8.37 | 8.55 | 8.73 | 8.9 | 9.08 | 9.26 | 9.43 | 9.61 | 9.79 | 9.96 | 10.14 | 10.32 | 10.49 | 10.67 | 10.84 | 11.02 | 11.2 | 11.37 | 11.55 | 11.73 | 11.9 | 12.08 | 12.26 | 12.43 | 12.61 | 12.79 | 12.96 | 13.14 | 13.31 | 13.49 | 13.67 | 13.84 | 14.02
3.42 | 6.96 | 7.14 | 7.31 | 7.49 | 7.67 | 7.84 | 8.02 | 8.2 | 8.37 | 8.55 | 8.72 | 8.9 | 9.08 | 9.25 | 9.43 | 9.61 | 9.78 | 9.96 | 10.14 | 10.31 | 10.49 | 10.67 | 10.84 | 11.02 | 11.19 | 11.37 | 11.55 | 11.72 | 11.9 | 12.08 | 12.25 | 12.43 | 12.61 | 12.78 | 12.96 | 13.13 | 13.31 | 13.49 | 13.66 | 13.84 | 14.02
3.43 | 6.96 | 7.13 | 7.31 | 7.49 | 7.66 | 7.84 | 8.02 | 8.19 | 8.37 | 8.55 | 8.72 | 8.9 | 9.08 | 9.25 | 9.43 | 9.6 | 9.78 | 9.96 | 10.13 | 10.31 | 10.49 | 10.66 | 10.84 | 11.02 | 11.19 | 11.37 | 11.54 | 11.72 | 11.9 | 12.07 | 12.25 | 12.43 | 12.6 | 12.78 | 12.95 | 13.13 | 13.31 | 13.48 | 13.66 | 13.84 | 14.01
3.44 | 6.96 | 7.13 | 7.31 | 7.49 | 7.66 | 7.84 | 8.01 | 8.19 | 8.37 | 8.54 | 8.72 | 8.9 | 9.07 | 9.25 | 9.43 | 9.6 | 9.78 | 9.95 | 10.13 | 10.31 | 10.48 | 10.66 | 10.84 | 11.01 | 11.19 | 11.36 | 11.54 | 11.72 | 11.89 | 12.07 | 12.25 | 12.42 | 12.6 | 12.78 | 12.95 | 13.13 | 13.3 | 13.48 | 13.66 | 13.83 | 14.01
3.45 | 6.95 | 7.13 | 7.31 | 7.48 | 7.66 | 7.84 | 8.01 | 8.19 | 8.37 | 8.54 | 8.72 | 8.89 | 9.07 | 9.25 | 9.42 | 9.6 | 9.78 | 9.95 | 10.13 | 10.3 | 10.48 | 10.66 | 10.83 | 11.01 | 11.19 | 11.36 | 11.54 | 11.71 | 11.89 | 12.07 | 12.24 | 12.42 | 12.6 | 12.77 | 12.95 | 13.12 | 13.3 | 13.48 | 13.65 | 13.83 | 14.01
3.46 | 6.95 | 7.13 | 7.31 | 7.48 | 7.66 | 7.83 | 8.01 | 8.19 | 8.36 | 8.54 | 8.72 | 8.89 | 9.07 | 9.24 | 9.42 | 9.6 | 9.77 | 9.95 | 10.13 | 10.3 | 10.48 | 10.65 | 10.83 | 11.01 | 11.18 | 11.36 | 11.54 | 11.71 | 11.89 | 12.06 | 12.24 | 12.42 | 12.59 | 12.77 | 12.94 | 13.12 | 13.3 | 13.47 | 13.65 | 13.83 | 14
3.47 | 6.95 | 7.13 | 7.3 | 7.48 | 7.66 | 7.83 | 8.01 | 8.18 | 8.36 | 8.54 | 8.71 | 8.89 | 9.07 | 9.24 | 9.42 | 9.59 | 9.77 | 9.95 | 10.12 | 10.3 | 10.47 | 10.65 | 10.83 | 11 | 11.18 | 11.36 | 11.53 | 11.71 | 11.88 | 12.06 | 12.24 | 12.41 | 12.59 | 12.76 | 12.94 | 13.12 | 13.29 | 13.47 | 13.65 | 13.82 | 14
3.48 | 6.95 | 7.12 | 7.3 | 7.48 | 7.65 | 7.83 | 8.01 | 8.18 | 8.36 | 8.53 | 8.71 | 8.89 | 9.06 | 9.24 | 9.42 | 9.59 | 9.77 | 9.94 | 10.12 | 10.3 | 10.47 | 10.65 | 10.82 | 11 | 11.18 | 11.35 | 11.53 | 11.71 | 11.88 | 12.06 | 12.23 | 12.41 | 12.59 | 12.76 | 12.94 | 13.11 | 13.29 | 13.47 | 13.64 | 13.82 | 13.99
3.49 | 6.95 | 7.12 | 7.3 | 7.48 | 7.65 | 7.83 | 8 | 8.18 | 8.36 | 8.53 | 8.71 | 8.88 | 9.06 | 9.24 | 9.41 | 9.59 | 9.77 | 9.94 | 10.12 | 10.29 | 10.47 | 10.65 | 10.82 | 11 | 11.17 | 11.35 | 11.53 | 11.7 | 11.88 | 12.05 | 12.23 | 12.41 | 12.58 | 12.76 | 12.93 | 13.11 | 13.29 | 13.46 | 13.64 | 13.81 | 13.99
3.5 | 6.94 | 7.12 | 7.3 | 7.47 | 7.65 | 7.83 | 8 | 8.18 | 8.35 | 8.53 | 8.71 | 8.88 | 9.06 | 9.23 | 9.41 | 9.59 | 9.76 | 9.94 | 10.11 | 10.29 | 10.47 | 10.64 | 10.82 | 10.99 | 11.17 | 11.35 | 11.52 | 11.7 | 11.87 | 12.05 | 12.23 | 12.4 | 12.58 | 12.75 | 12.93 | 13.11 | 13.28 | 13.46 | 13.63 | 13.81 | 13.99
3.51 | 6.94 | 7.12 | 7.3 | 7.47 | 7.65 | 7.82 | 8 | 8.18 | 8.35 | 8.53 | 8.7 | 8.88 | 9.06 | 9.23 | 9.41 | 9.58 | 9.76 | 9.94 | 10.11 | 10.29 | 10.46 | 10.64 | 10.82 | 10.99 | 11.17 | 11.34 | 11.52 | 11.7 | 11.87 | 12.05 | 12.22 | 12.4 | 12.58 | 12.75 | 12.93 | 13.1 | 13.28 | 13.46 | 13.63 | 13.81 | 13.98
3.52 | 6.94 | 7.12 | 7.29 | 7.47 | 7.65 | 7.82 | 8 | 8.17 | 8.35 | 8.53 | 8.7 | 8.88 | 9.05 | 9.23 | 9.41 | 9.58 | 9.76 | 9.93 | 10.11 | 10.29 | 10.46 | 10.64 | 10.81 | 10.99 | 11.16 | 11.34 | 11.52 | 11.69 | 11.87 | 12.04 | 12.22 | 12.4 | 12.57 | 12.75 | 12.92 | 13.1 | 13.28 | 13.45 | 13.63 | 13.8 | 13.98
3.53 | 6.94 | 7.12 | 7.29 | 7.47 | 7.64 | 7.82 | 8 | 8.17 | 8.35 | 8.52 | 8.7 | 8.87 | 9.05 | 9.23 | 9.4 | 9.58 | 9.75 | 9.93 | 10.11 | 10.28 | 10.46 | 10.63 | 10.81 | 10.99 | 11.16 | 11.34 | 11.51 | 11.69 | 11.87 | 12.04 | 12.22 | 12.39 | 12.57 | 12.74 | 12.92 | 13.1 | 13.27 | 13.45 | 13.62 | 13.8 | 13.98
3.54 | 6.94 | 7.11 | 7.29 | 7.47 | 7.64 | 7.82 | 7.99 | 8.17 | 8.34 | 8.52 | 8.7 | 8.87 | 9.05 | 9.22 | 9.4 | 9.58 | 9.75 | 9.93 | 10.1 | 10.28 | 10.46 | 10.63 | 10.81 | 10.98 | 11.16 | 11.33 | 11.51 | 11.69 | 11.86 | 12.04 | 12.21 | 12.39 | 12.57 | 12.74 | 12.92 | 13.09 | 13.27 | 13.44 | 13.62 | 13.8 | 13.97
3.55 | 6.94 | 7.11 | 7.29 | 7.46 | 7.64 | 7.81 | 7.99 | 8.17 | 8.34 | 8.52 | 8.69 | 8.87 | 9.05 | 9.22 | 9.4 | 9.57 | 9.75 | 9.93 | 10.1 | 10.28 | 10.45 | 10.63 | 10.8 | 10.98 | 11.16 | 11.33 | 11.51 | 11.68 | 11.86 | 12.03 | 12.21 | 12.39 | 12.56 | 12.74 | 12.91 | 13.09 | 13.27 | 13.44 | 13.62 | 13.79 | 13.97
3.56 | 6.93 | 7.11 | 7.29 | 7.46 | 7.64 | 7.81 | 7.99 | 8.16 | 8.34 | 8.52 | 8.69 | 8.87 | 9.04 | 9.22 | 9.4 | 9.57 | 9.75 | 9.92 | 10.1 | 10.27 | 10.45 | 10.63 | 10.8 | 10.98 | 11.15 | 11.33 | 11.5 | 11.68 | 11.86 | 12.03 | 12.21 | 12.38 | 12.56 | 12.73 | 12.91 | 13.09 | 13.26 | 13.44 | 13.61 | 13.79 | 13.96
3.57 | 6.93 | 7.11 | 7.28 | 7.46 | 7.63 | 7.81 | 7.99 | 8.16 | 8.34 | 8.51 | 8.69 | 8.87 | 9.04 | 9.22 | 9.39 | 9.57 | 9.74 | 9.92 | 10.1 | 10.27 | 10.45 | 10.62 | 10.8 | 10.97 | 11.15 | 11.33 | 11.5 | 11.68 | 11.85 | 12.03 | 12.2 | 12.38 | 12.56 | 12.73 | 12.91 | 13.08 | 13.26 | 13.43 | 13.61 | 13.79 | 13.96
3.58 | 6.93 | 7.11 | 7.28 | 7.46 | 7.63 | 7.81 | 7.98 | 8.16 | 8.34 | 8.51 | 8.69 | 8.86 | 9.04 | 9.21 | 9.39 | 9.57 | 9.74 | 9.92 | 10.09 | 10.27 | 10.44 | 10.62 | 10.8 | 10.97 | 11.15 | 11.32 | 11.5 | 11.67 | 11.85 | 12.03 | 12.2 | 12.38 | 12.55 | 12.73 | 12.9 | 13.08 | 13.25 | 13.43 | 13.61 | 13.78 | 13.96
3.59 | 6.93 | 7.1 | 7.28 | 7.46 | 7.63 | 7.81 | 7.98 | 8.16 | 8.33 | 8.51 | 8.68 | 8.86 | 9.04 | 9.21 | 9.39 | 9.56 | 9.74 | 9.91 | 10.09 | 10.27 | 10.44 | 10.62 | 10.79 | 10.97 | 11.14 | 11.32 | 11.5 | 11.67 | 11.85 | 12.02 | 12.2 | 12.37 | 12.55 | 12.72 | 12.9 | 13.08 | 13.25 | 13.43 | 13.6 | 13.78 | 13.95
3.6 | 6.93 | 7.1 | 7.28 | 7.45 | 7.63 | 7.8 | 7.98 | 8.16 | 8.33 | 8.51 | 8.68 | 8.86 | 9.03 | 9.21 | 9.39 | 9.56 | 9.74 | 9.91 | 10.09 | 10.26 | 10.44 | 10.61 | 10.79 | 10.97 | 11.14 | 11.32 | 11.49 | 11.67 | 11.84 | 12.02 | 12.19 | 12.37 | 12.55 | 12.72 | 12.9 | 13.07 | 13.25 | 13.42 | 13.6 | 13.77 | 13.95
3.61 | 6.92 | 7.1 | 7.28 | 7.45 | 7.63 | 7.8 | 7.98 | 8.15 | 8.33 | 8.5 | 8.68 | 8.86 | 9.03 | 9.21 | 9.38 | 9.56 | 9.73 | 9.91 | 10.08 | 10.26 | 10.44 | 10.61 | 10.79 | 10.96 | 11.14 | 11.31 | 11.49 | 11.66 | 11.84 | 12.02 | 12.19 | 12.37 | 12.54 | 12.72 | 12.89 | 13.07 | 13.24 | 13.42 | 13.6 | 13.77 | 13.95
3.62 | 6.92 | 7.1 | 7.27 | 7.45 | 7.62 | 7.8 | 7.98 | 8.15 | 8.33 | 8.5 | 8.68 | 8.85 | 9.03 | 9.2 | 9.38 | 9.56 | 9.73 | 9.91 | 10.08 | 10.26 | 10.43 | 10.61 | 10.78 | 10.96 | 11.14 | 11.31 | 11.49 | 11.66 | 11.84 | 12.01 | 12.19 | 12.36 | 12.54 | 12.71 | 12.89 | 13.07 | 13.24 | 13.42 | 13.59 | 13.77 | 13.94
3.63 | 6.92 | 7.1 | 7.27 | 7.45 | 7.62 | 7.8 | 7.97 | 8.15 | 8.32 | 8.5 | 8.68 | 8.85 | 9.03 | 9.2 | 9.38 | 9.55 | 9.73 | 9.9 | 10.08 | 10.25 | 10.43 | 10.61 | 10.78 | 10.96 | 11.13 | 11.31 | 11.48 | 11.66 | 11.83 | 12.01 | 12.18 | 12.36 | 12.54 | 12.71 | 12.89 | 13.06 | 13.24 | 13.41 | 13.59 | 13.76 | 13.94
3.64 | 6.92 | 7.09 | 7.27 | 7.45 | 7.62 | 7.8 | 7.97 | 8.15 | 8.32 | 8.5 | 8.67 | 8.85 | 9.02 | 9.2 | 9.38 | 9.55 | 9.73 | 9.9 | 10.08 | 10.25 | 10.43 | 10.6 | 10.78 | 10.95 | 11.13 | 11.3 | 11.48 | 11.66 | 11.83 | 12.01 | 12.18 | 12.36 | 12.53 | 12.71 | 12.88 | 13.06 | 13.23 | 13.41 | 13.58 | 13.76 | 13.93
3.65 | 6.92 | 7.09 | 7.27 | 7.44 | 7.62 | 7.79 | 7.97 | 8.14 | 8.32 | 8.5 | 8.67 | 8.85 | 9.02 | 9.2 | 9.37 | 9.55 | 9.72 | 9.9 | 10.07 | 10.25 | 10.42 | 10.6 | 10.78 | 10.95 | 11.13 | 11.3 | 11.48 | 11.65 | 11.83 | 12 | 12.18 | 12.35 | 12.53 | 12.7 | 12.88 | 13.05 | 13.23 | 13.41 | 13.58 | 13.76 | 13.93
3.66 | 6.91 | 7.09 | 7.27 | 7.44 | 7.62 | 7.79 | 7.97 | 8.14 | 8.32 | 8.49 | 8.67 | 8.84 | 9.02 | 9.19 | 9.37 | 9.55 | 9.72 | 9.9 | 10.07 | 10.25 | 10.42 | 10.6 | 10.77 | 10.95 | 11.12 | 11.3 | 11.47 | 11.65 | 11.82 | 12 | 12.17 | 12.35 | 12.53 | 12.7 | 12.88 | 13.05 | 13.23 | 13.4 | 13.58 | 13.75 | 13.93
3.67 | 6.91 | 7.09 | 7.26 | 7.44 | 7.61 | 7.79 | 7.97 | 8.14 | 8.32 | 8.49 | 8.67 | 8.84 | 9.02 | 9.19 | 9.37 | 9.54 | 9.72 | 9.89 | 10.07 | 10.24 | 10.42 | 10.59 | 10.77 | 10.95 | 11.12 | 11.3 | 11.47 | 11.65 | 11.82 | 12 | 12.17 | 12.35 | 12.52 | 12.7 | 12.87 | 13.05 | 13.22 | 13.4 | 13.57 | 13.75 | 13.92
3.68 | 6.91 | 7.09 | 7.26 | 7.44 | 7.61 | 7.79 | 7.96 | 8.14 | 8.31 | 8.49 | 8.66 | 8.84 | 9.01 | 9.19 | 9.37 | 9.54 | 9.72 | 9.89 | 10.07 | 10.24 | 10.42 | 10.59 | 10.77 | 10.94 | 11.12 | 11.29 | 11.47 | 11.64 | 11.82 | 11.99 | 12.17 | 12.34 | 12.52 | 12.69 | 12.87 | 13.04 | 13.22 | 13.39 | 13.57 | 13.74 | 13.92
3.69 | 6.91 | 7.08 | 7.26 | 7.44 | 7.61 | 7.79 | 7.96 | 8.14 | 8.31 | 8.49 | 8.66 | 8.84 | 9.01 | 9.19 | 9.36 | 9.54 | 9.71 | 9.89 | 10.06 | 10.24 | 10.41 | 10.59 | 10.76 | 10.94 | 11.11 | 11.29 | 11.46 | 11.64 | 11.82 | 11.99 | 12.17 | 12.34 | 12.52 | 12.69 | 12.87 | 13.04 | 13.22 | 13.39 | 13.57 | 13.74 | 13.92
3.7 | 6.91 | 7.08 | 7.26 | 7.43 | 7.61 | 7.78 | 7.96 | 8.13 | 8.31 | 8.48 | 8.66 | 8.83 | 9.01 | 9.18 | 9.36 | 9.54 | 9.71 | 9.89 | 10.06 | 10.24 | 10.41 | 10.59 | 10.76 | 10.94 | 11.11 | 11.29 | 11.46 | 11.64 | 11.81 | 11.99 | 12.16 | 12.34 | 12.51 | 12.69 | 12.86 | 13.04 | 13.21 | 13.39 | 13.56 | 13.74 | 13.91
3.71 | 6.91 | 7.08 | 7.26 | 7.43 | 7.61 | 7.78 | 7.96 | 8.13 | 8.31 | 8.48 | 8.66 | 8.83 | 9.01 | 9.18 | 9.36 | 9.53 | 9.71 | 9.88 | 10.06 | 10.23 | 10.41 | 10.58 | 10.76 | 10.93 | 11.11 | 11.28 | 11.46 | 11.63 | 11.81 | 11.98 | 12.16 | 12.33 | 12.51 | 12.68 | 12.86 | 13.03 | 13.21 | 13.38 | 13.56 | 13.73 | 13.91
3.72 | 6.9 | 7.08 | 7.25 | 7.43 | 7.6 | 7.78 | 7.95 | 8.13 | 8.3 | 8.48 | 8.65 | 8.83 | 9 | 9.18 | 9.36 | 9.53 | 9.71 | 9.88 | 10.06 | 10.23 | 10.41 | 10.58 | 10.76 | 10.93 | 11.11 | 11.28 | 11.46 | 11.63 | 11.81 | 11.98 | 12.16 | 12.33 | 12.51 | 12.68 | 12.86 | 13.03 | 13.21 | 13.38 | 13.56 | 13.73 | 13.91
3.73 | 6.9 | 7.08 | 7.25 | 7.43 | 7.6 | 7.78 | 7.95 | 8.13 | 8.3 | 8.48 | 8.65 | 8.83 | 9 | 9.18 | 9.35 | 9.53 | 9.7 | 9.88 | 10.05 | 10.23 | 10.4 | 10.58 | 10.75 | 10.93 | 11.1 | 11.28 | 11.45 | 11.63 | 11.8 | 11.98 | 12.15 | 12.33 | 12.5 | 12.68 | 12.85 | 13.03 | 13.2 | 13.38 | 13.55 | 13.73 | 13.9
3.74 | 6.9 | 7.08 | 7.25 | 7.43 | 7.6 | 7.78 | 7.95 | 8.13 | 8.3 | 8.48 | 8.65 | 8.83 | 9 | 9.18 | 9.35 | 9.53 | 9.7 | 9.88 | 10.05 | 10.22 | 10.4 | 10.57 | 10.75 | 10.92 | 11.1 | 11.27 | 11.45 | 11.62 | 11.8 | 11.97 | 12.15 | 12.32 | 12.5 | 12.67 | 12.85 | 13.02 | 13.2 | 13.37 | 13.55 | 13.72 | 13.9
3.75 | 6.9 | 7.07 | 7.25 | 7.42 | 7.6 | 7.77 | 7.95 | 8.12 | 8.3 | 8.47 | 8.65 | 8.82 | 9 | 9.17 | 9.35 | 9.52 | 9.7 | 9.87 | 10.05 | 10.22 | 10.4 | 10.57 | 10.75 | 10.92 | 11.1 | 11.27 | 11.45 | 11.62 | 11.8 | 11.97 | 12.15 | 12.32 | 12.5 | 12.67 | 12.85 | 13.02 | 13.2 | 13.37 | 13.54 | 13.72 | 13.89
3.76 | 6.9 | 7.07 | 7.25 | 7.42 | 7.6 | 7.77 | 7.95 | 8.12 | 8.3 | 8.47 | 8.65 | 8.82 | 9 | 9.17 | 9.35 | 9.52 | 9.69 | 9.87 | 10.04 | 10.22 | 10.39 | 10.57 | 10.74 | 10.92 | 11.09 | 11.27 | 11.44 | 11.62 | 11.79 | 11.97 | 12.14 | 12.32 | 12.49 | 12.67 | 12.84 | 13.02 | 13.19 | 13.37 | 13.54 | 13.72 | 13.89
3.77 | 6.89 | 7.07 | 7.24 | 7.42 | 7.59 | 7.77 | 7.94 | 8.12 | 8.29 | 8.47 | 8.64 | 8.82 | 8.99 | 9.17 | 9.34 | 9.52 | 9.69 | 9.87 | 10.04 | 10.22 | 10.39 | 10.57 | 10.74 | 10.92 | 11.09 | 11.27 | 11.44 | 11.62 | 11.79 | 11.96 | 12.14 | 12.31 | 12.49 | 12.66 | 12.84 | 13.01 | 13.19 | 13.36 | 13.54 | 13.71 | 13.89
3.78 | 6.89 | 7.07 | 7.24 | 7.42 | 7.59 | 7.77 | 7.94 | 8.12 | 8.29 | 8.47 | 8.64 | 8.82 | 8.99 | 9.17 | 9.34 | 9.51 | 9.69 | 9.86 | 10.04 | 10.21 | 10.39 | 10.56 | 10.74 | 10.91 | 11.09 | 11.26 | 11.44 | 11.61 | 11.79 | 11.96 | 12.14 | 12.31 | 12.49 | 12.66 | 12.84 | 13.01 | 13.18 | 13.36 | 13.53 | 13.71 | 13.88
3.79 | 6.89 | 7.07 | 7.24 | 7.42 | 7.59 | 7.76 | 7.94 | 8.11 | 8.29 | 8.46 | 8.64 | 8.81 | 8.99 | 9.16 | 9.34 | 9.51 | 9.69 | 9.86 | 10.04 | 10.21 | 10.39 | 10.56 | 10.74 | 10.91 | 11.08 | 11.26 | 11.43 | 11.61 | 11.78 | 11.96 | 12.13 | 12.31 | 12.48 | 12.66 | 12.83 | 13.01 | 13.18 | 13.36 | 13.53 | 13.71 | 13.88
3.8 | 6.89 | 7.06 | 7.24 | 7.41 | 7.59 | 7.76 | 7.94 | 8.11 | 8.29 | 8.46 | 8.64 | 8.81 | 8.99 | 9.16 | 9.34 | 9.51 | 9.68 | 9.86 | 10.03 | 10.21 | 10.38 | 10.56 | 10.73 | 10.91 | 11.08 | 11.26 | 11.43 | 11.61 | 11.78 | 11.96 | 12.13 | 12.3 | 12.48 | 12.65 | 12.83 | 13 | 13.18 | 13.35 | 13.53 | 13.7 | 13.88
3.81 | 6.89 | 7.06 | 7.24 | 7.41 | 7.59 | 7.76 | 7.94 | 8.11 | 8.28 | 8.46 | 8.63 | 8.81 | 8.98 | 9.16 | 9.33 | 9.51 | 9.68 | 9.86 | 10.03 | 10.21 | 10.38 | 10.56 | 10.73 | 10.9 | 11.08 | 11.25 | 11.43 | 11.6 | 11.78 | 11.95 | 12.13 | 12.3 | 12.48 | 12.65 | 12.83 | 13 | 13.17 | 13.35 | 13.52 | 13.7 | 13.87
3.82 | 6.89 | 7.06 | 7.23 | 7.41 | 7.58 | 7.76 | 7.93 | 8.11 | 8.28 | 8.46 | 8.63 | 8.81 | 8.98 | 9.16 | 9.33 | 9.5 | 9.68 | 9.85 | 10.03 | 10.2 | 10.38 | 10.55 | 10.73 | 10.9 | 11.08 | 11.25 | 11.43 | 11.6 | 11.77 | 11.95 | 12.12 | 12.3 | 12.47 | 12.65 | 12.82 | 13 | 13.17 | 13.35 | 13.52 | 13.69 | 13.87
3.83 | 6.88 | 7.06 | 7.23 | 7.41 | 7.58 | 7.76 | 7.93 | 8.11 | 8.28 | 8.45 | 8.63 | 8.8 | 8.98 | 9.15 | 9.33 | 9.5 | 9.68 | 9.85 | 10.03 | 10.2 | 10.38 | 10.55 | 10.72 | 10.9 | 11.07 | 11.25 | 11.42 | 11.6 | 11.77 | 11.95 | 12.12 | 12.29 | 12.47 | 12.64 | 12.82 | 12.99 | 13.17 | 13.34 | 13.52 | 13.69 | 13.87
3.84 | 6.88 | 7.06 | 7.23 | 7.41 | 7.58 | 7.75 | 7.93 | 8.1 | 8.28 | 8.45 | 8.63 | 8.8 | 8.98 | 9.15 | 9.33 | 9.5 | 9.67 | 9.85 | 10.02 | 10.2 | 10.37 | 10.55 | 10.72 | 10.9 | 11.07 | 11.24 | 11.42 | 11.59 | 11.77 | 11.94 | 12.12 | 12.29 | 12.47 | 12.64 | 12.81 | 12.99 | 13.16 | 13.34 | 13.51 | 13.69 | 13.86
3.85 | 6.88 | 7.05 | 7.23 | 7.4 | 7.58 | 7.75 | 7.93 | 8.1 | 8.28 | 8.45 | 8.62 | 8.8 | 8.97 | 9.15 | 9.32 | 9.5 | 9.67 | 9.85 | 10.02 | 10.2 | 10.37 | 10.54 | 10.72 | 10.89 | 11.07 | 11.24 | 11.42 | 11.59 | 11.77 | 11.94 | 12.11 | 12.29 | 12.46 | 12.64 | 12.81 | 12.99 | 13.16 | 13.33 | 13.51 | 13.68 | 13.86
3.86 | 6.88 | 7.05 | 7.23 | 7.4 | 7.58 | 7.75 | 7.92 | 8.1 | 8.27 | 8.45 | 8.62 | 8.8 | 8.97 | 9.15 | 9.32 | 9.49 | 9.67 | 9.84 | 10.02 | 10.19 | 10.37 | 10.54 | 10.72 | 10.89 | 11.06 | 11.24 | 11.41 | 11.59 | 11.76 | 11.94 | 12.11 | 12.29 | 12.46 | 12.63 | 12.81 | 12.98 | 13.16 | 13.33 | 13.51 | 13.68 | 13.85
3.87 | 6.88 | 7.05 | 7.22 | 7.4 | 7.57 | 7.75 | 7.92 | 8.1 | 8.27 | 8.45 | 8.62 | 8.79 | 8.97 | 9.14 | 9.32 | 9.49 | 9.67 | 9.84 | 10.02 | 10.19 | 10.36 | 10.54 | 10.71 | 10.89 | 11.06 | 11.24 | 11.41 | 11.58 | 11.76 | 11.93 | 12.11 | 12.28 | 12.46 | 12.63 | 12.8 | 12.98 | 13.15 | 13.33 | 13.5 | 13.68 | 13.85
3.88 | 6.87 | 7.05 | 7.22 | 7.4 | 7.57 | 7.75 | 7.92 | 8.09 | 8.27 | 8.44 | 8.62 | 8.79 | 8.97 | 9.14 | 9.32 | 9.49 | 9.66 | 9.84 | 10.01 | 10.19 | 10.36 | 10.54 | 10.71 | 10.88 | 11.06 | 11.23 | 11.41 | 11.58 | 11.76 | 11.93 | 12.1 | 12.28 | 12.45 | 12.63 | 12.8 | 12.98 | 13.15 | 13.32 | 13.5 | 13.67 | 13.85
3.89 | 6.87 | 7.05 | 7.22 | 7.4 | 7.57 | 7.74 | 7.92 | 8.09 | 8.27 | 8.44 | 8.62 | 8.79 | 8.96 | 9.14 | 9.31 | 9.49 | 9.66 | 9.84 | 10.01 | 10.18 | 10.36 | 10.53 | 10.71 | 10.88 | 11.06 | 11.23 | 11.4 | 11.58 | 11.75 | 11.93 | 12.1 | 12.28 | 12.45 | 12.62 | 12.8 | 12.97 | 13.15 | 13.32 | 13.49 | 13.67 | 13.84
3.9 | 6.87 | 7.04 | 7.22 | 7.39 | 7.57 | 7.74 | 7.92 | 8.09 | 8.26 | 8.44 | 8.61 | 8.79 | 8.96 | 9.14 | 9.31 | 9.48 | 9.66 | 9.83 | 10.01 | 10.18 | 10.36 | 10.53 | 10.7 | 10.88 | 11.05 | 11.23 | 11.4 | 11.58 | 11.75 | 11.92 | 12.1 | 12.27 | 12.45 | 12.62 | 12.79 | 12.97 | 13.14 | 13.32 | 13.49 | 13.67 | 13.84
3.91 | 6.87 | 7.04 | 7.22 | 7.39 | 7.57 | 7.74 | 7.91 | 8.09 | 8.26 | 8.44 | 8.61 | 8.79 | 8.96 | 9.13 | 9.31 | 9.48 | 9.66 | 9.83 | 10 | 10.18 | 10.35 | 10.53 | 10.7 | 10.88 | 11.05 | 11.22 | 11.4 | 11.57 | 11.75 | 11.92 | 12.09 | 12.27 | 12.44 | 12.62 | 12.79 | 12.97 | 13.14 | 13.31 | 13.49 | 13.66 | 13.84
3.92 | 6.87 | 7.04 | 7.22 | 7.39 | 7.56 | 7.74 | 7.91 | 8.09 | 8.26 | 8.43 | 8.61 | 8.78 | 8.96 | 9.13 | 9.31 | 9.48 | 9.65 | 9.83 | 10 | 10.18 | 10.35 | 10.52 | 10.7 | 10.87 | 11.05 | 11.22 | 11.4 | 11.57 | 11.74 | 11.92 | 12.09 | 12.27 | 12.44 | 12.61 | 12.79 | 12.96 | 13.14 | 13.31 | 13.48 | 13.66 | 13.83
3.93 | 6.86 | 7.04 | 7.21 | 7.39 | 7.56 | 7.74 | 7.91 | 8.08 | 8.26 | 8.43 | 8.61 | 8.78 | 8.95 | 9.13 | 9.3 | 9.48 | 9.65 | 9.83 | 10 | 10.17 | 10.35 | 10.52 | 10.7 | 10.87 | 11.04 | 11.22 | 11.39 | 11.57 | 11.74 | 11.91 | 12.09 | 12.26 | 12.44 | 12.61 | 12.78 | 12.96 | 13.13 | 13.31 | 13.48 | 13.65 | 13.83
3.94 | 6.86 | 7.04 | 7.21 | 7.39 | 7.56 | 7.73 | 7.91 | 8.08 | 8.26 | 8.43 | 8.6 | 8.78 | 8.95 | 9.13 | 9.3 | 9.47 | 9.65 | 9.82 | 10 | 10.17 | 10.34 | 10.52 | 10.69 | 10.87 | 11.04 | 11.22 | 11.39 | 11.56 | 11.74 | 11.91 | 12.09 | 12.26 | 12.43 | 12.61 | 12.78 | 12.96 | 13.13 | 13.3 | 13.48 | 13.65 | 13.82
3.95 | 6.86 | 7.04 | 7.21 | 7.38 | 7.56 | 7.73 | 7.91 | 8.08 | 8.25 | 8.43 | 8.6 | 8.78 | 8.95 | 9.12 | 9.3 | 9.47 | 9.65 | 9.82 | 9.99 | 10.17 | 10.34 | 10.52 | 10.69 | 10.86 | 11.04 | 11.21 | 11.39 | 11.56 | 11.73 | 11.91 | 12.08 | 12.26 | 12.43 | 12.6 | 12.78 | 12.95 | 13.13 | 13.3 | 13.47 | 13.65 | 13.82
3.96 | 6.86 | 7.03 | 7.21 | 7.38 | 7.56 | 7.73 | 7.9 | 8.08 | 8.25 | 8.43 | 8.6 | 8.77 | 8.95 | 9.12 | 9.3 | 9.47 | 9.64 | 9.82 | 9.99 | 10.17 | 10.34 | 10.51 | 10.69 | 10.86 | 11.04 | 11.21 | 11.38 | 11.56 | 11.73 | 11.9 | 12.08 | 12.25 | 12.43 | 12.6 | 12.77 | 12.95 | 13.12 | 13.3 | 13.47 | 13.64 | 13.82
3.97 | 6.86 | 7.03 | 7.21 | 7.38 | 7.55 | 7.73 | 7.9 | 8.08 | 8.25 | 8.42 | 8.6 | 8.77 | 8.95 | 9.12 | 9.29 | 9.47 | 9.64 | 9.81 | 9.99 | 10.16 | 10.34 | 10.51 | 10.68 | 10.86 | 11.03 | 11.21 | 11.38 | 11.55 | 11.73 | 11.9 | 12.08 | 12.25 | 12.42 | 12.6 | 12.77 | 12.94 | 13.12 | 13.29 | 13.47 | 13.64 | 13.81
3.98 | 6.86 | 7.03 | 7.2 | 7.38 | 7.55 | 7.73 | 7.9 | 8.07 | 8.25 | 8.42 | 8.59 | 8.77 | 8.94 | 9.12 | 9.29 | 9.46 | 9.64 | 9.81 | 9.99 | 10.16 | 10.33 | 10.51 | 10.68 | 10.86 | 11.03 | 11.2 | 11.38 | 11.55 | 11.72 | 11.9 | 12.07 | 12.25 | 12.42 | 12.59 | 12.77 | 12.94 | 13.12 | 13.29 | 13.46 | 13.64 | 13.81
3.99 | 6.85 | 7.03 | 7.2 | 7.38 | 7.55 | 7.72 | 7.9 | 8.07 | 8.24 | 8.42 | 8.59 | 8.77 | 8.94 | 9.11 | 9.29 | 9.46 | 9.64 | 9.81 | 9.98 | 10.16 | 10.33 | 10.5 | 10.68 | 10.85 | 11.03 | 11.2 | 11.37 | 11.55 | 11.72 | 11.9 | 12.07 | 12.24 | 12.42 | 12.59 | 12.76 | 12.94 | 13.11 | 13.29 | 13.46 | 13.63 | 13.81
4 | 6.85 | 7.03 | 7.2 | 7.37 | 7.55 | 7.72 | 7.9 | 8.07 | 8.24 | 8.42 | 8.59 | 8.76 | 8.94 | 9.11 | 9.29 | 9.46 | 9.63 | 9.81 | 9.98 | 10.15 | 10.33 | 10.5 | 10.68 | 10.85 | 11.02 | 11.2 | 11.37 | 11.54 | 11.72 | 11.89 | 12.07 | 12.24 | 12.41 | 12.59 | 12.76 | 12.93 | 13.11 | 13.28 | 13.46 | 13.63 | 13.8
4.01 | 6.85 | 7.02 | 7.2 | 7.37 | 7.55 | 7.72 | 7.89 | 8.07 | 8.24 | 8.41 | 8.59 | 8.76 | 8.94 | 9.11 | 9.28 | 9.46 | 9.63 | 9.8 | 9.98 | 10.15 | 10.33 | 10.5 | 10.67 | 10.85 | 11.02 | 11.19 | 11.37 | 11.54 | 11.72 | 11.89 | 12.06 | 12.24 | 12.41 | 12.58 | 12.76 | 12.93 | 13.1 | 13.28 | 13.45 | 13.63 | 13.8
4.02 | 6.85 | 7.02 | 7.2 | 7.37 | 7.54 | 7.72 | 7.89 | 8.06 | 8.24 | 8.41 | 8.59 | 8.76 | 8.93 | 9.11 | 9.28 | 9.45 | 9.63 | 9.8 | 9.98 | 10.15 | 10.32 | 10.5 | 10.67 | 10.84 | 11.02 | 11.19 | 11.37 | 11.54 | 11.71 | 11.89 | 12.06 | 12.23 | 12.41 | 12.58 | 12.75 | 12.93 | 13.1 | 13.28 | 13.45 | 13.62 | 13.8
4.03 | 6.85 | 7.02 | 7.19 | 7.37 | 7.54 | 7.71 | 7.89 | 8.06 | 8.24 | 8.41 | 8.58 | 8.76 | 8.93 | 9.1 | 9.28 | 9.45 | 9.63 | 9.8 | 9.97 | 10.15 | 10.32 | 10.49 | 10.67 | 10.84 | 11.01 | 11.19 | 11.36 | 11.54 | 11.71 | 11.88 | 12.06 | 12.23 | 12.4 | 12.58 | 12.75 | 12.92 | 13.1 | 13.27 | 13.45 | 13.62 | 13.79
4.04 | 6.84 | 7.02 | 7.19 | 7.37 | 7.54 | 7.71 | 7.89 | 8.06 | 8.23 | 8.41 | 8.58 | 8.75 | 8.93 | 9.1 | 9.28 | 9.45 | 9.62 | 9.8 | 9.97 | 10.14 | 10.32 | 10.49 | 10.66 | 10.84 | 11.01 | 11.19 | 11.36 | 11.53 | 11.71 | 11.88 | 12.05 | 12.23 | 12.4 | 12.57 | 12.75 | 12.92 | 13.09 | 13.27 | 13.44 | 13.62 | 13.79
4.05 | 6.84 | 7.02 | 7.19 | 7.36 | 7.54 | 7.71 | 7.88 | 8.06 | 8.23 | 8.41 | 8.58 | 8.75 | 8.93 | 9.1 | 9.27 | 9.45 | 9.62 | 9.79 | 9.97 | 10.14 | 10.31 | 10.49 | 10.66 | 10.84 | 11.01 | 11.18 | 11.36 | 11.53 | 11.7 | 11.88 | 12.05 | 12.22 | 12.4 | 12.57 | 12.74 | 12.92 | 13.09 | 13.26 | 13.44 | 13.61 | 13.79
4.06 | 6.84 | 7.01 | 7.19 | 7.36 | 7.54 | 7.71 | 7.88 | 8.06 | 8.23 | 8.4 | 8.58 | 8.75 | 8.92 | 9.1 | 9.27 | 9.44 | 9.62 | 9.79 | 9.96 | 10.14 | 10.31 | 10.49 | 10.66 | 10.83 | 11.01 | 11.18 | 11.35 | 11.53 | 11.7 | 11.87 | 12.05 | 12.22 | 12.39 | 12.57 | 12.74 | 12.91 | 13.09 | 13.26 | 13.43 | 13.61 | 13.78
4.07 | 6.84 | 7.01 | 7.19 | 7.36 | 7.53 | 7.71 | 7.88 | 8.05 | 8.23 | 8.4 | 8.57 | 8.75 | 8.92 | 9.09 | 9.27 | 9.44 | 9.62 | 9.79 | 9.96 | 10.14 | 10.31 | 10.48 | 10.66 | 10.83 | 11 | 11.18 | 11.35 | 11.52 | 11.7 | 11.87 | 12.04 | 12.22 | 12.39 | 12.56 | 12.74 | 12.91 | 13.08 | 13.26 | 13.43 | 13.6 | 13.78
4.08 | 6.84 | 7.01 | 7.18 | 7.36 | 7.53 | 7.7 | 7.88 | 8.05 | 8.23 | 8.4 | 8.57 | 8.75 | 8.92 | 9.09 | 9.27 | 9.44 | 9.61 | 9.79 | 9.96 | 10.13 | 10.31 | 10.48 | 10.65 | 10.83 | 11 | 11.17 | 11.35 | 11.52 | 11.69 | 11.87 | 12.04 | 12.21 | 12.39 | 12.56 | 12.73 | 12.91 | 13.08 | 13.25 | 13.43 | 13.6 | 13.77
4.09 | 6.84 | 7.01 | 7.18 | 7.36 | 7.53 | 7.7 | 7.88 | 8.05 | 8.22 | 8.4 | 8.57 | 8.74 | 8.92 | 9.09 | 9.26 | 9.44 | 9.61 | 9.78 | 9.96 | 10.13 | 10.3 | 10.48 | 10.65 | 10.82 | 11 | 11.17 | 11.34 | 11.52 | 11.69 | 11.86 | 12.04 | 12.21 | 12.38 | 12.56 | 12.73 | 12.9 | 13.08 | 13.25 | 13.42 | 13.6 | 13.77
4.1 | 6.83 | 7.01 | 7.18 | 7.35 | 7.53 | 7.7 | 7.87 | 8.05 | 8.22 | 8.39 | 8.57 | 8.74 | 8.91 | 9.09 | 9.26 | 9.43 | 9.61 | 9.78 | 9.95 | 10.13 | 10.3 | 10.47 | 10.65 | 10.82 | 10.99 | 11.17 | 11.34 | 11.51 | 11.69 | 11.86 | 12.03 | 12.21 | 12.38 | 12.55 | 12.73 | 12.9 | 13.07 | 13.25 | 13.42 | 13.59 | 13.77
4.11 | 6.83 | 7.01 | 7.18 | 7.35 | 7.53 | 7.7 | 7.87 | 8.05 | 8.22 | 8.39 | 8.57 | 8.74 | 8.91 | 9.09 | 9.26 | 9.43 | 9.61 | 9.78 | 9.95 | 10.13 | 10.3 | 10.47 | 10.64 | 10.82 | 10.99 | 11.16 | 11.34 | 11.51 | 11.68 | 11.86 | 12.03 | 12.2 | 12.38 | 12.55 | 12.72 | 12.9 | 13.07 | 13.24 | 13.42 | 13.59 | 13.76
4.12 | 6.83 | 7 | 7.18 | 7.35 | 7.52 | 7.7 | 7.87 | 8.04 | 8.22 | 8.39 | 8.56 | 8.74 | 8.91 | 9.08 | 9.26 | 9.43 | 9.6 | 9.78 | 9.95 | 10.12 | 10.3 | 10.47 | 10.64 | 10.82 | 10.99 | 11.16 | 11.34 | 11.51 | 11.68 | 11.85 | 12.03 | 12.2 | 12.37 | 12.55 | 12.72 | 12.89 | 13.07 | 13.24 | 13.41 | 13.59 | 13.76
4.13 | 6.83 | 7 | 7.17 | 7.35 | 7.52 | 7.69 | 7.87 | 8.04 | 8.21 | 8.39 | 8.56 | 8.73 | 8.91 | 9.08 | 9.25 | 9.43 | 9.6 | 9.77 | 9.95 | 10.12 | 10.29 | 10.47 | 10.64 | 10.81 | 10.99 | 11.16 | 11.33 | 11.51 | 11.68 | 11.85 | 12.02 | 12.2 | 12.37 | 12.54 | 12.72 | 12.89 | 13.06 | 13.24 | 13.41 | 13.58 | 13.76
4.14 | 6.83 | 7 | 7.17 | 7.35 | 7.52 | 7.69 | 7.87 | 8.04 | 8.21 | 8.39 | 8.56 | 8.73 | 8.9 | 9.08 | 9.25 | 9.42 | 9.6 | 9.77 | 9.94 | 10.12 | 10.29 | 10.46 | 10.64 | 10.81 | 10.98 | 11.16 | 11.33 | 11.5 | 11.68 | 11.85 | 12.02 | 12.19 | 12.37 | 12.54 | 12.71 | 12.89 | 13.06 | 13.23 | 13.41 | 13.58 | 13.75
4.15 | 6.82 | 7 | 7.17 | 7.34 | 7.52 | 7.69 | 7.86 | 8.04 | 8.21 | 8.38 | 8.56 | 8.73 | 8.9 | 9.08 | 9.25 | 9.42 | 9.6 | 9.77 | 9.94 | 10.11 | 10.29 | 10.46 | 10.63 | 10.81 | 10.98 | 11.15 | 11.33 | 11.5 | 11.67 | 11.85 | 12.02 | 12.19 | 12.36 | 12.54 | 12.71 | 12.88 | 13.06 | 13.23 | 13.4 | 13.58 | 13.75
4.16 | 6.82 | 7 | 7.17 | 7.34 | 7.52 | 7.69 | 7.86 | 8.03 | 8.21 | 8.38 | 8.55 | 8.73 | 8.9 | 9.07 | 9.25 | 9.42 | 9.59 | 9.77 | 9.94 | 10.11 | 10.28 | 10.46 | 10.63 | 10.8 | 10.98 | 11.15 | 11.32 | 11.5 | 11.67 | 11.84 | 12.02 | 12.19 | 12.36 | 12.53 | 12.71 | 12.88 | 13.05 | 13.23 | 13.4 | 13.57 | 13.75
4.17 | 6.82 | 6.99 | 7.17 | 7.34 | 7.51 | 7.69 | 7.86 | 8.03 | 8.21 | 8.38 | 8.55 | 8.72 | 8.9 | 9.07 | 9.24 | 9.42 | 9.59 | 9.76 | 9.94 | 10.11 | 10.28 | 10.46 | 10.63 | 10.8 | 10.97 | 11.15 | 11.32 | 11.49 | 11.67 | 11.84 | 12.01 | 12.19 | 12.36 | 12.53 | 12.7 | 12.88 | 13.05 | 13.22 | 13.4 | 13.57 | 13.74
4.18 | 6.82 | 6.99 | 7.17 | 7.34 | 7.51 | 7.68 | 7.86 | 8.03 | 8.2 | 8.38 | 8.55 | 8.72 | 8.9 | 9.07 | 9.24 | 9.41 | 9.59 | 9.76 | 9.93 | 10.11 | 10.28 | 10.45 | 10.63 | 10.8 | 10.97 | 11.14 | 11.32 | 11.49 | 11.66 | 11.84 | 12.01 | 12.18 | 12.35 | 12.53 | 12.7 | 12.87 | 13.05 | 13.22 | 13.39 | 13.57 | 13.74
4.19 | 6.82 | 6.99 | 7.16 | 7.34 | 7.51 | 7.68 | 7.86 | 8.03 | 8.2 | 8.37 | 8.55 | 8.72 | 8.89 | 9.07 | 9.24 | 9.41 | 9.58 | 9.76 | 9.93 | 10.1 | 10.28 | 10.45 | 10.62 | 10.8 | 10.97 | 11.14 | 11.31 | 11.49 | 11.66 | 11.83 | 12.01 | 12.18 | 12.35 | 12.52 | 12.7 | 12.87 | 13.04 | 13.22 | 13.39 | 13.56 | 13.73
4.2 | 6.82 | 6.99 | 7.16 | 7.33 | 7.51 | 7.68 | 7.85 | 8.03 | 8.2 | 8.37 | 8.54 | 8.72 | 8.89 | 9.06 | 9.24 | 9.41 | 9.58 | 9.76 | 9.93 | 10.1 | 10.27 | 10.45 | 10.62 | 10.79 | 10.97 | 11.14 | 11.31 | 11.48 | 11.66 | 11.83 | 12 | 12.18 | 12.35 | 12.52 | 12.69 | 12.87 | 13.04 | 13.21 | 13.39 | 13.56 | 13.73
4.21 | 6.81 | 6.99 | 7.16 | 7.33 | 7.51 | 7.68 | 7.85 | 8.02 | 8.2 | 8.37 | 8.54 | 8.72 | 8.89 | 9.06 | 9.23 | 9.41 | 9.58 | 9.75 | 9.93 | 10.1 | 10.27 | 10.44 | 10.62 | 10.79 | 10.96 | 11.14 | 11.31 | 11.48 | 11.65 | 11.83 | 12 | 12.17 | 12.35 | 12.52 | 12.69 | 12.86 | 13.04 | 13.21 | 13.38 | 13.55 | 13.73
4.22 | 6.81 | 6.98 | 7.16 | 7.33 | 7.5 | 7.68 | 7.85 | 8.02 | 8.19 | 8.37 | 8.54 | 8.71 | 8.89 | 9.06 | 9.23 | 9.4 | 9.58 | 9.75 | 9.92 | 10.1 | 10.27 | 10.44 | 10.61 | 10.79 | 10.96 | 11.13 | 11.31 | 11.48 | 11.65 | 11.82 | 12 | 12.17 | 12.34 | 12.51 | 12.69 | 12.86 | 13.03 | 13.21 | 13.38 | 13.55 | 13.72
4.23 | 6.81 | 6.98 | 7.16 | 7.33 | 7.5 | 7.67 | 7.85 | 8.02 | 8.19 | 8.37 | 8.54 | 8.71 | 8.88 | 9.06 | 9.23 | 9.4 | 9.57 | 9.75 | 9.92 | 10.09 | 10.27 | 10.44 | 10.61 | 10.78 | 10.96 | 11.13 | 11.3 | 11.48 | 11.65 | 11.82 | 11.99 | 12.17 | 12.34 | 12.51 | 12.68 | 12.86 | 13.03 | 13.2 | 13.37 | 13.55 | 13.72
4.24 | 6.81 | 6.98 | 7.15 | 7.33 | 7.5 | 7.67 | 7.84 | 8.02 | 8.19 | 8.36 | 8.54 | 8.71 | 8.88 | 9.05 | 9.23 | 9.4 | 9.57 | 9.74 | 9.92 | 10.09 | 10.26 | 10.44 | 10.61 | 10.78 | 10.95 | 11.13 | 11.3 | 11.47 | 11.64 | 11.82 | 11.99 | 12.16 | 12.34 | 12.51 | 12.68 | 12.85 | 13.03 | 13.2 | 13.37 | 13.54 | 13.72
4.25 | 6.81 | 6.98 | 7.15 | 7.32 | 7.5 | 7.67 | 7.84 | 8.02 | 8.19 | 8.36 | 8.53 | 8.71 | 8.88 | 9.05 | 9.22 | 9.4 | 9.57 | 9.74 | 9.92 | 10.09 | 10.26 | 10.43 | 10.61 | 10.78 | 10.95 | 11.12 | 11.3 | 11.47 | 11.64 | 11.81 | 11.99 | 12.16 | 12.33 | 12.5 | 12.68 | 12.85 | 13.02 | 13.2 | 13.37 | 13.54 | 13.71
4.26 | 6.8 | 6.98 | 7.15 | 7.32 | 7.5 | 7.67 | 7.84 | 8.01 | 8.19 | 8.36 | 8.53 | 8.7 | 8.88 | 9.05 | 9.22 | 9.39 | 9.57 | 9.74 | 9.91 | 10.09 | 10.26 | 10.43 | 10.6 | 10.78 | 10.95 | 11.12 | 11.29 | 11.47 | 11.64 | 11.81 | 11.98 | 12.16 | 12.33 | 12.5 | 12.67 | 12.85 | 13.02 | 13.19 | 13.36 | 13.54 | 13.71
4.27 | 6.8 | 6.98 | 7.15 | 7.32 | 7.49 | 7.67 | 7.84 | 8.01 | 8.18 | 8.36 | 8.53 | 8.7 | 8.87 | 9.05 | 9.22 | 9.39 | 9.56 | 9.74 | 9.91 | 10.08 | 10.26 | 10.43 | 10.6 | 10.77 | 10.95 | 11.12 | 11.29 | 11.46 | 11.64 | 11.81 | 11.98 | 12.15 | 12.33 | 12.5 | 12.67 | 12.84 | 13.02 | 13.19 | 13.36 | 13.53 | 13.71
4.28 | 6.8 | 6.97 | 7.15 | 7.32 | 7.49 | 7.66 | 7.84 | 8.01 | 8.18 | 8.35 | 8.53 | 8.7 | 8.87 | 9.04 | 9.22 | 9.39 | 9.56 | 9.73 | 9.91 | 10.08 | 10.25 | 10.42 | 10.6 | 10.77 | 10.94 | 11.11 | 11.29 | 11.46 | 11.63 | 11.8 | 11.98 | 12.15 | 12.32 | 12.49 | 12.67 | 12.84 | 13.01 | 13.18 | 13.36 | 13.53 | 13.7
4.29 | 6.8 | 6.97 | 7.14 | 7.32 | 7.49 | 7.66 | 7.83 | 8.01 | 8.18 | 8.35 | 8.52 | 8.7 | 8.87 | 9.04 | 9.21 | 9.39 | 9.56 | 9.73 | 9.9 | 10.08 | 10.25 | 10.42 | 10.59 | 10.77 | 10.94 | 11.11 | 11.28 | 11.46 | 11.63 | 11.8 | 11.97 | 12.15 | 12.32 | 12.49 | 12.66 | 12.84 | 13.01 | 13.18 | 13.35 | 13.53 | 13.7
4.3 | 6.8 | 6.97 | 7.14 | 7.31 | 7.49 | 7.66 | 7.83 | 8 | 8.18 | 8.35 | 8.52 | 8.69 | 8.87 | 9.04 | 9.21 | 9.38 | 9.56 | 9.73 | 9.9 | 10.07 | 10.25 | 10.42 | 10.59 | 10.76 | 10.94 | 11.11 | 11.28 | 11.45 | 11.63 | 11.8 | 11.97 | 12.14 | 12.32 | 12.49 | 12.66 | 12.83 | 13.01 | 13.18 | 13.35 | 13.52 | 13.7
4.31 | 6.8 | 6.97 | 7.14 | 7.31 | 7.49 | 7.66 | 7.83 | 8 | 8.18 | 8.35 | 8.52 | 8.69 | 8.86 | 9.04 | 9.21 | 9.38 | 9.55 | 9.73 | 9.9 | 10.07 | 10.24 | 10.42 | 10.59 | 10.76 | 10.93 | 11.11 | 11.28 | 11.45 | 11.62 | 11.8 | 11.97 | 12.14 | 12.31 | 12.49 | 12.66 | 12.83 | 13 | 13.17 | 13.35 | 13.52 | 13.69
4.32 | 6.79 | 6.97 | 7.14 | 7.31 | 7.48 | 7.66 | 7.83 | 8 | 8.17 | 8.35 | 8.52 | 8.69 | 8.86 | 9.03 | 9.21 | 9.38 | 9.55 | 9.72 | 9.9 | 10.07 | 10.24 | 10.41 | 10.59 | 10.76 | 10.93 | 11.1 | 11.28 | 11.45 | 11.62 | 11.79 | 11.96 | 12.14 | 12.31 | 12.48 | 12.65 | 12.83 | 13 | 13.17 | 13.34 | 13.52 | 13.69
4.33 | 6.79 | 6.96 | 7.14 | 7.31 | 7.48 | 7.65 | 7.83 | 8 | 8.17 | 8.34 | 8.52 | 8.69 | 8.86 | 9.03 | 9.2 | 9.38 | 9.55 | 9.72 | 9.89 | 10.07 | 10.24 | 10.41 | 10.58 | 10.76 | 10.93 | 11.1 | 11.27 | 11.44 | 11.62 | 11.79 | 11.96 | 12.13 | 12.31 | 12.48 | 12.65 | 12.82 | 13 | 13.17 | 13.34 | 13.51 | 13.68
4.34 | 6.79 | 6.96 | 7.13 | 7.31 | 7.48 | 7.65 | 7.82 | 8 | 8.17 | 8.34 | 8.51 | 8.69 | 8.86 | 9.03 | 9.2 | 9.37 | 9.55 | 9.72 | 9.89 | 10.06 | 10.24 | 10.41 | 10.58 | 10.75 | 10.93 | 11.1 | 11.27 | 11.44 | 11.61 | 11.79 | 11.96 | 12.13 | 12.3 | 12.48 | 12.65 | 12.82 | 12.99 | 13.16 | 13.34 | 13.51 | 13.68
4.35 | 6.79 | 6.96 | 7.13 | 7.3 | 7.48 | 7.65 | 7.82 | 7.99 | 8.17 | 8.34 | 8.51 | 8.68 | 8.86 | 9.03 | 9.2 | 9.37 | 9.54 | 9.72 | 9.89 | 10.06 | 10.23 | 10.41 | 10.58 | 10.75 | 10.92 | 11.09 | 11.27 | 11.44 | 11.61 | 11.78 | 11.96 | 12.13 | 12.3 | 12.47 | 12.64 | 12.82 | 12.99 | 13.16 | 13.33 | 13.51 | 13.68
4.36 | 6.79 | 6.96 | 7.13 | 7.3 | 7.48 | 7.65 | 7.82 | 7.99 | 8.16 | 8.34 | 8.51 | 8.68 | 8.85 | 9.03 | 9.2 | 9.37 | 9.54 | 9.71 | 9.89 | 10.06 | 10.23 | 10.4 | 10.58 | 10.75 | 10.92 | 11.09 | 11.26 | 11.44 | 11.61 | 11.78 | 11.95 | 12.12 | 12.3 | 12.47 | 12.64 | 12.81 | 12.99 | 13.16 | 13.33 | 13.5 | 13.67
4.37 | 6.78 | 6.96 | 7.13 | 7.3 | 7.47 | 7.65 | 7.82 | 7.99 | 8.16 | 8.33 | 8.51 | 8.68 | 8.85 | 9.02 | 9.2 | 9.37 | 9.54 | 9.71 | 9.88 | 10.06 | 10.23 | 10.4 | 10.57 | 10.74 | 10.92 | 11.09 | 11.26 | 11.43 | 11.61 | 11.78 | 11.95 | 12.12 | 12.29 | 12.47 | 12.64 | 12.81 | 12.98 | 13.15 | 13.33 | 13.5 | 13.67
4.38 | 6.78 | 6.95 | 7.13 | 7.3 | 7.47 | 7.64 | 7.82 | 7.99 | 8.16 | 8.33 | 8.5 | 8.68 | 8.85 | 9.02 | 9.19 | 9.36 | 9.54 | 9.71 | 9.88 | 10.05 | 10.23 | 10.4 | 10.57 | 10.74 | 10.91 | 11.09 | 11.26 | 11.43 | 11.6 | 11.77 | 11.95 | 12.12 | 12.29 | 12.46 | 12.63 | 12.81 | 12.98 | 13.15 | 13.32 | 13.49 | 13.67
4.39 | 6.78 | 6.95 | 7.13 | 7.3 | 7.47 | 7.64 | 7.81 | 7.99 | 8.16 | 8.33 | 8.5 | 8.67 | 8.85 | 9.02 | 9.19 | 9.36 | 9.53 | 9.71 | 9.88 | 10.05 | 10.22 | 10.39 | 10.57 | 10.74 | 10.91 | 11.08 | 11.25 | 11.43 | 11.6 | 11.77 | 11.94 | 12.12 | 12.29 | 12.46 | 12.63 | 12.8 | 12.98 | 13.15 | 13.32 | 13.49 | 13.66
4.4 | 6.78 | 6.95 | 7.12 | 7.3 | 7.47 | 7.64 | 7.81 | 7.98 | 8.16 | 8.33 | 8.5 | 8.67 | 8.84 | 9.02 | 9.19 | 9.36 | 9.53 | 9.7 | 9.88 | 10.05 | 10.22 | 10.39 | 10.56 | 10.74 | 10.91 | 11.08 | 11.25 | 11.42 | 11.6 | 11.77 | 11.94 | 12.11 | 12.28 | 12.46 | 12.63 | 12.8 | 12.97 | 13.14 | 13.32 | 13.49 | 13.66
4.41 | 6.78 | 6.95 | 7.12 | 7.29 | 7.47 | 7.64 | 7.81 | 7.98 | 8.15 | 8.33 | 8.5 | 8.67 | 8.84 | 9.01 | 9.19 | 9.36 | 9.53 | 9.7 | 9.87 | 10.05 | 10.22 | 10.39 | 10.56 | 10.73 | 10.91 | 11.08 | 11.25 | 11.42 | 11.59 | 11.76 | 11.94 | 12.11 | 12.28 | 12.45 | 12.62 | 12.8 | 12.97 | 13.14 | 13.31 | 13.48 | 13.66
4.42 | 6.78 | 6.95 | 7.12 | 7.29 | 7.46 | 7.64 | 7.81 | 7.98 | 8.15 | 8.32 | 8.5 | 8.67 | 8.84 | 9.01 | 9.18 | 9.35 | 9.53 | 9.7 | 9.87 | 10.04 | 10.21 | 10.39 | 10.56 | 10.73 | 10.9 | 11.07 | 11.25 | 11.42 | 11.59 | 11.76 | 11.93 | 12.11 | 12.28 | 12.45 | 12.62 | 12.79 | 12.96 | 13.14 | 13.31 | 13.48 | 13.65
4.43 | 6.77 | 6.95 | 7.12 | 7.29 | 7.46 | 7.63 | 7.81 | 7.98 | 8.15 | 8.32 | 8.49 | 8.66 | 8.84 | 9.01 | 9.18 | 9.35 | 9.52 | 9.7 | 9.87 | 10.04 | 10.21 | 10.38 | 10.56 | 10.73 | 10.9 | 11.07 | 11.24 | 11.41 | 11.59 | 11.76 | 11.93 | 12.1 | 12.27 | 12.45 | 12.62 | 12.79 | 12.96 | 13.13 | 13.31 | 13.48 | 13.65
4.44 | 6.77 | 6.94 | 7.12 | 7.29 | 7.46 | 7.63 | 7.8 | 7.98 | 8.15 | 8.32 | 8.49 | 8.66 | 8.83 | 9.01 | 9.18 | 9.35 | 9.52 | 9.69 | 9.87 | 10.04 | 10.21 | 10.38 | 10.55 | 10.72 | 10.9 | 11.07 | 11.24 | 11.41 | 11.58 | 11.76 | 11.93 | 12.1 | 12.27 | 12.44 | 12.61 | 12.79 | 12.96 | 13.13 | 13.3 | 13.47 | 13.65
4.45 | 6.77 | 6.94 | 7.11 | 7.29 | 7.46 | 7.63 | 7.8 | 7.97 | 8.14 | 8.32 | 8.49 | 8.66 | 8.83 | 9 | 9.18 | 9.35 | 9.52 | 9.69 | 9.86 | 10.03 | 10.21 | 10.38 | 10.55 | 10.72 | 10.89 | 11.07 | 11.24 | 11.41 | 11.58 | 11.75 | 11.92 | 12.1 | 12.27 | 12.44 | 12.61 | 12.78 | 12.95 | 13.13 | 13.3 | 13.47 | 13.64
4.46 | 6.77 | 6.94 | 7.11 | 7.28 | 7.46 | 7.63 | 7.8 | 7.97 | 8.14 | 8.31 | 8.49 | 8.66 | 8.83 | 9 | 9.17 | 9.35 | 9.52 | 9.69 | 9.86 | 10.03 | 10.2 | 10.38 | 10.55 | 10.72 | 10.89 | 11.06 | 11.23 | 11.41 | 11.58 | 11.75 | 11.92 | 12.09 | 12.26 | 12.44 | 12.61 | 12.78 | 12.95 | 13.12 | 13.29 | 13.47 | 13.64
4.47 | 6.77 | 6.94 | 7.11 | 7.28 | 7.45 | 7.63 | 7.8 | 7.97 | 8.14 | 8.31 | 8.48 | 8.66 | 8.83 | 9 | 9.17 | 9.34 | 9.51 | 9.69 | 9.86 | 10.03 | 10.2 | 10.37 | 10.54 | 10.72 | 10.89 | 11.06 | 11.23 | 11.4 | 11.57 | 11.75 | 11.92 | 12.09 | 12.26 | 12.43 | 12.6 | 12.78 | 12.95 | 13.12 | 13.29 | 13.46 | 13.63
4.48 | 6.76 | 6.94 | 7.11 | 7.28 | 7.45 | 7.62 | 7.79 | 7.97 | 8.14 | 8.31 | 8.48 | 8.65 | 8.83 | 9 | 9.17 | 9.34 | 9.51 | 9.68 | 9.86 | 10.03 | 10.2 | 10.37 | 10.54 | 10.71 | 10.89 | 11.06 | 11.23 | 11.4 | 11.57 | 11.74 | 11.91 | 12.09 | 12.26 | 12.43 | 12.6 | 12.77 | 12.94 | 13.12 | 13.29 | 13.46 | 13.63
4.49 | 6.76 | 6.93 | 7.11 | 7.28 | 7.45 | 7.62 | 7.79 | 7.96 | 8.14 | 8.31 | 8.48 | 8.65 | 8.82 | 8.99 | 9.17 | 9.34 | 9.51 | 9.68 | 9.85 | 10.02 | 10.2 | 10.37 | 10.54 | 10.71 | 10.88 | 11.05 | 11.23 | 11.4 | 11.57 | 11.74 | 11.91 | 12.08 | 12.25 | 12.43 | 12.6 | 12.77 | 12.94 | 13.11 | 13.28 | 13.46 | 13.63
4.5 | 6.76 | 6.93 | 7.1 | 7.28 | 7.45 | 7.62 | 7.79 | 7.96 | 8.13 | 8.31 | 8.48 | 8.65 | 8.82 | 8.99 | 9.16 | 9.34 | 9.51 | 9.68 | 9.85 | 10.02 | 10.19 | 10.36 | 10.54 | 10.71 | 10.88 | 11.05 | 11.22 | 11.39 | 11.57 | 11.74 | 11.91 | 12.08 | 12.25 | 12.42 | 12.59 | 12.77 | 12.94 | 13.11 | 13.28 | 13.45 | 13.62
4.51 | 6.76 | 6.93 | 7.1 | 7.27 | 7.45 | 7.62 | 7.79 | 7.96 | 8.13 | 8.3 | 8.48 | 8.65 | 8.82 | 8.99 | 9.16 | 9.33 | 9.5 | 9.68 | 9.85 | 10.02 | 10.19 | 10.36 | 10.53 | 10.71 | 10.88 | 11.05 | 11.22 | 11.39 | 11.56 | 11.73 | 11.91 | 12.08 | 12.25 | 12.42 | 12.59 | 12.76 | 12.93 | 13.11 | 13.28 | 13.45 | 13.62
4.52 | 6.76 | 6.93 | 7.1 | 7.27 | 7.44 | 7.62 | 7.79 | 7.96 | 8.13 | 8.3 | 8.47 | 8.64 | 8.82 | 8.99 | 9.16 | 9.33 | 9.5 | 9.67 | 9.84 | 10.02 | 10.19 | 10.36 | 10.53 | 10.7 | 10.87 | 11.05 | 11.22 | 11.39 | 11.56 | 11.73 | 11.9 | 12.07 | 12.25 | 12.42 | 12.59 | 12.76 | 12.93 | 13.1 | 13.27 | 13.45 | 13.62
4.53 | 6.76 | 6.93 | 7.1 | 7.27 | 7.44 | 7.61 | 7.78 | 7.96 | 8.13 | 8.3 | 8.47 | 8.64 | 8.81 | 8.98 | 9.16 | 9.33 | 9.5 | 9.67 | 9.84 | 10.01 | 10.19 | 10.36 | 10.53 | 10.7 | 10.87 | 11.04 | 11.21 | 11.39 | 11.56 | 11.73 | 11.9 | 12.07 | 12.24 | 12.41 | 12.58 | 12.76 | 12.93 | 13.1 | 13.27 | 13.44 | 13.61
4.54 | 6.75 | 6.93 | 7.1 | 7.27 | 7.44 | 7.61 | 7.78 | 7.95 | 8.13 | 8.3 | 8.47 | 8.64 | 8.81 | 8.98 | 9.15 | 9.33 | 9.5 | 9.67 | 9.84 | 10.01 | 10.18 | 10.35 | 10.53 | 10.7 | 10.87 | 11.04 | 11.21 | 11.38 | 11.55 | 11.72 | 11.9 | 12.07 | 12.24 | 12.41 | 12.58 | 12.75 | 12.92 | 13.1 | 13.27 | 13.44 | 13.61
4.55 | 6.75 | 6.92 | 7.09 | 7.27 | 7.44 | 7.61 | 7.78 | 7.95 | 8.12 | 8.29 | 8.47 | 8.64 | 8.81 | 8.98 | 9.15 | 9.32 | 9.49 | 9.67 | 9.84 | 10.01 | 10.18 | 10.35 | 10.52 | 10.69 | 10.87 | 11.04 | 11.21 | 11.38 | 11.55 | 11.72 | 11.89 | 12.06 | 12.24 | 12.41 | 12.58 | 12.75 | 12.92 | 13.09 | 13.26 | 13.43 | 13.61
4.56 | 6.75 | 6.92 | 7.09 | 7.26 | 7.44 | 7.61 | 7.78 | 7.95 | 8.12 | 8.29 | 8.46 | 8.64 | 8.81 | 8.98 | 9.15 | 9.32 | 9.49 | 9.66 | 9.83 | 10.01 | 10.18 | 10.35 | 10.52 | 10.69 | 10.86 | 11.03 | 11.2 | 11.38 | 11.55 | 11.72 | 11.89 | 12.06 | 12.23 | 12.4 | 12.58 | 12.75 | 12.92 | 13.09 | 13.26 | 13.43 | 13.6
4.57 | 6.75 | 6.92 | 7.09 | 7.26 | 7.43 | 7.61 | 7.78 | 7.95 | 8.12 | 8.29 | 8.46 | 8.63 | 8.8 | 8.98 | 9.15 | 9.32 | 9.49 | 9.66 | 9.83 | 10 | 10.17 | 10.35 | 10.52 | 10.69 | 10.86 | 11.03 | 11.2 | 11.37 | 11.54 | 11.72 | 11.89 | 12.06 | 12.23 | 12.4 | 12.57 | 12.74 | 12.91 | 13.09 | 13.26 | 13.43 | 13.6
4.58 | 6.75 | 6.92 | 7.09 | 7.26 | 7.43 | 7.6 | 7.77 | 7.95 | 8.12 | 8.29 | 8.46 | 8.63 | 8.8 | 8.97 | 9.14 | 9.32 | 9.49 | 9.66 | 9.83 | 10 | 10.17 | 10.34 | 10.51 | 10.69 | 10.86 | 11.03 | 11.2 | 11.37 | 11.54 | 11.71 | 11.88 | 12.06 | 12.23 | 12.4 | 12.57 | 12.74 | 12.91 | 13.08 | 13.25 | 13.42 | 13.6
4.59 | 6.74 | 6.92 | 7.09 | 7.26 | 7.43 | 7.6 | 7.77 | 7.94 | 8.11 | 8.29 | 8.46 | 8.63 | 8.8 | 8.97 | 9.14 | 9.31 | 9.48 | 9.66 | 9.83 | 10 | 10.17 | 10.34 | 10.51 | 10.68 | 10.85 | 11.03 | 11.2 | 11.37 | 11.54 | 11.71 | 11.88 | 12.05 | 12.22 | 12.39 | 12.57 | 12.74 | 12.91 | 13.08 | 13.25 | 13.42 | 13.59
4.6 | 6.74 | 6.91 | 7.09 | 7.26 | 7.43 | 7.6 | 7.77 | 7.94 | 8.11 | 8.28 | 8.45 | 8.63 | 8.8 | 8.97 | 9.14 | 9.31 | 9.48 | 9.65 | 9.82 | 10 | 10.17 | 10.34 | 10.51 | 10.68 | 10.85 | 11.02 | 11.19 | 11.36 | 11.54 | 11.71 | 11.88 | 12.05 | 12.22 | 12.39 | 12.56 | 12.73 | 12.9 | 13.08 | 13.25 | 13.42 | 13.59
4.61 | 6.74 | 6.91 | 7.08 | 7.25 | 7.43 | 7.6 | 7.77 | 7.94 | 8.11 | 8.28 | 8.45 | 8.62 | 8.79 | 8.97 | 9.14 | 9.31 | 9.48 | 9.65 | 9.82 | 9.99 | 10.16 | 10.33 | 10.51 | 10.68 | 10.85 | 11.02 | 11.19 | 11.36 | 11.53 | 11.7 | 11.87 | 12.05 | 12.22 | 12.39 | 12.56 | 12.73 | 12.9 | 13.07 | 13.24 | 13.41 | 13.58
4.62 | 6.74 | 6.91 | 7.08 | 7.25 | 7.42 | 7.59 | 7.77 | 7.94 | 8.11 | 8.28 | 8.45 | 8.62 | 8.79 | 8.96 | 9.13 | 9.31 | 9.48 | 9.65 | 9.82 | 9.99 | 10.16 | 10.33 | 10.5 | 10.67 | 10.85 | 11.02 | 11.19 | 11.36 | 11.53 | 11.7 | 11.87 | 12.04 | 12.21 | 12.38 | 12.56 | 12.73 | 12.9 | 13.07 | 13.24 | 13.41 | 13.58
4.63 | 6.74 | 6.91 | 7.08 | 7.25 | 7.42 | 7.59 | 7.76 | 7.94 | 8.11 | 8.28 | 8.45 | 8.62 | 8.79 | 8.96 | 9.13 | 9.3 | 9.47 | 9.65 | 9.82 | 9.99 | 10.16 | 10.33 | 10.5 | 10.67 | 10.84 | 11.01 | 11.18 | 11.36 | 11.53 | 11.7 | 11.87 | 12.04 | 12.21 | 12.38 | 12.55 | 12.72 | 12.89 | 13.07 | 13.24 | 13.41 | 13.58
4.64 | 6.74 | 6.91 | 7.08 | 7.25 | 7.42 | 7.59 | 7.76 | 7.93 | 8.1 | 8.28 | 8.45 | 8.62 | 8.79 | 8.96 | 9.13 | 9.3 | 9.47 | 9.64 | 9.81 | 9.98 | 10.16 | 10.33 | 10.5 | 10.67 | 10.84 | 11.01 | 11.18 | 11.35 | 11.52 | 11.69 | 11.87 | 12.04 | 12.21 | 12.38 | 12.55 | 12.72 | 12.89 | 13.06 | 13.23 | 13.4 | 13.57
4.65 | 6.73 | 6.91 | 7.08 | 7.25 | 7.42 | 7.59 | 7.76 | 7.93 | 8.1 | 8.27 | 8.44 | 8.61 | 8.79 | 8.96 | 9.13 | 9.3 | 9.47 | 9.64 | 9.81 | 9.98 | 10.15 | 10.32 | 10.5 | 10.67 | 10.84 | 11.01 | 11.18 | 11.35 | 11.52 | 11.69 | 11.86 | 12.03 | 12.2 | 12.37 | 12.55 | 12.72 | 12.89 | 13.06 | 13.23 | 13.4 | 13.57
4.66 | 6.73 | 6.9 | 7.07 | 7.25 | 7.42 | 7.59 | 7.76 | 7.93 | 8.1 | 8.27 | 8.44 | 8.61 | 8.78 | 8.95 | 9.13 | 9.3 | 9.47 | 9.64 | 9.81 | 9.98 | 10.15 | 10.32 | 10.49 | 10.66 | 10.83 | 11 | 11.18 | 11.35 | 11.52 | 11.69 | 11.86 | 12.03 | 12.2 | 12.37 | 12.54 | 12.71 | 12.88 | 13.05 | 13.23 | 13.4 | 13.57
4.67 | 6.73 | 6.9 | 7.07 | 7.24 | 7.41 | 7.58 | 7.76 | 7.93 | 8.1 | 8.27 | 8.44 | 8.61 | 8.78 | 8.95 | 9.12 | 9.29 | 9.46 | 9.64 | 9.81 | 9.98 | 10.15 | 10.32 | 10.49 | 10.66 | 10.83 | 11 | 11.17 | 11.34 | 11.51 | 11.69 | 11.86 | 12.03 | 12.2 | 12.37 | 12.54 | 12.71 | 12.88 | 13.05 | 13.22 | 13.39 | 13.56
4.68 | 6.73 | 6.9 | 7.07 | 7.24 | 7.41 | 7.58 | 7.75 | 7.92 | 8.1 | 8.27 | 8.44 | 8.61 | 8.78 | 8.95 | 9.12 | 9.29 | 9.46 | 9.63 | 9.8 | 9.97 | 10.15 | 10.32 | 10.49 | 10.66 | 10.83 | 11 | 11.17 | 11.34 | 11.51 | 11.68 | 11.85 | 12.02 | 12.19 | 12.37 | 12.54 | 12.71 | 12.88 | 13.05 | 13.22 | 13.39 | 13.56
4.69 | 6.73 | 6.9 | 7.07 | 7.24 | 7.41 | 7.58 | 7.75 | 7.92 | 8.09 | 8.26 | 8.43 | 8.61 | 8.78 | 8.95 | 9.12 | 9.29 | 9.46 | 9.63 | 9.8 | 9.97 | 10.14 | 10.31 | 10.48 | 10.65 | 10.83 | 11 | 11.17 | 11.34 | 11.51 | 11.68 | 11.85 | 12.02 | 12.19 | 12.36 | 12.53 | 12.7 | 12.87 | 13.04 | 13.22 | 13.39 | 13.56
4.7 | 6.73 | 6.9 | 7.07 | 7.24 | 7.41 | 7.58 | 7.75 | 7.92 | 8.09 | 8.26 | 8.43 | 8.6 | 8.77 | 8.94 | 9.12 | 9.29 | 9.46 | 9.63 | 9.8 | 9.97 | 10.14 | 10.31 | 10.48 | 10.65 | 10.82 | 10.99 | 11.16 | 11.33 | 11.51 | 11.68 | 11.85 | 12.02 | 12.19 | 12.36 | 12.53 | 12.7 | 12.87 | 13.04 | 13.21 | 13.38 | 13.55
4.71 | 6.72 | 6.89 | 7.06 | 7.24 | 7.41 | 7.58 | 7.75 | 7.92 | 8.09 | 8.26 | 8.43 | 8.6 | 8.77 | 8.94 | 9.11 | 9.28 | 9.45 | 9.63 | 9.8 | 9.97 | 10.14 | 10.31 | 10.48 | 10.65 | 10.82 | 10.99 | 11.16 | 11.33 | 11.5 | 11.67 | 11.84 | 12.01 | 12.18 | 12.36 | 12.53 | 12.7 | 12.87 | 13.04 | 13.21 | 13.38 | 13.55
4.72 | 6.72 | 6.89 | 7.06 | 7.23 | 7.4 | 7.57 | 7.75 | 7.92 | 8.09 | 8.26 | 8.43 | 8.6 | 8.77 | 8.94 | 9.11 | 9.28 | 9.45 | 9.62 | 9.79 | 9.96 | 10.13 | 10.31 | 10.48 | 10.65 | 10.82 | 10.99 | 11.16 | 11.33 | 11.5 | 11.67 | 11.84 | 12.01 | 12.18 | 12.35 | 12.52 | 12.69 | 12.86 | 13.03 | 13.21 | 13.38 | 13.55
4.73 | 6.72 | 6.89 | 7.06 | 7.23 | 7.4 | 7.57 | 7.74 | 7.91 | 8.08 | 8.26 | 8.43 | 8.6 | 8.77 | 8.94 | 9.11 | 9.28 | 9.45 | 9.62 | 9.79 | 9.96 | 10.13 | 10.3 | 10.47 | 10.64 | 10.81 | 10.98 | 11.16 | 11.33 | 11.5 | 11.67 | 11.84 | 12.01 | 12.18 | 12.35 | 12.52 | 12.69 | 12.86 | 13.03 | 13.2 | 13.37 | 13.54
4.74 | 6.72 | 6.89 | 7.06 | 7.23 | 7.4 | 7.57 | 7.74 | 7.91 | 8.08 | 8.25 | 8.42 | 8.59 | 8.76 | 8.94 | 9.11 | 9.28 | 9.45 | 9.62 | 9.79 | 9.96 | 10.13 | 10.3 | 10.47 | 10.64 | 10.81 | 10.98 | 11.15 | 11.32 | 11.49 | 11.66 | 11.83 | 12 | 12.18 | 12.35 | 12.52 | 12.69 | 12.86 | 13.03 | 13.2 | 13.37 | 13.54
4.75 | 6.72 | 6.89 | 7.06 | 7.23 | 7.4 | 7.57 | 7.74 | 7.91 | 8.08 | 8.25 | 8.42 | 8.59 | 8.76 | 8.93 | 9.1 | 9.27 | 9.44 | 9.62 | 9.79 | 9.96 | 10.13 | 10.3 | 10.47 | 10.64 | 10.81 | 10.98 | 11.15 | 11.32 | 11.49 | 11.66 | 11.83 | 12 | 12.17 | 12.34 | 12.51 | 12.68 | 12.85 | 13.02 | 13.19 | 13.37 | 13.54
4.76 | 6.71 | 6.88 | 7.06 | 7.23 | 7.4 | 7.57 | 7.74 | 7.91 | 8.08 | 8.25 | 8.42 | 8.59 | 8.76 | 8.93 | 9.1 | 9.27 | 9.44 | 9.61 | 9.78 | 9.95 | 10.12 | 10.29 | 10.46 | 10.64 | 10.81 | 10.98 | 11.15 | 11.32 | 11.49 | 11.66 | 11.83 | 12 | 12.17 | 12.34 | 12.51 | 12.68 | 12.85 | 13.02 | 13.19 | 13.36 | 13.53
4.77 | 6.71 | 6.88 | 7.05 | 7.22 | 7.39 | 7.56 | 7.74 | 7.91 | 8.08 | 8.25 | 8.42 | 8.59 | 8.76 | 8.93 | 9.1 | 9.27 | 9.44 | 9.61 | 9.78 | 9.95 | 10.12 | 10.29 | 10.46 | 10.63 | 10.8 | 10.97 | 11.14 | 11.31 | 11.48 | 11.65 | 11.83 | 12 | 12.17 | 12.34 | 12.51 | 12.68 | 12.85 | 13.02 | 13.19 | 13.36 | 13.53
4.78 | 6.71 | 6.88 | 7.05 | 7.22 | 7.39 | 7.56 | 7.73 | 7.9 | 8.07 | 8.24 | 8.42 | 8.59 | 8.76 | 8.93 | 9.1 | 9.27 | 9.44 | 9.61 | 9.78 | 9.95 | 10.12 | 10.29 | 10.46 | 10.63 | 10.8 | 10.97 | 11.14 | 11.31 | 11.48 | 11.65 | 11.82 | 11.99 | 12.16 | 12.33 | 12.5 | 12.67 | 12.84 | 13.01 | 13.18 | 13.35 | 13.53
4.79 | 6.71 | 6.88 | 7.05 | 7.22 | 7.39 | 7.56 | 7.73 | 7.9 | 8.07 | 8.24 | 8.41 | 8.58 | 8.75 | 8.92 | 9.09 | 9.26 | 9.43 | 9.61 | 9.78 | 9.95 | 10.12 | 10.29 | 10.46 | 10.63 | 10.8 | 10.97 | 11.14 | 11.31 | 11.48 | 11.65 | 11.82 | 11.99 | 12.16 | 12.33 | 12.5 | 12.67 | 12.84 | 13.01 | 13.18 | 13.35 | 13.52
4.8 | 6.71 | 6.88 | 7.05 | 7.22 | 7.39 | 7.56 | 7.73 | 7.9 | 8.07 | 8.24 | 8.41 | 8.58 | 8.75 | 8.92 | 9.09 | 9.26 | 9.43 | 9.6 | 9.77 | 9.94 | 10.11 | 10.28 | 10.45 | 10.62 | 10.79 | 10.96 | 11.14 | 11.31 | 11.48 | 11.65 | 11.82 | 11.99 | 12.16 | 12.33 | 12.5 | 12.67 | 12.84 | 13.01 | 13.18 | 13.35 | 13.52
4.81 | 6.71 | 6.88 | 7.05 | 7.22 | 7.39 | 7.56 | 7.73 | 7.9 | 8.07 | 8.24 | 8.41 | 8.58 | 8.75 | 8.92 | 9.09 | 9.26 | 9.43 | 9.6 | 9.77 | 9.94 | 10.11 | 10.28 | 10.45 | 10.62 | 10.79 | 10.96 | 11.13 | 11.3 | 11.47 | 11.64 | 11.81 | 11.98 | 12.15 | 12.32 | 12.49 | 12.66 | 12.83 | 13 | 13.17 | 13.34 | 13.51
4.82 | 6.7 | 6.87 | 7.04 | 7.21 | 7.38 | 7.56 | 7.73 | 7.9 | 8.07 | 8.24 | 8.41 | 8.58 | 8.75 | 8.92 | 9.09 | 9.26 | 9.43 | 9.6 | 9.77 | 9.94 | 10.11 | 10.28 | 10.45 | 10.62 | 10.79 | 10.96 | 11.13 | 11.3 | 11.47 | 11.64 | 11.81 | 11.98 | 12.15 | 12.32 | 12.49 | 12.66 | 12.83 | 13 | 13.17 | 13.34 | 13.51
4.83 | 6.7 | 6.87 | 7.04 | 7.21 | 7.38 | 7.55 | 7.72 | 7.89 | 8.06 | 8.23 | 8.4 | 8.57 | 8.74 | 8.91 | 9.08 | 9.25 | 9.42 | 9.6 | 9.77 | 9.94 | 10.11 | 10.28 | 10.45 | 10.62 | 10.79 | 10.96 | 11.13 | 11.3 | 11.47 | 11.64 | 11.81 | 11.98 | 12.15 | 12.32 | 12.49 | 12.66 | 12.83 | 13 | 13.17 | 13.34 | 13.51
4.84 | 6.7 | 6.87 | 7.04 | 7.21 | 7.38 | 7.55 | 7.72 | 7.89 | 8.06 | 8.23 | 8.4 | 8.57 | 8.74 | 8.91 | 9.08 | 9.25 | 9.42 | 9.59 | 9.76 | 9.93 | 10.1 | 10.27 | 10.44 | 10.61 | 10.78 | 10.95 | 11.12 | 11.29 | 11.46 | 11.63 | 11.8 | 11.97 | 12.14 | 12.31 | 12.48 | 12.65 | 12.82 | 12.99 | 13.16 | 13.33 | 13.5
4.85 | 6.7 | 6.87 | 7.04 | 7.21 | 7.38 | 7.55 | 7.72 | 7.89 | 8.06 | 8.23 | 8.4 | 8.57 | 8.74 | 8.91 | 9.08 | 9.25 | 9.42 | 9.59 | 9.76 | 9.93 | 10.1 | 10.27 | 10.44 | 10.61 | 10.78 | 10.95 | 11.12 | 11.29 | 11.46 | 11.63 | 11.8 | 11.97 | 12.14 | 12.31 | 12.48 | 12.65 | 12.82 | 12.99 | 13.16 | 13.33 | 13.5
4.86 | 6.7 | 6.87 | 7.04 | 7.21 | 7.38 | 7.55 | 7.72 | 7.89 | 8.06 | 8.23 | 8.4 | 8.57 | 8.74 | 8.91 | 9.08 | 9.25 | 9.42 | 9.59 | 9.76 | 9.93 | 10.1 | 10.27 | 10.44 | 10.61 | 10.78 | 10.95 | 11.12 | 11.29 | 11.46 | 11.63 | 11.8 | 11.97 | 12.14 | 12.31 | 12.48 | 12.65 | 12.82 | 12.99 | 13.16 | 13.33 | 13.5
4.87 | 6.69 | 6.86 | 7.03 | 7.21 | 7.38 | 7.55 | 7.72 | 7.89 | 8.06 | 8.23 | 8.4 | 8.57 | 8.74 | 8.91 | 9.08 | 9.25 | 9.42 | 9.59 | 9.76 | 9.93 | 10.1 | 10.27 | 10.43 | 10.6 | 10.77 | 10.94 | 11.11 | 11.28 | 11.45 | 11.62 | 11.79 | 11.96 | 12.13 | 12.3 | 12.47 | 12.64 | 12.81 | 12.98 | 13.15 | 13.32 | 13.49
4.88 | 6.69 | 6.86 | 7.03 | 7.2 | 7.37 | 7.54 | 7.71 | 7.88 | 8.05 | 8.22 | 8.39 | 8.56 | 8.73 | 8.9 | 9.07 | 9.24 | 9.41 | 9.58 | 9.75 | 9.92 | 10.09 | 10.26 | 10.43 | 10.6 | 10.77 | 10.94 | 11.11 | 11.28 | 11.45 | 11.62 | 11.79 | 11.96 | 12.13 | 12.3 | 12.47 | 12.64 | 12.81 | 12.98 | 13.15 | 13.32 | 13.49
4.89 | 6.69 | 6.86 | 7.03 | 7.2 | 7.37 | 7.54 | 7.71 | 7.88 | 8.05 | 8.22 | 8.39 | 8.56 | 8.73 | 8.9 | 9.07 | 9.24 | 9.41 | 9.58 | 9.75 | 9.92 | 10.09 | 10.26 | 10.43 | 10.6 | 10.77 | 10.94 | 11.11 | 11.28 | 11.45 | 11.62 | 11.79 | 11.96 | 12.13 | 12.3 | 12.47 | 12.64 | 12.81 | 12.98 | 13.15 | 13.32 | 13.49
4.9 | 6.69 | 6.86 | 7.03 | 7.2 | 7.37 | 7.54 | 7.71 | 7.88 | 8.05 | 8.22 | 8.39 | 8.56 | 8.73 | 8.9 | 9.07 | 9.24 | 9.41 | 9.58 | 9.75 | 9.92 | 10.09 | 10.26 | 10.43 | 10.6 | 10.77 | 10.94 | 11.11 | 11.28 | 11.45 | 11.62 | 11.79 | 11.96 | 12.12 | 12.29 | 12.46 | 12.63 | 12.8 | 12.97 | 13.14 | 13.31 | 13.48
4.91 | 6.69 | 6.86 | 7.03 | 7.2 | 7.37 | 7.54 | 7.71 | 7.88 | 8.05 | 8.22 | 8.39 | 8.56 | 8.73 | 8.9 | 9.07 | 9.24 | 9.41 | 9.58 | 9.74 | 9.91 | 10.08 | 10.25 | 10.42 | 10.59 | 10.76 | 10.93 | 11.1 | 11.27 | 11.44 | 11.61 | 11.78 | 11.95 | 12.12 | 12.29 | 12.46 | 12.63 | 12.8 | 12.97 | 13.14 | 13.31 | 13.48
4.92 | 6.69 | 6.86 | 7.03 | 7.2 | 7.37 | 7.54 | 7.7 | 7.87 | 8.04 | 8.21 | 8.38 | 8.55 | 8.72 | 8.89 | 9.06 | 9.23 | 9.4 | 9.57 | 9.74 | 9.91 | 10.08 | 10.25 | 10.42 | 10.59 | 10.76 | 10.93 | 11.1 | 11.27 | 11.44 | 11.61 | 11.78 | 11.95 | 12.12 | 12.29 | 12.46 | 12.63 | 12.8 | 12.97 | 13.14 | 13.31 | 13.48
4.93 | 6.68 | 6.85 | 7.02 | 7.19 | 7.36 | 7.53 | 7.7 | 7.87 | 8.04 | 8.21 | 8.38 | 8.55 | 8.72 | 8.89 | 9.06 | 9.23 | 9.4 | 9.57 | 9.74 | 9.91 | 10.08 | 10.25 | 10.42 | 10.59 | 10.76 | 10.93 | 11.1 | 11.27 | 11.44 | 11.61 | 11.78 | 11.95 | 12.12 | 12.29 | 12.45 | 12.62 | 12.79 | 12.96 | 13.13 | 13.3 | 13.47
4.94 | 6.68 | 6.85 | 7.02 | 7.19 | 7.36 | 7.53 | 7.7 | 7.87 | 8.04 | 8.21 | 8.38 | 8.55 | 8.72 | 8.89 | 9.06 | 9.23 | 9.4 | 9.57 | 9.74 | 9.91 | 10.08 | 10.25 | 10.42 | 10.59 | 10.76 | 10.92 | 11.09 | 11.26 | 11.43 | 11.6 | 11.77 | 11.94 | 12.11 | 12.28 | 12.45 | 12.62 | 12.79 | 12.96 | 13.13 | 13.3 | 13.47
4.95 | 6.68 | 6.85 | 7.02 | 7.19 | 7.36 | 7.53 | 7.7 | 7.87 | 8.04 | 8.21 | 8.38 | 8.55 | 8.72 | 8.89 | 9.06 | 9.23 | 9.4 | 9.57 | 9.73 | 9.9 | 10.07 | 10.24 | 10.41 | 10.58 | 10.75 | 10.92 | 11.09 | 11.26 | 11.43 | 11.6 | 11.77 | 11.94 | 12.11 | 12.28 | 12.45 | 12.62 | 12.79 | 12.96 | 13.13 | 13.3 | 13.47
4.96 | 6.68 | 6.85 | 7.02 | 7.19 | 7.36 | 7.53 | 7.7 | 7.87 | 8.04 | 8.21 | 8.38 | 8.55 | 8.71 | 8.88 | 9.05 | 9.22 | 9.39 | 9.56 | 9.73 | 9.9 | 10.07 | 10.24 | 10.41 | 10.58 | 10.75 | 10.92 | 11.09 | 11.26 | 11.43 | 11.6 | 11.77 | 11.94 | 12.11 | 12.28 | 12.45 | 12.61 | 12.78 | 12.95 | 13.12 | 13.29 | 13.46
4.97 | 6.68 | 6.85 | 7.02 | 7.19 | 7.36 | 7.53 | 7.69 | 7.86 | 8.03 | 8.2 | 8.37 | 8.54 | 8.71 | 8.88 | 9.05 | 9.22 | 9.39 | 9.56 | 9.73 | 9.9 | 10.07 | 10.24 | 10.41 | 10.58 | 10.75 | 10.92 | 11.09 | 11.26 | 11.42 | 11.59 | 11.76 | 11.93 | 12.1 | 12.27 | 12.44 | 12.61 | 12.78 | 12.95 | 13.12 | 13.29 | 13.46
4.98 | 6.68 | 6.85 | 7.01 | 7.18 | 7.35 | 7.52 | 7.69 | 7.86 | 8.03 | 8.2 | 8.37 | 8.54 | 8.71 | 8.88 | 9.05 | 9.22 | 9.39 | 9.56 | 9.73 | 9.9 | 10.07 | 10.24 | 10.41 | 10.57 | 10.74 | 10.91 | 11.08 | 11.25 | 11.42 | 11.59 | 11.76 | 11.93 | 12.1 | 12.27 | 12.44 | 12.61 | 12.78 | 12.95 | 13.12 | 13.29 | 13.46
4.99 | 6.67 | 6.84 | 7.01 | 7.18 | 7.35 | 7.52 | 7.69 | 7.86 | 8.03 | 8.2 | 8.37 | 8.54 | 8.71 | 8.88 | 9.05 | 9.22 | 9.39 | 9.56 | 9.72 | 9.89 | 10.06 | 10.23 | 10.4 | 10.57 | 10.74 | 10.91 | 11.08 | 11.25 | 11.42 | 11.59 | 11.76 | 11.93 | 12.1 | 12.27 | 12.44 | 12.6 | 12.77 | 12.94 | 13.11 | 13.28 | 13.45
5 | 6.67 | 6.84 | 7.01 | 7.18 | 7.35 | 7.52 | 7.69 | 7.86 | 8.03 | 8.2 | 8.37 | 8.54 | 8.71 | 8.87 | 9.04 | 9.21 | 9.38 | 9.55 | 9.72 | 9.89 | 10.06 | 10.23 | 10.4 | 10.57 | 10.74 | 10.91 | 11.08 | 11.25 | 11.42 | 11.59 | 11.75 | 11.92 | 12.09 | 12.26 | 12.43 | 12.6 | 12.77 | 12.94 | 13.11 | 13.28 | 13.45
5.01 | 6.67 | 6.84 | 7.01 | 7.18 | 7.35 | 7.52 | 7.69 | 7.86 | 8.03 | 8.2 | 8.36 | 8.53 | 8.7 | 8.87 | 9.04 | 9.21 | 9.38 | 9.55 | 9.72 | 9.89 | 10.06 | 10.23 | 10.4 | 10.57 | 10.74 | 10.91 | 11.07 | 11.24 | 11.41 | 11.58 | 11.75 | 11.92 | 12.09 | 12.26 | 12.43 | 12.6 | 12.77 | 12.94 | 13.11 | 13.28 | 13.44
5.02 | 6.67 | 6.84 | 7.01 | 7.18 | 7.35 | 7.52 | 7.68 | 7.85 | 8.02 | 8.19 | 8.36 | 8.53 | 8.7 | 8.87 | 9.04 | 9.21 | 9.38 | 9.55 | 9.72 | 9.89 | 10.06 | 10.23 | 10.39 | 10.56 | 10.73 | 10.9 | 11.07 | 11.24 | 11.41 | 11.58 | 11.75 | 11.92 | 12.09 | 12.26 | 12.43 | 12.6 | 12.76 | 12.93 | 13.1 | 13.27 | 13.44
5.03 | 6.67 | 6.84 | 7.01 | 7.17 | 7.34 | 7.51 | 7.68 | 7.85 | 8.02 | 8.19 | 8.36 | 8.53 | 8.7 | 8.87 | 9.04 | 9.21 | 9.38 | 9.55 | 9.71 | 9.88 | 10.05 | 10.22 | 10.39 | 10.56 | 10.73 | 10.9 | 11.07 | 11.24 | 11.41 | 11.58 | 11.75 | 11.91 | 12.08 | 12.25 | 12.42 | 12.59 | 12.76 | 12.93 | 13.1 | 13.27 | 13.44
5.04 | 6.66 | 6.83 | 7 | 7.17 | 7.34 | 7.51 | 7.68 | 7.85 | 8.02 | 8.19 | 8.36 | 8.53 | 8.7 | 8.87 | 9.03 | 9.2 | 9.37 | 9.54 | 9.71 | 9.88 | 10.05 | 10.22 | 10.39 | 10.56 | 10.73 | 10.9 | 11.07 | 11.24 | 11.4 | 11.57 | 11.74 | 11.91 | 12.08 | 12.25 | 12.42 | 12.59 | 12.76 | 12.93 | 13.1 | 13.27 | 13.43
5.05 | 6.66 | 6.83 | 7 | 7.17 | 7.34 | 7.51 | 7.68 | 7.85 | 8.02 | 8.19 | 8.36 | 8.52 | 8.69 | 8.86 | 9.03 | 9.2 | 9.37 | 9.54 | 9.71 | 9.88 | 10.05 | 10.22 | 10.39 | 10.56 | 10.72 | 10.89 | 11.06 | 11.23 | 11.4 | 11.57 | 11.74 | 11.91 | 12.08 | 12.25 | 12.42 | 12.59 | 12.75 | 12.92 | 13.09 | 13.26 | 13.43
5.06 | 6.66 | 6.83 | 7 | 7.17 | 7.34 | 7.51 | 7.68 | 7.85 | 8.02 | 8.18 | 8.35 | 8.52 | 8.69 | 8.86 | 9.03 | 9.2 | 9.37 | 9.54 | 9.71 | 9.88 | 10.05 | 10.21 | 10.38 | 10.55 | 10.72 | 10.89 | 11.06 | 11.23 | 11.4 | 11.57 | 11.74 | 11.91 | 12.07 | 12.24 | 12.41 | 12.58 | 12.75 | 12.92 | 13.09 | 13.26 | 13.43
5.07 | 6.66 | 6.83 | 7 | 7.17 | 7.34 | 7.51 | 7.67 | 7.84 | 8.01 | 8.18 | 8.35 | 8.52 | 8.69 | 8.86 | 9.03 | 9.2 | 9.37 | 9.54 | 9.7 | 9.87 | 10.04 | 10.21 | 10.38 | 10.55 | 10.72 | 10.89 | 11.06 | 11.23 | 11.4 | 11.56 | 11.73 | 11.9 | 12.07 | 12.24 | 12.41 | 12.58 | 12.75 | 12.92 | 13.09 | 13.26 | 13.42
5.08 | 6.66 | 6.83 | 7 | 7.17 | 7.33 | 7.5 | 7.67 | 7.84 | 8.01 | 8.18 | 8.35 | 8.52 | 8.69 | 8.86 | 9.03 | 9.19 | 9.36 | 9.53 | 9.7 | 9.87 | 10.04 | 10.21 | 10.38 | 10.55 | 10.72 | 10.89 | 11.05 | 11.22 | 11.39 | 11.56 | 11.73 | 11.9 | 12.07 | 12.24 | 12.41 | 12.58 | 12.74 | 12.91 | 13.08 | 13.25 | 13.42
5.09 | 6.66 | 6.83 | 6.99 | 7.16 | 7.33 | 7.5 | 7.67 | 7.84 | 8.01 | 8.18 | 8.35 | 8.52 | 8.69 | 8.85 | 9.02 | 9.19 | 9.36 | 9.53 | 9.7 | 9.87 | 10.04 | 10.21 | 10.38 | 10.54 | 10.71 | 10.88 | 11.05 | 11.22 | 11.39 | 11.56 | 11.73 | 11.9 | 12.07 | 12.23 | 12.4 | 12.57 | 12.74 | 12.91 | 13.08 | 13.25 | 13.42
5.1 | 6.65 | 6.82 | 6.99 | 7.16 | 7.33 | 7.5 | 7.67 | 7.84 | 8.01 | 8.18 | 8.34 | 8.51 | 8.68 | 8.85 | 9.02 | 9.19 | 9.36 | 9.53 | 9.7 | 9.87 | 10.03 | 10.2 | 10.37 | 10.54 | 10.71 | 10.88 | 11.05 | 11.22 | 11.39 | 11.56 | 11.72 | 11.89 | 12.06 | 12.23 | 12.4 | 12.57 | 12.74 | 12.91 | 13.08 | 13.24 | 13.41
5.11 | 6.65 | 6.82 | 6.99 | 7.16 | 7.33 | 7.5 | 7.67 | 7.84 | 8 | 8.17 | 8.34 | 8.51 | 8.68 | 8.85 | 9.02 | 9.19 | 9.36 | 9.53 | 9.69 | 9.86 | 10.03 | 10.2 | 10.37 | 10.54 | 10.71 | 10.88 | 11.05 | 11.21 | 11.38 | 11.55 | 11.72 | 11.89 | 12.06 | 12.23 | 12.4 | 12.57 | 12.73 | 12.9 | 13.07 | 13.24 | 13.41
5.12 | 6.65 | 6.82 | 6.99 | 7.16 | 7.33 | 7.5 | 7.66 | 7.83 | 8 | 8.17 | 8.34 | 8.51 | 8.68 | 8.85 | 9.02 | 9.19 | 9.35 | 9.52 | 9.69 | 9.86 | 10.03 | 10.2 | 10.37 | 10.54 | 10.71 | 10.87 | 11.04 | 11.21 | 11.38 | 11.55 | 11.72 | 11.89 | 12.06 | 12.22 | 12.39 | 12.56 | 12.73 | 12.9 | 13.07 | 13.24 | 13.41
5.13 | 6.65 | 6.82 | 6.99 | 7.16 | 7.32 | 7.49 | 7.66 | 7.83 | 8 | 8.17 | 8.34 | 8.51 | 8.68 | 8.84 | 9.01 | 9.18 | 9.35 | 9.52 | 9.69 | 9.86 | 10.03 | 10.2 | 10.36 | 10.53 | 10.7 | 10.87 | 11.04 | 11.21 | 11.38 | 11.55 | 11.72 | 11.88 | 12.05 | 12.22 | 12.39 | 12.56 | 12.73 | 12.9 | 13.07 | 13.23 | 13.4
5.14 | 6.65 | 6.82 | 6.99 | 7.15 | 7.32 | 7.49 | 7.66 | 7.83 | 8 | 8.17 | 8.34 | 8.5 | 8.67 | 8.84 | 9.01 | 9.18 | 9.35 | 9.52 | 9.69 | 9.86 | 10.02 | 10.19 | 10.36 | 10.53 | 10.7 | 10.87 | 11.04 | 11.21 | 11.37 | 11.54 | 11.71 | 11.88 | 12.05 | 12.22 | 12.39 | 12.56 | 12.72 | 12.89 | 13.06 | 13.23 | 13.4
5.15 | 6.65 | 6.81 | 6.98 | 7.15 | 7.32 | 7.49 | 7.66 | 7.83 | 8 | 8.17 | 8.33 | 8.5 | 8.67 | 8.84 | 9.01 | 9.18 | 9.35 | 9.52 | 9.68 | 9.85 | 10.02 | 10.19 | 10.36 | 10.53 | 10.7 | 10.87 | 11.03 | 11.2 | 11.37 | 11.54 | 11.71 | 11.88 | 12.05 | 12.22 | 12.38 | 12.55 | 12.72 | 12.89 | 13.06 | 13.23 | 13.4
5.16 | 6.64 | 6.81 | 6.98 | 7.15 | 7.32 | 7.49 | 7.66 | 7.83 | 7.99 | 8.16 | 8.33 | 8.5 | 8.67 | 8.84 | 9.01 | 9.18 | 9.34 | 9.51 | 9.68 | 9.85 | 10.02 | 10.19 | 10.36 | 10.53 | 10.69 | 10.86 | 11.03 | 11.2 | 11.37 | 11.54 | 11.71 | 11.88 | 12.04 | 12.21 | 12.38 | 12.55 | 12.72 | 12.89 | 13.06 | 13.22 | 13.39
5.17 | 6.64 | 6.81 | 6.98 | 7.15 | 7.32 | 7.49 | 7.65 | 7.82 | 7.99 | 8.16 | 8.33 | 8.5 | 8.67 | 8.84 | 9 | 9.17 | 9.34 | 9.51 | 9.68 | 9.85 | 10.02 | 10.19 | 10.35 | 10.52 | 10.69 | 10.86 | 11.03 | 11.2 | 11.37 | 11.53 | 11.7 | 11.87 | 12.04 | 12.21 | 12.38 | 12.55 | 12.72 | 12.88 | 13.05 | 13.22 | 13.39
5.18 | 6.64 | 6.81 | 6.98 | 7.15 | 7.32 | 7.48 | 7.65 | 7.82 | 7.99 | 8.16 | 8.33 | 8.5 | 8.66 | 8.83 | 9 | 9.17 | 9.34 | 9.51 | 9.68 | 9.85 | 10.01 | 10.18 | 10.35 | 10.52 | 10.69 | 10.86 | 11.03 | 11.19 | 11.36 | 11.53 | 11.7 | 11.87 | 12.04 | 12.21 | 12.37 | 12.54 | 12.71 | 12.88 | 13.05 | 13.22 | 13.39
5.19 | 6.64 | 6.81 | 6.98 | 7.14 | 7.31 | 7.48 | 7.65 | 7.82 | 7.99 | 8.16 | 8.33 | 8.49 | 8.66 | 8.83 | 9 | 9.17 | 9.34 | 9.51 | 9.67 | 9.84 | 10.01 | 10.18 | 10.35 | 10.52 | 10.69 | 10.85 | 11.02 | 11.19 | 11.36 | 11.53 | 11.7 | 11.87 | 12.03 | 12.2 | 12.37 | 12.54 | 12.71 | 12.88 | 13.05 | 13.21 | 13.38
5.2 | 6.64 | 6.81 | 6.97 | 7.14 | 7.31 | 7.48 | 7.65 | 7.82 | 7.99 | 8.15 | 8.32 | 8.49 | 8.66 | 8.83 | 9 | 9.17 | 9.33 | 9.5 | 9.67 | 9.84 | 10.01 | 10.18 | 10.35 | 10.51 | 10.68 | 10.85 | 11.02 | 11.19 | 11.36 | 11.53 | 11.69 | 11.86 | 12.03 | 12.2 | 12.37 | 12.54 | 12.71 | 12.87 | 13.04 | 13.21 | 13.38
5.21 | 6.64 | 6.8 | 6.97 | 7.14 | 7.31 | 7.48 | 7.65 | 7.82 | 7.98 | 8.15 | 8.32 | 8.49 | 8.66 | 8.83 | 9 | 9.16 | 9.33 | 9.5 | 9.67 | 9.84 | 10.01 | 10.17 | 10.34 | 10.51 | 10.68 | 10.85 | 11.02 | 11.19 | 11.35 | 11.52 | 11.69 | 11.86 | 12.03 | 12.2 | 12.37 | 12.53 | 12.7 | 12.87 | 13.04 | 13.21 | 13.38
5.22 | 6.63 | 6.8 | 6.97 | 7.14 | 7.31 | 7.48 | 7.64 | 7.81 | 7.98 | 8.15 | 8.32 | 8.49 | 8.66 | 8.82 | 8.99 | 9.16 | 9.33 | 9.5 | 9.67 | 9.84 | 10 | 10.17 | 10.34 | 10.51 | 10.68 | 10.85 | 11.01 | 11.18 | 11.35 | 11.52 | 11.69 | 11.86 | 12.03 | 12.19 | 12.36 | 12.53 | 12.7 | 12.87 | 13.04 | 13.2 | 13.37
5.23 | 6.63 | 6.8 | 6.97 | 7.14 | 7.31 | 7.47 | 7.64 | 7.81 | 7.98 | 8.15 | 8.32 | 8.49 | 8.65 | 8.82 | 8.99 | 9.16 | 9.33 | 9.5 | 9.66 | 9.83 | 10 | 10.17 | 10.34 | 10.51 | 10.67 | 10.84 | 11.01 | 11.18 | 11.35 | 11.52 | 11.69 | 11.85 | 12.02 | 12.19 | 12.36 | 12.53 | 12.7 | 12.86 | 13.03 | 13.2 | 13.37
5.24 | 6.63 | 6.8 | 6.97 | 7.14 | 7.3 | 7.47 | 7.64 | 7.81 | 7.98 | 8.15 | 8.31 | 8.48 | 8.65 | 8.82 | 8.99 | 9.16 | 9.32 | 9.49 | 9.66 | 9.83 | 10 | 10.17 | 10.34 | 10.5 | 10.67 | 10.84 | 11.01 | 11.18 | 11.35 | 11.51 | 11.68 | 11.85 | 12.02 | 12.19 | 12.36 | 12.52 | 12.69 | 12.86 | 13.03 | 13.2 | 13.37
5.25 | 6.63 | 6.8 | 6.96 | 7.13 | 7.3 | 7.47 | 7.64 | 7.81 | 7.98 | 8.14 | 8.31 | 8.48 | 8.65 | 8.82 | 8.99 | 9.15 | 9.32 | 9.49 | 9.66 | 9.83 | 10 | 10.16 | 10.33 | 10.5 | 10.67 | 10.84 | 11.01 | 11.17 | 11.34 | 11.51 | 11.68 | 11.85 | 12.02 | 12.18 | 12.35 | 12.52 | 12.69 | 12.86 | 13.03 | 13.19 | 13.36
5.26 | 6.63 | 6.79 | 6.96 | 7.13 | 7.3 | 7.47 | 7.64 | 7.8 | 7.97 | 8.14 | 8.31 | 8.48 | 8.65 | 8.82 | 8.98 | 9.15 | 9.32 | 9.49 | 9.66 | 9.83 | 9.99 | 10.16 | 10.33 | 10.5 | 10.67 | 10.83 | 11 | 11.17 | 11.34 | 11.51 | 11.68 | 11.84 | 12.01 | 12.18 | 12.35 | 12.52 | 12.69 | 12.85 | 13.02 | 13.19 | 13.36
5.27 | 6.62 | 6.79 | 6.96 | 7.13 | 7.3 | 7.47 | 7.63 | 7.8 | 7.97 | 8.14 | 8.31 | 8.48 | 8.64 | 8.81 | 8.98 | 9.15 | 9.32 | 9.49 | 9.65 | 9.82 | 9.99 | 10.16 | 10.33 | 10.5 | 10.66 | 10.83 | 11 | 11.17 | 11.34 | 11.5 | 11.67 | 11.84 | 12.01 | 12.18 | 12.35 | 12.51 | 12.68 | 12.85 | 13.02 | 13.19 | 13.36
5.28 | 6.62 | 6.79 | 6.96 | 7.13 | 7.3 | 7.46 | 7.63 | 7.8 | 7.97 | 8.14 | 8.31 | 8.47 | 8.64 | 8.81 | 8.98 | 9.15 | 9.32 | 9.48 | 9.65 | 9.82 | 9.99 | 10.16 | 10.32 | 10.49 | 10.66 | 10.83 | 11 | 11.17 | 11.33 | 11.5 | 11.67 | 11.84 | 12.01 | 12.17 | 12.34 | 12.51 | 12.68 | 12.85 | 13.02 | 13.18 | 13.35
5.29 | 6.62 | 6.79 | 6.96 | 7.13 | 7.29 | 7.46 | 7.63 | 7.8 | 7.97 | 8.14 | 8.3 | 8.47 | 8.64 | 8.81 | 8.98 | 9.14 | 9.31 | 9.48 | 9.65 | 9.82 | 9.99 | 10.15 | 10.32 | 10.49 | 10.66 | 10.83 | 10.99 | 11.16 | 11.33 | 11.5 | 11.67 | 11.84 | 12 | 12.17 | 12.34 | 12.51 | 12.68 | 12.84 | 13.01 | 13.18 | 13.35
5.3 | 6.62 | 6.79 | 6.96 | 7.12 | 7.29 | 7.46 | 7.63 | 7.8 | 7.97 | 8.13 | 8.3 | 8.47 | 8.64 | 8.81 | 8.97 | 9.14 | 9.31 | 9.48 | 9.65 | 9.81 | 9.98 | 10.15 | 10.32 | 10.49 | 10.66 | 10.82 | 10.99 | 11.16 | 11.33 | 11.5 | 11.66 | 11.83 | 12 | 12.17 | 12.34 | 12.5 | 12.67 | 12.84 | 13.01 | 13.18 | 13.34
5.31 | 6.62 | 6.79 | 6.95 | 7.12 | 7.29 | 7.46 | 7.63 | 7.79 | 7.96 | 8.13 | 8.3 | 8.47 | 8.64 | 8.8 | 8.97 | 9.14 | 9.31 | 9.48 | 9.64 | 9.81 | 9.98 | 10.15 | 10.32 | 10.48 | 10.65 | 10.82 | 10.99 | 11.16 | 11.33 | 11.49 | 11.66 | 11.83 | 12 | 12.17 | 12.33 | 12.5 | 12.67 | 12.84 | 13.01 | 13.17 | 13.34
5.32 | 6.62 | 6.78 | 6.95 | 7.12 | 7.29 | 7.46 | 7.62 | 7.79 | 7.96 | 8.13 | 8.3 | 8.47 | 8.63 | 8.8 | 8.97 | 9.14 | 9.31 | 9.47 | 9.64 | 9.81 | 9.98 | 10.15 | 10.31 | 10.48 | 10.65 | 10.82 | 10.99 | 11.15 | 11.32 | 11.49 | 11.66 | 11.83 | 11.99 | 12.16 | 12.33 | 12.5 | 12.67 | 12.83 | 13 | 13.17 | 13.34
5.33 | 6.61 | 6.78 | 6.95 | 7.12 | 7.29 | 7.45 | 7.62 | 7.79 | 7.96 | 8.13 | 8.29 | 8.46 | 8.63 | 8.8 | 8.97 | 9.14 | 9.3 | 9.47 | 9.64 | 9.81 | 9.98 | 10.14 | 10.31 | 10.48 | 10.65 | 10.82 | 10.98 | 11.15 | 11.32 | 11.49 | 11.66 | 11.82 | 11.99 | 12.16 | 12.33 | 12.49 | 12.66 | 12.83 | 13 | 13.17 | 13.33
5.34 | 6.61 | 6.78 | 6.95 | 7.12 | 7.28 | 7.45 | 7.62 | 7.79 | 7.96 | 8.12 | 8.29 | 8.46 | 8.63 | 8.8 | 8.96 | 9.13 | 9.3 | 9.47 | 9.64 | 9.8 | 9.97 | 10.14 | 10.31 | 10.48 | 10.64 | 10.81 | 10.98 | 11.15 | 11.32 | 11.48 | 11.65 | 11.82 | 11.99 | 12.16 | 12.32 | 12.49 | 12.66 | 12.83 | 13 | 13.16 | 13.33
5.35 | 6.61 | 6.78 | 6.95 | 7.11 | 7.28 | 7.45 | 7.62 | 7.79 | 7.95 | 8.12 | 8.29 | 8.46 | 8.63 | 8.79 | 8.96 | 9.13 | 9.3 | 9.47 | 9.63 | 9.8 | 9.97 | 10.14 | 10.31 | 10.47 | 10.64 | 10.81 | 10.98 | 11.15 | 11.31 | 11.48 | 11.65 | 11.82 | 11.98 | 12.15 | 12.32 | 12.49 | 12.66 | 12.82 | 12.99 | 13.16 | 13.33
5.36 | 6.61 | 6.78 | 6.94 | 7.11 | 7.28 | 7.45 | 7.62 | 7.78 | 7.95 | 8.12 | 8.29 | 8.46 | 8.62 | 8.79 | 8.96 | 9.13 | 9.3 | 9.46 | 9.63 | 9.8 | 9.97 | 10.14 | 10.3 | 10.47 | 10.64 | 10.81 | 10.97 | 11.14 | 11.31 | 11.48 | 11.65 | 11.81 | 11.98 | 12.15 | 12.32 | 12.49 | 12.65 | 12.82 | 12.99 | 13.16 | 13.32
5.37 | 6.61 | 6.78 | 6.94 | 7.11 | 7.28 | 7.45 | 7.61 | 7.78 | 7.95 | 8.12 | 8.29 | 8.45 | 8.62 | 8.79 | 8.96 | 9.13 | 9.29 | 9.46 | 9.63 | 9.8 | 9.96 | 10.13 | 10.3 | 10.47 | 10.64 | 10.8 | 10.97 | 11.14 | 11.31 | 11.48 | 11.64 | 11.81 | 11.98 | 12.15 | 12.31 | 12.48 | 12.65 | 12.82 | 12.99 | 13.15 | 13.32
5.38 | 6.61 | 6.77 | 6.94 | 7.11 | 7.28 | 7.44 | 7.61 | 7.78 | 7.95 | 8.12 | 8.28 | 8.45 | 8.62 | 8.79 | 8.96 | 9.12 | 9.29 | 9.46 | 9.63 | 9.79 | 9.96 | 10.13 | 10.3 | 10.47 | 10.63 | 10.8 | 10.97 | 11.14 | 11.3 | 11.47 | 11.64 | 11.81 | 11.98 | 12.14 | 12.31 | 12.48 | 12.65 | 12.81 | 12.98 | 13.15 | 13.32
5.39 | 6.6 | 6.77 | 6.94 | 7.11 | 7.28 | 7.44 | 7.61 | 7.78 | 7.95 | 8.11 | 8.28 | 8.45 | 8.62 | 8.79 | 8.95 | 9.12 | 9.29 | 9.46 | 9.62 | 9.79 | 9.96 | 10.13 | 10.3 | 10.46 | 10.63 | 10.8 | 10.97 | 11.13 | 11.3 | 11.47 | 11.64 | 11.8 | 11.97 | 12.14 | 12.31 | 12.48 | 12.64 | 12.81 | 12.98 | 13.15 | 13.31
5.4 | 6.6 | 6.77 | 6.94 | 7.11 | 7.27 | 7.44 | 7.61 | 7.78 | 7.94 | 8.11 | 8.28 | 8.45 | 8.62 | 8.78 | 8.95 | 9.12 | 9.29 | 9.45 | 9.62 | 9.79 | 9.96 | 10.12 | 10.29 | 10.46 | 10.63 | 10.8 | 10.96 | 11.13 | 11.3 | 11.47 | 11.63 | 11.8 | 11.97 | 12.14 | 12.3 | 12.47 | 12.64 | 12.81 | 12.98 | 13.14 | 13.31
5.41 | 6.6 | 6.77 | 6.94 | 7.1 | 7.27 | 7.44 | 7.61 | 7.77 | 7.94 | 8.11 | 8.28 | 8.45 | 8.61 | 8.78 | 8.95 | 9.12 | 9.28 | 9.45 | 9.62 | 9.79 | 9.95 | 10.12 | 10.29 | 10.46 | 10.63 | 10.79 | 10.96 | 11.13 | 11.3 | 11.46 | 11.63 | 11.8 | 11.97 | 12.13 | 12.3 | 12.47 | 12.64 | 12.8 | 12.97 | 13.14 | 13.31
5.42 | 6.6 | 6.77 | 6.93 | 7.1 | 7.27 | 7.44 | 7.6 | 7.77 | 7.94 | 8.11 | 8.28 | 8.44 | 8.61 | 8.78 | 8.95 | 9.11 | 9.28 | 9.45 | 9.62 | 9.78 | 9.95 | 10.12 | 10.29 | 10.45 | 10.62 | 10.79 | 10.96 | 11.13 | 11.29 | 11.46 | 11.63 | 11.8 | 11.96 | 12.13 | 12.3 | 12.47 | 12.63 | 12.8 | 12.97 | 13.14 | 13.3
5.43 | 6.6 | 6.76 | 6.93 | 7.1 | 7.27 | 7.44 | 7.6 | 7.77 | 7.94 | 8.11 | 8.27 | 8.44 | 8.61 | 8.78 | 8.94 | 9.11 | 9.28 | 9.45 | 9.61 | 9.78 | 9.95 | 10.12 | 10.28 | 10.45 | 10.62 | 10.79 | 10.95 | 11.12 | 11.29 | 11.46 | 11.63 | 11.79 | 11.96 | 12.13 | 12.3 | 12.46 | 12.63 | 12.8 | 12.97 | 13.13 | 13.3
5.44 | 6.6 | 6.76 | 6.93 | 7.1 | 7.27 | 7.43 | 7.6 | 7.77 | 7.94 | 8.1 | 8.27 | 8.44 | 8.61 | 8.77 | 8.94 | 9.11 | 9.28 | 9.44 | 9.61 | 9.78 | 9.95 | 10.11 | 10.28 | 10.45 | 10.62 | 10.78 | 10.95 | 11.12 | 11.29 | 11.45 | 11.62 | 11.79 | 11.96 | 12.12 | 12.29 | 12.46 | 12.63 | 12.79 | 12.96 | 13.13 | 13.3
5.45 | 6.59 | 6.76 | 6.93 | 7.1 | 7.26 | 7.43 | 7.6 | 7.77 | 7.93 | 8.1 | 8.27 | 8.44 | 8.6 | 8.77 | 8.94 | 9.11 | 9.27 | 9.44 | 9.61 | 9.78 | 9.94 | 10.11 | 10.28 | 10.45 | 10.61 | 10.78 | 10.95 | 11.12 | 11.28 | 11.45 | 11.62 | 11.79 | 11.95 | 12.12 | 12.29 | 12.46 | 12.62 | 12.79 | 12.96 | 13.13 | 13.29
5.46 | 6.59 | 6.76 | 6.93 | 7.09 | 7.26 | 7.43 | 7.6 | 7.76 | 7.93 | 8.1 | 8.27 | 8.43 | 8.6 | 8.77 | 8.94 | 9.1 | 9.27 | 9.44 | 9.61 | 9.77 | 9.94 | 10.11 | 10.28 | 10.44 | 10.61 | 10.78 | 10.95 | 11.11 | 11.28 | 11.45 | 11.62 | 11.78 | 11.95 | 12.12 | 12.29 | 12.45 | 12.62 | 12.79 | 12.96 | 13.12 | 13.29
5.47 | 6.59 | 6.76 | 6.92 | 7.09 | 7.26 | 7.43 | 7.59 | 7.76 | 7.93 | 8.1 | 8.26 | 8.43 | 8.6 | 8.77 | 8.93 | 9.1 | 9.27 | 9.44 | 9.6 | 9.77 | 9.94 | 10.11 | 10.27 | 10.44 | 10.61 | 10.78 | 10.94 | 11.11 | 11.28 | 11.45 | 11.61 | 11.78 | 11.95 | 12.12 | 12.28 | 12.45 | 12.62 | 12.78 | 12.95 | 13.12 | 13.29
5.48 | 6.59 | 6.76 | 6.92 | 7.09 | 7.26 | 7.43 | 7.59 | 7.76 | 7.93 | 8.1 | 8.26 | 8.43 | 8.6 | 8.76 | 8.93 | 9.1 | 9.27 | 9.43 | 9.6 | 9.77 | 9.94 | 10.1 | 10.27 | 10.44 | 10.61 | 10.77 | 10.94 | 11.11 | 11.28 | 11.44 | 11.61 | 11.78 | 11.94 | 12.11 | 12.28 | 12.45 | 12.61 | 12.78 | 12.95 | 13.12 | 13.28
5.49 | 6.59 | 6.75 | 6.92 | 7.09 | 7.26 | 7.42 | 7.59 | 7.76 | 7.93 | 8.09 | 8.26 | 8.43 | 8.6 | 8.76 | 8.93 | 9.1 | 9.26 | 9.43 | 9.6 | 9.77 | 9.93 | 10.1 | 10.27 | 10.44 | 10.6 | 10.77 | 10.94 | 11.11 | 11.27 | 11.44 | 11.61 | 11.77 | 11.94 | 12.11 | 12.28 | 12.44 | 12.61 | 12.78 | 12.95 | 13.11 | 13.28
5.5 | 6.58 | 6.75 | 6.92 | 7.09 | 7.25 | 7.42 | 7.59 | 7.76 | 7.92 | 8.09 | 8.26 | 8.43 | 8.59 | 8.76 | 8.93 | 9.1 | 9.26 | 9.43 | 9.6 | 9.76 | 9.93 | 10.1 | 10.27 | 10.43 | 10.6 | 10.77 | 10.94 | 11.1 | 11.27 | 11.44 | 11.6 | 11.77 | 11.94 | 12.11 | 12.27 | 12.44 | 12.61 | 12.78 | 12.94 | 13.11 | 13.28
5.51 | 6.58 | 6.75 | 6.92 | 7.09 | 7.25 | 7.42 | 7.59 | 7.75 | 7.92 | 8.09 | 8.26 | 8.42 | 8.59 | 8.76 | 8.93 | 9.09 | 9.26 | 9.43 | 9.59 | 9.76 | 9.93 | 10.1 | 10.26 | 10.43 | 10.6 | 10.77 | 10.93 | 11.1 | 11.27 | 11.43 | 11.6 | 11.77 | 11.94 | 12.1 | 12.27 | 12.44 | 12.6 | 12.77 | 12.94 | 13.11 | 13.27
5.52 | 6.58 | 6.75 | 6.92 | 7.08 | 7.25 | 7.42 | 7.59 | 7.75 | 7.92 | 8.09 | 8.25 | 8.42 | 8.59 | 8.76 | 8.92 | 9.09 | 9.26 | 9.42 | 9.59 | 9.76 | 9.93 | 10.09 | 10.26 | 10.43 | 10.6 | 10.76 | 10.93 | 11.1 | 11.26 | 11.43 | 11.6 | 11.77 | 11.93 | 12.1 | 12.27 | 12.43 | 12.6 | 12.77 | 12.94 | 13.1 | 13.27
5.53 | 6.58 | 6.75 | 6.91 | 7.08 | 7.25 | 7.42 | 7.58 | 7.75 | 7.92 | 8.08 | 8.25 | 8.42 | 8.59 | 8.75 | 8.92 | 9.09 | 9.26 | 9.42 | 9.59 | 9.76 | 9.92 | 10.09 | 10.26 | 10.43 | 10.59 | 10.76 | 10.93 | 11.09 | 11.26 | 11.43 | 11.6 | 11.76 | 11.93 | 12.1 | 12.26 | 12.43 | 12.6 | 12.77 | 12.93 | 13.1 | 13.27
5.54 | 6.58 | 6.75 | 6.91 | 7.08 | 7.25 | 7.41 | 7.58 | 7.75 | 7.92 | 8.08 | 8.25 | 8.42 | 8.58 | 8.75 | 8.92 | 9.09 | 9.25 | 9.42 | 9.59 | 9.75 | 9.92 | 10.09 | 10.26 | 10.42 | 10.59 | 10.76 | 10.92 | 11.09 | 11.26 | 11.43 | 11.59 | 11.76 | 11.93 | 12.09 | 12.26 | 12.43 | 12.59 | 12.76 | 12.93 | 13.1 | 13.26
5.55 | 6.58 | 6.74 | 6.91 | 7.08 | 7.24 | 7.41 | 7.58 | 7.75 | 7.91 | 8.08 | 8.25 | 8.41 | 8.58 | 8.75 | 8.92 | 9.08 | 9.25 | 9.42 | 9.58 | 9.75 | 9.92 | 10.09 | 10.25 | 10.42 | 10.59 | 10.75 | 10.92 | 11.09 | 11.26 | 11.42 | 11.59 | 11.76 | 11.92 | 12.09 | 12.26 | 12.42 | 12.59 | 12.76 | 12.93 | 13.09 | 13.26
5.56 | 6.57 | 6.74 | 6.91 | 7.08 | 7.24 | 7.41 | 7.58 | 7.74 | 7.91 | 8.08 | 8.25 | 8.41 | 8.58 | 8.75 | 8.91 | 9.08 | 9.25 | 9.42 | 9.58 | 9.75 | 9.92 | 10.08 | 10.25 | 10.42 | 10.58 | 10.75 | 10.92 | 11.09 | 11.25 | 11.42 | 11.59 | 11.75 | 11.92 | 12.09 | 12.25 | 12.42 | 12.59 | 12.76 | 12.92 | 13.09 | 13.26
5.57 | 6.57 | 6.74 | 6.91 | 7.07 | 7.24 | 7.41 | 7.58 | 7.74 | 7.91 | 8.08 | 8.24 | 8.41 | 8.58 | 8.74 | 8.91 | 9.08 | 9.25 | 9.41 | 9.58 | 9.75 | 9.91 | 10.08 | 10.25 | 10.41 | 10.58 | 10.75 | 10.92 | 11.08 | 11.25 | 11.42 | 11.58 | 11.75 | 11.92 | 12.08 | 12.25 | 12.42 | 12.59 | 12.75 | 12.92 | 13.09 | 13.25
5.58 | 6.57 | 6.74 | 6.91 | 7.07 | 7.24 | 7.41 | 7.57 | 7.74 | 7.91 | 8.07 | 8.24 | 8.41 | 8.58 | 8.74 | 8.91 | 9.08 | 9.24 | 9.41 | 9.58 | 9.74 | 9.91 | 10.08 | 10.25 | 10.41 | 10.58 | 10.75 | 10.91 | 11.08 | 11.25 | 11.41 | 11.58 | 11.75 | 11.91 | 12.08 | 12.25 | 12.42 | 12.58 | 12.75 | 12.92 | 13.08 | 13.25
5.59 | 6.57 | 6.74 | 6.9 | 7.07 | 7.24 | 7.4 | 7.57 | 7.74 | 7.91 | 8.07 | 8.24 | 8.41 | 8.57 | 8.74 | 8.91 | 9.07 | 9.24 | 9.41 | 9.57 | 9.74 | 9.91 | 10.08 | 10.24 | 10.41 | 10.58 | 10.74 | 10.91 | 11.08 | 11.24 | 11.41 | 11.58 | 11.74 | 11.91 | 12.08 | 12.25 | 12.41 | 12.58 | 12.75 | 12.91 | 13.08 | 13.25
5.6 | 6.57 | 6.73 | 6.9 | 7.07 | 7.24 | 7.4 | 7.57 | 7.74 | 7.9 | 8.07 | 8.24 | 8.4 | 8.57 | 8.74 | 8.9 | 9.07 | 9.24 | 9.41 | 9.57 | 9.74 | 9.91 | 10.07 | 10.24 | 10.41 | 10.57 | 10.74 | 10.91 | 11.07 | 11.24 | 11.41 | 11.57 | 11.74 | 11.91 | 12.08 | 12.24 | 12.41 | 12.58 | 12.74 | 12.91 | 13.08 | 13.24
5.61 | 6.57 | 6.73 | 6.9 | 7.07 | 7.23 | 7.4 | 7.57 | 7.73 | 7.9 | 8.07 | 8.23 | 8.4 | 8.57 | 8.74 | 8.9 | 9.07 | 9.24 | 9.4 | 9.57 | 9.74 | 9.9 | 10.07 | 10.24 | 10.4 | 10.57 | 10.74 | 10.9 | 11.07 | 11.24 | 11.4 | 11.57 | 11.74 | 11.91 | 12.07 | 12.24 | 12.41 | 12.57 | 12.74 | 12.91 | 13.07 | 13.24
5.62 | 6.56 | 6.73 | 6.9 | 7.06 | 7.23 | 7.4 | 7.57 | 7.73 | 7.9 | 8.07 | 8.23 | 8.4 | 8.57 | 8.73 | 8.9 | 9.07 | 9.23 | 9.4 | 9.57 | 9.73 | 9.9 | 10.07 | 10.23 | 10.4 | 10.57 | 10.73 | 10.9 | 11.07 | 11.24 | 11.4 | 11.57 | 11.74 | 11.9 | 12.07 | 12.24 | 12.4 | 12.57 | 12.74 | 12.9 | 13.07 | 13.24
5.63 | 6.56 | 6.73 | 6.9 | 7.06 | 7.23 | 7.4 | 7.56 | 7.73 | 7.9 | 8.06 | 8.23 | 8.4 | 8.56 | 8.73 | 8.9 | 9.06 | 9.23 | 9.4 | 9.56 | 9.73 | 9.9 | 10.07 | 10.23 | 10.4 | 10.57 | 10.73 | 10.9 | 11.07 | 11.23 | 11.4 | 11.57 | 11.73 | 11.9 | 12.07 | 12.23 | 12.4 | 12.57 | 12.73 | 12.9 | 13.07 | 13.23
5.64 | 6.56 | 6.73 | 6.89 | 7.06 | 7.23 | 7.39 | 7.56 | 7.73 | 7.89 | 8.06 | 8.23 | 8.4 | 8.56 | 8.73 | 8.9 | 9.06 | 9.23 | 9.4 | 9.56 | 9.73 | 9.9 | 10.06 | 10.23 | 10.4 | 10.56 | 10.73 | 10.9 | 11.06 | 11.23 | 11.4 | 11.56 | 11.73 | 11.9 | 12.06 | 12.23 | 12.4 | 12.56 | 12.73 | 12.9 | 13.06 | 13.23
5.65 | 6.56 | 6.73 | 6.89 | 7.06 | 7.23 | 7.39 | 7.56 | 7.73 | 7.89 | 8.06 | 8.23 | 8.39 | 8.56 | 8.73 | 8.89 | 9.06 | 9.23 | 9.39 | 9.56 | 9.73 | 9.89 | 10.06 | 10.23 | 10.39 | 10.56 | 10.73 | 10.89 | 11.06 | 11.23 | 11.39 | 11.56 | 11.73 | 11.89 | 12.06 | 12.23 | 12.39 | 12.56 | 12.73 | 12.89 | 13.06 | 13.23
5.66 | 6.56 | 6.72 | 6.89 | 7.06 | 7.22 | 7.39 | 7.56 | 7.72 | 7.89 | 8.06 | 8.22 | 8.39 | 8.56 | 8.72 | 8.89 | 9.06 | 9.22 | 9.39 | 9.56 | 9.72 | 9.89 | 10.06 | 10.22 | 10.39 | 10.56 | 10.72 | 10.89 | 11.06 | 11.22 | 11.39 | 11.56 | 11.72 | 11.89 | 12.06 | 12.22 | 12.39 | 12.56 | 12.72 | 12.89 | 13.06 | 13.22
5.67 | 6.56 | 6.72 | 6.89 | 7.06 | 7.22 | 7.39 | 7.56 | 7.72 | 7.89 | 8.06 | 8.22 | 8.39 | 8.56 | 8.72 | 8.89 | 9.06 | 9.22 | 9.39 | 9.56 | 9.72 | 9.89 | 10.05 | 10.22 | 10.39 | 10.55 | 10.72 | 10.89 | 11.05 | 11.22 | 11.39 | 11.55 | 11.72 | 11.89 | 12.05 | 12.22 | 12.39 | 12.55 | 12.72 | 12.89 | 13.05 | 13.22
5.68 | 6.55 | 6.72 | 6.89 | 7.05 | 7.22 | 7.39 | 7.55 | 7.72 | 7.89 | 8.05 | 8.22 | 8.39 | 8.55 | 8.72 | 8.89 | 9.05 | 9.22 | 9.39 | 9.55 | 9.72 | 9.89 | 10.05 | 10.22 | 10.39 | 10.55 | 10.72 | 10.88 | 11.05 | 11.22 | 11.38 | 11.55 | 11.72 | 11.88 | 12.05 | 12.22 | 12.38 | 12.55 | 12.72 | 12.88 | 13.05 | 13.22
5.69 | 6.55 | 6.72 | 6.89 | 7.05 | 7.22 | 7.39 | 7.55 | 7.72 | 7.88 | 8.05 | 8.22 | 8.38 | 8.55 | 8.72 | 8.88 | 9.05 | 9.22 | 9.38 | 9.55 | 9.72 | 9.88 | 10.05 | 10.22 | 10.38 | 10.55 | 10.72 | 10.88 | 11.05 | 11.22 | 11.38 | 11.55 | 11.71 | 11.88 | 12.05 | 12.21 | 12.38 | 12.55 | 12.71 | 12.88 | 13.05 | 13.21
5.7 | 6.55 | 6.72 | 6.88 | 7.05 | 7.22 | 7.38 | 7.55 | 7.72 | 7.88 | 8.05 | 8.22 | 8.38 | 8.55 | 8.72 | 8.88 | 9.05 | 9.21 | 9.38 | 9.55 | 9.71 | 9.88 | 10.05 | 10.21 | 10.38 | 10.55 | 10.71 | 10.88 | 11.05 | 11.21 | 11.38 | 11.55 | 11.71 | 11.88 | 12.04 | 12.21 | 12.38 | 12.54 | 12.71 | 12.88 | 13.04 | 13.21
5.71 | 6.55 | 6.72 | 6.88 | 7.05 | 7.21 | 7.38 | 7.55 | 7.71 | 7.88 | 8.05 | 8.21 | 8.38 | 8.55 | 8.71 | 8.88 | 9.05 | 9.21 | 9.38 | 9.55 | 9.71 | 9.88 | 10.04 | 10.21 | 10.38 | 10.54 | 10.71 | 10.88 | 11.04 | 11.21 | 11.38 | 11.54 | 11.71 | 11.87 | 12.04 | 12.21 | 12.37 | 12.54 | 12.71 | 12.87 | 13.04 | 13.21
5.72 | 6.55 | 6.71 | 6.88 | 7.05 | 7.21 | 7.38 | 7.55 | 7.71 | 7.88 | 8.05 | 8.21 | 8.38 | 8.54 | 8.71 | 8.88 | 9.04 | 9.21 | 9.38 | 9.54 | 9.71 | 9.88 | 10.04 | 10.21 | 10.37 | 10.54 | 10.71 | 10.87 | 11.04 | 11.21 | 11.37 | 11.54 | 11.71 | 11.87 | 12.04 | 12.2 | 12.37 | 12.54 | 12.7 | 12.87 | 13.04 | 13.2
5.73 | 6.55 | 6.71 | 6.88 | 7.04 | 7.21 | 7.38 | 7.54 | 7.71 | 7.88 | 8.04 | 8.21 | 8.38 | 8.54 | 8.71 | 8.87 | 9.04 | 9.21 | 9.37 | 9.54 | 9.71 | 9.87 | 10.04 | 10.21 | 10.37 | 10.54 | 10.7 | 10.87 | 11.04 | 11.2 | 11.37 | 11.54 | 11.7 | 11.87 | 12.04 | 12.2 | 12.37 | 12.53 | 12.7 | 12.87 | 13.03 | 13.2
5.74 | 6.54 | 6.71 | 6.88 | 7.04 | 7.21 | 7.38 | 7.54 | 7.71 | 7.87 | 8.04 | 8.21 | 8.37 | 8.54 | 8.71 | 8.87 | 9.04 | 9.21 | 9.37 | 9.54 | 9.7 | 9.87 | 10.04 | 10.2 | 10.37 | 10.54 | 10.7 | 10.87 | 11.03 | 11.2 | 11.37 | 11.53 | 11.7 | 11.87 | 12.03 | 12.2 | 12.36 | 12.53 | 12.7 | 12.86 | 13.03 | 13.2
5.75 | 6.54 | 6.71 | 6.87 | 7.04 | 7.21 | 7.37 | 7.54 | 7.71 | 7.87 | 8.04 | 8.21 | 8.37 | 8.54 | 8.7 | 8.87 | 9.04 | 9.2 | 9.37 | 9.54 | 9.7 | 9.87 | 10.03 | 10.2 | 10.37 | 10.53 | 10.7 | 10.87 | 11.03 | 11.2 | 11.36 | 11.53 | 11.7 | 11.86 | 12.03 | 12.2 | 12.36 | 12.53 | 12.69 | 12.86 | 13.03 | 13.19
5.76 | 6.54 | 6.71 | 6.87 | 7.04 | 7.21 | 7.37 | 7.54 | 7.7 | 7.87 | 8.04 | 8.2 | 8.37 | 8.54 | 8.7 | 8.87 | 9.03 | 9.2 | 9.37 | 9.53 | 9.7 | 9.87 | 10.03 | 10.2 | 10.36 | 10.53 | 10.7 | 10.86 | 11.03 | 11.2 | 11.36 | 11.53 | 11.69 | 11.86 | 12.03 | 12.19 | 12.36 | 12.52 | 12.69 | 12.86 | 13.02 | 13.19
5.77 | 6.54 | 6.7 | 6.87 | 7.04 | 7.2 | 7.37 | 7.54 | 7.7 | 7.87 | 8.03 | 8.2 | 8.37 | 8.53 | 8.7 | 8.87 | 9.03 | 9.2 | 9.36 | 9.53 | 9.7 | 9.86 | 10.03 | 10.2 | 10.36 | 10.53 | 10.69 | 10.86 | 11.03 | 11.19 | 11.36 | 11.52 | 11.69 | 11.86 | 12.02 | 12.19 | 12.36 | 12.52 | 12.69 | 12.85 | 13.02 | 13.19
5.78 | 6.54 | 6.7 | 6.87 | 7.04 | 7.2 | 7.37 | 7.53 | 7.7 | 7.87 | 8.03 | 8.2 | 8.36 | 8.53 | 8.7 | 8.86 | 9.03 | 9.2 | 9.36 | 9.53 | 9.69 | 9.86 | 10.03 | 10.19 | 10.36 | 10.52 | 10.69 | 10.86 | 11.02 | 11.19 | 11.36 | 11.52 | 11.69 | 11.85 | 12.02 | 12.19 | 12.35 | 12.52 | 12.68 | 12.85 | 13.02 | 13.18
5.79 | 6.54 | 6.7 | 6.87 | 7.03 | 7.2 | 7.37 | 7.53 | 7.7 | 7.86 | 8.03 | 8.2 | 8.36 | 8.53 | 8.7 | 8.86 | 9.03 | 9.19 | 9.36 | 9.53 | 9.69 | 9.86 | 10.02 | 10.19 | 10.36 | 10.52 | 10.69 | 10.85 | 11.02 | 11.19 | 11.35 | 11.52 | 11.68 | 11.85 | 12.02 | 12.18 | 12.35 | 12.51 | 12.68 | 12.85 | 13.01 | 13.18
5.8 | 6.53 | 6.7 | 6.87 | 7.03 | 7.2 | 7.36 | 7.53 | 7.7 | 7.86 | 8.03 | 8.19 | 8.36 | 8.53 | 8.69 | 8.86 | 9.02 | 9.19 | 9.36 | 9.52 | 9.69 | 9.86 | 10.02 | 10.19 | 10.35 | 10.52 | 10.69 | 10.85 | 11.02 | 11.18 | 11.35 | 11.52 | 11.68 | 11.85 | 12.01 | 12.18 | 12.35 | 12.51 | 12.68 | 12.84 | 13.01 | 13.18
5.81 | 6.53 | 6.7 | 6.86 | 7.03 | 7.2 | 7.36 | 7.53 | 7.69 | 7.86 | 8.03 | 8.19 | 8.36 | 8.52 | 8.69 | 8.86 | 9.02 | 9.19 | 9.35 | 9.52 | 9.69 | 9.85 | 10.02 | 10.18 | 10.35 | 10.52 | 10.68 | 10.85 | 11.01 | 11.18 | 11.35 | 11.51 | 11.68 | 11.84 | 12.01 | 12.18 | 12.34 | 12.51 | 12.67 | 12.84 | 13.01 | 13.17
5.82 | 6.53 | 6.7 | 6.86 | 7.03 | 7.19 | 7.36 | 7.53 | 7.69 | 7.86 | 8.02 | 8.19 | 8.36 | 8.52 | 8.69 | 8.85 | 9.02 | 9.19 | 9.35 | 9.52 | 9.68 | 9.85 | 10.02 | 10.18 | 10.35 | 10.51 | 10.68 | 10.85 | 11.01 | 11.18 | 11.34 | 11.51 | 11.68 | 11.84 | 12.01 | 12.17 | 12.34 | 12.51 | 12.67 | 12.84 | 13 | 13.17
5.83 | 6.53 | 6.69 | 6.86 | 7.03 | 7.19 | 7.36 | 7.52 | 7.69 | 7.86 | 8.02 | 8.19 | 8.35 | 8.52 | 8.69 | 8.85 | 9.02 | 9.18 | 9.35 | 9.52 | 9.68 | 9.85 | 10.01 | 10.18 | 10.35 | 10.51 | 10.68 | 10.84 | 11.01 | 11.18 | 11.34 | 11.51 | 11.67 | 11.84 | 12 | 12.17 | 12.34 | 12.5 | 12.67 | 12.83 | 13 | 13.17
5.84 | 6.53 | 6.69 | 6.86 | 7.02 | 7.19 | 7.36 | 7.52 | 7.69 | 7.85 | 8.02 | 8.19 | 8.35 | 8.52 | 8.68 | 8.85 | 9.02 | 9.18 | 9.35 | 9.51 | 9.68 | 9.85 | 10.01 | 10.18 | 10.34 | 10.51 | 10.67 | 10.84 | 11.01 | 11.17 | 11.34 | 11.5 | 11.67 | 11.84 | 12 | 12.17 | 12.33 | 12.5 | 12.66 | 12.83 | 13 | 13.16
5.85 | 6.52 | 6.69 | 6.86 | 7.02 | 7.19 | 7.35 | 7.52 | 7.69 | 7.85 | 8.02 | 8.18 | 8.35 | 8.52 | 8.68 | 8.85 | 9.01 | 9.18 | 9.35 | 9.51 | 9.68 | 9.84 | 10.01 | 10.17 | 10.34 | 10.51 | 10.67 | 10.84 | 11 | 11.17 | 11.34 | 11.5 | 11.67 | 11.83 | 12 | 12.16 | 12.33 | 12.5 | 12.66 | 12.83 | 12.99 | 13.16
5.86 | 6.52 | 6.69 | 6.85 | 7.02 | 7.19 | 7.35 | 7.52 | 7.68 | 7.85 | 8.02 | 8.18 | 8.35 | 8.51 | 8.68 | 8.85 | 9.01 | 9.18 | 9.34 | 9.51 | 9.67 | 9.84 | 10.01 | 10.17 | 10.34 | 10.5 | 10.67 | 10.83 | 11 | 11.17 | 11.33 | 11.5 | 11.66 | 11.83 | 12 | 12.16 | 12.33 | 12.49 | 12.66 | 12.82 | 12.99 | 13.16
5.87 | 6.52 | 6.69 | 6.85 | 7.02 | 7.18 | 7.35 | 7.52 | 7.68 | 7.85 | 8.01 | 8.18 | 8.35 | 8.51 | 8.68 | 8.84 | 9.01 | 9.17 | 9.34 | 9.51 | 9.67 | 9.84 | 10 | 10.17 | 10.33 | 10.5 | 10.67 | 10.83 | 11 | 11.16 | 11.33 | 11.5 | 11.66 | 11.83 | 11.99 | 12.16 | 12.32 | 12.49 | 12.66 | 12.82 | 12.99 | 13.15
5.88 | 6.52 | 6.69 | 6.85 | 7.02 | 7.18 | 7.35 | 7.51 | 7.68 | 7.85 | 8.01 | 8.18 | 8.34 | 8.51 | 8.67 | 8.84 | 9.01 | 9.17 | 9.34 | 9.5 | 9.67 | 9.84 | 10 | 10.17 | 10.33 | 10.5 | 10.66 | 10.83 | 11 | 11.16 | 11.33 | 11.49 | 11.66 | 11.82 | 11.99 | 12.15 | 12.32 | 12.49 | 12.65 | 12.82 | 12.98 | 13.15
5.89 | 6.52 | 6.68 | 6.85 | 7.02 | 7.18 | 7.35 | 7.51 | 7.68 | 7.84 | 8.01 | 8.18 | 8.34 | 8.51 | 8.67 | 8.84 | 9 | 9.17 | 9.34 | 9.5 | 9.67 | 9.83 | 10 | 10.16 | 10.33 | 10.5 | 10.66 | 10.83 | 10.99 | 11.16 | 11.32 | 11.49 | 11.65 | 11.82 | 11.99 | 12.15 | 12.32 | 12.48 | 12.65 | 12.81 | 12.98 | 13.15
5.9 | 6.52 | 6.68 | 6.85 | 7.01 | 7.18 | 7.34 | 7.51 | 7.68 | 7.84 | 8.01 | 8.17 | 8.34 | 8.5 | 8.67 | 8.84 | 9 | 9.17 | 9.33 | 9.5 | 9.66 | 9.83 | 10 | 10.16 | 10.33 | 10.49 | 10.66 | 10.82 | 10.99 | 11.16 | 11.32 | 11.49 | 11.65 | 11.82 | 11.98 | 12.15 | 12.31 | 12.48 | 12.65 | 12.81 | 12.98 | 13.14
5.91 | 6.51 | 6.68 | 6.85 | 7.01 | 7.18 | 7.34 | 7.51 | 7.67 | 7.84 | 8.01 | 8.17 | 8.34 | 8.5 | 8.67 | 8.83 | 9 | 9.17 | 9.33 | 9.5 | 9.66 | 9.83 | 9.99 | 10.16 | 10.32 | 10.49 | 10.66 | 10.82 | 10.99 | 11.15 | 11.32 | 11.48 | 11.65 | 11.81 | 11.98 | 12.15 | 12.31 | 12.48 | 12.64 | 12.81 | 12.97 | 13.14
5.92 | 6.51 | 6.68 | 6.84 | 7.01 | 7.18 | 7.34 | 7.51 | 7.67 | 7.84 | 8 | 8.17 | 8.33 | 8.5 | 8.67 | 8.83 | 9 | 9.16 | 9.33 | 9.49 | 9.66 | 9.83 | 9.99 | 10.16 | 10.32 | 10.49 | 10.65 | 10.82 | 10.98 | 11.15 | 11.31 | 11.48 | 11.65 | 11.81 | 11.98 | 12.14 | 12.31 | 12.47 | 12.64 | 12.8 | 12.97 | 13.14
5.93 | 6.51 | 6.68 | 6.84 | 7.01 | 7.17 | 7.34 | 7.5 | 7.67 | 7.84 | 8 | 8.17 | 8.33 | 8.5 | 8.66 | 8.83 | 8.99 | 9.16 | 9.33 | 9.49 | 9.66 | 9.82 | 9.99 | 10.15 | 10.32 | 10.48 | 10.65 | 10.82 | 10.98 | 11.15 | 11.31 | 11.48 | 11.64 | 11.81 | 11.97 | 12.14 | 12.3 | 12.47 | 12.64 | 12.8 | 12.97 | 13.13
5.94 | 6.51 | 6.68 | 6.84 | 7.01 | 7.17 | 7.34 | 7.5 | 7.67 | 7.83 | 8 | 8.17 | 8.33 | 8.5 | 8.66 | 8.83 | 8.99 | 9.16 | 9.32 | 9.49 | 9.65 | 9.82 | 9.99 | 10.15 | 10.32 | 10.48 | 10.65 | 10.81 | 10.98 | 11.14 | 11.31 | 11.47 | 11.64 | 11.81 | 11.97 | 12.14 | 12.3 | 12.47 | 12.63 | 12.8 | 12.96 | 13.13
5.95 | 6.51 | 6.67 | 6.84 | 7 | 7.17 | 7.34 | 7.5 | 7.67 | 7.83 | 8 | 8.16 | 8.33 | 8.49 | 8.66 | 8.82 | 8.99 | 9.16 | 9.32 | 9.49 | 9.65 | 9.82 | 9.98 | 10.15 | 10.31 | 10.48 | 10.64 | 10.81 | 10.98 | 11.14 | 11.31 | 11.47 | 11.64 | 11.8 | 11.97 | 12.13 | 12.3 | 12.46 | 12.63 | 12.79 | 12.96 | 13.13
5.96 | 6.51 | 6.67 | 6.84 | 7 | 7.17 | 7.33 | 7.5 | 7.66 | 7.83 | 8 | 8.16 | 8.33 | 8.49 | 8.66 | 8.82 | 8.99 | 9.15 | 9.32 | 9.48 | 9.65 | 9.82 | 9.98 | 10.15 | 10.31 | 10.48 | 10.64 | 10.81 | 10.97 | 11.14 | 11.3 | 11.47 | 11.63 | 11.8 | 11.96 | 12.13 | 12.3 | 12.46 | 12.63 | 12.79 | 12.96 | 13.12
5.97 | 6.5 | 6.67 | 6.84 | 7 | 7.17 | 7.33 | 7.5 | 7.66 | 7.83 | 7.99 | 8.16 | 8.32 | 8.49 | 8.65 | 8.82 | 8.99 | 9.15 | 9.32 | 9.48 | 9.65 | 9.81 | 9.98 | 10.14 | 10.31 | 10.47 | 10.64 | 10.8 | 10.97 | 11.14 | 11.3 | 11.47 | 11.63 | 11.8 | 11.96 | 12.13 | 12.29 | 12.46 | 12.62 | 12.79 | 12.95 | 13.12
5.98 | 6.5 | 6.67 | 6.83 | 7 | 7.16 | 7.33 | 7.5 | 7.66 | 7.83 | 7.99 | 8.16 | 8.32 | 8.49 | 8.65 | 8.82 | 8.98 | 9.15 | 9.31 | 9.48 | 9.64 | 9.81 | 9.98 | 10.14 | 10.31 | 10.47 | 10.64 | 10.8 | 10.97 | 11.13 | 11.3 | 11.46 | 11.63 | 11.79 | 11.96 | 12.12 | 12.29 | 12.45 | 12.62 | 12.79 | 12.95 | 13.12
5.99 | 6.5 | 6.67 | 6.83 | 7 | 7.16 | 7.33 | 7.49 | 7.66 | 7.82 | 7.99 | 8.15 | 8.32 | 8.49 | 8.65 | 8.82 | 8.98 | 9.15 | 9.31 | 9.48 | 9.64 | 9.81 | 9.97 | 10.14 | 10.3 | 10.47 | 10.63 | 10.8 | 10.96 | 11.13 | 11.29 | 11.46 | 11.63 | 11.79 | 11.96 | 12.12 | 12.29 | 12.45 | 12.62 | 12.78 | 12.95 | 13.11
6 | 6.5 | 6.66 | 6.83 | 7 | 7.16 | 7.33 | 7.49 | 7.66 | 7.82 | 7.99 | 8.15 | 8.32 | 8.48 | 8.65 | 8.81 | 8.98 | 9.14 | 9.31 | 9.47 | 9.64 | 9.8 | 9.97 | 10.14 | 10.3 | 10.47 | 10.63 | 10.8 | 10.96 | 11.13 | 11.29 | 11.46 | 11.62 | 11.79 | 11.95 | 12.12 | 12.28 | 12.45 | 12.61 | 12.78 | 12.94 | 13.11
6.01 | 6.5 | 6.66 | 6.83 | 6.99 | 7.16 | 7.32 | 7.49 | 7.65 | 7.82 | 7.99 | 8.15 | 8.32 | 8.48 | 8.65 | 8.81 | 8.98 | 9.14 | 9.31 | 9.47 | 9.64 | 9.8 | 9.97 | 10.13 | 10.3 | 10.46 | 10.63 | 10.79 | 10.96 | 11.12 | 11.29 | 11.45 | 11.62 | 11.78 | 11.95 | 12.11 | 12.28 | 12.45 | 12.61 | 12.78 | 12.94 | 13.11
6.02 | 6.5 | 6.66 | 6.83 | 6.99 | 7.16 | 7.32 | 7.49 | 7.65 | 7.82 | 7.98 | 8.15 | 8.31 | 8.48 | 8.64 | 8.81 | 8.97 | 9.14 | 9.3 | 9.47 | 9.63 | 9.8 | 9.97 | 10.13 | 10.3 | 10.46 | 10.63 | 10.79 | 10.96 | 11.12 | 11.29 | 11.45 | 11.62 | 11.78 | 11.95 | 12.11 | 12.28 | 12.44 | 12.61 | 12.77 | 12.94 | 13.1
6.03 | 6.49 | 6.66 | 6.82 | 6.99 | 7.16 | 7.32 | 7.49 | 7.65 | 7.82 | 7.98 | 8.15 | 8.31 | 8.48 | 8.64 | 8.81 | 8.97 | 9.14 | 9.3 | 9.47 | 9.63 | 9.8 | 9.96 | 10.13 | 10.29 | 10.46 | 10.62 | 10.79 | 10.95 | 11.12 | 11.28 | 11.45 | 11.61 | 11.78 | 11.94 | 12.11 | 12.27 | 12.44 | 12.6 | 12.77 | 12.93 | 13.1
6.04 | 6.49 | 6.66 | 6.82 | 6.99 | 7.15 | 7.32 | 7.48 | 7.65 | 7.81 | 7.98 | 8.14 | 8.31 | 8.47 | 8.64 | 8.8 | 8.97 | 9.13 | 9.3 | 9.46 | 9.63 | 9.79 | 9.96 | 10.13 | 10.29 | 10.46 | 10.62 | 10.79 | 10.95 | 11.12 | 11.28 | 11.45 | 11.61 | 11.78 | 11.94 | 12.11 | 12.27 | 12.44 | 12.6 | 12.77 | 12.93 | 13.1
6.05 | 6.49 | 6.66 | 6.82 | 6.99 | 7.15 | 7.32 | 7.48 | 7.65 | 7.81 | 7.98 | 8.14 | 8.31 | 8.47 | 8.64 | 8.8 | 8.97 | 9.13 | 9.3 | 9.46 | 9.63 | 9.79 | 9.96 | 10.12 | 10.29 | 10.45 | 10.62 | 10.78 | 10.95 | 11.11 | 11.28 | 11.44 | 11.61 | 11.77 | 11.94 | 12.1 | 12.27 | 12.43 | 12.6 | 12.76 | 12.93 | 13.09
6.06 | 6.49 | 6.65 | 6.82 | 6.98 | 7.15 | 7.31 | 7.48 | 7.64 | 7.81 | 7.97 | 8.14 | 8.3 | 8.47 | 8.63 | 8.8 | 8.96 | 9.13 | 9.29 | 9.46 | 9.62 | 9.79 | 9.95 | 10.12 | 10.28 | 10.45 | 10.61 | 10.78 | 10.94 | 11.11 | 11.27 | 11.44 | 11.6 | 11.77 | 11.93 | 12.1 | 12.26 | 12.43 | 12.59 | 12.76 | 12.92 | 13.09
6.07 | 6.49 | 6.65 | 6.82 | 6.98 | 7.15 | 7.31 | 7.48 | 7.64 | 7.81 | 7.97 | 8.14 | 8.3 | 8.47 | 8.63 | 8.8 | 8.96 | 9.13 | 9.29 | 9.46 | 9.62 | 9.79 | 9.95 | 10.12 | 10.28 | 10.45 | 10.61 | 10.78 | 10.94 | 11.11 | 11.27 | 11.44 | 11.6 | 11.77 | 11.93 | 12.1 | 12.26 | 12.43 | 12.59 | 12.76 | 12.92 | 13.09
6.08 | 6.49 | 6.65 | 6.82 | 6.98 | 7.15 | 7.31 | 7.48 | 7.64 | 7.81 | 7.97 | 8.14 | 8.3 | 8.47 | 8.63 | 8.8 | 8.96 | 9.13 | 9.29 | 9.46 | 9.62 | 9.78 | 9.95 | 10.11 | 10.28 | 10.44 | 10.61 | 10.77 | 10.94 | 11.1 | 11.27 | 11.43 | 11.6 | 11.76 | 11.93 | 12.09 | 12.26 | 12.42 | 12.59 | 12.75 | 12.92 | 13.08
6.09 | 6.48 | 6.65 | 6.81 | 6.98 | 7.14 | 7.31 | 7.47 | 7.64 | 7.8 | 7.97 | 8.13 | 8.3 | 8.46 | 8.63 | 8.79 | 8.96 | 9.12 | 9.29 | 9.45 | 9.62 | 9.78 | 9.95 | 10.11 | 10.28 | 10.44 | 10.61 | 10.77 | 10.94 | 11.1 | 11.27 | 11.43 | 11.6 | 11.76 | 11.93 | 12.09 | 12.26 | 12.42 | 12.58 | 12.75 | 12.91 | 13.08
6.1 | 6.48 | 6.65 | 6.81 | 6.98 | 7.14 | 7.31 | 7.47 | 7.64 | 7.8 | 7.97 | 8.13 | 8.3 | 8.46 | 8.63 | 8.79 | 8.96 | 9.12 | 9.29 | 9.45 | 9.62 | 9.78 | 9.94 | 10.11 | 10.27 | 10.44 | 10.6 | 10.77 | 10.93 | 11.1 | 11.26 | 11.43 | 11.59 | 11.76 | 11.92 | 12.09 | 12.25 | 12.42 | 12.58 | 12.75 | 12.91 | 13.08
6.11 | 6.48 | 6.65 | 6.81 | 6.98 | 7.14 | 7.31 | 7.47 | 7.63 | 7.8 | 7.96 | 8.13 | 8.29 | 8.46 | 8.62 | 8.79 | 8.95 | 9.12 | 9.28 | 9.45 | 9.61 | 9.78 | 9.94 | 10.11 | 10.27 | 10.44 | 10.6 | 10.77 | 10.93 | 11.1 | 11.26 | 11.43 | 11.59 | 11.75 | 11.92 | 12.08 | 12.25 | 12.41 | 12.58 | 12.74 | 12.91 | 13.07
6.12 | 6.48 | 6.64 | 6.81 | 6.97 | 7.14 | 7.3 | 7.47 | 7.63 | 7.8 | 7.96 | 8.13 | 8.29 | 8.46 | 8.62 | 8.79 | 8.95 | 9.12 | 9.28 | 9.45 | 9.61 | 9.78 | 9.94 | 10.1 | 10.27 | 10.43 | 10.6 | 10.76 | 10.93 | 11.09 | 11.26 | 11.42 | 11.59 | 11.75 | 11.92 | 12.08 | 12.25 | 12.41 | 12.58 | 12.74 | 12.9 | 13.07
6.13 | 6.48 | 6.64 | 6.81 | 6.97 | 7.14 | 7.3 | 7.47 | 7.63 | 7.8 | 7.96 | 8.13 | 8.29 | 8.45 | 8.62 | 8.78 | 8.95 | 9.11 | 9.28 | 9.44 | 9.61 | 9.77 | 9.94 | 10.1 | 10.27 | 10.43 | 10.6 | 10.76 | 10.93 | 11.09 | 11.25 | 11.42 | 11.58 | 11.75 | 11.91 | 12.08 | 12.24 | 12.41 | 12.57 | 12.74 | 12.9 | 13.07
6.14 | 6.48 | 6.64 | 6.81 | 6.97 | 7.13 | 7.3 | 7.46 | 7.63 | 7.79 | 7.96 | 8.12 | 8.29 | 8.45 | 8.62 | 8.78 | 8.95 | 9.11 | 9.28 | 9.44 | 9.61 | 9.77 | 9.93 | 10.1 | 10.26 | 10.43 | 10.59 | 10.76 | 10.92 | 11.09 | 11.25 | 11.42 | 11.58 | 11.75 | 11.91 | 12.07 | 12.24 | 12.4 | 12.57 | 12.73 | 12.9 | 13.06
6.15 | 6.47 | 6.64 | 6.8 | 6.97 | 7.13 | 7.3 | 7.46 | 7.63 | 7.79 | 7.96 | 8.12 | 8.29 | 8.45 | 8.62 | 8.78 | 8.94 | 9.11 | 9.27 | 9.44 | 9.6 | 9.77 | 9.93 | 10.1 | 10.26 | 10.43 | 10.59 | 10.76 | 10.92 | 11.08 | 11.25 | 11.41 | 11.58 | 11.74 | 11.91 | 12.07 | 12.24 | 12.4 | 12.57 | 12.73 | 12.89 | 13.06
6.16 | 6.47 | 6.64 | 6.8 | 6.97 | 7.13 | 7.3 | 7.46 | 7.63 | 7.79 | 7.95 | 8.12 | 8.28 | 8.45 | 8.61 | 8.78 | 8.94 | 9.11 | 9.27 | 9.44 | 9.6 | 9.77 | 9.93 | 10.09 | 10.26 | 10.42 | 10.59 | 10.75 | 10.92 | 11.08 | 11.25 | 11.41 | 11.58 | 11.74 | 11.9 | 12.07 | 12.23 | 12.4 | 12.56 | 12.73 | 12.89 | 13.06
6.17 | 6.47 | 6.64 | 6.8 | 6.96 | 7.13 | 7.29 | 7.46 | 7.62 | 7.79 | 7.95 | 8.12 | 8.28 | 8.45 | 8.61 | 8.78 | 8.94 | 9.1 | 9.27 | 9.43 | 9.6 | 9.76 | 9.93 | 10.09 | 10.26 | 10.42 | 10.59 | 10.75 | 10.91 | 11.08 | 11.24 | 11.41 | 11.57 | 11.74 | 11.9 | 12.07 | 12.23 | 12.39 | 12.56 | 12.72 | 12.89 | 13.05
6.18 | 6.47 | 6.63 | 6.8 | 6.96 | 7.13 | 7.29 | 7.46 | 7.62 | 7.79 | 7.95 | 8.11 | 8.28 | 8.44 | 8.61 | 8.77 | 8.94 | 9.1 | 9.27 | 9.43 | 9.6 | 9.76 | 9.92 | 10.09 | 10.25 | 10.42 | 10.58 | 10.75 | 10.91 | 11.08 | 11.24 | 11.4 | 11.57 | 11.73 | 11.9 | 12.06 | 12.23 | 12.39 | 12.56 | 12.72 | 12.88 | 13.05
6.19 | 6.47 | 6.63 | 6.8 | 6.96 | 7.13 | 7.29 | 7.45 | 7.62 | 7.78 | 7.95 | 8.11 | 8.28 | 8.44 | 8.61 | 8.77 | 8.94 | 9.1 | 9.26 | 9.43 | 9.59 | 9.76 | 9.92 | 10.09 | 10.25 | 10.42 | 10.58 | 10.74 | 10.91 | 11.07 | 11.24 | 11.4 | 11.57 | 11.73 | 11.9 | 12.06 | 12.22 | 12.39 | 12.55 | 12.72 | 12.88 | 13.05
6.2 | 6.47 | 6.63 | 6.79 | 6.96 | 7.12 | 7.29 | 7.45 | 7.62 | 7.78 | 7.95 | 8.11 | 8.28 | 8.44 | 8.6 | 8.77 | 8.93 | 9.1 | 9.26 | 9.43 | 9.59 | 9.76 | 9.92 | 10.08 | 10.25 | 10.41 | 10.58 | 10.74 | 10.91 | 11.07 | 11.23 | 11.4 | 11.56 | 11.73 | 11.89 | 12.06 | 12.22 | 12.39 | 12.55 | 12.71 | 12.88 | 13.04
6.21 | 6.46 | 6.63 | 6.79 | 6.96 | 7.12 | 7.29 | 7.45 | 7.62 | 7.78 | 7.94 | 8.11 | 8.27 | 8.44 | 8.6 | 8.77 | 8.93 | 9.1 | 9.26 | 9.42 | 9.59 | 9.75 | 9.92 | 10.08 | 10.25 | 10.41 | 10.57 | 10.74 | 10.9 | 11.07 | 11.23 | 11.4 | 11.56 | 11.72 | 11.89 | 12.05 | 12.22 | 12.38 | 12.55 | 12.71 | 12.88 | 13.04
6.22 | 6.46 | 6.63 | 6.79 | 6.96 | 7.12 | 7.28 | 7.45 | 7.61 | 7.78 | 7.94 | 8.11 | 8.27 | 8.44 | 8.6 | 8.76 | 8.93 | 9.09 | 9.26 | 9.42 | 9.59 | 9.75 | 9.91 | 10.08 | 10.24 | 10.41 | 10.57 | 10.74 | 10.9 | 11.06 | 11.23 | 11.39 | 11.56 | 11.72 | 11.89 | 12.05 | 12.21 | 12.38 | 12.54 | 12.71 | 12.87 | 13.04
6.23 | 6.46 | 6.63 | 6.79 | 6.95 | 7.12 | 7.28 | 7.45 | 7.61 | 7.78 | 7.94 | 8.1 | 8.27 | 8.43 | 8.6 | 8.76 | 8.93 | 9.09 | 9.25 | 9.42 | 9.58 | 9.75 | 9.91 | 10.08 | 10.24 | 10.4 | 10.57 | 10.73 | 10.9 | 11.06 | 11.23 | 11.39 | 11.55 | 11.72 | 11.88 | 12.05 | 12.21 | 12.38 | 12.54 | 12.7 | 12.87 | 13.03
6.24 | 6.46 | 6.62 | 6.79 | 6.95 | 7.12 | 7.28 | 7.45 | 7.61 | 7.77 | 7.94 | 8.1 | 8.27 | 8.43 | 8.6 | 8.76 | 8.92 | 9.09 | 9.25 | 9.42 | 9.58 | 9.75 | 9.91 | 10.07 | 10.24 | 10.4 | 10.57 | 10.73 | 10.89 | 11.06 | 11.22 | 11.39 | 11.55 | 11.72 | 11.88 | 12.04 | 12.21 | 12.37 | 12.54 | 12.7 | 12.87 | 13.03
6.25 | 6.46 | 6.62 | 6.79 | 6.95 | 7.11 | 7.28 | 7.44 | 7.61 | 7.77 | 7.94 | 8.1 | 8.26 | 8.43 | 8.59 | 8.76 | 8.92 | 9.09 | 9.25 | 9.41 | 9.58 | 9.74 | 9.91 | 10.07 | 10.24 | 10.4 | 10.56 | 10.73 | 10.89 | 11.06 | 11.22 | 11.38 | 11.55 | 11.71 | 11.88 | 12.04 | 12.21 | 12.37 | 12.53 | 12.7 | 12.86 | 13.03
6.26 | 6.46 | 6.62 | 6.78 | 6.95 | 7.11 | 7.28 | 7.44 | 7.61 | 7.77 | 7.93 | 8.1 | 8.26 | 8.43 | 8.59 | 8.76 | 8.92 | 9.08 | 9.25 | 9.41 | 9.58 | 9.74 | 9.9 | 10.07 | 10.23 | 10.4 | 10.56 | 10.73 | 10.89 | 11.05 | 11.22 | 11.38 | 11.55 | 11.71 | 11.87 | 12.04 | 12.2 | 12.37 | 12.53 | 12.69 | 12.86 | 13.02
6.27 | 6.45 | 6.62 | 6.78 | 6.95 | 7.11 | 7.28 | 7.44 | 7.6 | 7.77 | 7.93 | 8.1 | 8.26 | 8.42 | 8.59 | 8.75 | 8.92 | 9.08 | 9.25 | 9.41 | 9.57 | 9.74 | 9.9 | 10.07 | 10.23 | 10.39 | 10.56 | 10.72 | 10.89 | 11.05 | 11.21 | 11.38 | 11.54 | 11.71 | 11.87 | 12.04 | 12.2 | 12.36 | 12.53 | 12.69 | 12.86 | 13.02
6.28 | 6.45 | 6.62 | 6.78 | 6.95 | 7.11 | 7.27 | 7.44 | 7.6 | 7.77 | 7.93 | 8.09 | 8.26 | 8.42 | 8.59 | 8.75 | 8.91 | 9.08 | 9.24 | 9.41 | 9.57 | 9.74 | 9.9 | 10.06 | 10.23 | 10.39 | 10.56 | 10.72 | 10.88 | 11.05 | 11.21 | 11.38 | 11.54 | 11.7 | 11.87 | 12.03 | 12.2 | 12.36 | 12.52 | 12.69 | 12.85 | 13.02
6.29 | 6.45 | 6.62 | 6.78 | 6.94 | 7.11 | 7.27 | 7.44 | 7.6 | 7.76 | 7.93 | 8.09 | 8.26 | 8.42 | 8.58 | 8.75 | 8.91 | 9.08 | 9.24 | 9.4 | 9.57 | 9.73 | 9.9 | 10.06 | 10.22 | 10.39 | 10.55 | 10.72 | 10.88 | 11.05 | 11.21 | 11.37 | 11.54 | 11.7 | 11.87 | 12.03 | 12.19 | 12.36 | 12.52 | 12.69 | 12.85 | 13.01
6.3 | 6.45 | 6.61 | 6.78 | 6.94 | 7.11 | 7.27 | 7.43 | 7.6 | 7.76 | 7.93 | 8.09 | 8.25 | 8.42 | 8.58 | 8.75 | 8.91 | 9.07 | 9.24 | 9.4 | 9.57 | 9.73 | 9.89 | 10.06 | 10.22 | 10.39 | 10.55 | 10.71 | 10.88 | 11.04 | 11.21 | 11.37 | 11.53 | 11.7 | 11.86 | 12.03 | 12.19 | 12.35 | 12.52 | 12.68 | 12.85 | 13.01
6.31 | 6.45 | 6.61 | 6.78 | 6.94 | 7.1 | 7.27 | 7.43 | 7.6 | 7.76 | 7.92 | 8.09 | 8.25 | 8.42 | 8.58 | 8.74 | 8.91 | 9.07 | 9.24 | 9.4 | 9.56 | 9.73 | 9.89 | 10.06 | 10.22 | 10.38 | 10.55 | 10.71 | 10.88 | 11.04 | 11.2 | 11.37 | 11.53 | 11.7 | 11.86 | 12.02 | 12.19 | 12.35 | 12.51 | 12.68 | 12.84 | 13.01
6.32 | 6.45 | 6.61 | 6.77 | 6.94 | 7.1 | 7.27 | 7.43 | 7.59 | 7.76 | 7.92 | 8.09 | 8.25 | 8.41 | 8.58 | 8.74 | 8.91 | 9.07 | 9.23 | 9.4 | 9.56 | 9.73 | 9.89 | 10.05 | 10.22 | 10.38 | 10.55 | 10.71 | 10.87 | 11.04 | 11.2 | 11.36 | 11.53 | 11.69 | 11.86 | 12.02 | 12.18 | 12.35 | 12.51 | 12.68 | 12.84 | 13
6.33 | 6.44 | 6.61 | 6.77 | 6.94 | 7.1 | 7.26 | 7.43 | 7.59 | 7.76 | 7.92 | 8.08 | 8.25 | 8.41 | 8.58 | 8.74 | 8.9 | 9.07 | 9.23 | 9.4 | 9.56 | 9.72 | 9.89 | 10.05 | 10.21 | 10.38 | 10.54 | 10.71 | 10.87 | 11.03 | 11.2 | 11.36 | 11.53 | 11.69 | 11.85 | 12.02 | 12.18 | 12.34 | 12.51 | 12.67 | 12.84 | 13
6.34 | 6.44 | 6.61 | 6.77 | 6.93 | 7.1 | 7.26 | 7.43 | 7.59 | 7.75 | 7.92 | 8.08 | 8.25 | 8.41 | 8.57 | 8.74 | 8.9 | 9.06 | 9.23 | 9.39 | 9.56 | 9.72 | 9.88 | 10.05 | 10.21 | 10.38 | 10.54 | 10.7 | 10.87 | 11.03 | 11.19 | 11.36 | 11.52 | 11.69 | 11.85 | 12.01 | 12.18 | 12.34 | 12.51 | 12.67 | 12.83 | 13
6.35 | 6.44 | 6.6 | 6.77 | 6.93 | 7.1 | 7.26 | 7.42 | 7.59 | 7.75 | 7.92 | 8.08 | 8.24 | 8.41 | 8.57 | 8.74 | 8.9 | 9.06 | 9.23 | 9.39 | 9.55 | 9.72 | 9.88 | 10.05 | 10.21 | 10.37 | 10.54 | 10.7 | 10.86 | 11.03 | 11.19 | 11.36 | 11.52 | 11.68 | 11.85 | 12.01 | 12.17 | 12.34 | 12.5 | 12.67 | 12.83 | 12.99
6.36 | 6.44 | 6.6 | 6.77 | 6.93 | 7.09 | 7.26 | 7.42 | 7.59 | 7.75 | 7.91 | 8.08 | 8.24 | 8.41 | 8.57 | 8.73 | 8.9 | 9.06 | 9.22 | 9.39 | 9.55 | 9.72 | 9.88 | 10.04 | 10.21 | 10.37 | 10.53 | 10.7 | 10.86 | 11.03 | 11.19 | 11.35 | 11.52 | 11.68 | 11.84 | 12.01 | 12.17 | 12.34 | 12.5 | 12.66 | 12.83 | 12.99
6.37 | 6.44 | 6.6 | 6.77 | 6.93 | 7.09 | 7.26 | 7.42 | 7.58 | 7.75 | 7.91 | 8.08 | 8.24 | 8.4 | 8.57 | 8.73 | 8.89 | 9.06 | 9.22 | 9.39 | 9.55 | 9.71 | 9.88 | 10.04 | 10.2 | 10.37 | 10.53 | 10.7 | 10.86 | 11.02 | 11.19 | 11.35 | 11.51 | 11.68 | 11.84 | 12 | 12.17 | 12.33 | 12.5 | 12.66 | 12.82 | 12.99
6.38 | 6.44 | 6.6 | 6.76 | 6.93 | 7.09 | 7.25 | 7.42 | 7.58 | 7.75 | 7.91 | 8.07 | 8.24 | 8.4 | 8.56 | 8.73 | 8.89 | 9.06 | 9.22 | 9.38 | 9.55 | 9.71 | 9.87 | 10.04 | 10.2 | 10.37 | 10.53 | 10.69 | 10.86 | 11.02 | 11.18 | 11.35 | 11.51 | 11.67 | 11.84 | 12 | 12.17 | 12.33 | 12.49 | 12.66 | 12.82 | 12.98
6.39 | 6.43 | 6.6 | 6.76 | 6.93 | 7.09 | 7.25 | 7.42 | 7.58 | 7.74 | 7.91 | 8.07 | 8.24 | 8.4 | 8.56 | 8.73 | 8.89 | 9.05 | 9.22 | 9.38 | 9.54 | 9.71 | 9.87 | 10.04 | 10.2 | 10.36 | 10.53 | 10.69 | 10.85 | 11.02 | 11.18 | 11.34 | 11.51 | 11.67 | 11.84 | 12 | 12.16 | 12.33 | 12.49 | 12.65 | 12.82 | 12.98
6.4 | 6.43 | 6.6 | 6.76 | 6.92 | 7.09 | 7.25 | 7.41 | 7.58 | 7.74 | 7.91 | 8.07 | 8.23 | 8.4 | 8.56 | 8.72 | 8.89 | 9.05 | 9.21 | 9.38 | 9.54 | 9.71 | 9.87 | 10.03 | 10.2 | 10.36 | 10.52 | 10.69 | 10.85 | 11.01 | 11.18 | 11.34 | 11.51 | 11.67 | 11.83 | 12 | 12.16 | 12.32 | 12.49 | 12.65 | 12.81 | 12.98
6.41 | 6.43 | 6.59 | 6.76 | 6.92 | 7.09 | 7.25 | 7.41 | 7.58 | 7.74 | 7.9 | 8.07 | 8.23 | 8.39 | 8.56 | 8.72 | 8.89 | 9.05 | 9.21 | 9.38 | 9.54 | 9.7 | 9.87 | 10.03 | 10.19 | 10.36 | 10.52 | 10.68 | 10.85 | 11.01 | 11.18 | 11.34 | 11.5 | 11.67 | 11.83 | 11.99 | 12.16 | 12.32 | 12.48 | 12.65 | 12.81 | 12.97
6.42 | 6.43 | 6.59 | 6.76 | 6.92 | 7.08 | 7.25 | 7.41 | 7.57 | 7.74 | 7.9 | 8.07 | 8.23 | 8.39 | 8.56 | 8.72 | 8.88 | 9.05 | 9.21 | 9.37 | 9.54 | 9.7 | 9.86 | 10.03 | 10.19 | 10.35 | 10.52 | 10.68 | 10.85 | 11.01 | 11.17 | 11.34 | 11.5 | 11.66 | 11.83 | 11.99 | 12.15 | 12.32 | 12.48 | 12.64 | 12.81 | 12.97
6.43 | 6.43 | 6.59 | 6.75 | 6.92 | 7.08 | 7.25 | 7.41 | 7.57 | 7.74 | 7.9 | 8.06 | 8.23 | 8.39 | 8.55 | 8.72 | 8.88 | 9.04 | 9.21 | 9.37 | 9.53 | 9.7 | 9.86 | 10.03 | 10.19 | 10.35 | 10.52 | 10.68 | 10.84 | 11.01 | 11.17 | 11.33 | 11.5 | 11.66 | 11.82 | 11.99 | 12.15 | 12.31 | 12.48 | 12.64 | 12.8 | 12.97
6.44 | 6.43 | 6.59 | 6.75 | 6.92 | 7.08 | 7.24 | 7.41 | 7.57 | 7.73 | 7.9 | 8.06 | 8.22 | 8.39 | 8.55 | 8.72 | 8.88 | 9.04 | 9.21 | 9.37 | 9.53 | 9.7 | 9.86 | 10.02 | 10.19 | 10.35 | 10.51 | 10.68 | 10.84 | 11 | 11.17 | 11.33 | 11.49 | 11.66 | 11.82 | 11.98 | 12.15 | 12.31 | 12.47 | 12.64 | 12.8 | 12.96
6.45 | 6.42 | 6.59 | 6.75 | 6.91 | 7.08 | 7.24 | 7.41 | 7.57 | 7.73 | 7.9 | 8.06 | 8.22 | 8.39 | 8.55 | 8.71 | 8.88 | 9.04 | 9.2 | 9.37 | 9.53 | 9.69 | 9.86 | 10.02 | 10.18 | 10.35 | 10.51 | 10.67 | 10.84 | 11 | 11.16 | 11.33 | 11.49 | 11.65 | 11.82 | 11.98 | 12.14 | 12.31 | 12.47 | 12.63 | 12.8 | 12.96
6.46 | 6.42 | 6.59 | 6.75 | 6.91 | 7.08 | 7.24 | 7.4 | 7.57 | 7.73 | 7.89 | 8.06 | 8.22 | 8.38 | 8.55 | 8.71 | 8.87 | 9.04 | 9.2 | 9.36 | 9.53 | 9.69 | 9.85 | 10.02 | 10.18 | 10.34 | 10.51 | 10.67 | 10.83 | 11 | 11.16 | 11.32 | 11.49 | 11.65 | 11.81 | 11.98 | 12.14 | 12.3 | 12.47 | 12.63 | 12.79 | 12.96
6.47 | 6.42 | 6.58 | 6.75 | 6.91 | 7.07 | 7.24 | 7.4 | 7.56 | 7.73 | 7.89 | 8.05 | 8.22 | 8.38 | 8.55 | 8.71 | 8.87 | 9.04 | 9.2 | 9.36 | 9.53 | 9.69 | 9.85 | 10.02 | 10.18 | 10.34 | 10.51 | 10.67 | 10.83 | 11 | 11.16 | 11.32 | 11.48 | 11.65 | 11.81 | 11.97 | 12.14 | 12.3 | 12.46 | 12.63 | 12.79 | 12.95
6.48 | 6.42 | 6.58 | 6.75 | 6.91 | 7.07 | 7.24 | 7.4 | 7.56 | 7.73 | 7.89 | 8.05 | 8.22 | 8.38 | 8.54 | 8.71 | 8.87 | 9.03 | 9.2 | 9.36 | 9.52 | 9.69 | 9.85 | 10.01 | 10.18 | 10.34 | 10.5 | 10.67 | 10.83 | 10.99 | 11.16 | 11.32 | 11.48 | 11.65 | 11.81 | 11.97 | 12.13 | 12.3 | 12.46 | 12.62 | 12.79 | 12.95
6.49 | 6.42 | 6.58 | 6.74 | 6.91 | 7.07 | 7.23 | 7.4 | 7.56 | 7.72 | 7.89 | 8.05 | 8.21 | 8.38 | 8.54 | 8.7 | 8.87 | 9.03 | 9.19 | 9.36 | 9.52 | 9.68 | 9.85 | 10.01 | 10.17 | 10.34 | 10.5 | 10.66 | 10.83 | 10.99 | 11.15 | 11.32 | 11.48 | 11.64 | 11.81 | 11.97 | 12.13 | 12.3 | 12.46 | 12.62 | 12.78 | 12.95
6.5 | 6.42 | 6.58 | 6.74 | 6.91 | 7.07 | 7.23 | 7.4 | 7.56 | 7.72 | 7.89 | 8.05 | 8.21 | 8.38 | 8.54 | 8.7 | 8.86 | 9.03 | 9.19 | 9.35 | 9.52 | 9.68 | 9.84 | 10.01 | 10.17 | 10.33 | 10.5 | 10.66 | 10.82 | 10.99 | 11.15 | 11.31 | 11.48 | 11.64 | 11.8 | 11.97 | 12.13 | 12.29 | 12.46 | 12.62 | 12.78 | 12.94
6.51 | 6.41 | 6.58 | 6.74 | 6.9 | 7.07 | 7.23 | 7.39 | 7.56 | 7.72 | 7.88 | 8.05 | 8.21 | 8.37 | 8.54 | 8.7 | 8.86 | 9.03 | 9.19 | 9.35 | 9.52 | 9.68 | 9.84 | 10 | 10.17 | 10.33 | 10.49 | 10.66 | 10.82 | 10.98 | 11.15 | 11.31 | 11.47 | 11.64 | 11.8 | 11.96 | 12.13 | 12.29 | 12.45 | 12.62 | 12.78 | 12.94
6.52 | 6.41 | 6.58 | 6.74 | 6.9 | 7.07 | 7.23 | 7.39 | 7.56 | 7.72 | 7.88 | 8.04 | 8.21 | 8.37 | 8.53 | 8.7 | 8.86 | 9.02 | 9.19 | 9.35 | 9.51 | 9.68 | 9.84 | 10 | 10.17 | 10.33 | 10.49 | 10.65 | 10.82 | 10.98 | 11.14 | 11.31 | 11.47 | 11.63 | 11.8 | 11.96 | 12.12 | 12.29 | 12.45 | 12.61 | 12.78 | 12.94
6.53 | 6.41 | 6.57 | 6.74 | 6.9 | 7.06 | 7.23 | 7.39 | 7.55 | 7.72 | 7.88 | 8.04 | 8.21 | 8.37 | 8.53 | 8.7 | 8.86 | 9.02 | 9.18 | 9.35 | 9.51 | 9.67 | 9.84 | 10 | 10.16 | 10.33 | 10.49 | 10.65 | 10.82 | 10.98 | 11.14 | 11.3 | 11.47 | 11.63 | 11.79 | 11.96 | 12.12 | 12.28 | 12.45 | 12.61 | 12.77 | 12.93
6.54 | 6.41 | 6.57 | 6.74 | 6.9 | 7.06 | 7.23 | 7.39 | 7.55 | 7.71 | 7.88 | 8.04 | 8.2 | 8.37 | 8.53 | 8.69 | 8.86 | 9.02 | 9.18 | 9.35 | 9.51 | 9.67 | 9.83 | 10 | 10.16 | 10.32 | 10.49 | 10.65 | 10.81 | 10.98 | 11.14 | 11.3 | 11.46 | 11.63 | 11.79 | 11.95 | 12.12 | 12.28 | 12.44 | 12.61 | 12.77 | 12.93
6.55 | 6.41 | 6.57 | 6.73 | 6.9 | 7.06 | 7.22 | 7.39 | 7.55 | 7.71 | 7.88 | 8.04 | 8.2 | 8.36 | 8.53 | 8.69 | 8.85 | 9.02 | 9.18 | 9.34 | 9.51 | 9.67 | 9.83 | 9.99 | 10.16 | 10.32 | 10.48 | 10.65 | 10.81 | 10.97 | 11.14 | 11.3 | 11.46 | 11.62 | 11.79 | 11.95 | 12.11 | 12.28 | 12.44 | 12.6 | 12.77 | 12.93
6.56 | 6.41 | 6.57 | 6.73 | 6.9 | 7.06 | 7.22 | 7.38 | 7.55 | 7.71 | 7.87 | 8.04 | 8.2 | 8.36 | 8.53 | 8.69 | 8.85 | 9.01 | 9.18 | 9.34 | 9.5 | 9.67 | 9.83 | 9.99 | 10.16 | 10.32 | 10.48 | 10.64 | 10.81 | 10.97 | 11.13 | 11.3 | 11.46 | 11.62 | 11.78 | 11.95 | 12.11 | 12.27 | 12.44 | 12.6 | 12.76 | 12.93
6.57 | 6.4 | 6.57 | 6.73 | 6.89 | 7.06 | 7.22 | 7.38 | 7.55 | 7.71 | 7.87 | 8.03 | 8.2 | 8.36 | 8.52 | 8.69 | 8.85 | 9.01 | 9.18 | 9.34 | 9.5 | 9.66 | 9.83 | 9.99 | 10.15 | 10.32 | 10.48 | 10.64 | 10.8 | 10.97 | 11.13 | 11.29 | 11.46 | 11.62 | 11.78 | 11.94 | 12.11 | 12.27 | 12.43 | 12.6 | 12.76 | 12.92
6.58 | 6.4 | 6.57 | 6.73 | 6.89 | 7.05 | 7.22 | 7.38 | 7.54 | 7.71 | 7.87 | 8.03 | 8.2 | 8.36 | 8.52 | 8.68 | 8.85 | 9.01 | 9.17 | 9.34 | 9.5 | 9.66 | 9.82 | 9.99 | 10.15 | 10.31 | 10.48 | 10.64 | 10.8 | 10.96 | 11.13 | 11.29 | 11.45 | 11.62 | 11.78 | 11.94 | 12.1 | 12.27 | 12.43 | 12.59 | 12.76 | 12.92
6.59 | 6.4 | 6.56 | 6.73 | 6.89 | 7.05 | 7.22 | 7.38 | 7.54 | 7.7 | 7.87 | 8.03 | 8.19 | 8.36 | 8.52 | 8.68 | 8.84 | 9.01 | 9.17 | 9.33 | 9.5 | 9.66 | 9.82 | 9.98 | 10.15 | 10.31 | 10.47 | 10.64 | 10.8 | 10.96 | 11.12 | 11.29 | 11.45 | 11.61 | 11.78 | 11.94 | 12.1 | 12.26 | 12.43 | 12.59 | 12.75 | 12.92
6.6 | 6.4 | 6.56 | 6.73 | 6.89 | 7.05 | 7.21 | 7.38 | 7.54 | 7.7 | 7.87 | 8.03 | 8.19 | 8.35 | 8.52 | 8.68 | 8.84 | 9.01 | 9.17 | 9.33 | 9.49 | 9.66 | 9.82 | 9.98 | 10.15 | 10.31 | 10.47 | 10.63 | 10.8 | 10.96 | 11.12 | 11.28 | 11.45 | 11.61 | 11.77 | 11.94 | 12.1 | 12.26 | 12.42 | 12.59 | 12.75 | 12.91
6.61 | 6.4 | 6.56 | 6.72 | 6.89 | 7.05 | 7.21 | 7.37 | 7.54 | 7.7 | 7.86 | 8.03 | 8.19 | 8.35 | 8.51 | 8.68 | 8.84 | 9 | 9.17 | 9.33 | 9.49 | 9.65 | 9.82 | 9.98 | 10.14 | 10.31 | 10.47 | 10.63 | 10.79 | 10.96 | 11.12 | 11.28 | 11.44 | 11.61 | 11.77 | 11.93 | 12.1 | 12.26 | 12.42 | 12.58 | 12.75 | 12.91
6.62 | 6.4 | 6.56 | 6.72 | 6.88 | 7.05 | 7.21 | 7.37 | 7.54 | 7.7 | 7.86 | 8.02 | 8.19 | 8.35 | 8.51 | 8.68 | 8.84 | 9 | 9.16 | 9.33 | 9.49 | 9.65 | 9.81 | 9.98 | 10.14 | 10.3 | 10.47 | 10.63 | 10.79 | 10.95 | 11.12 | 11.28 | 11.44 | 11.6 | 11.77 | 11.93 | 12.09 | 12.26 | 12.42 | 12.58 | 12.74 | 12.91
6.63 | 6.39 | 6.56 | 6.72 | 6.88 | 7.05 | 7.21 | 7.37 | 7.53 | 7.7 | 7.86 | 8.02 | 8.18 | 8.35 | 8.51 | 8.67 | 8.84 | 9 | 9.16 | 9.32 | 9.49 | 9.65 | 9.81 | 9.97 | 10.14 | 10.3 | 10.46 | 10.63 | 10.79 | 10.95 | 11.11 | 11.28 | 11.44 | 11.6 | 11.76 | 11.93 | 12.09 | 12.25 | 12.41 | 12.58 | 12.74 | 12.9
6.64 | 6.39 | 6.56 | 6.72 | 6.88 | 7.04 | 7.21 | 7.37 | 7.53 | 7.69 | 7.86 | 8.02 | 8.18 | 8.35 | 8.51 | 8.67 | 8.83 | 9 | 9.16 | 9.32 | 9.48 | 9.65 | 9.81 | 9.97 | 10.13 | 10.3 | 10.46 | 10.62 | 10.79 | 10.95 | 11.11 | 11.27 | 11.44 | 11.6 | 11.76 | 11.92 | 12.09 | 12.25 | 12.41 | 12.57 | 12.74 | 12.9
6.65 | 6.39 | 6.55 | 6.72 | 6.88 | 7.04 | 7.2 | 7.37 | 7.53 | 7.69 | 7.86 | 8.02 | 8.18 | 8.34 | 8.51 | 8.67 | 8.83 | 8.99 | 9.16 | 9.32 | 9.48 | 9.64 | 9.81 | 9.97 | 10.13 | 10.29 | 10.46 | 10.62 | 10.78 | 10.95 | 11.11 | 11.27 | 11.43 | 11.6 | 11.76 | 11.92 | 12.08 | 12.25 | 12.41 | 12.57 | 12.73 | 12.9
6.66 | 6.39 | 6.55 | 6.72 | 6.88 | 7.04 | 7.2 | 7.37 | 7.53 | 7.69 | 7.85 | 8.02 | 8.18 | 8.34 | 8.5 | 8.67 | 8.83 | 8.99 | 9.15 | 9.32 | 9.48 | 9.64 | 9.8 | 9.97 | 10.13 | 10.29 | 10.45 | 10.62 | 10.78 | 10.94 | 11.11 | 11.27 | 11.43 | 11.59 | 11.76 | 11.92 | 12.08 | 12.24 | 12.41 | 12.57 | 12.73 | 12.89
6.67 | 6.39 | 6.55 | 6.71 | 6.88 | 7.04 | 7.2 | 7.36 | 7.53 | 7.69 | 7.85 | 8.01 | 8.18 | 8.34 | 8.5 | 8.66 | 8.83 | 8.99 | 9.15 | 9.31 | 9.48 | 9.64 | 9.8 | 9.96 | 10.13 | 10.29 | 10.45 | 10.61 | 10.78 | 10.94 | 11.1 | 11.26 | 11.43 | 11.59 | 11.75 | 11.91 | 12.08 | 12.24 | 12.4 | 12.56 | 12.73 | 12.89
6.68 | 6.39 | 6.55 | 6.71 | 6.87 | 7.04 | 7.2 | 7.36 | 7.52 | 7.69 | 7.85 | 8.01 | 8.17 | 8.34 | 8.5 | 8.66 | 8.82 | 8.99 | 9.15 | 9.31 | 9.47 | 9.64 | 9.8 | 9.96 | 10.12 | 10.29 | 10.45 | 10.61 | 10.77 | 10.94 | 11.1 | 11.26 | 11.42 | 11.59 | 11.75 | 11.91 | 12.07 | 12.24 | 12.4 | 12.56 | 12.72 | 12.89
6.69 | 6.38 | 6.55 | 6.71 | 6.87 | 7.03 | 7.2 | 7.36 | 7.52 | 7.68 | 7.85 | 8.01 | 8.17 | 8.33 | 8.5 | 8.66 | 8.82 | 8.98 | 9.15 | 9.31 | 9.47 | 9.63 | 9.8 | 9.96 | 10.12 | 10.28 | 10.45 | 10.61 | 10.77 | 10.93 | 11.1 | 11.26 | 11.42 | 11.58 | 11.75 | 11.91 | 12.07 | 12.23 | 12.4 | 12.56 | 12.72 | 12.88
6.7 | 6.38 | 6.55 | 6.71 | 6.87 | 7.03 | 7.2 | 7.36 | 7.52 | 7.68 | 7.85 | 8.01 | 8.17 | 8.33 | 8.5 | 8.66 | 8.82 | 8.98 | 9.15 | 9.31 | 9.47 | 9.63 | 9.79 | 9.96 | 10.12 | 10.28 | 10.44 | 10.61 | 10.77 | 10.93 | 11.09 | 11.26 | 11.42 | 11.58 | 11.74 | 11.91 | 12.07 | 12.23 | 12.39 | 12.56 | 12.72 | 12.88
6.71 | 6.38 | 6.54 | 6.71 | 6.87 | 7.03 | 7.19 | 7.36 | 7.52 | 7.68 | 7.84 | 8.01 | 8.17 | 8.33 | 8.49 | 8.66 | 8.82 | 8.98 | 9.14 | 9.31 | 9.47 | 9.63 | 9.79 | 9.95 | 10.12 | 10.28 | 10.44 | 10.6 | 10.77 | 10.93 | 11.09 | 11.25 | 11.42 | 11.58 | 11.74 | 11.9 | 12.07 | 12.23 | 12.39 | 12.55 | 12.71 | 12.88
6.72 | 6.38 | 6.54 | 6.7 | 6.87 | 7.03 | 7.19 | 7.35 | 7.52 | 7.68 | 7.84 | 8 | 8.17 | 8.33 | 8.49 | 8.65 | 8.82 | 8.98 | 9.14 | 9.3 | 9.47 | 9.63 | 9.79 | 9.95 | 10.11 | 10.28 | 10.44 | 10.6 | 10.76 | 10.93 | 11.09 | 11.25 | 11.41 | 11.58 | 11.74 | 11.9 | 12.06 | 12.22 | 12.39 | 12.55 | 12.71 | 12.87
6.73 | 6.38 | 6.54 | 6.7 | 6.87 | 7.03 | 7.19 | 7.35 | 7.51 | 7.68 | 7.84 | 8 | 8.16 | 8.33 | 8.49 | 8.65 | 8.81 | 8.98 | 9.14 | 9.3 | 9.46 | 9.63 | 9.79 | 9.95 | 10.11 | 10.27 | 10.44 | 10.6 | 10.76 | 10.92 | 11.09 | 11.25 | 11.41 | 11.57 | 11.73 | 11.9 | 12.06 | 12.22 | 12.38 | 12.55 | 12.71 | 12.87
6.74 | 6.38 | 6.54 | 6.7 | 6.86 | 7.03 | 7.19 | 7.35 | 7.51 | 7.68 | 7.84 | 8 | 8.16 | 8.32 | 8.49 | 8.65 | 8.81 | 8.97 | 9.14 | 9.3 | 9.46 | 9.62 | 9.78 | 9.95 | 10.11 | 10.27 | 10.43 | 10.6 | 10.76 | 10.92 | 11.08 | 11.25 | 11.41 | 11.57 | 11.73 | 11.89 | 12.06 | 12.22 | 12.38 | 12.54 | 12.7 | 12.87
6.75 | 6.37 | 6.54 | 6.7 | 6.86 | 7.02 | 7.19 | 7.35 | 7.51 | 7.67 | 7.84 | 8 | 8.16 | 8.32 | 8.48 | 8.65 | 8.81 | 8.97 | 9.13 | 9.3 | 9.46 | 9.62 | 9.78 | 9.94 | 10.11 | 10.27 | 10.43 | 10.59 | 10.76 | 10.92 | 11.08 | 11.24 | 11.4 | 11.57 | 11.73 | 11.89 | 12.05 | 12.22 | 12.38 | 12.54 | 12.7 | 12.86
6.76 | 6.37 | 6.54 | 6.7 | 6.86 | 7.02 | 7.18 | 7.35 | 7.51 | 7.67 | 7.83 | 8 | 8.16 | 8.32 | 8.48 | 8.64 | 8.81 | 8.97 | 9.13 | 9.29 | 9.46 | 9.62 | 9.78 | 9.94 | 10.1 | 10.27 | 10.43 | 10.59 | 10.75 | 10.92 | 11.08 | 11.24 | 11.4 | 11.56 | 11.73 | 11.89 | 12.05 | 12.21 | 12.37 | 12.54 | 12.7 | 12.86
6.77 | 6.37 | 6.53 | 6.7 | 6.86 | 7.02 | 7.18 | 7.34 | 7.51 | 7.67 | 7.83 | 7.99 | 8.16 | 8.32 | 8.48 | 8.64 | 8.8 | 8.97 | 9.13 | 9.29 | 9.45 | 9.62 | 9.78 | 9.94 | 10.1 | 10.26 | 10.43 | 10.59 | 10.75 | 10.91 | 11.07 | 11.24 | 11.4 | 11.56 | 11.72 | 11.88 | 12.05 | 12.21 | 12.37 | 12.53 | 12.7 | 12.86
6.78 | 6.37 | 6.53 | 6.69 | 6.86 | 7.02 | 7.18 | 7.34 | 7.51 | 7.67 | 7.83 | 7.99 | 8.15 | 8.32 | 8.48 | 8.64 | 8.8 | 8.96 | 9.13 | 9.29 | 9.45 | 9.61 | 9.77 | 9.94 | 10.1 | 10.26 | 10.42 | 10.59 | 10.75 | 10.91 | 11.07 | 11.23 | 11.4 | 11.56 | 11.72 | 11.88 | 12.04 | 12.21 | 12.37 | 12.53 | 12.69 | 12.85
6.79 | 6.37 | 6.53 | 6.69 | 6.85 | 7.02 | 7.18 | 7.34 | 7.5 | 7.67 | 7.83 | 7.99 | 8.15 | 8.31 | 8.48 | 8.64 | 8.8 | 8.96 | 9.12 | 9.29 | 9.45 | 9.61 | 9.77 | 9.93 | 10.1 | 10.26 | 10.42 | 10.58 | 10.74 | 10.91 | 11.07 | 11.23 | 11.39 | 11.55 | 11.72 | 11.88 | 12.04 | 12.2 | 12.37 | 12.53 | 12.69 | 12.85
6.8 | 6.37 | 6.53 | 6.69 | 6.85 | 7.02 | 7.18 | 7.34 | 7.5 | 7.66 | 7.83 | 7.99 | 8.15 | 8.31 | 8.47 | 8.64 | 8.8 | 8.96 | 9.12 | 9.28 | 9.45 | 9.61 | 9.77 | 9.93 | 10.09 | 10.26 | 10.42 | 10.58 | 10.74 | 10.9 | 11.07 | 11.23 | 11.39 | 11.55 | 11.71 | 11.88 | 12.04 | 12.2 | 12.36 | 12.52 | 12.69 | 12.85
6.81 | 6.37 | 6.53 | 6.69 | 6.85 | 7.01 | 7.18 | 7.34 | 7.5 | 7.66 | 7.82 | 7.99 | 8.15 | 8.31 | 8.47 | 8.63 | 8.8 | 8.96 | 9.12 | 9.28 | 9.44 | 9.61 | 9.77 | 9.93 | 10.09 | 10.25 | 10.42 | 10.58 | 10.74 | 10.9 | 11.06 | 11.23 | 11.39 | 11.55 | 11.71 | 11.87 | 12.03 | 12.2 | 12.36 | 12.52 | 12.68 | 12.84
6.82 | 6.36 | 6.53 | 6.69 | 6.85 | 7.01 | 7.17 | 7.34 | 7.5 | 7.66 | 7.82 | 7.98 | 8.15 | 8.31 | 8.47 | 8.63 | 8.79 | 8.96 | 9.12 | 9.28 | 9.44 | 9.6 | 9.77 | 9.93 | 10.09 | 10.25 | 10.41 | 10.57 | 10.74 | 10.9 | 11.06 | 11.22 | 11.38 | 11.55 | 11.71 | 11.87 | 12.03 | 12.19 | 12.36 | 12.52 | 12.68 | 12.84
6.83 | 6.36 | 6.52 | 6.69 | 6.85 | 7.01 | 7.17 | 7.33 | 7.5 | 7.66 | 7.82 | 7.98 | 8.14 | 8.31 | 8.47 | 8.63 | 8.79 | 8.95 | 9.12 | 9.28 | 9.44 | 9.6 | 9.76 | 9.92 | 10.09 | 10.25 | 10.41 | 10.57 | 10.73 | 10.9 | 11.06 | 11.22 | 11.38 | 11.54 | 11.71 | 11.87 | 12.03 | 12.19 | 12.35 | 12.51 | 12.68 | 12.84
6.84 | 6.36 | 6.52 | 6.68 | 6.85 | 7.01 | 7.17 | 7.33 | 7.49 | 7.66 | 7.82 | 7.98 | 8.14 | 8.3 | 8.47 | 8.63 | 8.79 | 8.95 | 9.11 | 9.27 | 9.44 | 9.6 | 9.76 | 9.92 | 10.08 | 10.25 | 10.41 | 10.57 | 10.73 | 10.89 | 11.06 | 11.22 | 11.38 | 11.54 | 11.7 | 11.86 | 12.03 | 12.19 | 12.35 | 12.51 | 12.67 | 12.83
6.85 | 6.36 | 6.52 | 6.68 | 6.84 | 7.01 | 7.17 | 7.33 | 7.49 | 7.65 | 7.82 | 7.98 | 8.14 | 8.3 | 8.46 | 8.62 | 8.79 | 8.95 | 9.11 | 9.27 | 9.43 | 9.6 | 9.76 | 9.92 | 10.08 | 10.24 | 10.41 | 10.57 | 10.73 | 10.89 | 11.05 | 11.21 | 11.38 | 11.54 | 11.7 | 11.86 | 12.02 | 12.18 | 12.35 | 12.51 | 12.67 | 12.83
6.86 | 6.36 | 6.52 | 6.68 | 6.84 | 7 | 7.17 | 7.33 | 7.49 | 7.65 | 7.81 | 7.98 | 8.14 | 8.3 | 8.46 | 8.62 | 8.78 | 8.95 | 9.11 | 9.27 | 9.43 | 9.59 | 9.76 | 9.92 | 10.08 | 10.24 | 10.4 | 10.56 | 10.73 | 10.89 | 11.05 | 11.21 | 11.37 | 11.53 | 11.7 | 11.86 | 12.02 | 12.18 | 12.34 | 12.51 | 12.67 | 12.83
6.87 | 6.36 | 6.52 | 6.68 | 6.84 | 7 | 7.16 | 7.33 | 7.49 | 7.65 | 7.81 | 7.97 | 8.14 | 8.3 | 8.46 | 8.62 | 8.78 | 8.94 | 9.11 | 9.27 | 9.43 | 9.59 | 9.75 | 9.91 | 10.08 | 10.24 | 10.4 | 10.56 | 10.72 | 10.89 | 11.05 | 11.21 | 11.37 | 11.53 | 11.69 | 11.86 | 12.02 | 12.18 | 12.34 | 12.5 | 12.66 | 12.83
6.88 | 6.35 | 6.52 | 6.68 | 6.84 | 7 | 7.16 | 7.32 | 7.49 | 7.65 | 7.81 | 7.97 | 8.13 | 8.29 | 8.46 | 8.62 | 8.78 | 8.94 | 9.1 | 9.27 | 9.43 | 9.59 | 9.75 | 9.91 | 10.07 | 10.24 | 10.4 | 10.56 | 10.72 | 10.88 | 11.04 | 11.21 | 11.37 | 11.53 | 11.69 | 11.85 | 12.01 | 12.18 | 12.34 | 12.5 | 12.66 | 12.82
6.89 | 6.35 | 6.51 | 6.68 | 6.84 | 7 | 7.16 | 7.32 | 7.48 | 7.65 | 7.81 | 7.97 | 8.13 | 8.29 | 8.45 | 8.62 | 8.78 | 8.94 | 9.1 | 9.26 | 9.42 | 9.59 | 9.75 | 9.91 | 10.07 | 10.23 | 10.39 | 10.56 | 10.72 | 10.88 | 11.04 | 11.2 | 11.36 | 11.53 | 11.69 | 11.85 | 12.01 | 12.17 | 12.33 | 12.5 | 12.66 | 12.82
6.9 | 6.35 | 6.51 | 6.67 | 6.84 | 7 | 7.16 | 7.32 | 7.48 | 7.64 | 7.81 | 7.97 | 8.13 | 8.29 | 8.45 | 8.61 | 8.78 | 8.94 | 9.1 | 9.26 | 9.42 | 9.58 | 9.75 | 9.91 | 10.07 | 10.23 | 10.39 | 10.55 | 10.72 | 10.88 | 11.04 | 11.2 | 11.36 | 11.52 | 11.68 | 11.85 | 12.01 | 12.17 | 12.33 | 12.49 | 12.65 | 12.82
6.91 | 6.35 | 6.51 | 6.67 | 6.83 | 7 | 7.16 | 7.32 | 7.48 | 7.64 | 7.8 | 7.97 | 8.13 | 8.29 | 8.45 | 8.61 | 8.77 | 8.94 | 9.1 | 9.26 | 9.42 | 9.58 | 9.74 | 9.9 | 10.07 | 10.23 | 10.39 | 10.55 | 10.71 | 10.87 | 11.04 | 11.2 | 11.36 | 11.52 | 11.68 | 11.84 | 12 | 12.17 | 12.33 | 12.49 | 12.65 | 12.81
6.92 | 6.35 | 6.51 | 6.67 | 6.83 | 6.99 | 7.16 | 7.32 | 7.48 | 7.64 | 7.8 | 7.96 | 8.12 | 8.29 | 8.45 | 8.61 | 8.77 | 8.93 | 9.09 | 9.26 | 9.42 | 9.58 | 9.74 | 9.9 | 10.06 | 10.23 | 10.39 | 10.55 | 10.71 | 10.87 | 11.03 | 11.19 | 11.36 | 11.52 | 11.68 | 11.84 | 12 | 12.16 | 12.32 | 12.49 | 12.65 | 12.81
6.93 | 6.35 | 6.51 | 6.67 | 6.83 | 6.99 | 7.15 | 7.32 | 7.48 | 7.64 | 7.8 | 7.96 | 8.12 | 8.28 | 8.45 | 8.61 | 8.77 | 8.93 | 9.09 | 9.25 | 9.42 | 9.58 | 9.74 | 9.9 | 10.06 | 10.22 | 10.38 | 10.55 | 10.71 | 10.87 | 11.03 | 11.19 | 11.35 | 11.51 | 11.68 | 11.84 | 12 | 12.16 | 12.32 | 12.48 | 12.64 | 12.81
6.94 | 6.34 | 6.51 | 6.67 | 6.83 | 6.99 | 7.15 | 7.31 | 7.47 | 7.64 | 7.8 | 7.96 | 8.12 | 8.28 | 8.44 | 8.61 | 8.77 | 8.93 | 9.09 | 9.25 | 9.41 | 9.57 | 9.74 | 9.9 | 10.06 | 10.22 | 10.38 | 10.54 | 10.7 | 10.87 | 11.03 | 11.19 | 11.35 | 11.51 | 11.67 | 11.83 | 12 | 12.16 | 12.32 | 12.48 | 12.64 | 12.8
6.95 | 6.34 | 6.5 | 6.67 | 6.83 | 6.99 | 7.15 | 7.31 | 7.47 | 7.63 | 7.8 | 7.96 | 8.12 | 8.28 | 8.44 | 8.6 | 8.76 | 8.93 | 9.09 | 9.25 | 9.41 | 9.57 | 9.73 | 9.89 | 10.06 | 10.22 | 10.38 | 10.54 | 10.7 | 10.86 | 11.02 | 11.19 | 11.35 | 11.51 | 11.67 | 11.83 | 11.99 | 12.15 | 12.32 | 12.48 | 12.64 | 12.8
6.96 | 6.34 | 6.5 | 6.66 | 6.83 | 6.99 | 7.15 | 7.31 | 7.47 | 7.63 | 7.79 | 7.96 | 8.12 | 8.28 | 8.44 | 8.6 | 8.76 | 8.92 | 9.09 | 9.25 | 9.41 | 9.57 | 9.73 | 9.89 | 10.05 | 10.22 | 10.38 | 10.54 | 10.7 | 10.86 | 11.02 | 11.18 | 11.34 | 11.51 | 11.67 | 11.83 | 11.99 | 12.15 | 12.31 | 12.47 | 12.64 | 12.8
6.97 | 6.34 | 6.5 | 6.66 | 6.82 | 6.98 | 7.15 | 7.31 | 7.47 | 7.63 | 7.79 | 7.95 | 8.11 | 8.28 | 8.44 | 8.6 | 8.76 | 8.92 | 9.08 | 9.24 | 9.41 | 9.57 | 9.73 | 9.89 | 10.05 | 10.21 | 10.37 | 10.54 | 10.7 | 10.86 | 11.02 | 11.18 | 11.34 | 11.5 | 11.66 | 11.83 | 11.99 | 12.15 | 12.31 | 12.47 | 12.63 | 12.79
6.98 | 6.34 | 6.5 | 6.66 | 6.82 | 6.98 | 7.14 | 7.31 | 7.47 | 7.63 | 7.79 | 7.95 | 8.11 | 8.27 | 8.44 | 8.6 | 8.76 | 8.92 | 9.08 | 9.24 | 9.4 | 9.56 | 9.73 | 9.89 | 10.05 | 10.21 | 10.37 | 10.53 | 10.69 | 10.86 | 11.02 | 11.18 | 11.34 | 11.5 | 11.66 | 11.82 | 11.98 | 12.15 | 12.31 | 12.47 | 12.63 | 12.79
6.99 | 6.34 | 6.5 | 6.66 | 6.82 | 6.98 | 7.14 | 7.3 | 7.47 | 7.63 | 7.79 | 7.95 | 8.11 | 8.27 | 8.43 | 8.59 | 8.76 | 8.92 | 9.08 | 9.24 | 9.4 | 9.56 | 9.72 | 9.88 | 10.05 | 10.21 | 10.37 | 10.53 | 10.69 | 10.85 | 11.01 | 11.17 | 11.34 | 11.5 | 11.66 | 11.82 | 11.98 | 12.14 | 12.3 | 12.46 | 12.63 | 12.79
7 | 6.33 | 6.5 | 6.66 | 6.82 | 6.98 | 7.14 | 7.3 | 7.46 | 7.62 | 7.79 | 7.95 | 8.11 | 8.27 | 8.43 | 8.59 | 8.75 | 8.91 | 9.08 | 9.24 | 9.4 | 9.56 | 9.72 | 9.88 | 10.04 | 10.2 | 10.37 | 10.53 | 10.69 | 10.85 | 11.01 | 11.17 | 11.33 | 11.49 | 11.66 | 11.82 | 11.98 | 12.14 | 12.3 | 12.46 | 12.62 | 12.78
7.01 | 6.33 | 6.49 | 6.66 | 6.82 | 6.98 | 7.14 | 7.3 | 7.46 | 7.62 | 7.78 | 7.95 | 8.11 | 8.27 | 8.43 | 8.59 | 8.75 | 8.91 | 9.07 | 9.24 | 9.4 | 9.56 | 9.72 | 9.88 | 10.04 | 10.2 | 10.36 | 10.52 | 10.69 | 10.85 | 11.01 | 11.17 | 11.33 | 11.49 | 11.65 | 11.81 | 11.98 | 12.14 | 12.3 | 12.46 | 12.62 | 12.78
7.02 | 6.33 | 6.49 | 6.65 | 6.81 | 6.98 | 7.14 | 7.3 | 7.46 | 7.62 | 7.78 | 7.94 | 8.1 | 8.27 | 8.43 | 8.59 | 8.75 | 8.91 | 9.07 | 9.23 | 9.39 | 9.56 | 9.72 | 9.88 | 10.04 | 10.2 | 10.36 | 10.52 | 10.68 | 10.84 | 11.01 | 11.17 | 11.33 | 11.49 | 11.65 | 11.81 | 11.97 | 12.13 | 12.29 | 12.46 | 12.62 | 12.78
7.03 | 6.33 | 6.49 | 6.65 | 6.81 | 6.97 | 7.14 | 7.3 | 7.46 | 7.62 | 7.78 | 7.94 | 8.1 | 8.26 | 8.42 | 8.59 | 8.75 | 8.91 | 9.07 | 9.23 | 9.39 | 9.55 | 9.71 | 9.87 | 10.04 | 10.2 | 10.36 | 10.52 | 10.68 | 10.84 | 11 | 11.16 | 11.32 | 11.49 | 11.65 | 11.81 | 11.97 | 12.13 | 12.29 | 12.45 | 12.61 | 12.77
7.04 | 6.33 | 6.49 | 6.65 | 6.81 | 6.97 | 7.13 | 7.29 | 7.46 | 7.62 | 7.78 | 7.94 | 8.1 | 8.26 | 8.42 | 8.58 | 8.74 | 8.91 | 9.07 | 9.23 | 9.39 | 9.55 | 9.71 | 9.87 | 10.03 | 10.19 | 10.36 | 10.52 | 10.68 | 10.84 | 11 | 11.16 | 11.32 | 11.48 | 11.64 | 11.81 | 11.97 | 12.13 | 12.29 | 12.45 | 12.61 | 12.77
7.05 | 6.33 | 6.49 | 6.65 | 6.81 | 6.97 | 7.13 | 7.29 | 7.45 | 7.62 | 7.78 | 7.94 | 8.1 | 8.26 | 8.42 | 8.58 | 8.74 | 8.9 | 9.06 | 9.23 | 9.39 | 9.55 | 9.71 | 9.87 | 10.03 | 10.19 | 10.35 | 10.51 | 10.68 | 10.84 | 11 | 11.16 | 11.32 | 11.48 | 11.64 | 11.8 | 11.96 | 12.12 | 12.29 | 12.45 | 12.61 | 12.77
7.06 | 6.32 | 6.49 | 6.65 | 6.81 | 6.97 | 7.13 | 7.29 | 7.45 | 7.61 | 7.77 | 7.94 | 8.1 | 8.26 | 8.42 | 8.58 | 8.74 | 8.9 | 9.06 | 9.22 | 9.38 | 9.55 | 9.71 | 9.87 | 10.03 | 10.19 | 10.35 | 10.51 | 10.67 | 10.83 | 10.99 | 11.16 | 11.32 | 11.48 | 11.64 | 11.8 | 11.96 | 12.12 | 12.28 | 12.44 | 12.6 | 12.76
7.07 | 6.32 | 6.48 | 6.65 | 6.81 | 6.97 | 7.13 | 7.29 | 7.45 | 7.61 | 7.77 | 7.93 | 8.09 | 8.26 | 8.42 | 8.58 | 8.74 | 8.9 | 9.06 | 9.22 | 9.38 | 9.54 | 9.7 | 9.87 | 10.03 | 10.19 | 10.35 | 10.51 | 10.67 | 10.83 | 10.99 | 11.15 | 11.31 | 11.47 | 11.64 | 11.8 | 11.96 | 12.12 | 12.28 | 12.44 | 12.6 | 12.76
7.08 | 6.32 | 6.48 | 6.64 | 6.8 | 6.97 | 7.13 | 7.29 | 7.45 | 7.61 | 7.77 | 7.93 | 8.09 | 8.25 | 8.41 | 8.58 | 8.74 | 8.9 | 9.06 | 9.22 | 9.38 | 9.54 | 9.7 | 9.86 | 10.02 | 10.18 | 10.35 | 10.51 | 10.67 | 10.83 | 10.99 | 11.15 | 11.31 | 11.47 | 11.63 | 11.79 | 11.95 | 12.12 | 12.28 | 12.44 | 12.6 | 12.76
7.09 | 6.32 | 6.48 | 6.64 | 6.8 | 6.96 | 7.12 | 7.29 | 7.45 | 7.61 | 7.77 | 7.93 | 8.09 | 8.25 | 8.41 | 8.57 | 8.73 | 8.89 | 9.06 | 9.22 | 9.38 | 9.54 | 9.7 | 9.86 | 10.02 | 10.18 | 10.34 | 10.5 | 10.66 | 10.83 | 10.99 | 11.15 | 11.31 | 11.47 | 11.63 | 11.79 | 11.95 | 12.11 | 12.27 | 12.43 | 12.59 | 12.76
7.1 | 6.32 | 6.48 | 6.64 | 6.8 | 6.96 | 7.12 | 7.28 | 7.44 | 7.61 | 7.77 | 7.93 | 8.09 | 8.25 | 8.41 | 8.57 | 8.73 | 8.89 | 9.05 | 9.21 | 9.38 | 9.54 | 9.7 | 9.86 | 10.02 | 10.18 | 10.34 | 10.5 | 10.66 | 10.82 | 10.98 | 11.14 | 11.31 | 11.47 | 11.63 | 11.79 | 11.95 | 12.11 | 12.27 | 12.43 | 12.59 | 12.75
7.11 | 6.32 | 6.48 | 6.64 | 6.8 | 6.96 | 7.12 | 7.28 | 7.44 | 7.6 | 7.76 | 7.93 | 8.09 | 8.25 | 8.41 | 8.57 | 8.73 | 8.89 | 9.05 | 9.21 | 9.37 | 9.53 | 9.69 | 9.86 | 10.02 | 10.18 | 10.34 | 10.5 | 10.66 | 10.82 | 10.98 | 11.14 | 11.3 | 11.46 | 11.62 | 11.78 | 11.95 | 12.11 | 12.27 | 12.43 | 12.59 | 12.75
7.12 | 6.31 | 6.48 | 6.64 | 6.8 | 6.96 | 7.12 | 7.28 | 7.44 | 7.6 | 7.76 | 7.92 | 8.08 | 8.24 | 8.41 | 8.57 | 8.73 | 8.89 | 9.05 | 9.21 | 9.37 | 9.53 | 9.69 | 9.85 | 10.01 | 10.17 | 10.33 | 10.5 | 10.66 | 10.82 | 10.98 | 11.14 | 11.3 | 11.46 | 11.62 | 11.78 | 11.94 | 12.1 | 12.26 | 12.42 | 12.59 | 12.75
7.13 | 6.31 | 6.47 | 6.63 | 6.8 | 6.96 | 7.12 | 7.28 | 7.44 | 7.6 | 7.76 | 7.92 | 8.08 | 8.24 | 8.4 | 8.56 | 8.73 | 8.89 | 9.05 | 9.21 | 9.37 | 9.53 | 9.69 | 9.85 | 10.01 | 10.17 | 10.33 | 10.49 | 10.65 | 10.81 | 10.98 | 11.14 | 11.3 | 11.46 | 11.62 | 11.78 | 11.94 | 12.1 | 12.26 | 12.42 | 12.58 | 12.74
7.14 | 6.31 | 6.47 | 6.63 | 6.79 | 6.95 | 7.12 | 7.28 | 7.44 | 7.6 | 7.76 | 7.92 | 8.08 | 8.24 | 8.4 | 8.56 | 8.72 | 8.88 | 9.04 | 9.2 | 9.37 | 9.53 | 9.69 | 9.85 | 10.01 | 10.17 | 10.33 | 10.49 | 10.65 | 10.81 | 10.97 | 11.13 | 11.29 | 11.45 | 11.62 | 11.78 | 11.94 | 12.1 | 12.26 | 12.42 | 12.58 | 12.74
7.15 | 6.31 | 6.47 | 6.63 | 6.79 | 6.95 | 7.11 | 7.27 | 7.44 | 7.6 | 7.76 | 7.92 | 8.08 | 8.24 | 8.4 | 8.56 | 8.72 | 8.88 | 9.04 | 9.2 | 9.36 | 9.52 | 9.68 | 9.85 | 10.01 | 10.17 | 10.33 | 10.49 | 10.65 | 10.81 | 10.97 | 11.13 | 11.29 | 11.45 | 11.61 | 11.77 | 11.93 | 12.09 | 12.25 | 12.42 | 12.58 | 12.74
7.16 | 6.31 | 6.47 | 6.63 | 6.79 | 6.95 | 7.11 | 7.27 | 7.43 | 7.59 | 7.75 | 7.92 | 8.08 | 8.24 | 8.4 | 8.56 | 8.72 | 8.88 | 9.04 | 9.2 | 9.36 | 9.52 | 9.68 | 9.84 | 10 | 10.16 | 10.32 | 10.49 | 10.65 | 10.81 | 10.97 | 11.13 | 11.29 | 11.45 | 11.61 | 11.77 | 11.93 | 12.09 | 12.25 | 12.41 | 12.57 | 12.73
7.17 | 6.31 | 6.47 | 6.63 | 6.79 | 6.95 | 7.11 | 7.27 | 7.43 | 7.59 | 7.75 | 7.91 | 8.07 | 8.23 | 8.4 | 8.56 | 8.72 | 8.88 | 9.04 | 9.2 | 9.36 | 9.52 | 9.68 | 9.84 | 10 | 10.16 | 10.32 | 10.48 | 10.64 | 10.8 | 10.96 | 11.12 | 11.29 | 11.45 | 11.61 | 11.77 | 11.93 | 12.09 | 12.25 | 12.41 | 12.57 | 12.73
7.18 | 6.31 | 6.47 | 6.63 | 6.79 | 6.95 | 7.11 | 7.27 | 7.43 | 7.59 | 7.75 | 7.91 | 8.07 | 8.23 | 8.39 | 8.55 | 8.71 | 8.87 | 9.04 | 9.2 | 9.36 | 9.52 | 9.68 | 9.84 | 10 | 10.16 | 10.32 | 10.48 | 10.64 | 10.8 | 10.96 | 11.12 | 11.28 | 11.44 | 11.6 | 11.76 | 11.92 | 12.09 | 12.25 | 12.41 | 12.57 | 12.73
7.19 | 6.3 | 6.46 | 6.62 | 6.79 | 6.95 | 7.11 | 7.27 | 7.43 | 7.59 | 7.75 | 7.91 | 8.07 | 8.23 | 8.39 | 8.55 | 8.71 | 8.87 | 9.03 | 9.19 | 9.35 | 9.51 | 9.67 | 9.84 | 10 | 10.16 | 10.32 | 10.48 | 10.64 | 10.8 | 10.96 | 11.12 | 11.28 | 11.44 | 11.6 | 11.76 | 11.92 | 12.08 | 12.24 | 12.4 | 12.56 | 12.72
7.2 | 6.3 | 6.46 | 6.62 | 6.78 | 6.94 | 7.1 | 7.27 | 7.43 | 7.59 | 7.75 | 7.91 | 8.07 | 8.23 | 8.39 | 8.55 | 8.71 | 8.87 | 9.03 | 9.19 | 9.35 | 9.51 | 9.67 | 9.83 | 9.99 | 10.15 | 10.31 | 10.47 | 10.64 | 10.8 | 10.96 | 11.12 | 11.28 | 11.44 | 11.6 | 11.76 | 11.92 | 12.08 | 12.24 | 12.4 | 12.56 | 12.72
7.21 | 6.3 | 6.46 | 6.62 | 6.78 | 6.94 | 7.1 | 7.26 | 7.42 | 7.58 | 7.74 | 7.91 | 8.07 | 8.23 | 8.39 | 8.55 | 8.71 | 8.87 | 9.03 | 9.19 | 9.35 | 9.51 | 9.67 | 9.83 | 9.99 | 10.15 | 10.31 | 10.47 | 10.63 | 10.79 | 10.95 | 11.11 | 11.27 | 11.43 | 11.59 | 11.76 | 11.92 | 12.08 | 12.24 | 12.4 | 12.56 | 12.72
7.22 | 6.3 | 6.46 | 6.62 | 6.78 | 6.94 | 7.1 | 7.26 | 7.42 | 7.58 | 7.74 | 7.9 | 8.06 | 8.22 | 8.38 | 8.54 | 8.71 | 8.87 | 9.03 | 9.19 | 9.35 | 9.51 | 9.67 | 9.83 | 9.99 | 10.15 | 10.31 | 10.47 | 10.63 | 10.79 | 10.95 | 11.11 | 11.27 | 11.43 | 11.59 | 11.75 | 11.91 | 12.07 | 12.23 | 12.39 | 12.55 | 12.71
7.23 | 6.3 | 6.46 | 6.62 | 6.78 | 6.94 | 7.1 | 7.26 | 7.42 | 7.58 | 7.74 | 7.9 | 8.06 | 8.22 | 8.38 | 8.54 | 8.7 | 8.86 | 9.02 | 9.18 | 9.34 | 9.5 | 9.67 | 9.83 | 9.99 | 10.15 | 10.31 | 10.47 | 10.63 | 10.79 | 10.95 | 11.11 | 11.27 | 11.43 | 11.59 | 11.75 | 11.91 | 12.07 | 12.23 | 12.39 | 12.55 | 12.71
7.24 | 6.3 | 6.46 | 6.62 | 6.78 | 6.94 | 7.1 | 7.26 | 7.42 | 7.58 | 7.74 | 7.9 | 8.06 | 8.22 | 8.38 | 8.54 | 8.7 | 8.86 | 9.02 | 9.18 | 9.34 | 9.5 | 9.66 | 9.82 | 9.98 | 10.14 | 10.3 | 10.46 | 10.62 | 10.78 | 10.95 | 11.11 | 11.27 | 11.43 | 11.59 | 11.75 | 11.91 | 12.07 | 12.23 | 12.39 | 12.55 | 12.71
7.25 | 6.29 | 6.45 | 6.61 | 6.78 | 6.94 | 7.1 | 7.26 | 7.42 | 7.58 | 7.74 | 7.9 | 8.06 | 8.22 | 8.38 | 8.54 | 8.7 | 8.86 | 9.02 | 9.18 | 9.34 | 9.5 | 9.66 | 9.82 | 9.98 | 10.14 | 10.3 | 10.46 | 10.62 | 10.78 | 10.94 | 11.1 | 11.26 | 11.42 | 11.58 | 11.74 | 11.9 | 12.06 | 12.22 | 12.38 | 12.54 | 12.7
7.26 | 6.29 | 6.45 | 6.61 | 6.77 | 6.93 | 7.09 | 7.25 | 7.41 | 7.57 | 7.74 | 7.9 | 8.06 | 8.22 | 8.38 | 8.54 | 8.7 | 8.86 | 9.02 | 9.18 | 9.34 | 9.5 | 9.66 | 9.82 | 9.98 | 10.14 | 10.3 | 10.46 | 10.62 | 10.78 | 10.94 | 11.1 | 11.26 | 11.42 | 11.58 | 11.74 | 11.9 | 12.06 | 12.22 | 12.38 | 12.54 | 12.7
7.27 | 6.29 | 6.45 | 6.61 | 6.77 | 6.93 | 7.09 | 7.25 | 7.41 | 7.57 | 7.73 | 7.89 | 8.05 | 8.21 | 8.37 | 8.53 | 8.69 | 8.85 | 9.01 | 9.18 | 9.34 | 9.5 | 9.66 | 9.82 | 9.98 | 10.14 | 10.3 | 10.46 | 10.62 | 10.78 | 10.94 | 11.1 | 11.26 | 11.42 | 11.58 | 11.74 | 11.9 | 12.06 | 12.22 | 12.38 | 12.54 | 12.7
7.28 | 6.29 | 6.45 | 6.61 | 6.77 | 6.93 | 7.09 | 7.25 | 7.41 | 7.57 | 7.73 | 7.89 | 8.05 | 8.21 | 8.37 | 8.53 | 8.69 | 8.85 | 9.01 | 9.17 | 9.33 | 9.49 | 9.65 | 9.81 | 9.97 | 10.13 | 10.29 | 10.45 | 10.61 | 10.77 | 10.93 | 11.09 | 11.25 | 11.41 | 11.57 | 11.73 | 11.89 | 12.06 | 12.22 | 12.38 | 12.54 | 12.7
7.29 | 6.29 | 6.45 | 6.61 | 6.77 | 6.93 | 7.09 | 7.25 | 7.41 | 7.57 | 7.73 | 7.89 | 8.05 | 8.21 | 8.37 | 8.53 | 8.69 | 8.85 | 9.01 | 9.17 | 9.33 | 9.49 | 9.65 | 9.81 | 9.97 | 10.13 | 10.29 | 10.45 | 10.61 | 10.77 | 10.93 | 11.09 | 11.25 | 11.41 | 11.57 | 11.73 | 11.89 | 12.05 | 12.21 | 12.37 | 12.53 | 12.69
7.3 | 6.29 | 6.45 | 6.61 | 6.77 | 6.93 | 7.09 | 7.25 | 7.41 | 7.57 | 7.73 | 7.89 | 8.05 | 8.21 | 8.37 | 8.53 | 8.69 | 8.85 | 9.01 | 9.17 | 9.33 | 9.49 | 9.65 | 9.81 | 9.97 | 10.13 | 10.29 | 10.45 | 10.61 | 10.77 | 10.93 | 11.09 | 11.25 | 11.41 | 11.57 | 11.73 | 11.89 | 12.05 | 12.21 | 12.37 | 12.53 | 12.69
7.31 | 6.28 | 6.44 | 6.6 | 6.76 | 6.92 | 7.08 | 7.25 | 7.41 | 7.57 | 7.73 | 7.89 | 8.05 | 8.21 | 8.37 | 8.53 | 8.69 | 8.85 | 9.01 | 9.17 | 9.33 | 9.49 | 9.65 | 9.81 | 9.97 | 10.13 | 10.29 | 10.45 | 10.61 | 10.77 | 10.93 | 11.09 | 11.25 | 11.41 | 11.57 | 11.73 | 11.89 | 12.05 | 12.21 | 12.37 | 12.53 | 12.69
7.32 | 6.28 | 6.44 | 6.6 | 6.76 | 6.92 | 7.08 | 7.24 | 7.4 | 7.56 | 7.72 | 7.88 | 8.04 | 8.2 | 8.36 | 8.52 | 8.68 | 8.84 | 9 | 9.16 | 9.32 | 9.48 | 9.64 | 9.8 | 9.96 | 10.12 | 10.28 | 10.44 | 10.6 | 10.76 | 10.92 | 11.08 | 11.24 | 11.4 | 11.56 | 11.72 | 11.88 | 12.04 | 12.2 | 12.36 | 12.52 | 12.68
7.33 | 6.28 | 6.44 | 6.6 | 6.76 | 6.92 | 7.08 | 7.24 | 7.4 | 7.56 | 7.72 | 7.88 | 8.04 | 8.2 | 8.36 | 8.52 | 8.68 | 8.84 | 9 | 9.16 | 9.32 | 9.48 | 9.64 | 9.8 | 9.96 | 10.12 | 10.28 | 10.44 | 10.6 | 10.76 | 10.92 | 11.08 | 11.24 | 11.4 | 11.56 | 11.72 | 11.88 | 12.04 | 12.2 | 12.36 | 12.52 | 12.68
7.34 | 6.28 | 6.44 | 6.6 | 6.76 | 6.92 | 7.08 | 7.24 | 7.4 | 7.56 | 7.72 | 7.88 | 8.04 | 8.2 | 8.36 | 8.52 | 8.68 | 8.84 | 9 | 9.16 | 9.32 | 9.48 | 9.64 | 9.8 | 9.96 | 10.12 | 10.28 | 10.44 | 10.6 | 10.76 | 10.92 | 11.08 | 11.24 | 11.4 | 11.56 | 11.72 | 11.88 | 12.04 | 12.2 | 12.36 | 12.52 | 12.68
7.35 | 6.28 | 6.44 | 6.6 | 6.76 | 6.92 | 7.08 | 7.24 | 7.4 | 7.56 | 7.72 | 7.88 | 8.04 | 8.2 | 8.36 | 8.52 | 8.68 | 8.84 | 9 | 9.16 | 9.32 | 9.48 | 9.64 | 9.8 | 9.96 | 10.12 | 10.28 | 10.44 | 10.6 | 10.76 | 10.92 | 11.08 | 11.24 | 11.39 | 11.55 | 11.71 | 11.87 | 12.03 | 12.19 | 12.35 | 12.51 | 12.67
7.36 | 6.28 | 6.44 | 6.6 | 6.76 | 6.92 | 7.08 | 7.24 | 7.4 | 7.56 | 7.72 | 7.88 | 8.04 | 8.2 | 8.36 | 8.52 | 8.67 | 8.83 | 8.99 | 9.15 | 9.31 | 9.47 | 9.63 | 9.79 | 9.95 | 10.11 | 10.27 | 10.43 | 10.59 | 10.75 | 10.91 | 11.07 | 11.23 | 11.39 | 11.55 | 11.71 | 11.87 | 12.03 | 12.19 | 12.35 | 12.51 | 12.67
7.37 | 6.27 | 6.43 | 6.59 | 6.75 | 6.91 | 7.07 | 7.23 | 7.39 | 7.55 | 7.71 | 7.87 | 8.03 | 8.19 | 8.35 | 8.51 | 8.67 | 8.83 | 8.99 | 9.15 | 9.31 | 9.47 | 9.63 | 9.79 | 9.95 | 10.11 | 10.27 | 10.43 | 10.59 | 10.75 | 10.91 | 11.07 | 11.23 | 11.39 | 11.55 | 11.71 | 11.87 | 12.03 | 12.19 | 12.35 | 12.51 | 12.67
7.38 | 6.27 | 6.43 | 6.59 | 6.75 | 6.91 | 7.07 | 7.23 | 7.39 | 7.55 | 7.71 | 7.87 | 8.03 | 8.19 | 8.35 | 8.51 | 8.67 | 8.83 | 8.99 | 9.15 | 9.31 | 9.47 | 9.63 | 9.79 | 9.95 | 10.11 | 10.27 | 10.43 | 10.59 | 10.75 | 10.91 | 11.07 | 11.23 | 11.39 | 11.55 | 11.71 | 11.87 | 12.03 | 12.18 | 12.34 | 12.5 | 12.66
7.39 | 6.27 | 6.43 | 6.59 | 6.75 | 6.91 | 7.07 | 7.23 | 7.39 | 7.55 | 7.71 | 7.87 | 8.03 | 8.19 | 8.35 | 8.51 | 8.67 | 8.83 | 8.99 | 9.15 | 9.31 | 9.47 | 9.63 | 9.79 | 9.95 | 10.11 | 10.27 | 10.43 | 10.59 | 10.74 | 10.9 | 11.06 | 11.22 | 11.38 | 11.54 | 11.7 | 11.86 | 12.02 | 12.18 | 12.34 | 12.5 | 12.66
7.4 | 6.27 | 6.43 | 6.59 | 6.75 | 6.91 | 7.07 | 7.23 | 7.39 | 7.55 | 7.71 | 7.87 | 8.03 | 8.19 | 8.35 | 8.51 | 8.67 | 8.83 | 8.99 | 9.15 | 9.31 | 9.46 | 9.62 | 9.78 | 9.94 | 10.1 | 10.26 | 10.42 | 10.58 | 10.74 | 10.9 | 11.06 | 11.22 | 11.38 | 11.54 | 11.7 | 11.86 | 12.02 | 12.18 | 12.34 | 12.5 | 12.66
7.41 | 6.27 | 6.43 | 6.59 | 6.75 | 6.91 | 7.07 | 7.23 | 7.39 | 7.55 | 7.71 | 7.87 | 8.03 | 8.19 | 8.34 | 8.5 | 8.66 | 8.82 | 8.98 | 9.14 | 9.3 | 9.46 | 9.62 | 9.78 | 9.94 | 10.1 | 10.26 | 10.42 | 10.58 | 10.74 | 10.9 | 11.06 | 11.22 | 11.38 | 11.54 | 11.7 | 11.86 | 12.02 | 12.18 | 12.34 | 12.5 | 12.65
7.42 | 6.27 | 6.43 | 6.59 | 6.75 | 6.91 | 7.07 | 7.23 | 7.38 | 7.54 | 7.7 | 7.86 | 8.02 | 8.18 | 8.34 | 8.5 | 8.66 | 8.82 | 8.98 | 9.14 | 9.3 | 9.46 | 9.62 | 9.78 | 9.94 | 10.1 | 10.26 | 10.42 | 10.58 | 10.74 | 10.9 | 11.06 | 11.22 | 11.38 | 11.53 | 11.69 | 11.85 | 12.01 | 12.17 | 12.33 | 12.49 | 12.65
7.43 | 6.27 | 6.43 | 6.58 | 6.74 | 6.9 | 7.06 | 7.22 | 7.38 | 7.54 | 7.7 | 7.86 | 8.02 | 8.18 | 8.34 | 8.5 | 8.66 | 8.82 | 8.98 | 9.14 | 9.3 | 9.46 | 9.62 | 9.78 | 9.94 | 10.1 | 10.26 | 10.42 | 10.57 | 10.73 | 10.89 | 11.05 | 11.21 | 11.37 | 11.53 | 11.69 | 11.85 | 12.01 | 12.17 | 12.33 | 12.49 | 12.65
7.44 | 6.26 | 6.42 | 6.58 | 6.74 | 6.9 | 7.06 | 7.22 | 7.38 | 7.54 | 7.7 | 7.86 | 8.02 | 8.18 | 8.34 | 8.5 | 8.66 | 8.82 | 8.98 | 9.14 | 9.3 | 9.46 | 9.61 | 9.77 | 9.93 | 10.09 | 10.25 | 10.41 | 10.57 | 10.73 | 10.89 | 11.05 | 11.21 | 11.37 | 11.53 | 11.69 | 11.85 | 12.01 | 12.17 | 12.33 | 12.49 | 12.65
7.45 | 6.26 | 6.42 | 6.58 | 6.74 | 6.9 | 7.06 | 7.22 | 7.38 | 7.54 | 7.7 | 7.86 | 8.02 | 8.18 | 8.34 | 8.5 | 8.66 | 8.81 | 8.97 | 9.13 | 9.29 | 9.45 | 9.61 | 9.77 | 9.93 | 10.09 | 10.25 | 10.41 | 10.57 | 10.73 | 10.89 | 11.05 | 11.21 | 11.37 | 11.53 | 11.69 | 11.85 | 12 | 12.16 | 12.32 | 12.48 | 12.64
7.46 | 6.26 | 6.42 | 6.58 | 6.74 | 6.9 | 7.06 | 7.22 | 7.38 | 7.54 | 7.7 | 7.86 | 8.02 | 8.17 | 8.33 | 8.49 | 8.65 | 8.81 | 8.97 | 9.13 | 9.29 | 9.45 | 9.61 | 9.77 | 9.93 | 10.09 | 10.25 | 10.41 | 10.57 | 10.73 | 10.89 | 11.05 | 11.2 | 11.36 | 11.52 | 11.68 | 11.84 | 12 | 12.16 | 12.32 | 12.48 | 12.64
7.47 | 6.26 | 6.42 | 6.58 | 6.74 | 6.9 | 7.06 | 7.22 | 7.38 | 7.53 | 7.69 | 7.85 | 8.01 | 8.17 | 8.33 | 8.49 | 8.65 | 8.81 | 8.97 | 9.13 | 9.29 | 9.45 | 9.61 | 9.77 | 9.93 | 10.09 | 10.25 | 10.4 | 10.56 | 10.72 | 10.88 | 11.04 | 11.2 | 11.36 | 11.52 | 11.68 | 11.84 | 12 | 12.16 | 12.32 | 12.48 | 12.64
7.48 | 6.26 | 6.42 | 6.58 | 6.74 | 6.9 | 7.05 | 7.21 | 7.37 | 7.53 | 7.69 | 7.85 | 8.01 | 8.17 | 8.33 | 8.49 | 8.65 | 8.81 | 8.97 | 9.13 | 9.29 | 9.45 | 9.61 | 9.76 | 9.92 | 10.08 | 10.24 | 10.4 | 10.56 | 10.72 | 10.88 | 11.04 | 11.2 | 11.36 | 11.52 | 11.68 | 11.84 | 12 | 12.15 | 12.31 | 12.47 | 12.63
7.49 | 6.26 | 6.42 | 6.57 | 6.73 | 6.89 | 7.05 | 7.21 | 7.37 | 7.53 | 7.69 | 7.85 | 8.01 | 8.17 | 8.33 | 8.49 | 8.65 | 8.81 | 8.97 | 9.12 | 9.28 | 9.44 | 9.6 | 9.76 | 9.92 | 10.08 | 10.24 | 10.4 | 10.56 | 10.72 | 10.88 | 11.04 | 11.2 | 11.36 | 11.51 | 11.67 | 11.83 | 11.99 | 12.15 | 12.31 | 12.47 | 12.63
7.5 | 6.25 | 6.41 | 6.57 | 6.73 | 6.89 | 7.05 | 7.21 | 7.37 | 7.53 | 7.69 | 7.85 | 8.01 | 8.17 | 8.33 | 8.49 | 8.64 | 8.8 | 8.96 | 9.12 | 9.28 | 9.44 | 9.6 | 9.76 | 9.92 | 10.08 | 10.24 | 10.4 | 10.56 | 10.72 | 10.87 | 11.03 | 11.19 | 11.35 | 11.51 | 11.67 | 11.83 | 11.99 | 12.15 | 12.31 | 12.47 | 12.63
7.51 | 6.25 | 6.41 | 6.57 | 6.73 | 6.89 | 7.05 | 7.21 | 7.37 | 7.53 | 7.69 | 7.85 | 8.01 | 8.16 | 8.32 | 8.48 | 8.64 | 8.8 | 8.96 | 9.12 | 9.28 | 9.44 | 9.6 | 9.76 | 9.92 | 10.08 | 10.24 | 10.39 | 10.55 | 10.71 | 10.87 | 11.03 | 11.19 | 11.35 | 11.51 | 11.67 | 11.83 | 11.99 | 12.15 | 12.31 | 12.46 | 12.62
7.52 | 6.25 | 6.41 | 6.57 | 6.73 | 6.89 | 7.05 | 7.21 | 7.37 | 7.53 | 7.68 | 7.84 | 8 | 8.16 | 8.32 | 8.48 | 8.64 | 8.8 | 8.96 | 9.12 | 9.28 | 9.44 | 9.6 | 9.75 | 9.91 | 10.07 | 10.23 | 10.39 | 10.55 | 10.71 | 10.87 | 11.03 | 11.19 | 11.35 | 11.51 | 11.67 | 11.82 | 11.98 | 12.14 | 12.3 | 12.46 | 12.62
7.53 | 6.25 | 6.41 | 6.57 | 6.73 | 6.89 | 7.05 | 7.21 | 7.36 | 7.52 | 7.68 | 7.84 | 8 | 8.16 | 8.32 | 8.48 | 8.64 | 8.8 | 8.96 | 9.12 | 9.27 | 9.43 | 9.59 | 9.75 | 9.91 | 10.07 | 10.23 | 10.39 | 10.55 | 10.71 | 10.87 | 11.03 | 11.19 | 11.34 | 11.5 | 11.66 | 11.82 | 11.98 | 12.14 | 12.3 | 12.46 | 12.62
7.54 | 6.25 | 6.41 | 6.57 | 6.73 | 6.88 | 7.04 | 7.2 | 7.36 | 7.52 | 7.68 | 7.84 | 8 | 8.16 | 8.32 | 8.48 | 8.64 | 8.8 | 8.95 | 9.11 | 9.27 | 9.43 | 9.59 | 9.75 | 9.91 | 10.07 | 10.23 | 10.39 | 10.55 | 10.7 | 10.86 | 11.02 | 11.18 | 11.34 | 11.5 | 11.66 | 11.82 | 11.98 | 12.14 | 12.3 | 12.46 | 12.61
7.55 | 6.25 | 6.41 | 6.56 | 6.72 | 6.88 | 7.04 | 7.2 | 7.36 | 7.52 | 7.68 | 7.84 | 8 | 8.16 | 8.32 | 8.47 | 8.63 | 8.79 | 8.95 | 9.11 | 9.27 | 9.43 | 9.59 | 9.75 | 9.91 | 10.07 | 10.23 | 10.38 | 10.54 | 10.7 | 10.86 | 11.02 | 11.18 | 11.34 | 11.5 | 11.66 | 11.82 | 11.97 | 12.13 | 12.29 | 12.45 | 12.61
7.56 | 6.24 | 6.4 | 6.56 | 6.72 | 6.88 | 7.04 | 7.2 | 7.36 | 7.52 | 7.68 | 7.84 | 8 | 8.15 | 8.31 | 8.47 | 8.63 | 8.79 | 8.95 | 9.11 | 9.27 | 9.43 | 9.59 | 9.75 | 9.9 | 10.06 | 10.22 | 10.38 | 10.54 | 10.7 | 10.86 | 11.02 | 11.18 | 11.34 | 11.49 | 11.65 | 11.81 | 11.97 | 12.13 | 12.29 | 12.45 | 12.61
7.57 | 6.24 | 6.4 | 6.56 | 6.72 | 6.88 | 7.04 | 7.2 | 7.36 | 7.52 | 7.68 | 7.83 | 7.99 | 8.15 | 8.31 | 8.47 | 8.63 | 8.79 | 8.95 | 9.11 | 9.27 | 9.42 | 9.58 | 9.74 | 9.9 | 10.06 | 10.22 | 10.38 | 10.54 | 10.7 | 10.86 | 11.02 | 11.17 | 11.33 | 11.49 | 11.65 | 11.81 | 11.97 | 12.13 | 12.29 | 12.45 | 12.6
7.58 | 6.24 | 6.4 | 6.56 | 6.72 | 6.88 | 7.04 | 7.2 | 7.36 | 7.51 | 7.67 | 7.83 | 7.99 | 8.15 | 8.31 | 8.47 | 8.63 | 8.79 | 8.95 | 9.1 | 9.26 | 9.42 | 9.58 | 9.74 | 9.9 | 10.06 | 10.22 | 10.38 | 10.54 | 10.69 | 10.85 | 11.01 | 11.17 | 11.33 | 11.49 | 11.65 | 11.81 | 11.97 | 12.13 | 12.28 | 12.44 | 12.6
7.59 | 6.24 | 6.4 | 6.56 | 6.72 | 6.88 | 7.04 | 7.19 | 7.35 | 7.51 | 7.67 | 7.83 | 7.99 | 8.15 | 8.31 | 8.47 | 8.63 | 8.78 | 8.94 | 9.1 | 9.26 | 9.42 | 9.58 | 9.74 | 9.9 | 10.06 | 10.21 | 10.37 | 10.53 | 10.69 | 10.85 | 11.01 | 11.17 | 11.33 | 11.49 | 11.65 | 11.8 | 11.96 | 12.12 | 12.28 | 12.44 | 12.6
7.6 | 6.24 | 6.4 | 6.56 | 6.72 | 6.87 | 7.03 | 7.19 | 7.35 | 7.51 | 7.67 | 7.83 | 7.99 | 8.15 | 8.31 | 8.46 | 8.62 | 8.78 | 8.94 | 9.1 | 9.26 | 9.42 | 9.58 | 9.74 | 9.89 | 10.05 | 10.21 | 10.37 | 10.53 | 10.69 | 10.85 | 11.01 | 11.17 | 11.32 | 11.48 | 11.64 | 11.8 | 11.96 | 12.12 | 12.28 | 12.44 | 12.6
7.61 | 6.24 | 6.4 | 6.55 | 6.71 | 6.87 | 7.03 | 7.19 | 7.35 | 7.51 | 7.67 | 7.83 | 7.99 | 8.14 | 8.3 | 8.46 | 8.62 | 8.78 | 8.94 | 9.1 | 9.26 | 9.42 | 9.57 | 9.73 | 9.89 | 10.05 | 10.21 | 10.37 | 10.53 | 10.69 | 10.85 | 11 | 11.16 | 11.32 | 11.48 | 11.64 | 11.8 | 11.96 | 12.12 | 12.27 | 12.43 | 12.59
7.62 | 6.24 | 6.39 | 6.55 | 6.71 | 6.87 | 7.03 | 7.19 | 7.35 | 7.51 | 7.67 | 7.82 | 7.98 | 8.14 | 8.3 | 8.46 | 8.62 | 8.78 | 8.94 | 9.1 | 9.25 | 9.41 | 9.57 | 9.73 | 9.89 | 10.05 | 10.21 | 10.37 | 10.52 | 10.68 | 10.84 | 11 | 11.16 | 11.32 | 11.48 | 11.64 | 11.8 | 11.95 | 12.11 | 12.27 | 12.43 | 12.59
7.63 | 6.23 | 6.39 | 6.55 | 6.71 | 6.87 | 7.03 | 7.19 | 7.35 | 7.5 | 7.66 | 7.82 | 7.98 | 8.14 | 8.3 | 8.46 | 8.62 | 8.78 | 8.93 | 9.09 | 9.25 | 9.41 | 9.57 | 9.73 | 9.89 | 10.05 | 10.2 | 10.36 | 10.52 | 10.68 | 10.84 | 11 | 11.16 | 11.32 | 11.48 | 11.63 | 11.79 | 11.95 | 12.11 | 12.27 | 12.43 | 12.59
7.64 | 6.23 | 6.39 | 6.55 | 6.71 | 6.87 | 7.03 | 7.19 | 7.34 | 7.5 | 7.66 | 7.82 | 7.98 | 8.14 | 8.3 | 8.46 | 8.61 | 8.77 | 8.93 | 9.09 | 9.25 | 9.41 | 9.57 | 9.73 | 9.88 | 10.04 | 10.2 | 10.36 | 10.52 | 10.68 | 10.84 | 11 | 11.15 | 11.31 | 11.47 | 11.63 | 11.79 | 11.95 | 12.11 | 12.27 | 12.42 | 12.58
7.65 | 6.23 | 6.39 | 6.55 | 6.71 | 6.87 | 7.02 | 7.18 | 7.34 | 7.5 | 7.66 | 7.82 | 7.98 | 8.14 | 8.29 | 8.45 | 8.61 | 8.77 | 8.93 | 9.09 | 9.25 | 9.41 | 9.56 | 9.72 | 9.88 | 10.04 | 10.2 | 10.36 | 10.52 | 10.68 | 10.83 | 10.99 | 11.15 | 11.31 | 11.47 | 11.63 | 11.79 | 11.95 | 12.1 | 12.26 | 12.42 | 12.58
7.66 | 6.23 | 6.39 | 6.55 | 6.71 | 6.86 | 7.02 | 7.18 | 7.34 | 7.5 | 7.66 | 7.82 | 7.98 | 8.13 | 8.29 | 8.45 | 8.61 | 8.77 | 8.93 | 9.09 | 9.25 | 9.4 | 9.56 | 9.72 | 9.88 | 10.04 | 10.2 | 10.36 | 10.51 | 10.67 | 10.83 | 10.99 | 11.15 | 11.31 | 11.47 | 11.63 | 11.78 | 11.94 | 12.1 | 12.26 | 12.42 | 12.58
7.67 | 6.23 | 6.39 | 6.54 | 6.7 | 6.86 | 7.02 | 7.18 | 7.34 | 7.5 | 7.66 | 7.81 | 7.97 | 8.13 | 8.29 | 8.45 | 8.61 | 8.77 | 8.93 | 9.08 | 9.24 | 9.4 | 9.56 | 9.72 | 9.88 | 10.04 | 10.19 | 10.35 | 10.51 | 10.67 | 10.83 | 10.99 | 11.15 | 11.31 | 11.46 | 11.62 | 11.78 | 11.94 | 12.1 | 12.26 | 12.42 | 12.57
7.68 | 6.23 | 6.38 | 6.54 | 6.7 | 6.86 | 7.02 | 7.18 | 7.34 | 7.5 | 7.65 | 7.81 | 7.97 | 8.13 | 8.29 | 8.45 | 8.61 | 8.76 | 8.92 | 9.08 | 9.24 | 9.4 | 9.56 | 9.72 | 9.87 | 10.03 | 10.19 | 10.35 | 10.51 | 10.67 | 10.83 | 10.99 | 11.14 | 11.3 | 11.46 | 11.62 | 11.78 | 11.94 | 12.1 | 12.25 | 12.41 | 12.57
7.69 | 6.22 | 6.38 | 6.54 | 6.7 | 6.86 | 7.02 | 7.18 | 7.33 | 7.49 | 7.65 | 7.81 | 7.97 | 8.13 | 8.29 | 8.45 | 8.6 | 8.76 | 8.92 | 9.08 | 9.24 | 9.4 | 9.56 | 9.71 | 9.87 | 10.03 | 10.19 | 10.35 | 10.51 | 10.67 | 10.82 | 10.98 | 11.14 | 11.3 | 11.46 | 11.62 | 11.78 | 11.93 | 12.09 | 12.25 | 12.41 | 12.57
7.7 | 6.22 | 6.38 | 6.54 | 6.7 | 6.86 | 7.02 | 7.17 | 7.33 | 7.49 | 7.65 | 7.81 | 7.97 | 8.13 | 8.28 | 8.44 | 8.6 | 8.76 | 8.92 | 9.08 | 9.24 | 9.39 | 9.55 | 9.71 | 9.87 | 10.03 | 10.19 | 10.35 | 10.5 | 10.66 | 10.82 | 10.98 | 11.14 | 11.3 | 11.46 | 11.61 | 11.77 | 11.93 | 12.09 | 12.25 | 12.41 | 12.56
7.71 | 6.22 | 6.38 | 6.54 | 6.7 | 6.86 | 7.01 | 7.17 | 7.33 | 7.49 | 7.65 | 7.81 | 7.97 | 8.12 | 8.28 | 8.44 | 8.6 | 8.76 | 8.92 | 9.08 | 9.23 | 9.39 | 9.55 | 9.71 | 9.87 | 10.03 | 10.18 | 10.34 | 10.5 | 10.66 | 10.82 | 10.98 | 11.14 | 11.29 | 11.45 | 11.61 | 11.77 | 11.93 | 12.09 | 12.24 | 12.4 | 12.56
7.72 | 6.22 | 6.38 | 6.54 | 6.7 | 6.85 | 7.01 | 7.17 | 7.33 | 7.49 | 7.65 | 7.8 | 7.96 | 8.12 | 8.28 | 8.44 | 8.6 | 8.76 | 8.91 | 9.07 | 9.23 | 9.39 | 9.55 | 9.71 | 9.87 | 10.02 | 10.18 | 10.34 | 10.5 | 10.66 | 10.82 | 10.97 | 11.13 | 11.29 | 11.45 | 11.61 | 11.77 | 11.92 | 12.08 | 12.24 | 12.4 | 12.56
7.73 | 6.22 | 6.38 | 6.53 | 6.69 | 6.85 | 7.01 | 7.17 | 7.33 | 7.49 | 7.64 | 7.8 | 7.96 | 8.12 | 8.28 | 8.44 | 8.6 | 8.75 | 8.91 | 9.07 | 9.23 | 9.39 | 9.55 | 9.7 | 9.86 | 10.02 | 10.18 | 10.34 | 10.5 | 10.65 | 10.81 | 10.97 | 11.13 | 11.29 | 11.45 | 11.61 | 11.76 | 11.92 | 12.08 | 12.24 | 12.4 | 12.56
7.74 | 6.22 | 6.37 | 6.53 | 6.69 | 6.85 | 7.01 | 7.17 | 7.33 | 7.48 | 7.64 | 7.8 | 7.96 | 8.12 | 8.28 | 8.43 | 8.59 | 8.75 | 8.91 | 9.07 | 9.23 | 9.39 | 9.54 | 9.7 | 9.86 | 10.02 | 10.18 | 10.34 | 10.49 | 10.65 | 10.81 | 10.97 | 11.13 | 11.29 | 11.44 | 11.6 | 11.76 | 11.92 | 12.08 | 12.24 | 12.39 | 12.55
7.75 | 6.21 | 6.37 | 6.53 | 6.69 | 6.85 | 7.01 | 7.17 | 7.32 | 7.48 | 7.64 | 7.8 | 7.96 | 8.12 | 8.27 | 8.43 | 8.59 | 8.75 | 8.91 | 9.07 | 9.22 | 9.38 | 9.54 | 9.7 | 9.86 | 10.02 | 10.17 | 10.33 | 10.49 | 10.65 | 10.81 | 10.97 | 11.12 | 11.28 | 11.44 | 11.6 | 11.76 | 11.92 | 12.07 | 12.23 | 12.39 | 12.55
7.76 | 6.21 | 6.37 | 6.53 | 6.69 | 6.85 | 7.01 | 7.16 | 7.32 | 7.48 | 7.64 | 7.8 | 7.96 | 8.11 | 8.27 | 8.43 | 8.59 | 8.75 | 8.91 | 9.06 | 9.22 | 9.38 | 9.54 | 9.7 | 9.86 | 10.01 | 10.17 | 10.33 | 10.49 | 10.65 | 10.81 | 10.96 | 11.12 | 11.28 | 11.44 | 11.6 | 11.75 | 11.91 | 12.07 | 12.23 | 12.39 | 12.55
7.77 | 6.21 | 6.37 | 6.53 | 6.69 | 6.85 | 7 | 7.16 | 7.32 | 7.48 | 7.64 | 7.8 | 7.95 | 8.11 | 8.27 | 8.43 | 8.59 | 8.74 | 8.9 | 9.06 | 9.22 | 9.38 | 9.54 | 9.69 | 9.85 | 10.01 | 10.17 | 10.33 | 10.49 | 10.64 | 10.8 | 10.96 | 11.12 | 11.28 | 11.44 | 11.59 | 11.75 | 11.91 | 12.07 | 12.23 | 12.38 | 12.54
7.78 | 6.21 | 6.37 | 6.53 | 6.69 | 6.84 | 7 | 7.16 | 7.32 | 7.48 | 7.63 | 7.79 | 7.95 | 8.11 | 8.27 | 8.43 | 8.58 | 8.74 | 8.9 | 9.06 | 9.22 | 9.38 | 9.53 | 9.69 | 9.85 | 10.01 | 10.17 | 10.33 | 10.48 | 10.64 | 10.8 | 10.96 | 11.12 | 11.27 | 11.43 | 11.59 | 11.75 | 11.91 | 12.07 | 12.22 | 12.38 | 12.54
7.79 | 6.21 | 6.37 | 6.53 | 6.68 | 6.84 | 7 | 7.16 | 7.32 | 7.47 | 7.63 | 7.79 | 7.95 | 8.11 | 8.27 | 8.42 | 8.58 | 8.74 | 8.9 | 9.06 | 9.22 | 9.37 | 9.53 | 9.69 | 9.85 | 10.01 | 10.16 | 10.32 | 10.48 | 10.64 | 10.8 | 10.96 | 11.11 | 11.27 | 11.43 | 11.59 | 11.75 | 11.9 | 12.06 | 12.22 | 12.38 | 12.54
7.8 | 6.21 | 6.37 | 6.52 | 6.68 | 6.84 | 7 | 7.16 | 7.31 | 7.47 | 7.63 | 7.79 | 7.95 | 8.11 | 8.26 | 8.42 | 8.58 | 8.74 | 8.9 | 9.05 | 9.21 | 9.37 | 9.53 | 9.69 | 9.85 | 10 | 10.16 | 10.32 | 10.48 | 10.64 | 10.79 | 10.95 | 11.11 | 11.27 | 11.43 | 11.59 | 11.74 | 11.9 | 12.06 | 12.22 | 12.38 | 12.53
7.81 | 6.21 | 6.36 | 6.52 | 6.68 | 6.84 | 7 | 7.15 | 7.31 | 7.47 | 7.63 | 7.79 | 7.95 | 8.1 | 8.26 | 8.42 | 8.58 | 8.74 | 8.89 | 9.05 | 9.21 | 9.37 | 9.53 | 9.69 | 9.84 | 10 | 10.16 | 10.32 | 10.48 | 10.63 | 10.79 | 10.95 | 11.11 | 11.27 | 11.42 | 11.58 | 11.74 | 11.9 | 12.06 | 12.21 | 12.37 | 12.53
7.82 | 6.2 | 6.36 | 6.52 | 6.68 | 6.84 | 6.99 | 7.15 | 7.31 | 7.47 | 7.63 | 7.79 | 7.94 | 8.1 | 8.26 | 8.42 | 8.58 | 8.73 | 8.89 | 9.05 | 9.21 | 9.37 | 9.52 | 9.68 | 9.84 | 10 | 10.16 | 10.32 | 10.47 | 10.63 | 10.79 | 10.95 | 11.11 | 11.26 | 11.42 | 11.58 | 11.74 | 11.9 | 12.05 | 12.21 | 12.37 | 12.53
7.83 | 6.2 | 6.36 | 6.52 | 6.68 | 6.83 | 6.99 | 7.15 | 7.31 | 7.47 | 7.63 | 7.78 | 7.94 | 8.1 | 8.26 | 8.42 | 8.57 | 8.73 | 8.89 | 9.05 | 9.21 | 9.36 | 9.52 | 9.68 | 9.84 | 10 | 10.15 | 10.31 | 10.47 | 10.63 | 10.79 | 10.94 | 11.1 | 11.26 | 11.42 | 11.58 | 11.73 | 11.89 | 12.05 | 12.21 | 12.37 | 12.52
7.84 | 6.2 | 6.36 | 6.52 | 6.68 | 6.83 | 6.99 | 7.15 | 7.31 | 7.47 | 7.62 | 7.78 | 7.94 | 8.1 | 8.26 | 8.41 | 8.57 | 8.73 | 8.89 | 9.05 | 9.2 | 9.36 | 9.52 | 9.68 | 9.84 | 9.99 | 10.15 | 10.31 | 10.47 | 10.63 | 10.78 | 10.94 | 11.1 | 11.26 | 11.42 | 11.57 | 11.73 | 11.89 | 12.05 | 12.21 | 12.36 | 12.52
7.85 | 6.2 | 6.36 | 6.52 | 6.67 | 6.83 | 6.99 | 7.15 | 7.31 | 7.46 | 7.62 | 7.78 | 7.94 | 8.1 | 8.25 | 8.41 | 8.57 | 8.73 | 8.89 | 9.04 | 9.2 | 9.36 | 9.52 | 9.68 | 9.83 | 9.99 | 10.15 | 10.31 | 10.47 | 10.62 | 10.78 | 10.94 | 11.1 | 11.26 | 11.41 | 11.57 | 11.73 | 11.89 | 12.04 | 12.2 | 12.36 | 12.52
7.86 | 6.2 | 6.36 | 6.51 | 6.67 | 6.83 | 6.99 | 7.15 | 7.3 | 7.46 | 7.62 | 7.78 | 7.94 | 8.09 | 8.25 | 8.41 | 8.57 | 8.73 | 8.88 | 9.04 | 9.2 | 9.36 | 9.52 | 9.67 | 9.83 | 9.99 | 10.15 | 10.3 | 10.46 | 10.62 | 10.78 | 10.94 | 11.09 | 11.25 | 11.41 | 11.57 | 11.73 | 11.88 | 12.04 | 12.2 | 12.36 | 12.52
7.87 | 6.2 | 6.35 | 6.51 | 6.67 | 6.83 | 6.99 | 7.14 | 7.3 | 7.46 | 7.62 | 7.78 | 7.93 | 8.09 | 8.25 | 8.41 | 8.57 | 8.72 | 8.88 | 9.04 | 9.2 | 9.35 | 9.51 | 9.67 | 9.83 | 9.99 | 10.14 | 10.3 | 10.46 | 10.62 | 10.78 | 10.93 | 11.09 | 11.25 | 11.41 | 11.57 | 11.72 | 11.88 | 12.04 | 12.2 | 12.35 | 12.51
7.88 | 6.19 | 6.35 | 6.51 | 6.67 | 6.83 | 6.98 | 7.14 | 7.3 | 7.46 | 7.62 | 7.77 | 7.93 | 8.09 | 8.25 | 8.41 | 8.56 | 8.72 | 8.88 | 9.04 | 9.19 | 9.35 | 9.51 | 9.67 | 9.83 | 9.98 | 10.14 | 10.3 | 10.46 | 10.62 | 10.77 | 10.93 | 11.09 | 11.25 | 11.4 | 11.56 | 11.72 | 11.88 | 12.04 | 12.19 | 12.35 | 12.51
7.89 | 6.19 | 6.35 | 6.51 | 6.67 | 6.82 | 6.98 | 7.14 | 7.3 | 7.46 | 7.61 | 7.77 | 7.93 | 8.09 | 8.25 | 8.4 | 8.56 | 8.72 | 8.88 | 9.03 | 9.19 | 9.35 | 9.51 | 9.67 | 9.82 | 9.98 | 10.14 | 10.3 | 10.46 | 10.61 | 10.77 | 10.93 | 11.09 | 11.24 | 11.4 | 11.56 | 11.72 | 11.88 | 12.03 | 12.19 | 12.35 | 12.51
7.9 | 6.19 | 6.35 | 6.51 | 6.66 | 6.82 | 6.98 | 7.14 | 7.3 | 7.45 | 7.61 | 7.77 | 7.93 | 8.09 | 8.24 | 8.4 | 8.56 | 8.72 | 8.87 | 9.03 | 9.19 | 9.35 | 9.51 | 9.66 | 9.82 | 9.98 | 10.14 | 10.29 | 10.45 | 10.61 | 10.77 | 10.93 | 11.08 | 11.24 | 11.4 | 11.56 | 11.71 | 11.87 | 12.03 | 12.19 | 12.35 | 12.5
7.91 | 6.19 | 6.35 | 6.51 | 6.66 | 6.82 | 6.98 | 7.14 | 7.29 | 7.45 | 7.61 | 7.77 | 7.93 | 8.08 | 8.24 | 8.4 | 8.56 | 8.71 | 8.87 | 9.03 | 9.19 | 9.35 | 9.5 | 9.66 | 9.82 | 9.98 | 10.13 | 10.29 | 10.45 | 10.61 | 10.77 | 10.92 | 11.08 | 11.24 | 11.4 | 11.55 | 11.71 | 11.87 | 12.03 | 12.18 | 12.34 | 12.5
7.92 | 6.19 | 6.35 | 6.5 | 6.66 | 6.82 | 6.98 | 7.13 | 7.29 | 7.45 | 7.61 | 7.77 | 7.92 | 8.08 | 8.24 | 8.4 | 8.55 | 8.71 | 8.87 | 9.03 | 9.19 | 9.34 | 9.5 | 9.66 | 9.82 | 9.97 | 10.13 | 10.29 | 10.45 | 10.61 | 10.76 | 10.92 | 11.08 | 11.24 | 11.39 | 11.55 | 11.71 | 11.87 | 12.02 | 12.18 | 12.34 | 12.5
7.93 | 6.19 | 6.34 | 6.5 | 6.66 | 6.82 | 6.98 | 7.13 | 7.29 | 7.45 | 7.61 | 7.76 | 7.92 | 8.08 | 8.24 | 8.39 | 8.55 | 8.71 | 8.87 | 9.03 | 9.18 | 9.34 | 9.5 | 9.66 | 9.81 | 9.97 | 10.13 | 10.29 | 10.44 | 10.6 | 10.76 | 10.92 | 11.08 | 11.23 | 11.39 | 11.55 | 11.71 | 11.86 | 12.02 | 12.18 | 12.34 | 12.49
7.94 | 6.19 | 6.34 | 6.5 | 6.66 | 6.82 | 6.97 | 7.13 | 7.29 | 7.45 | 7.6 | 7.76 | 7.92 | 8.08 | 8.24 | 8.39 | 8.55 | 8.71 | 8.87 | 9.02 | 9.18 | 9.34 | 9.5 | 9.65 | 9.81 | 9.97 | 10.13 | 10.28 | 10.44 | 10.6 | 10.76 | 10.92 | 11.07 | 11.23 | 11.39 | 11.55 | 11.7 | 11.86 | 12.02 | 12.18 | 12.33 | 12.49
7.95 | 6.18 | 6.34 | 6.5 | 6.66 | 6.81 | 6.97 | 7.13 | 7.29 | 7.44 | 7.6 | 7.76 | 7.92 | 8.08 | 8.23 | 8.39 | 8.55 | 8.71 | 8.86 | 9.02 | 9.18 | 9.34 | 9.49 | 9.65 | 9.81 | 9.97 | 10.12 | 10.28 | 10.44 | 10.6 | 10.75 | 10.91 | 11.07 | 11.23 | 11.39 | 11.54 | 11.7 | 11.86 | 12.02 | 12.17 | 12.33 | 12.49
7.96 | 6.18 | 6.34 | 6.5 | 6.65 | 6.81 | 6.97 | 7.13 | 7.29 | 7.44 | 7.6 | 7.76 | 7.92 | 8.07 | 8.23 | 8.39 | 8.55 | 8.7 | 8.86 | 9.02 | 9.18 | 9.33 | 9.49 | 9.65 | 9.81 | 9.96 | 10.12 | 10.28 | 10.44 | 10.59 | 10.75 | 10.91 | 11.07 | 11.22 | 11.38 | 11.54 | 11.7 | 11.85 | 12.01 | 12.17 | 12.33 | 12.48
7.97 | 6.18 | 6.34 | 6.5 | 6.65 | 6.81 | 6.97 | 7.13 | 7.28 | 7.44 | 7.6 | 7.76 | 7.91 | 8.07 | 8.23 | 8.39 | 8.54 | 8.7 | 8.86 | 9.02 | 9.17 | 9.33 | 9.49 | 9.65 | 9.8 | 9.96 | 10.12 | 10.28 | 10.43 | 10.59 | 10.75 | 10.91 | 11.06 | 11.22 | 11.38 | 11.54 | 11.69 | 11.85 | 12.01 | 12.17 | 12.32 | 12.48
7.98 | 6.18 | 6.34 | 6.49 | 6.65 | 6.81 | 6.97 | 7.12 | 7.28 | 7.44 | 7.6 | 7.75 | 7.91 | 8.07 | 8.23 | 8.38 | 8.54 | 8.7 | 8.86 | 9.01 | 9.17 | 9.33 | 9.49 | 9.64 | 9.8 | 9.96 | 10.12 | 10.27 | 10.43 | 10.59 | 10.75 | 10.9 | 11.06 | 11.22 | 11.38 | 11.53 | 11.69 | 11.85 | 12.01 | 12.16 | 12.32 | 12.48
7.99 | 6.18 | 6.33 | 6.49 | 6.65 | 6.81 | 6.97 | 7.12 | 7.28 | 7.44 | 7.6 | 7.75 | 7.91 | 8.07 | 8.23 | 8.38 | 8.54 | 8.7 | 8.85 | 9.01 | 9.17 | 9.33 | 9.48 | 9.64 | 9.8 | 9.96 | 10.11 | 10.27 | 10.43 | 10.59 | 10.74 | 10.9 | 11.06 | 11.22 | 11.37 | 11.53 | 11.69 | 11.85 | 12 | 12.16 | 12.32 | 12.48
8 | 6.18 | 6.33 | 6.49 | 6.65 | 6.81 | 6.96 | 7.12 | 7.28 | 7.44 | 7.59 | 7.75 | 7.91 | 8.07 | 8.22 | 8.38 | 8.54 | 8.7 | 8.85 | 9.01 | 9.17 | 9.33 | 9.48 | 9.64 | 9.8 | 9.95 | 10.11 | 10.27 | 10.43 | 10.58 | 10.74 | 10.9 | 11.06 | 11.21 | 11.37 | 11.53 | 11.69 | 11.84 | 12 | 12.16 | 12.32 | 12.47
8.01 | 6.17 | 6.33 | 6.49 | 6.65 | 6.8 | 6.96 | 7.12 | 7.28 | 7.43 | 7.59 | 7.75 | 7.91 | 8.06 | 8.22 | 8.38 | 8.54 | 8.69 | 8.85 | 9.01 | 9.17 | 9.32 | 9.48 | 9.64 | 9.79 | 9.95 | 10.11 | 10.27 | 10.42 | 10.58 | 10.74 | 10.9 | 11.05 | 11.21 | 11.37 | 11.53 | 11.68 | 11.84 | 12 | 12.16 | 12.31 | 12.47
8.02 | 6.17 | 6.33 | 6.49 | 6.64 | 6.8 | 6.96 | 7.12 | 7.27 | 7.43 | 7.59 | 7.75 | 7.9 | 8.06 | 8.22 | 8.38 | 8.53 | 8.69 | 8.85 | 9.01 | 9.16 | 9.32 | 9.48 | 9.64 | 9.79 | 9.95 | 10.11 | 10.26 | 10.42 | 10.58 | 10.74 | 10.89 | 11.05 | 11.21 | 11.37 | 11.52 | 11.68 | 11.84 | 11.99 | 12.15 | 12.31 | 12.47
8.03 | 6.17 | 6.33 | 6.49 | 6.64 | 6.8 | 6.96 | 7.12 | 7.27 | 7.43 | 7.59 | 7.74 | 7.9 | 8.06 | 8.22 | 8.37 | 8.53 | 8.69 | 8.85 | 9 | 9.16 | 9.32 | 9.48 | 9.63 | 9.79 | 9.95 | 10.1 | 10.26 | 10.42 | 10.58 | 10.73 | 10.89 | 11.05 | 11.21 | 11.36 | 11.52 | 11.68 | 11.83 | 11.99 | 12.15 | 12.31 | 12.46
8.04 | 6.17 | 6.33 | 6.48 | 6.64 | 6.8 | 6.96 | 7.11 | 7.27 | 7.43 | 7.59 | 7.74 | 7.9 | 8.06 | 8.21 | 8.37 | 8.53 | 8.69 | 8.84 | 9 | 9.16 | 9.32 | 9.47 | 9.63 | 9.79 | 9.94 | 10.1 | 10.26 | 10.42 | 10.57 | 10.73 | 10.89 | 11.05 | 11.2 | 11.36 | 11.52 | 11.67 | 11.83 | 11.99 | 12.15 | 12.3 | 12.46
8.05 | 6.17 | 6.33 | 6.48 | 6.64 | 6.8 | 6.95 | 7.11 | 7.27 | 7.43 | 7.58 | 7.74 | 7.9 | 8.06 | 8.21 | 8.37 | 8.53 | 8.68 | 8.84 | 9 | 9.16 | 9.31 | 9.47 | 9.63 | 9.79 | 9.94 | 10.1 | 10.26 | 10.41 | 10.57 | 10.73 | 10.89 | 11.04 | 11.2 | 11.36 | 11.51 | 11.67 | 11.83 | 11.99 | 12.14 | 12.3 | 12.46
8.06 | 6.17 | 6.32 | 6.48 | 6.64 | 6.8 | 6.95 | 7.11 | 7.27 | 7.42 | 7.58 | 7.74 | 7.9 | 8.05 | 8.21 | 8.37 | 8.53 | 8.68 | 8.84 | 9 | 9.15 | 9.31 | 9.47 | 9.63 | 9.78 | 9.94 | 10.1 | 10.25 | 10.41 | 10.57 | 10.73 | 10.88 | 11.04 | 11.2 | 11.35 | 11.51 | 11.67 | 11.83 | 11.98 | 12.14 | 12.3 | 12.45
8.07 | 6.16 | 6.32 | 6.48 | 6.64 | 6.79 | 6.95 | 7.11 | 7.27 | 7.42 | 7.58 | 7.74 | 7.89 | 8.05 | 8.21 | 8.37 | 8.52 | 8.68 | 8.84 | 8.99 | 9.15 | 9.31 | 9.47 | 9.62 | 9.78 | 9.94 | 10.09 | 10.25 | 10.41 | 10.57 | 10.72 | 10.88 | 11.04 | 11.19 | 11.35 | 11.51 | 11.67 | 11.82 | 11.98 | 12.14 | 12.29 | 12.45
8.08 | 6.16 | 6.32 | 6.48 | 6.64 | 6.79 | 6.95 | 7.11 | 7.26 | 7.42 | 7.58 | 7.74 | 7.89 | 8.05 | 8.21 | 8.36 | 8.52 | 8.68 | 8.84 | 8.99 | 9.15 | 9.31 | 9.46 | 9.62 | 9.78 | 9.94 | 10.09 | 10.25 | 10.41 | 10.56 | 10.72 | 10.88 | 11.03 | 11.19 | 11.35 | 11.51 | 11.66 | 11.82 | 11.98 | 12.13 | 12.29 | 12.45
8.09 | 6.16 | 6.32 | 6.48 | 6.63 | 6.79 | 6.95 | 7.1 | 7.26 | 7.42 | 7.58 | 7.73 | 7.89 | 8.05 | 8.2 | 8.36 | 8.52 | 8.68 | 8.83 | 8.99 | 9.15 | 9.3 | 9.46 | 9.62 | 9.78 | 9.93 | 10.09 | 10.25 | 10.4 | 10.56 | 10.72 | 10.88 | 11.03 | 11.19 | 11.35 | 11.5 | 11.66 | 11.82 | 11.97 | 12.13 | 12.29 | 12.45
8.1 | 6.16 | 6.32 | 6.47 | 6.63 | 6.79 | 6.95 | 7.1 | 7.26 | 7.42 | 7.57 | 7.73 | 7.89 | 8.05 | 8.2 | 8.36 | 8.52 | 8.67 | 8.83 | 8.99 | 9.15 | 9.3 | 9.46 | 9.62 | 9.77 | 9.93 | 10.09 | 10.24 | 10.4 | 10.56 | 10.72 | 10.87 | 11.03 | 11.19 | 11.34 | 11.5 | 11.66 | 11.81 | 11.97 | 12.13 | 12.29 | 12.44
8.11 | 6.16 | 6.32 | 6.47 | 6.63 | 6.79 | 6.94 | 7.1 | 7.26 | 7.42 | 7.57 | 7.73 | 7.89 | 8.04 | 8.2 | 8.36 | 8.51 | 8.67 | 8.83 | 8.99 | 9.14 | 9.3 | 9.46 | 9.61 | 9.77 | 9.93 | 10.08 | 10.24 | 10.4 | 10.56 | 10.71 | 10.87 | 11.03 | 11.18 | 11.34 | 11.5 | 11.65 | 11.81 | 11.97 | 12.13 | 12.28 | 12.44
8.12 | 6.16 | 6.31 | 6.47 | 6.63 | 6.79 | 6.94 | 7.1 | 7.26 | 7.41 | 7.57 | 7.73 | 7.88 | 8.04 | 8.2 | 8.36 | 8.51 | 8.67 | 8.83 | 8.98 | 9.14 | 9.3 | 9.45 | 9.61 | 9.77 | 9.93 | 10.08 | 10.24 | 10.4 | 10.55 | 10.71 | 10.87 | 11.02 | 11.18 | 11.34 | 11.49 | 11.65 | 11.81 | 11.97 | 12.12 | 12.28 | 12.44
8.13 | 6.16 | 6.31 | 6.47 | 6.63 | 6.78 | 6.94 | 7.1 | 7.25 | 7.41 | 7.57 | 7.73 | 7.88 | 8.04 | 8.2 | 8.35 | 8.51 | 8.67 | 8.82 | 8.98 | 9.14 | 9.3 | 9.45 | 9.61 | 9.77 | 9.92 | 10.08 | 10.24 | 10.39 | 10.55 | 10.71 | 10.86 | 11.02 | 11.18 | 11.34 | 11.49 | 11.65 | 11.81 | 11.96 | 12.12 | 12.28 | 12.43
8.14 | 6.15 | 6.31 | 6.47 | 6.63 | 6.78 | 6.94 | 7.1 | 7.25 | 7.41 | 7.57 | 7.72 | 7.88 | 8.04 | 8.19 | 8.35 | 8.51 | 8.67 | 8.82 | 8.98 | 9.14 | 9.29 | 9.45 | 9.61 | 9.76 | 9.92 | 10.08 | 10.23 | 10.39 | 10.55 | 10.7 | 10.86 | 11.02 | 11.18 | 11.33 | 11.49 | 11.65 | 11.8 | 11.96 | 12.12 | 12.27 | 12.43
8.15 | 6.15 | 6.31 | 6.47 | 6.62 | 6.78 | 6.94 | 7.09 | 7.25 | 7.41 | 7.56 | 7.72 | 7.88 | 8.04 | 8.19 | 8.35 | 8.51 | 8.66 | 8.82 | 8.98 | 9.13 | 9.29 | 9.45 | 9.6 | 9.76 | 9.92 | 10.07 | 10.23 | 10.39 | 10.55 | 10.7 | 10.86 | 11.02 | 11.17 | 11.33 | 11.49 | 11.64 | 11.8 | 11.96 | 12.11 | 12.27 | 12.43
8.16 | 6.15 | 6.31 | 6.46 | 6.62 | 6.78 | 6.94 | 7.09 | 7.25 | 7.41 | 7.56 | 7.72 | 7.88 | 8.03 | 8.19 | 8.35 | 8.5 | 8.66 | 8.82 | 8.97 | 9.13 | 9.29 | 9.45 | 9.6 | 9.76 | 9.92 | 10.07 | 10.23 | 10.39 | 10.54 | 10.7 | 10.86 | 11.01 | 11.17 | 11.33 | 11.48 | 11.64 | 11.8 | 11.95 | 12.11 | 12.27 | 12.42
8.17 | 6.15 | 6.31 | 6.46 | 6.62 | 6.78 | 6.93 | 7.09 | 7.25 | 7.4 | 7.56 | 7.72 | 7.87 | 8.03 | 8.19 | 8.35 | 8.5 | 8.66 | 8.82 | 8.97 | 9.13 | 9.29 | 9.44 | 9.6 | 9.76 | 9.91 | 10.07 | 10.23 | 10.38 | 10.54 | 10.7 | 10.85 | 11.01 | 11.17 | 11.32 | 11.48 | 11.64 | 11.79 | 11.95 | 12.11 | 12.26 | 12.42
8.18 | 6.15 | 6.3 | 6.46 | 6.62 | 6.78 | 6.93 | 7.09 | 7.25 | 7.4 | 7.56 | 7.72 | 7.87 | 8.03 | 8.19 | 8.34 | 8.5 | 8.66 | 8.81 | 8.97 | 9.13 | 9.28 | 9.44 | 9.6 | 9.75 | 9.91 | 10.07 | 10.22 | 10.38 | 10.54 | 10.69 | 10.85 | 11.01 | 11.16 | 11.32 | 11.48 | 11.63 | 11.79 | 11.95 | 12.1 | 12.26 | 12.42
8.19 | 6.15 | 6.3 | 6.46 | 6.62 | 6.77 | 6.93 | 7.09 | 7.24 | 7.4 | 7.56 | 7.71 | 7.87 | 8.03 | 8.18 | 8.34 | 8.5 | 8.65 | 8.81 | 8.97 | 9.12 | 9.28 | 9.44 | 9.59 | 9.75 | 9.91 | 10.07 | 10.22 | 10.38 | 10.54 | 10.69 | 10.85 | 11.01 | 11.16 | 11.32 | 11.48 | 11.63 | 11.79 | 11.95 | 12.1 | 12.26 | 12.42
8.2 | 6.14 | 6.3 | 6.46 | 6.62 | 6.77 | 6.93 | 7.09 | 7.24 | 7.4 | 7.56 | 7.71 | 7.87 | 8.03 | 8.18 | 8.34 | 8.5 | 8.65 | 8.81 | 8.97 | 9.12 | 9.28 | 9.44 | 9.59 | 9.75 | 9.91 | 10.06 | 10.22 | 10.38 | 10.53 | 10.69 | 10.85 | 11 | 11.16 | 11.32 | 11.47 | 11.63 | 11.79 | 11.94 | 12.1 | 12.26 | 12.41
8.21 | 6.14 | 6.3 | 6.46 | 6.61 | 6.77 | 6.93 | 7.08 | 7.24 | 7.4 | 7.55 | 7.71 | 7.87 | 8.02 | 8.18 | 8.34 | 8.49 | 8.65 | 8.81 | 8.96 | 9.12 | 9.28 | 9.43 | 9.59 | 9.75 | 9.9 | 10.06 | 10.22 | 10.37 | 10.53 | 10.69 | 10.84 | 11 | 11.16 | 11.31 | 11.47 | 11.63 | 11.78 | 11.94 | 12.1 | 12.25 | 12.41
8.22 | 6.14 | 6.3 | 6.46 | 6.61 | 6.77 | 6.93 | 7.08 | 7.24 | 7.4 | 7.55 | 7.71 | 7.87 | 8.02 | 8.18 | 8.33 | 8.49 | 8.65 | 8.8 | 8.96 | 9.12 | 9.27 | 9.43 | 9.59 | 9.74 | 9.9 | 10.06 | 10.21 | 10.37 | 10.53 | 10.68 | 10.84 | 11 | 11.15 | 11.31 | 11.47 | 11.62 | 11.78 | 11.94 | 12.09 | 12.25 | 12.41
8.23 | 6.14 | 6.3 | 6.45 | 6.61 | 6.77 | 6.92 | 7.08 | 7.24 | 7.39 | 7.55 | 7.71 | 7.86 | 8.02 | 8.18 | 8.33 | 8.49 | 8.65 | 8.8 | 8.96 | 9.12 | 9.27 | 9.43 | 9.59 | 9.74 | 9.9 | 10.06 | 10.21 | 10.37 | 10.52 | 10.68 | 10.84 | 10.99 | 11.15 | 11.31 | 11.46 | 11.62 | 11.78 | 11.93 | 12.09 | 12.25 | 12.4
8.24 | 6.14 | 6.3 | 6.45 | 6.61 | 6.77 | 6.92 | 7.08 | 7.23 | 7.39 | 7.55 | 7.7 | 7.86 | 8.02 | 8.17 | 8.33 | 8.49 | 8.64 | 8.8 | 8.96 | 9.11 | 9.27 | 9.43 | 9.58 | 9.74 | 9.9 | 10.05 | 10.21 | 10.37 | 10.52 | 10.68 | 10.84 | 10.99 | 11.15 | 11.3 | 11.46 | 11.62 | 11.77 | 11.93 | 12.09 | 12.24 | 12.4
8.25 | 6.14 | 6.29 | 6.45 | 6.61 | 6.76 | 6.92 | 7.08 | 7.23 | 7.39 | 7.55 | 7.7 | 7.86 | 8.02 | 8.17 | 8.33 | 8.49 | 8.64 | 8.8 | 8.95 | 9.11 | 9.27 | 9.42 | 9.58 | 9.74 | 9.89 | 10.05 | 10.21 | 10.36 | 10.52 | 10.68 | 10.83 | 10.99 | 11.15 | 11.3 | 11.46 | 11.61 | 11.77 | 11.93 | 12.08 | 12.24 | 12.4
8.26 | 6.14 | 6.29 | 6.45 | 6.61 | 6.76 | 6.92 | 7.07 | 7.23 | 7.39 | 7.54 | 7.7 | 7.86 | 8.01 | 8.17 | 8.33 | 8.48 | 8.64 | 8.8 | 8.95 | 9.11 | 9.27 | 9.42 | 9.58 | 9.73 | 9.89 | 10.05 | 10.2 | 10.36 | 10.52 | 10.67 | 10.83 | 10.99 | 11.14 | 11.3 | 11.46 | 11.61 | 11.77 | 11.92 | 12.08 | 12.24 | 12.39
8.27 | 6.13 | 6.29 | 6.45 | 6.6 | 6.76 | 6.92 | 7.07 | 7.23 | 7.39 | 7.54 | 7.7 | 7.86 | 8.01 | 8.17 | 8.32 | 8.48 | 8.64 | 8.79 | 8.95 | 9.11 | 9.26 | 9.42 | 9.58 | 9.73 | 9.89 | 10.05 | 10.2 | 10.36 | 10.51 | 10.67 | 10.83 | 10.98 | 11.14 | 11.3 | 11.45 | 11.61 | 11.77 | 11.92 | 12.08 | 12.23 | 12.39
8.28 | 6.13 | 6.29 | 6.45 | 6.6 | 6.76 | 6.91 | 7.07 | 7.23 | 7.38 | 7.54 | 7.7 | 7.85 | 8.01 | 8.17 | 8.32 | 8.48 | 8.64 | 8.79 | 8.95 | 9.1 | 9.26 | 9.42 | 9.57 | 9.73 | 9.89 | 10.04 | 10.2 | 10.36 | 10.51 | 10.67 | 10.82 | 10.98 | 11.14 | 11.29 | 11.45 | 11.61 | 11.76 | 11.92 | 12.08 | 12.23 | 12.39
8.29 | 6.13 | 6.29 | 6.44 | 6.6 | 6.76 | 6.91 | 7.07 | 7.23 | 7.38 | 7.54 | 7.7 | 7.85 | 8.01 | 8.16 | 8.32 | 8.48 | 8.63 | 8.79 | 8.95 | 9.1 | 9.26 | 9.42 | 9.57 | 9.73 | 9.88 | 10.04 | 10.2 | 10.35 | 10.51 | 10.67 | 10.82 | 10.98 | 11.13 | 11.29 | 11.45 | 11.6 | 11.76 | 11.92 | 12.07 | 12.23 | 12.38
8.3 | 6.13 | 6.29 | 6.44 | 6.6 | 6.76 | 6.91 | 7.07 | 7.22 | 7.38 | 7.54 | 7.69 | 7.85 | 8.01 | 8.16 | 8.32 | 8.47 | 8.63 | 8.79 | 8.94 | 9.1 | 9.26 | 9.41 | 9.57 | 9.73 | 9.88 | 10.04 | 10.19 | 10.35 | 10.51 | 10.66 | 10.82 | 10.98 | 11.13 | 11.29 | 11.44 | 11.6 | 11.76 | 11.91 | 12.07 | 12.23 | 12.38
8.31 | 6.13 | 6.28 | 6.44 | 6.6 | 6.75 | 6.91 | 7.07 | 7.22 | 7.38 | 7.53 | 7.69 | 7.85 | 8 | 8.16 | 8.32 | 8.47 | 8.63 | 8.79 | 8.94 | 9.1 | 9.25 | 9.41 | 9.57 | 9.72 | 9.88 | 10.04 | 10.19 | 10.35 | 10.5 | 10.66 | 10.82 | 10.97 | 11.13 | 11.29 | 11.44 | 11.6 | 11.75 | 11.91 | 12.07 | 12.22 | 12.38
8.32 | 6.13 | 6.28 | 6.44 | 6.6 | 6.75 | 6.91 | 7.06 | 7.22 | 7.38 | 7.53 | 7.69 | 7.85 | 8 | 8.16 | 8.31 | 8.47 | 8.63 | 8.78 | 8.94 | 9.1 | 9.25 | 9.41 | 9.56 | 9.72 | 9.88 | 10.03 | 10.19 | 10.35 | 10.5 | 10.66 | 10.81 | 10.97 | 11.13 | 11.28 | 11.44 | 11.6 | 11.75 | 11.91 | 12.06 | 12.22 | 12.38
8.33 | 6.12 | 6.28 | 6.44 | 6.59 | 6.75 | 6.91 | 7.06 | 7.22 | 7.38 | 7.53 | 7.69 | 7.84 | 8 | 8.16 | 8.31 | 8.47 | 8.62 | 8.78 | 8.94 | 9.09 | 9.25 | 9.41 | 9.56 | 9.72 | 9.87 | 10.03 | 10.19 | 10.34 | 10.5 | 10.66 | 10.81 | 10.97 | 11.12 | 11.28 | 11.44 | 11.59 | 11.75 | 11.9 | 12.06 | 12.22 | 12.37
8.34 | 6.12 | 6.28 | 6.44 | 6.59 | 6.75 | 6.9 | 7.06 | 7.22 | 7.37 | 7.53 | 7.69 | 7.84 | 8 | 8.15 | 8.31 | 8.47 | 8.62 | 8.78 | 8.94 | 9.09 | 9.25 | 9.4 | 9.56 | 9.72 | 9.87 | 10.03 | 10.18 | 10.34 | 10.5 | 10.65 | 10.81 | 10.96 | 11.12 | 11.28 | 11.43 | 11.59 | 11.75 | 11.9 | 12.06 | 12.21 | 12.37
8.35 | 6.12 | 6.28 | 6.43 | 6.59 | 6.75 | 6.9 | 7.06 | 7.22 | 7.37 | 7.53 | 7.68 | 7.84 | 8 | 8.15 | 8.31 | 8.46 | 8.62 | 8.78 | 8.93 | 9.09 | 9.25 | 9.4 | 9.56 | 9.71 | 9.87 | 10.03 | 10.18 | 10.34 | 10.49 | 10.65 | 10.81 | 10.96 | 11.12 | 11.27 | 11.43 | 11.59 | 11.74 | 11.9 | 12.05 | 12.21 | 12.37
8.36 | 6.12 | 6.28 | 6.43 | 6.59 | 6.74 | 6.9 | 7.06 | 7.21 | 7.37 | 7.53 | 7.68 | 7.84 | 7.99 | 8.15 | 8.31 | 8.46 | 8.62 | 8.77 | 8.93 | 9.09 | 9.24 | 9.4 | 9.56 | 9.71 | 9.87 | 10.02 | 10.18 | 10.34 | 10.49 | 10.65 | 10.8 | 10.96 | 11.12 | 11.27 | 11.43 | 11.58 | 11.74 | 11.9 | 12.05 | 12.21 | 12.36
8.37 | 6.12 | 6.27 | 6.43 | 6.59 | 6.74 | 6.9 | 7.06 | 7.21 | 7.37 | 7.52 | 7.68 | 7.84 | 7.99 | 8.15 | 8.3 | 8.46 | 8.62 | 8.77 | 8.93 | 9.08 | 9.24 | 9.4 | 9.55 | 9.71 | 9.86 | 10.02 | 10.18 | 10.33 | 10.49 | 10.64 | 10.8 | 10.96 | 11.11 | 11.27 | 11.43 | 11.58 | 11.74 | 11.89 | 12.05 | 12.2 | 12.36
8.38 | 6.12 | 6.27 | 6.43 | 6.59 | 6.74 | 6.9 | 7.05 | 7.21 | 7.37 | 7.52 | 7.68 | 7.83 | 7.99 | 8.15 | 8.3 | 8.46 | 8.61 | 8.77 | 8.93 | 9.08 | 9.24 | 9.39 | 9.55 | 9.71 | 9.86 | 10.02 | 10.17 | 10.33 | 10.49 | 10.64 | 10.8 | 10.95 | 11.11 | 11.27 | 11.42 | 11.58 | 11.73 | 11.89 | 12.05 | 12.2 | 12.36
8.39 | 6.12 | 6.27 | 6.43 | 6.58 | 6.74 | 6.9 | 7.05 | 7.21 | 7.36 | 7.52 | 7.68 | 7.83 | 7.99 | 8.14 | 8.3 | 8.46 | 8.61 | 8.77 | 8.92 | 9.08 | 9.24 | 9.39 | 9.55 | 9.7 | 9.86 | 10.02 | 10.17 | 10.33 | 10.48 | 10.64 | 10.8 | 10.95 | 11.11 | 11.26 | 11.42 | 11.58 | 11.73 | 11.89 | 12.04 | 12.2 | 12.35
8.4 | 6.11 | 6.27 | 6.43 | 6.58 | 6.74 | 6.89 | 7.05 | 7.21 | 7.36 | 7.52 | 7.67 | 7.83 | 7.99 | 8.14 | 8.3 | 8.45 | 8.61 | 8.77 | 8.92 | 9.08 | 9.23 | 9.39 | 9.55 | 9.7 | 9.86 | 10.01 | 10.17 | 10.33 | 10.48 | 10.64 | 10.79 | 10.95 | 11.1 | 11.26 | 11.42 | 11.57 | 11.73 | 11.88 | 12.04 | 12.2 | 12.35
8.41 | 6.11 | 6.27 | 6.42 | 6.58 | 6.74 | 6.89 | 7.05 | 7.2 | 7.36 | 7.52 | 7.67 | 7.83 | 7.98 | 8.14 | 8.3 | 8.45 | 8.61 | 8.76 | 8.92 | 9.08 | 9.23 | 9.39 | 9.54 | 9.7 | 9.86 | 10.01 | 10.17 | 10.32 | 10.48 | 10.63 | 10.79 | 10.95 | 11.1 | 11.26 | 11.41 | 11.57 | 11.73 | 11.88 | 12.04 | 12.19 | 12.35
8.42 | 6.11 | 6.27 | 6.42 | 6.58 | 6.73 | 6.89 | 7.05 | 7.2 | 7.36 | 7.51 | 7.67 | 7.83 | 7.98 | 8.14 | 8.29 | 8.45 | 8.61 | 8.76 | 8.92 | 9.07 | 9.23 | 9.39 | 9.54 | 9.7 | 9.85 | 10.01 | 10.16 | 10.32 | 10.48 | 10.63 | 10.79 | 10.94 | 11.1 | 11.26 | 11.41 | 11.57 | 11.72 | 11.88 | 12.03 | 12.19 | 12.35
8.43 | 6.11 | 6.27 | 6.42 | 6.58 | 6.73 | 6.89 | 7.05 | 7.2 | 7.36 | 7.51 | 7.67 | 7.82 | 7.98 | 8.14 | 8.29 | 8.45 | 8.6 | 8.76 | 8.92 | 9.07 | 9.23 | 9.38 | 9.54 | 9.69 | 9.85 | 10.01 | 10.16 | 10.32 | 10.47 | 10.63 | 10.79 | 10.94 | 11.1 | 11.25 | 11.41 | 11.56 | 11.72 | 11.88 | 12.03 | 12.19 | 12.34
8.44 | 6.11 | 6.26 | 6.42 | 6.58 | 6.73 | 6.89 | 7.04 | 7.2 | 7.35 | 7.51 | 7.67 | 7.82 | 7.98 | 8.13 | 8.29 | 8.45 | 8.6 | 8.76 | 8.91 | 9.07 | 9.22 | 9.38 | 9.54 | 9.69 | 9.85 | 10 | 10.16 | 10.32 | 10.47 | 10.63 | 10.78 | 10.94 | 11.09 | 11.25 | 11.41 | 11.56 | 11.72 | 11.87 | 12.03 | 12.18 | 12.34
8.45 | 6.11 | 6.26 | 6.42 | 6.57 | 6.73 | 6.89 | 7.04 | 7.2 | 7.35 | 7.51 | 7.66 | 7.82 | 7.98 | 8.13 | 8.29 | 8.44 | 8.6 | 8.76 | 8.91 | 9.07 | 9.22 | 9.38 | 9.53 | 9.69 | 9.85 | 10 | 10.16 | 10.31 | 10.47 | 10.62 | 10.78 | 10.94 | 11.09 | 11.25 | 11.4 | 11.56 | 11.71 | 11.87 | 12.03 | 12.18 | 12.34
8.46 | 6.11 | 6.26 | 6.42 | 6.57 | 6.73 | 6.88 | 7.04 | 7.2 | 7.35 | 7.51 | 7.66 | 7.82 | 7.97 | 8.13 | 8.29 | 8.44 | 8.6 | 8.75 | 8.91 | 9.06 | 9.22 | 9.38 | 9.53 | 9.69 | 9.84 | 10 | 10.15 | 10.31 | 10.47 | 10.62 | 10.78 | 10.93 | 11.09 | 11.24 | 11.4 | 11.56 | 11.71 | 11.87 | 12.02 | 12.18 | 12.33
8.47 | 6.1 | 6.26 | 6.42 | 6.57 | 6.73 | 6.88 | 7.04 | 7.19 | 7.35 | 7.51 | 7.66 | 7.82 | 7.97 | 8.13 | 8.28 | 8.44 | 8.6 | 8.75 | 8.91 | 9.06 | 9.22 | 9.37 | 9.53 | 9.69 | 9.84 | 10 | 10.15 | 10.31 | 10.46 | 10.62 | 10.77 | 10.93 | 11.09 | 11.24 | 11.4 | 11.55 | 11.71 | 11.86 | 12.02 | 12.18 | 12.33
8.48 | 6.1 | 6.26 | 6.41 | 6.57 | 6.72 | 6.88 | 7.04 | 7.19 | 7.35 | 7.5 | 7.66 | 7.81 | 7.97 | 8.13 | 8.28 | 8.44 | 8.59 | 8.75 | 8.9 | 9.06 | 9.22 | 9.37 | 9.53 | 9.68 | 9.84 | 9.99 | 10.15 | 10.31 | 10.46 | 10.62 | 10.77 | 10.93 | 11.08 | 11.24 | 11.39 | 11.55 | 11.71 | 11.86 | 12.02 | 12.17 | 12.33
8.49 | 6.1 | 6.26 | 6.41 | 6.57 | 6.72 | 6.88 | 7.03 | 7.19 | 7.35 | 7.5 | 7.66 | 7.81 | 7.97 | 8.12 | 8.28 | 8.44 | 8.59 | 8.75 | 8.9 | 9.06 | 9.21 | 9.37 | 9.52 | 9.68 | 9.84 | 9.99 | 10.15 | 10.3 | 10.46 | 10.61 | 10.77 | 10.93 | 11.08 | 11.24 | 11.39 | 11.55 | 11.7 | 11.86 | 12.01 | 12.17 | 12.33
8.5 | 6.1 | 6.25 | 6.41 | 6.57 | 6.72 | 6.88 | 7.03 | 7.19 | 7.34 | 7.5 | 7.66 | 7.81 | 7.97 | 8.12 | 8.28 | 8.43 | 8.59 | 8.74 | 8.9 | 9.06 | 9.21 | 9.37 | 9.52 | 9.68 | 9.83 | 9.99 | 10.14 | 10.3 | 10.46 | 10.61 | 10.77 | 10.92 | 11.08 | 11.23 | 11.39 | 11.54 | 11.7 | 11.86 | 12.01 | 12.17 | 12.32
8.51 | 6.1 | 6.25 | 6.41 | 6.56 | 6.72 | 6.88 | 7.03 | 7.19 | 7.34 | 7.5 | 7.65 | 7.81 | 7.96 | 8.12 | 8.28 | 8.43 | 8.59 | 8.74 | 8.9 | 9.05 | 9.21 | 9.36 | 9.52 | 9.68 | 9.83 | 9.99 | 10.14 | 10.3 | 10.45 | 10.61 | 10.76 | 10.92 | 11.08 | 11.23 | 11.39 | 11.54 | 11.7 | 11.85 | 12.01 | 12.16 | 12.32
8.52 | 6.1 | 6.25 | 6.41 | 6.56 | 6.72 | 6.87 | 7.03 | 7.18 | 7.34 | 7.5 | 7.65 | 7.81 | 7.96 | 8.12 | 8.27 | 8.43 | 8.58 | 8.74 | 8.9 | 9.05 | 9.21 | 9.36 | 9.52 | 9.67 | 9.83 | 9.98 | 10.14 | 10.3 | 10.45 | 10.61 | 10.76 | 10.92 | 11.07 | 11.23 | 11.38 | 11.54 | 11.69 | 11.85 | 12.01 | 12.16 | 12.32
8.53 | 6.09 | 6.25 | 6.41 | 6.56 | 6.72 | 6.87 | 7.03 | 7.18 | 7.34 | 7.49 | 7.65 | 7.81 | 7.96 | 8.12 | 8.27 | 8.43 | 8.58 | 8.74 | 8.89 | 9.05 | 9.2 | 9.36 | 9.52 | 9.67 | 9.83 | 9.98 | 10.14 | 10.29 | 10.45 | 10.6 | 10.76 | 10.91 | 11.07 | 11.23 | 11.38 | 11.54 | 11.69 | 11.85 | 12 | 12.16 | 12.31
8.54 | 6.09 | 6.25 | 6.4 | 6.56 | 6.71 | 6.87 | 7.03 | 7.18 | 7.34 | 7.49 | 7.65 | 7.8 | 7.96 | 8.11 | 8.27 | 8.43 | 8.58 | 8.74 | 8.89 | 9.05 | 9.2 | 9.36 | 9.51 | 9.67 | 9.82 | 9.98 | 10.13 | 10.29 | 10.45 | 10.6 | 10.76 | 10.91 | 11.07 | 11.22 | 11.38 | 11.53 | 11.69 | 11.84 | 12 | 12.15 | 12.31
8.55 | 6.09 | 6.25 | 6.4 | 6.56 | 6.71 | 6.87 | 7.02 | 7.18 | 7.34 | 7.49 | 7.65 | 7.8 | 7.96 | 8.11 | 8.27 | 8.42 | 8.58 | 8.73 | 8.89 | 9.04 | 9.2 | 9.36 | 9.51 | 9.67 | 9.82 | 9.98 | 10.13 | 10.29 | 10.44 | 10.6 | 10.75 | 10.91 | 11.06 | 11.22 | 11.38 | 11.53 | 11.69 | 11.84 | 12 | 12.15 | 12.31
8.56 | 6.09 | 6.25 | 6.4 | 6.56 | 6.71 | 6.87 | 7.02 | 7.18 | 7.33 | 7.49 | 7.64 | 7.8 | 7.95 | 8.11 | 8.27 | 8.42 | 8.58 | 8.73 | 8.89 | 9.04 | 9.2 | 9.35 | 9.51 | 9.66 | 9.82 | 9.97 | 10.13 | 10.29 | 10.44 | 10.6 | 10.75 | 10.91 | 11.06 | 11.22 | 11.37 | 11.53 | 11.68 | 11.84 | 11.99 | 12.15 | 12.3
8.57 | 6.09 | 6.24 | 6.4 | 6.55 | 6.71 | 6.87 | 7.02 | 7.18 | 7.33 | 7.49 | 7.64 | 7.8 | 7.95 | 8.11 | 8.26 | 8.42 | 8.57 | 8.73 | 8.88 | 9.04 | 9.2 | 9.35 | 9.51 | 9.66 | 9.82 | 9.97 | 10.13 | 10.28 | 10.44 | 10.59 | 10.75 | 10.9 | 11.06 | 11.21 | 11.37 | 11.52 | 11.68 | 11.84 | 11.99 | 12.15 | 12.3
8.58 | 6.09 | 6.24 | 6.4 | 6.55 | 6.71 | 6.86 | 7.02 | 7.17 | 7.33 | 7.48 | 7.64 | 7.8 | 7.95 | 8.11 | 8.26 | 8.42 | 8.57 | 8.73 | 8.88 | 9.04 | 9.19 | 9.35 | 9.5 | 9.66 | 9.81 | 9.97 | 10.12 | 10.28 | 10.44 | 10.59 | 10.75 | 10.9 | 11.06 | 11.21 | 11.37 | 11.52 | 11.68 | 11.83 | 11.99 | 12.14 | 12.3
8.59 | 6.09 | 6.24 | 6.4 | 6.55 | 6.71 | 6.86 | 7.02 | 7.17 | 7.33 | 7.48 | 7.64 | 7.79 | 7.95 | 8.1 | 8.26 | 8.41 | 8.57 | 8.73 | 8.88 | 9.04 | 9.19 | 9.35 | 9.5 | 9.66 | 9.81 | 9.97 | 10.12 | 10.28 | 10.43 | 10.59 | 10.74 | 10.9 | 11.05 | 11.21 | 11.36 | 11.52 | 11.67 | 11.83 | 11.98 | 12.14 | 12.3
8.6 | 6.08 | 6.24 | 6.39 | 6.55 | 6.7 | 6.86 | 7.02 | 7.17 | 7.33 | 7.48 | 7.64 | 7.79 | 7.95 | 8.1 | 8.26 | 8.41 | 8.57 | 8.72 | 8.88 | 9.03 | 9.19 | 9.34 | 9.5 | 9.65 | 9.81 | 9.96 | 10.12 | 10.28 | 10.43 | 10.59 | 10.74 | 10.9 | 11.05 | 11.21 | 11.36 | 11.52 | 11.67 | 11.83 | 11.98 | 12.14 | 12.29
8.61 | 6.08 | 6.24 | 6.39 | 6.55 | 6.7 | 6.86 | 7.01 | 7.17 | 7.32 | 7.48 | 7.63 | 7.79 | 7.95 | 8.1 | 8.26 | 8.41 | 8.57 | 8.72 | 8.88 | 9.03 | 9.19 | 9.34 | 9.5 | 9.65 | 9.81 | 9.96 | 10.12 | 10.27 | 10.43 | 10.58 | 10.74 | 10.89 | 11.05 | 11.2 | 11.36 | 11.51 | 11.67 | 11.82 | 11.98 | 12.13 | 12.29
8.62 | 6.08 | 6.24 | 6.39 | 6.55 | 6.7 | 6.86 | 7.01 | 7.17 | 7.32 | 7.48 | 7.63 | 7.79 | 7.94 | 8.1 | 8.25 | 8.41 | 8.56 | 8.72 | 8.87 | 9.03 | 9.18 | 9.34 | 9.49 | 9.65 | 9.8 | 9.96 | 10.12 | 10.27 | 10.43 | 10.58 | 10.74 | 10.89 | 11.05 | 11.2 | 11.36 | 11.51 | 11.67 | 11.82 | 11.98 | 12.13 | 12.29
8.63 | 6.08 | 6.23 | 6.39 | 6.54 | 6.7 | 6.86 | 7.01 | 7.17 | 7.32 | 7.48 | 7.63 | 7.79 | 7.94 | 8.1 | 8.25 | 8.41 | 8.56 | 8.72 | 8.87 | 9.03 | 9.18 | 9.34 | 9.49 | 9.65 | 9.8 | 9.96 | 10.11 | 10.27 | 10.42 | 10.58 | 10.73 | 10.89 | 11.04 | 11.2 | 11.35 | 11.51 | 11.66 | 11.82 | 11.97 | 12.13 | 12.28
8.64 | 6.08 | 6.23 | 6.39 | 6.54 | 6.7 | 6.85 | 7.01 | 7.16 | 7.32 | 7.47 | 7.63 | 7.78 | 7.94 | 8.09 | 8.25 | 8.4 | 8.56 | 8.71 | 8.87 | 9.02 | 9.18 | 9.33 | 9.49 | 9.65 | 9.8 | 9.96 | 10.11 | 10.27 | 10.42 | 10.58 | 10.73 | 10.89 | 11.04 | 11.2 | 11.35 | 11.51 | 11.66 | 11.82 | 11.97 | 12.13 | 12.28
8.65 | 6.08 | 6.23 | 6.39 | 6.54 | 6.7 | 6.85 | 7.01 | 7.16 | 7.32 | 7.47 | 7.63 | 7.78 | 7.94 | 8.09 | 8.25 | 8.4 | 8.56 | 8.71 | 8.87 | 9.02 | 9.18 | 9.33 | 9.49 | 9.64 | 9.8 | 9.95 | 10.11 | 10.26 | 10.42 | 10.57 | 10.73 | 10.88 | 11.04 | 11.19 | 11.35 | 11.5 | 11.66 | 11.81 | 11.97 | 12.12 | 12.28
8.66 | 6.07 | 6.23 | 6.38 | 6.54 | 6.7 | 6.85 | 7.01 | 7.16 | 7.32 | 7.47 | 7.63 | 7.78 | 7.94 | 8.09 | 8.25 | 8.4 | 8.56 | 8.71 | 8.87 | 9.02 | 9.18 | 9.33 | 9.49 | 9.64 | 9.8 | 9.95 | 10.11 | 10.26 | 10.42 | 10.57 | 10.73 | 10.88 | 11.04 | 11.19 | 11.34 | 11.5 | 11.65 | 11.81 | 11.96 | 12.12 | 12.27
8.67 | 6.07 | 6.23 | 6.38 | 6.54 | 6.69 | 6.85 | 7 | 7.16 | 7.31 | 7.47 | 7.62 | 7.78 | 7.93 | 8.09 | 8.24 | 8.4 | 8.55 | 8.71 | 8.86 | 9.02 | 9.17 | 9.33 | 9.48 | 9.64 | 9.79 | 9.95 | 10.1 | 10.26 | 10.41 | 10.57 | 10.72 | 10.88 | 11.03 | 11.19 | 11.34 | 11.5 | 11.65 | 11.81 | 11.96 | 12.12 | 12.27
8.68 | 6.07 | 6.23 | 6.38 | 6.54 | 6.69 | 6.85 | 7 | 7.16 | 7.31 | 7.47 | 7.62 | 7.78 | 7.93 | 8.09 | 8.24 | 8.4 | 8.55 | 8.71 | 8.86 | 9.02 | 9.17 | 9.33 | 9.48 | 9.64 | 9.79 | 9.95 | 10.1 | 10.26 | 10.41 | 10.57 | 10.72 | 10.87 | 11.03 | 11.18 | 11.34 | 11.49 | 11.65 | 11.8 | 11.96 | 12.11 | 12.27
8.69 | 6.07 | 6.23 | 6.38 | 6.54 | 6.69 | 6.84 | 7 | 7.15 | 7.31 | 7.46 | 7.62 | 7.77 | 7.93 | 8.08 | 8.24 | 8.39 | 8.55 | 8.7 | 8.86 | 9.01 | 9.17 | 9.32 | 9.48 | 9.63 | 9.79 | 9.94 | 10.1 | 10.25 | 10.41 | 10.56 | 10.72 | 10.87 | 11.03 | 11.18 | 11.34 | 11.49 | 11.65 | 11.8 | 11.96 | 12.11 | 12.27
8.7 | 6.07 | 6.22 | 6.38 | 6.53 | 6.69 | 6.84 | 7 | 7.15 | 7.31 | 7.46 | 7.62 | 7.77 | 7.93 | 8.08 | 8.24 | 8.39 | 8.55 | 8.7 | 8.86 | 9.01 | 9.17 | 9.32 | 9.48 | 9.63 | 9.79 | 9.94 | 10.1 | 10.25 | 10.41 | 10.56 | 10.71 | 10.87 | 11.02 | 11.18 | 11.33 | 11.49 | 11.64 | 11.8 | 11.95 | 12.11 | 12.26
8.71 | 6.07 | 6.22 | 6.38 | 6.53 | 6.69 | 6.84 | 7 | 7.15 | 7.31 | 7.46 | 7.62 | 7.77 | 7.93 | 8.08 | 8.24 | 8.39 | 8.54 | 8.7 | 8.85 | 9.01 | 9.16 | 9.32 | 9.47 | 9.63 | 9.78 | 9.94 | 10.09 | 10.25 | 10.4 | 10.56 | 10.71 | 10.87 | 11.02 | 11.18 | 11.33 | 11.49 | 11.64 | 11.8 | 11.95 | 12.11 | 12.26
8.72 | 6.07 | 6.22 | 6.38 | 6.53 | 6.69 | 6.84 | 6.99 | 7.15 | 7.3 | 7.46 | 7.61 | 7.77 | 7.92 | 8.08 | 8.23 | 8.39 | 8.54 | 8.7 | 8.85 | 9.01 | 9.16 | 9.32 | 9.47 | 9.63 | 9.78 | 9.94 | 10.09 | 10.25 | 10.4 | 10.55 | 10.71 | 10.86 | 11.02 | 11.17 | 11.33 | 11.48 | 11.64 | 11.79 | 11.95 | 12.1 | 12.26
8.73 | 6.06 | 6.22 | 6.37 | 6.53 | 6.68 | 6.84 | 6.99 | 7.15 | 7.3 | 7.46 | 7.61 | 7.77 | 7.92 | 8.08 | 8.23 | 8.39 | 8.54 | 8.7 | 8.85 | 9.01 | 9.16 | 9.31 | 9.47 | 9.62 | 9.78 | 9.93 | 10.09 | 10.24 | 10.4 | 10.55 | 10.71 | 10.86 | 11.02 | 11.17 | 11.33 | 11.48 | 11.64 | 11.79 | 11.94 | 12.1 | 12.25
8.74 | 6.06 | 6.22 | 6.37 | 6.53 | 6.68 | 6.84 | 6.99 | 7.15 | 7.3 | 7.46 | 7.61 | 7.77 | 7.92 | 8.07 | 8.23 | 8.38 | 8.54 | 8.69 | 8.85 | 9 | 9.16 | 9.31 | 9.47 | 9.62 | 9.78 | 9.93 | 10.09 | 10.24 | 10.4 | 10.55 | 10.7 | 10.86 | 11.01 | 11.17 | 11.32 | 11.48 | 11.63 | 11.79 | 11.94 | 12.1 | 12.25
8.75 | 6.06 | 6.22 | 6.37 | 6.53 | 6.68 | 6.83 | 6.99 | 7.14 | 7.3 | 7.45 | 7.61 | 7.76 | 7.92 | 8.07 | 8.23 | 8.38 | 8.54 | 8.69 | 8.85 | 9 | 9.16 | 9.31 | 9.46 | 9.62 | 9.77 | 9.93 | 10.08 | 10.24 | 10.39 | 10.55 | 10.7 | 10.86 | 11.01 | 11.17 | 11.32 | 11.47 | 11.63 | 11.78 | 11.94 | 12.09 | 12.25
8.76 | 6.06 | 6.21 | 6.37 | 6.52 | 6.68 | 6.83 | 6.99 | 7.14 | 7.3 | 7.45 | 7.61 | 7.76 | 7.92 | 8.07 | 8.23 | 8.38 | 8.53 | 8.69 | 8.84 | 9 | 9.15 | 9.31 | 9.46 | 9.62 | 9.77 | 9.93 | 10.08 | 10.24 | 10.39 | 10.54 | 10.7 | 10.85 | 11.01 | 11.16 | 11.32 | 11.47 | 11.63 | 11.78 | 11.94 | 12.09 | 12.25
8.77 | 6.06 | 6.21 | 6.37 | 6.52 | 6.68 | 6.83 | 6.99 | 7.14 | 7.3 | 7.45 | 7.6 | 7.76 | 7.91 | 8.07 | 8.22 | 8.38 | 8.53 | 8.69 | 8.84 | 9 | 9.15 | 9.31 | 9.46 | 9.61 | 9.77 | 9.92 | 10.08 | 10.23 | 10.39 | 10.54 | 10.7 | 10.85 | 11.01 | 11.16 | 11.31 | 11.47 | 11.62 | 11.78 | 11.93 | 12.09 | 12.24
8.78 | 6.06 | 6.21 | 6.37 | 6.52 | 6.68 | 6.83 | 6.98 | 7.14 | 7.29 | 7.45 | 7.6 | 7.76 | 7.91 | 8.07 | 8.22 | 8.38 | 8.53 | 8.68 | 8.84 | 8.99 | 9.15 | 9.3 | 9.46 | 9.61 | 9.77 | 9.92 | 10.08 | 10.23 | 10.38 | 10.54 | 10.69 | 10.85 | 11 | 11.16 | 11.31 | 11.47 | 11.62 | 11.78 | 11.93 | 12.08 | 12.24
8.79 | 6.06 | 6.21 | 6.36 | 6.52 | 6.67 | 6.83 | 6.98 | 7.14 | 7.29 | 7.45 | 7.6 | 7.76 | 7.91 | 8.06 | 8.22 | 8.37 | 8.53 | 8.68 | 8.84 | 8.99 | 9.15 | 9.3 | 9.46 | 9.61 | 9.76 | 9.92 | 10.07 | 10.23 | 10.38 | 10.54 | 10.69 | 10.85 | 11 | 11.15 | 11.31 | 11.46 | 11.62 | 11.77 | 11.93 | 12.08 | 12.24
8.8 | 6.05 | 6.21 | 6.36 | 6.52 | 6.67 | 6.83 | 6.98 | 7.14 | 7.29 | 7.44 | 7.6 | 7.75 | 7.91 | 8.06 | 8.22 | 8.37 | 8.53 | 8.68 | 8.84 | 8.99 | 9.14 | 9.3 | 9.45 | 9.61 | 9.76 | 9.92 | 10.07 | 10.23 | 10.38 | 10.53 | 10.69 | 10.84 | 11 | 11.15 | 11.31 | 11.46 | 11.62 | 11.77 | 11.92 | 12.08 | 12.23
8.81 | 6.05 | 6.21 | 6.36 | 6.52 | 6.67 | 6.82 | 6.98 | 7.13 | 7.29 | 7.44 | 7.6 | 7.75 | 7.91 | 8.06 | 8.22 | 8.37 | 8.52 | 8.68 | 8.83 | 8.99 | 9.14 | 9.3 | 9.45 | 9.61 | 9.76 | 9.91 | 10.07 | 10.22 | 10.38 | 10.53 | 10.69 | 10.84 | 11 | 11.15 | 11.3 | 11.46 | 11.61 | 11.77 | 11.92 | 12.08 | 12.23
8.82 | 6.05 | 6.21 | 6.36 | 6.51 | 6.67 | 6.82 | 6.98 | 7.13 | 7.29 | 7.44 | 7.6 | 7.75 | 7.9 | 8.06 | 8.21 | 8.37 | 8.52 | 8.68 | 8.83 | 8.99 | 9.14 | 9.29 | 9.45 | 9.6 | 9.76 | 9.91 | 10.07 | 10.22 | 10.37 | 10.53 | 10.68 | 10.84 | 10.99 | 11.15 | 11.3 | 11.46 | 11.61 | 11.76 | 11.92 | 12.07 | 12.23
8.83 | 6.05 | 6.2 | 6.36 | 6.51 | 6.67 | 6.82 | 6.98 | 7.13 | 7.28 | 7.44 | 7.59 | 7.75 | 7.9 | 8.06 | 8.21 | 8.37 | 8.52 | 8.67 | 8.83 | 8.98 | 9.14 | 9.29 | 9.45 | 9.6 | 9.75 | 9.91 | 10.06 | 10.22 | 10.37 | 10.53 | 10.68 | 10.84 | 10.99 | 11.14 | 11.3 | 11.45 | 11.61 | 11.76 | 11.92 | 12.07 | 12.22
8.84 | 6.05 | 6.2 | 6.36 | 6.51 | 6.67 | 6.82 | 6.97 | 7.13 | 7.28 | 7.44 | 7.59 | 7.75 | 7.9 | 8.05 | 8.21 | 8.36 | 8.52 | 8.67 | 8.83 | 8.98 | 9.14 | 9.29 | 9.44 | 9.6 | 9.75 | 9.91 | 10.06 | 10.22 | 10.37 | 10.52 | 10.68 | 10.83 | 10.99 | 11.14 | 11.3 | 11.45 | 11.6 | 11.76 | 11.91 | 12.07 | 12.22
8.85 | 6.05 | 6.2 | 6.35 | 6.51 | 6.66 | 6.82 | 6.97 | 7.13 | 7.28 | 7.44 | 7.59 | 7.74 | 7.9 | 8.05 | 8.21 | 8.36 | 8.52 | 8.67 | 8.82 | 8.98 | 9.13 | 9.29 | 9.44 | 9.6 | 9.75 | 9.9 | 10.06 | 10.21 | 10.37 | 10.52 | 10.68 | 10.83 | 10.98 | 11.14 | 11.29 | 11.45 | 11.6 | 11.76 | 11.91 | 12.06 | 12.22
8.86 | 6.04 | 6.2 | 6.35 | 6.51 | 6.66 | 6.82 | 6.97 | 7.13 | 7.28 | 7.43 | 7.59 | 7.74 | 7.9 | 8.05 | 8.21 | 8.36 | 8.51 | 8.67 | 8.82 | 8.98 | 9.13 | 9.29 | 9.44 | 9.59 | 9.75 | 9.9 | 10.06 | 10.21 | 10.36 | 10.52 | 10.67 | 10.83 | 10.98 | 11.14 | 11.29 | 11.44 | 11.6 | 11.75 | 11.91 | 12.06 | 12.22
8.87 | 6.04 | 6.2 | 6.35 | 6.51 | 6.66 | 6.81 | 6.97 | 7.12 | 7.28 | 7.43 | 7.59 | 7.74 | 7.89 | 8.05 | 8.2 | 8.36 | 8.51 | 8.67 | 8.82 | 8.97 | 9.13 | 9.28 | 9.44 | 9.59 | 9.75 | 9.9 | 10.05 | 10.21 | 10.36 | 10.52 | 10.67 | 10.83 | 10.98 | 11.13 | 11.29 | 11.44 | 11.6 | 11.75 | 11.9 | 12.06 | 12.21
8.88 | 6.04 | 6.2 | 6.35 | 6.5 | 6.66 | 6.81 | 6.97 | 7.12 | 7.28 | 7.43 | 7.58 | 7.74 | 7.89 | 8.05 | 8.2 | 8.36 | 8.51 | 8.66 | 8.82 | 8.97 | 9.13 | 9.28 | 9.43 | 9.59 | 9.74 | 9.9 | 10.05 | 10.21 | 10.36 | 10.51 | 10.67 | 10.82 | 10.98 | 11.13 | 11.28 | 11.44 | 11.59 | 11.75 | 11.9 | 12.06 | 12.21
8.89 | 6.04 | 6.19 | 6.35 | 6.5 | 6.66 | 6.81 | 6.97 | 7.12 | 7.27 | 7.43 | 7.58 | 7.74 | 7.89 | 8.04 | 8.2 | 8.35 | 8.51 | 8.66 | 8.82 | 8.97 | 9.12 | 9.28 | 9.43 | 9.59 | 9.74 | 9.89 | 10.05 | 10.2 | 10.36 | 10.51 | 10.67 | 10.82 | 10.97 | 11.13 | 11.28 | 11.44 | 11.59 | 11.74 | 11.9 | 12.05 | 12.21
8.9 | 6.04 | 6.19 | 6.35 | 6.5 | 6.66 | 6.81 | 6.96 | 7.12 | 7.27 | 7.43 | 7.58 | 7.73 | 7.89 | 8.04 | 8.2 | 8.35 | 8.51 | 8.66 | 8.81 | 8.97 | 9.12 | 9.28 | 9.43 | 9.58 | 9.74 | 9.89 | 10.05 | 10.2 | 10.35 | 10.51 | 10.66 | 10.82 | 10.97 | 11.13 | 11.28 | 11.43 | 11.59 | 11.74 | 11.9 | 12.05 | 12.2
8.91 | 6.04 | 6.19 | 6.35 | 6.5 | 6.65 | 6.81 | 6.96 | 7.12 | 7.27 | 7.42 | 7.58 | 7.73 | 7.89 | 8.04 | 8.2 | 8.35 | 8.5 | 8.66 | 8.81 | 8.97 | 9.12 | 9.27 | 9.43 | 9.58 | 9.74 | 9.89 | 10.04 | 10.2 | 10.35 | 10.51 | 10.66 | 10.81 | 10.97 | 11.12 | 11.28 | 11.43 | 11.58 | 11.74 | 11.89 | 12.05 | 12.2
8.92 | 6.04 | 6.19 | 6.34 | 6.5 | 6.65 | 6.81 | 6.96 | 7.11 | 7.27 | 7.42 | 7.58 | 7.73 | 7.88 | 8.04 | 8.19 | 8.35 | 8.5 | 8.66 | 8.81 | 8.96 | 9.12 | 9.27 | 9.43 | 9.58 | 9.73 | 9.89 | 10.04 | 10.2 | 10.35 | 10.5 | 10.66 | 10.81 | 10.97 | 11.12 | 11.27 | 11.43 | 11.58 | 11.74 | 11.89 | 12.04 | 12.2
8.93 | 6.03 | 6.19 | 6.34 | 6.5 | 6.65 | 6.8 | 6.96 | 7.11 | 7.27 | 7.42 | 7.57 | 7.73 | 7.88 | 8.04 | 8.19 | 8.35 | 8.5 | 8.65 | 8.81 | 8.96 | 9.12 | 9.27 | 9.42 | 9.58 | 9.73 | 9.89 | 10.04 | 10.19 | 10.35 | 10.5 | 10.66 | 10.81 | 10.96 | 11.12 | 11.27 | 11.43 | 11.58 | 11.73 | 11.89 | 12.04 | 12.2
8.94 | 6.03 | 6.19 | 6.34 | 6.49 | 6.65 | 6.8 | 6.96 | 7.11 | 7.27 | 7.42 | 7.57 | 7.73 | 7.88 | 8.04 | 8.19 | 8.34 | 8.5 | 8.65 | 8.81 | 8.96 | 9.11 | 9.27 | 9.42 | 9.58 | 9.73 | 9.88 | 10.04 | 10.19 | 10.34 | 10.5 | 10.65 | 10.81 | 10.96 | 11.11 | 11.27 | 11.42 | 11.58 | 11.73 | 11.88 | 12.04 | 12.19
8.95 | 6.03 | 6.19 | 6.34 | 6.49 | 6.65 | 6.8 | 6.96 | 7.11 | 7.26 | 7.42 | 7.57 | 7.73 | 7.88 | 8.03 | 8.19 | 8.34 | 8.5 | 8.65 | 8.8 | 8.96 | 9.11 | 9.26 | 9.42 | 9.57 | 9.73 | 9.88 | 10.03 | 10.19 | 10.34 | 10.5 | 10.65 | 10.8 | 10.96 | 11.11 | 11.27 | 11.42 | 11.57 | 11.73 | 11.88 | 12.04 | 12.19
8.96 | 6.03 | 6.18 | 6.34 | 6.49 | 6.65 | 6.8 | 6.95 | 7.11 | 7.26 | 7.42 | 7.57 | 7.72 | 7.88 | 8.03 | 8.19 | 8.34 | 8.49 | 8.65 | 8.8 | 8.95 | 9.11 | 9.26 | 9.42 | 9.57 | 9.72 | 9.88 | 10.03 | 10.19 | 10.34 | 10.49 | 10.65 | 10.8 | 10.96 | 11.11 | 11.26 | 11.42 | 11.57 | 11.72 | 11.88 | 12.03 | 12.19
8.97 | 6.03 | 6.18 | 6.34 | 6.49 | 6.64 | 6.8 | 6.95 | 7.11 | 7.26 | 7.41 | 7.57 | 7.72 | 7.88 | 8.03 | 8.18 | 8.34 | 8.49 | 8.64 | 8.8 | 8.95 | 9.11 | 9.26 | 9.41 | 9.57 | 9.72 | 9.88 | 10.03 | 10.18 | 10.34 | 10.49 | 10.65 | 10.8 | 10.95 | 11.11 | 11.26 | 11.41 | 11.57 | 11.72 | 11.88 | 12.03 | 12.18
8.98 | 6.03 | 6.18 | 6.33 | 6.49 | 6.64 | 6.8 | 6.95 | 7.1 | 7.26 | 7.41 | 7.57 | 7.72 | 7.87 | 8.03 | 8.18 | 8.34 | 8.49 | 8.64 | 8.8 | 8.95 | 9.1 | 9.26 | 9.41 | 9.57 | 9.72 | 9.87 | 10.03 | 10.18 | 10.33 | 10.49 | 10.64 | 10.8 | 10.95 | 11.1 | 11.26 | 11.41 | 11.57 | 11.72 | 11.87 | 12.03 | 12.18
8.99 | 6.03 | 6.18 | 6.33 | 6.49 | 6.64 | 6.79 | 6.95 | 7.1 | 7.26 | 7.41 | 7.56 | 7.72 | 7.87 | 8.03 | 8.18 | 8.33 | 8.49 | 8.64 | 8.79 | 8.95 | 9.1 | 9.26 | 9.41 | 9.56 | 9.72 | 9.87 | 10.02 | 10.18 | 10.33 | 10.49 | 10.64 | 10.79 | 10.95 | 11.1 | 11.26 | 11.41 | 11.56 | 11.72 | 11.87 | 12.02 | 12.18
9 | 6.02 | 6.18 | 6.33 | 6.49 | 6.64 | 6.79 | 6.95 | 7.1 | 7.25 | 7.41 | 7.56 | 7.72 | 7.87 | 8.02 | 8.18 | 8.33 | 8.48 | 8.64 | 8.79 | 8.95 | 9.1 | 9.25 | 9.41 | 9.56 | 9.71 | 9.87 | 10.02 | 10.18 | 10.33 | 10.48 | 10.64 | 10.79 | 10.94 | 11.1 | 11.25 | 11.41 | 11.56 | 11.71 | 11.87 | 12.02 | 12.17
9.01 | 6.02 | 6.18 | 6.33 | 6.48 | 6.64 | 6.79 | 6.95 | 7.1 | 7.25 | 7.41 | 7.56 | 7.71 | 7.87 | 8.02 | 8.18 | 8.33 | 8.48 | 8.64 | 8.79 | 8.94 | 9.1 | 9.25 | 9.41 | 9.56 | 9.71 | 9.87 | 10.02 | 10.17 | 10.33 | 10.48 | 10.63 | 10.79 | 10.94 | 11.1 | 11.25 | 11.4 | 11.56 | 11.71 | 11.86 | 12.02 | 12.17
9.02 | 6.02 | 6.17 | 6.33 | 6.48 | 6.64 | 6.79 | 6.94 | 7.1 | 7.25 | 7.4 | 7.56 | 7.71 | 7.87 | 8.02 | 8.17 | 8.33 | 8.48 | 8.63 | 8.79 | 8.94 | 9.1 | 9.25 | 9.4 | 9.56 | 9.71 | 9.86 | 10.02 | 10.17 | 10.32 | 10.48 | 10.63 | 10.79 | 10.94 | 11.09 | 11.25 | 11.4 | 11.55 | 11.71 | 11.86 | 12.02 | 12.17
9.03 | 6.02 | 6.17 | 6.33 | 6.48 | 6.63 | 6.79 | 6.94 | 7.1 | 7.25 | 7.4 | 7.56 | 7.71 | 7.86 | 8.02 | 8.17 | 8.32 | 8.48 | 8.63 | 8.79 | 8.94 | 9.09 | 9.25 | 9.4 | 9.55 | 9.71 | 9.86 | 10.02 | 10.17 | 10.32 | 10.48 | 10.63 | 10.78 | 10.94 | 11.09 | 11.24 | 11.4 | 11.55 | 11.71 | 11.86 | 12.01 | 12.17
9.04 | 6.02 | 6.17 | 6.33 | 6.48 | 6.63 | 6.79 | 6.94 | 7.09 | 7.25 | 7.4 | 7.55 | 7.71 | 7.86 | 8.02 | 8.17 | 8.32 | 8.48 | 8.63 | 8.78 | 8.94 | 9.09 | 9.24 | 9.4 | 9.55 | 9.71 | 9.86 | 10.01 | 10.17 | 10.32 | 10.47 | 10.63 | 10.78 | 10.93 | 11.09 | 11.24 | 11.4 | 11.55 | 11.7 | 11.86 | 12.01 | 12.16
9.05 | 6.02 | 6.17 | 6.32 | 6.48 | 6.63 | 6.78 | 6.94 | 7.09 | 7.25 | 7.4 | 7.55 | 7.71 | 7.86 | 8.01 | 8.17 | 8.32 | 8.47 | 8.63 | 8.78 | 8.94 | 9.09 | 9.24 | 9.4 | 9.55 | 9.7 | 9.86 | 10.01 | 10.16 | 10.32 | 10.47 | 10.62 | 10.78 | 10.93 | 11.09 | 11.24 | 11.39 | 11.55 | 11.7 | 11.85 | 12.01 | 12.16
9.06 | 6.01 | 6.17 | 6.32 | 6.48 | 6.63 | 6.78 | 6.94 | 7.09 | 7.24 | 7.4 | 7.55 | 7.7 | 7.86 | 8.01 | 8.17 | 8.32 | 8.47 | 8.63 | 8.78 | 8.93 | 9.09 | 9.24 | 9.39 | 9.55 | 9.7 | 9.85 | 10.01 | 10.16 | 10.32 | 10.47 | 10.62 | 10.78 | 10.93 | 11.08 | 11.24 | 11.39 | 11.54 | 11.7 | 11.85 | 12 | 12.16
9.07 | 6.01 | 6.17 | 6.32 | 6.47 | 6.63 | 6.78 | 6.93 | 7.09 | 7.24 | 7.4 | 7.55 | 7.7 | 7.86 | 8.01 | 8.16 | 8.32 | 8.47 | 8.62 | 8.78 | 8.93 | 9.08 | 9.24 | 9.39 | 9.55 | 9.7 | 9.85 | 10.01 | 10.16 | 10.31 | 10.47 | 10.62 | 10.77 | 10.93 | 11.08 | 11.23 | 11.39 | 11.54 | 11.69 | 11.85 | 12 | 12.15
9.08 | 6.01 | 6.17 | 6.32 | 6.47 | 6.63 | 6.78 | 6.93 | 7.09 | 7.24 | 7.39 | 7.55 | 7.7 | 7.85 | 8.01 | 8.16 | 8.31 | 8.47 | 8.62 | 8.78 | 8.93 | 9.08 | 9.24 | 9.39 | 9.54 | 9.7 | 9.85 | 10 | 10.16 | 10.31 | 10.46 | 10.62 | 10.77 | 10.92 | 11.08 | 11.23 | 11.38 | 11.54 | 11.69 | 11.84 | 12 | 12.15
9.09 | 6.01 | 6.16 | 6.32 | 6.47 | 6.62 | 6.78 | 6.93 | 7.08 | 7.24 | 7.39 | 7.55 | 7.7 | 7.85 | 8.01 | 8.16 | 8.31 | 8.47 | 8.62 | 8.77 | 8.93 | 9.08 | 9.23 | 9.39 | 9.54 | 9.69 | 9.85 | 10 | 10.15 | 10.31 | 10.46 | 10.61 | 10.77 | 10.92 | 11.07 | 11.23 | 11.38 | 11.53 | 11.69 | 11.84 | 12 | 12.15
9.1 | 6.01 | 6.16 | 6.32 | 6.47 | 6.62 | 6.78 | 6.93 | 7.08 | 7.24 | 7.39 | 7.54 | 7.7 | 7.85 | 8 | 8.16 | 8.31 | 8.46 | 8.62 | 8.77 | 8.92 | 9.08 | 9.23 | 9.38 | 9.54 | 9.69 | 9.84 | 10 | 10.15 | 10.31 | 10.46 | 10.61 | 10.77 | 10.92 | 11.07 | 11.23 | 11.38 | 11.53 | 11.69 | 11.84 | 11.99 | 12.15
9.11 | 6.01 | 6.16 | 6.31 | 6.47 | 6.62 | 6.77 | 6.93 | 7.08 | 7.23 | 7.39 | 7.54 | 7.7 | 7.85 | 8 | 8.16 | 8.31 | 8.46 | 8.62 | 8.77 | 8.92 | 9.08 | 9.23 | 9.38 | 9.54 | 9.69 | 9.84 | 10 | 10.15 | 10.3 | 10.46 | 10.61 | 10.76 | 10.92 | 11.07 | 11.22 | 11.38 | 11.53 | 11.68 | 11.84 | 11.99 | 12.14
9.12 | 6.01 | 6.16 | 6.31 | 6.47 | 6.62 | 6.77 | 6.93 | 7.08 | 7.23 | 7.39 | 7.54 | 7.69 | 7.85 | 8 | 8.15 | 8.31 | 8.46 | 8.61 | 8.77 | 8.92 | 9.07 | 9.23 | 9.38 | 9.53 | 9.69 | 9.84 | 9.99 | 10.15 | 10.3 | 10.45 | 10.61 | 10.76 | 10.91 | 11.07 | 11.22 | 11.37 | 11.53 | 11.68 | 11.83 | 11.99 | 12.14
9.13 | 6 | 6.16 | 6.31 | 6.46 | 6.62 | 6.77 | 6.92 | 7.08 | 7.23 | 7.38 | 7.54 | 7.69 | 7.84 | 8 | 8.15 | 8.3 | 8.46 | 8.61 | 8.76 | 8.92 | 9.07 | 9.22 | 9.38 | 9.53 | 9.68 | 9.84 | 9.99 | 10.14 | 10.3 | 10.45 | 10.6 | 10.76 | 10.91 | 11.06 | 11.22 | 11.37 | 11.52 | 11.68 | 11.83 | 11.98 | 12.14
9.14 | 6 | 6.16 | 6.31 | 6.46 | 6.62 | 6.77 | 6.92 | 7.08 | 7.23 | 7.38 | 7.54 | 7.69 | 7.84 | 8 | 8.15 | 8.3 | 8.46 | 8.61 | 8.76 | 8.92 | 9.07 | 9.22 | 9.38 | 9.53 | 9.68 | 9.84 | 9.99 | 10.14 | 10.3 | 10.45 | 10.6 | 10.75 | 10.91 | 11.06 | 11.21 | 11.37 | 11.52 | 11.67 | 11.83 | 11.98 | 12.13
9.15 | 6 | 6.15 | 6.31 | 6.46 | 6.61 | 6.77 | 6.92 | 7.07 | 7.23 | 7.38 | 7.53 | 7.69 | 7.84 | 7.99 | 8.15 | 8.3 | 8.45 | 8.61 | 8.76 | 8.91 | 9.07 | 9.22 | 9.37 | 9.53 | 9.68 | 9.83 | 9.99 | 10.14 | 10.29 | 10.45 | 10.6 | 10.75 | 10.91 | 11.06 | 11.21 | 11.37 | 11.52 | 11.67 | 11.82 | 11.98 | 12.13
9.16 | 6 | 6.15 | 6.31 | 6.46 | 6.61 | 6.77 | 6.92 | 7.07 | 7.23 | 7.38 | 7.53 | 7.69 | 7.84 | 7.99 | 8.15 | 8.3 | 8.45 | 8.61 | 8.76 | 8.91 | 9.06 | 9.22 | 9.37 | 9.52 | 9.68 | 9.83 | 9.98 | 10.14 | 10.29 | 10.44 | 10.6 | 10.75 | 10.9 | 11.06 | 11.21 | 11.36 | 11.52 | 11.67 | 11.82 | 11.97 | 12.13
9.17 | 6 | 6.15 | 6.31 | 6.46 | 6.61 | 6.76 | 6.92 | 7.07 | 7.22 | 7.38 | 7.53 | 7.68 | 7.84 | 7.99 | 8.14 | 8.3 | 8.45 | 8.6 | 8.76 | 8.91 | 9.06 | 9.22 | 9.37 | 9.52 | 9.68 | 9.83 | 9.98 | 10.13 | 10.29 | 10.44 | 10.59 | 10.75 | 10.9 | 11.05 | 11.21 | 11.36 | 11.51 | 11.67 | 11.82 | 11.97 | 12.13
9.18 | 6 | 6.15 | 6.3 | 6.46 | 6.61 | 6.76 | 6.92 | 7.07 | 7.22 | 7.38 | 7.53 | 7.68 | 7.84 | 7.99 | 8.14 | 8.29 | 8.45 | 8.6 | 8.75 | 8.91 | 9.06 | 9.21 | 9.37 | 9.52 | 9.67 | 9.83 | 9.98 | 10.13 | 10.29 | 10.44 | 10.59 | 10.74 | 10.9 | 11.05 | 11.2 | 11.36 | 11.51 | 11.66 | 11.82 | 11.97 | 12.12
9.19 | 6 | 6.15 | 6.3 | 6.46 | 6.61 | 6.76 | 6.91 | 7.07 | 7.22 | 7.37 | 7.53 | 7.68 | 7.83 | 7.99 | 8.14 | 8.29 | 8.45 | 8.6 | 8.75 | 8.91 | 9.06 | 9.21 | 9.36 | 9.52 | 9.67 | 9.82 | 9.98 | 10.13 | 10.28 | 10.44 | 10.59 | 10.74 | 10.9 | 11.05 | 11.2 | 11.35 | 11.51 | 11.66 | 11.81 | 11.97 | 12.12
9.2 | 5.99 | 6.15 | 6.3 | 6.45 | 6.61 | 6.76 | 6.91 | 7.07 | 7.22 | 7.37 | 7.53 | 7.68 | 7.83 | 7.98 | 8.14 | 8.29 | 8.44 | 8.6 | 8.75 | 8.9 | 9.06 | 9.21 | 9.36 | 9.52 | 9.67 | 9.82 | 9.97 | 10.13 | 10.28 | 10.43 | 10.59 | 10.74 | 10.89 | 11.05 | 11.2 | 11.35 | 11.5 | 11.66 | 11.81 | 11.96 | 12.12
9.21 | 5.99 | 6.15 | 6.3 | 6.45 | 6.6 | 6.76 | 6.91 | 7.06 | 7.22 | 7.37 | 7.52 | 7.68 | 7.83 | 7.98 | 8.14 | 8.29 | 8.44 | 8.59 | 8.75 | 8.9 | 9.05 | 9.21 | 9.36 | 9.51 | 9.67 | 9.82 | 9.97 | 10.12 | 10.28 | 10.43 | 10.58 | 10.74 | 10.89 | 11.04 | 11.2 | 11.35 | 11.5 | 11.65 | 11.81 | 11.96 | 12.11
9.22 | 5.99 | 6.14 | 6.3 | 6.45 | 6.6 | 6.76 | 6.91 | 7.06 | 7.22 | 7.37 | 7.52 | 7.67 | 7.83 | 7.98 | 8.13 | 8.29 | 8.44 | 8.59 | 8.75 | 8.9 | 9.05 | 9.2 | 9.36 | 9.51 | 9.66 | 9.82 | 9.97 | 10.12 | 10.28 | 10.43 | 10.58 | 10.73 | 10.89 | 11.04 | 11.19 | 11.35 | 11.5 | 11.65 | 11.8 | 11.96 | 12.11
9.23 | 5.99 | 6.14 | 6.3 | 6.45 | 6.6 | 6.75 | 6.91 | 7.06 | 7.21 | 7.37 | 7.52 | 7.67 | 7.83 | 7.98 | 8.13 | 8.28 | 8.44 | 8.59 | 8.74 | 8.9 | 9.05 | 9.2 | 9.36 | 9.51 | 9.66 | 9.81 | 9.97 | 10.12 | 10.27 | 10.43 | 10.58 | 10.73 | 10.88 | 11.04 | 11.19 | 11.34 | 11.5 | 11.65 | 11.8 | 11.95 | 12.11
9.24 | 5.99 | 6.14 | 6.29 | 6.45 | 6.6 | 6.75 | 6.91 | 7.06 | 7.21 | 7.36 | 7.52 | 7.67 | 7.82 | 7.98 | 8.13 | 8.28 | 8.44 | 8.59 | 8.74 | 8.89 | 9.05 | 9.2 | 9.35 | 9.51 | 9.66 | 9.81 | 9.96 | 10.12 | 10.27 | 10.42 | 10.58 | 10.73 | 10.88 | 11.03 | 11.19 | 11.34 | 11.49 | 11.65 | 11.8 | 11.95 | 12.1
9.25 | 5.99 | 6.14 | 6.29 | 6.45 | 6.6 | 6.75 | 6.9 | 7.06 | 7.21 | 7.36 | 7.52 | 7.67 | 7.82 | 7.97 | 8.13 | 8.28 | 8.43 | 8.59 | 8.74 | 8.89 | 9.05 | 9.2 | 9.35 | 9.5 | 9.66 | 9.81 | 9.96 | 10.12 | 10.27 | 10.42 | 10.57 | 10.73 | 10.88 | 11.03 | 11.19 | 11.34 | 11.49 | 11.64 | 11.8 | 11.95 | 12.1
9.26 | 5.99 | 6.14 | 6.29 | 6.44 | 6.6 | 6.75 | 6.9 | 7.06 | 7.21 | 7.36 | 7.51 | 7.67 | 7.82 | 7.97 | 8.13 | 8.28 | 8.43 | 8.58 | 8.74 | 8.89 | 9.04 | 9.2 | 9.35 | 9.5 | 9.65 | 9.81 | 9.96 | 10.11 | 10.27 | 10.42 | 10.57 | 10.72 | 10.88 | 11.03 | 11.18 | 11.34 | 11.49 | 11.64 | 11.79 | 11.95 | 12.1
9.27 | 5.98 | 6.14 | 6.29 | 6.44 | 6.6 | 6.75 | 6.9 | 7.05 | 7.21 | 7.36 | 7.51 | 7.67 | 7.82 | 7.97 | 8.12 | 8.28 | 8.43 | 8.58 | 8.74 | 8.89 | 9.04 | 9.19 | 9.35 | 9.5 | 9.65 | 9.8 | 9.96 | 10.11 | 10.26 | 10.42 | 10.57 | 10.72 | 10.87 | 11.03 | 11.18 | 11.33 | 11.49 | 11.64 | 11.79 | 11.94 | 12.1
9.28 | 5.98 | 6.14 | 6.29 | 6.44 | 6.59 | 6.75 | 6.9 | 7.05 | 7.2 | 7.36 | 7.51 | 7.66 | 7.82 | 7.97 | 8.12 | 8.27 | 8.43 | 8.58 | 8.73 | 8.89 | 9.04 | 9.19 | 9.34 | 9.5 | 9.65 | 9.8 | 9.96 | 10.11 | 10.26 | 10.41 | 10.57 | 10.72 | 10.87 | 11.02 | 11.18 | 11.33 | 11.48 | 11.64 | 11.79 | 11.94 | 12.09
9.29 | 5.98 | 6.13 | 6.29 | 6.44 | 6.59 | 6.74 | 6.9 | 7.05 | 7.2 | 7.36 | 7.51 | 7.66 | 7.81 | 7.97 | 8.12 | 8.27 | 8.43 | 8.58 | 8.73 | 8.88 | 9.04 | 9.19 | 9.34 | 9.49 | 9.65 | 9.8 | 9.95 | 10.11 | 10.26 | 10.41 | 10.56 | 10.72 | 10.87 | 11.02 | 11.17 | 11.33 | 11.48 | 11.63 | 11.79 | 11.94 | 12.09
9.3 | 5.98 | 6.13 | 6.28 | 6.44 | 6.59 | 6.74 | 6.9 | 7.05 | 7.2 | 7.35 | 7.51 | 7.66 | 7.81 | 7.97 | 8.12 | 8.27 | 8.42 | 8.58 | 8.73 | 8.88 | 9.03 | 9.19 | 9.34 | 9.49 | 9.64 | 9.8 | 9.95 | 10.1 | 10.26 | 10.41 | 10.56 | 10.71 | 10.87 | 11.02 | 11.17 | 11.32 | 11.48 | 11.63 | 11.78 | 11.93 | 12.09
9.31 | 5.98 | 6.13 | 6.28 | 6.44 | 6.59 | 6.74 | 6.89 | 7.05 | 7.2 | 7.35 | 7.51 | 7.66 | 7.81 | 7.96 | 8.12 | 8.27 | 8.42 | 8.57 | 8.73 | 8.88 | 9.03 | 9.18 | 9.34 | 9.49 | 9.64 | 9.8 | 9.95 | 10.1 | 10.25 | 10.41 | 10.56 | 10.71 | 10.86 | 11.02 | 11.17 | 11.32 | 11.47 | 11.63 | 11.78 | 11.93 | 12.08
9.32 | 5.98 | 6.13 | 6.28 | 6.43 | 6.59 | 6.74 | 6.89 | 7.05 | 7.2 | 7.35 | 7.5 | 7.66 | 7.81 | 7.96 | 8.11 | 8.27 | 8.42 | 8.57 | 8.72 | 8.88 | 9.03 | 9.18 | 9.34 | 9.49 | 9.64 | 9.79 | 9.95 | 10.1 | 10.25 | 10.4 | 10.56 | 10.71 | 10.86 | 11.01 | 11.17 | 11.32 | 11.47 | 11.62 | 11.78 | 11.93 | 12.08
9.33 | 5.97 | 6.13 | 6.28 | 6.43 | 6.59 | 6.74 | 6.89 | 7.04 | 7.2 | 7.35 | 7.5 | 7.65 | 7.81 | 7.96 | 8.11 | 8.26 | 8.42 | 8.57 | 8.72 | 8.88 | 9.03 | 9.18 | 9.33 | 9.49 | 9.64 | 9.79 | 9.94 | 10.1 | 10.25 | 10.4 | 10.55 | 10.71 | 10.86 | 11.01 | 11.16 | 11.32 | 11.47 | 11.62 | 11.77 | 11.93 | 12.08
9.34 | 5.97 | 6.13 | 6.28 | 6.43 | 6.58 | 6.74 | 6.89 | 7.04 | 7.19 | 7.35 | 7.5 | 7.65 | 7.8 | 7.96 | 8.11 | 8.26 | 8.42 | 8.57 | 8.72 | 8.87 | 9.03 | 9.18 | 9.33 | 9.48 | 9.64 | 9.79 | 9.94 | 10.09 | 10.25 | 10.4 | 10.55 | 10.7 | 10.86 | 11.01 | 11.16 | 11.31 | 11.47 | 11.62 | 11.77 | 11.92 | 12.08
9.35 | 5.97 | 6.12 | 6.28 | 6.43 | 6.58 | 6.73 | 6.89 | 7.04 | 7.19 | 7.35 | 7.5 | 7.65 | 7.8 | 7.96 | 8.11 | 8.26 | 8.41 | 8.57 | 8.72 | 8.87 | 9.02 | 9.18 | 9.33 | 9.48 | 9.63 | 9.79 | 9.94 | 10.09 | 10.24 | 10.4 | 10.55 | 10.7 | 10.85 | 11.01 | 11.16 | 11.31 | 11.46 | 11.62 | 11.77 | 11.92 | 12.07
9.36 | 5.97 | 6.12 | 6.28 | 6.43 | 6.58 | 6.73 | 6.89 | 7.04 | 7.19 | 7.34 | 7.5 | 7.65 | 7.8 | 7.95 | 8.11 | 8.26 | 8.41 | 8.56 | 8.72 | 8.87 | 9.02 | 9.17 | 9.33 | 9.48 | 9.63 | 9.78 | 9.94 | 10.09 | 10.24 | 10.39 | 10.55 | 10.7 | 10.85 | 11 | 11.16 | 11.31 | 11.46 | 11.61 | 11.77 | 11.92 | 12.07
9.37 | 5.97 | 6.12 | 6.27 | 6.43 | 6.58 | 6.73 | 6.88 | 7.04 | 7.19 | 7.34 | 7.49 | 7.65 | 7.8 | 7.95 | 8.1 | 8.26 | 8.41 | 8.56 | 8.71 | 8.87 | 9.02 | 9.17 | 9.32 | 9.48 | 9.63 | 9.78 | 9.93 | 10.09 | 10.24 | 10.39 | 10.54 | 10.7 | 10.85 | 11 | 11.15 | 11.31 | 11.46 | 11.61 | 11.76 | 11.92 | 12.07
9.38 | 5.97 | 6.12 | 6.27 | 6.43 | 6.58 | 6.73 | 6.88 | 7.03 | 7.19 | 7.34 | 7.49 | 7.64 | 7.8 | 7.95 | 8.1 | 8.25 | 8.41 | 8.56 | 8.71 | 8.86 | 9.02 | 9.17 | 9.32 | 9.47 | 9.63 | 9.78 | 9.93 | 10.08 | 10.24 | 10.39 | 10.54 | 10.69 | 10.85 | 11 | 11.15 | 11.3 | 11.46 | 11.61 | 11.76 | 11.91 | 12.06
9.39 | 5.97 | 6.12 | 6.27 | 6.42 | 6.58 | 6.73 | 6.88 | 7.03 | 7.19 | 7.34 | 7.49 | 7.64 | 7.8 | 7.95 | 8.1 | 8.25 | 8.41 | 8.56 | 8.71 | 8.86 | 9.01 | 9.17 | 9.32 | 9.47 | 9.62 | 9.78 | 9.93 | 10.08 | 10.23 | 10.39 | 10.54 | 10.69 | 10.84 | 11 | 11.15 | 11.3 | 11.45 | 11.6 | 11.76 | 11.91 | 12.06
9.4 | 5.96 | 6.12 | 6.27 | 6.42 | 6.57 | 6.73 | 6.88 | 7.03 | 7.18 | 7.34 | 7.49 | 7.64 | 7.79 | 7.95 | 8.1 | 8.25 | 8.4 | 8.56 | 8.71 | 8.86 | 9.01 | 9.16 | 9.32 | 9.47 | 9.62 | 9.77 | 9.93 | 10.08 | 10.23 | 10.38 | 10.54 | 10.69 | 10.84 | 10.99 | 11.15 | 11.3 | 11.45 | 11.6 | 11.75 | 11.91 | 12.06
9.41 | 5.96 | 6.12 | 6.27 | 6.42 | 6.57 | 6.73 | 6.88 | 7.03 | 7.18 | 7.33 | 7.49 | 7.64 | 7.79 | 7.94 | 8.1 | 8.25 | 8.4 | 8.55 | 8.71 | 8.86 | 9.01 | 9.16 | 9.31 | 9.47 | 9.62 | 9.77 | 9.92 | 10.08 | 10.23 | 10.38 | 10.53 | 10.69 | 10.84 | 10.99 | 11.14 | 11.29 | 11.45 | 11.6 | 11.75 | 11.9 | 12.06
9.42 | 5.96 | 6.11 | 6.27 | 6.42 | 6.57 | 6.72 | 6.88 | 7.03 | 7.18 | 7.33 | 7.49 | 7.64 | 7.79 | 7.94 | 8.09 | 8.25 | 8.4 | 8.55 | 8.7 | 8.86 | 9.01 | 9.16 | 9.31 | 9.46 | 9.62 | 9.77 | 9.92 | 10.07 | 10.23 | 10.38 | 10.53 | 10.68 | 10.84 | 10.99 | 11.14 | 11.29 | 11.44 | 11.6 | 11.75 | 11.9 | 12.05
9.43 | 5.96 | 6.11 | 6.26 | 6.42 | 6.57 | 6.72 | 6.87 | 7.03 | 7.18 | 7.33 | 7.48 | 7.64 | 7.79 | 7.94 | 8.09 | 8.24 | 8.4 | 8.55 | 8.7 | 8.85 | 9.01 | 9.16 | 9.31 | 9.46 | 9.61 | 9.77 | 9.92 | 10.07 | 10.22 | 10.38 | 10.53 | 10.68 | 10.83 | 10.98 | 11.14 | 11.29 | 11.44 | 11.59 | 11.75 | 11.9 | 12.05
9.44 | 5.96 | 6.11 | 6.26 | 6.42 | 6.57 | 6.72 | 6.87 | 7.02 | 7.18 | 7.33 | 7.48 | 7.63 | 7.79 | 7.94 | 8.09 | 8.24 | 8.39 | 8.55 | 8.7 | 8.85 | 9 | 9.16 | 9.31 | 9.46 | 9.61 | 9.76 | 9.92 | 10.07 | 10.22 | 10.37 | 10.53 | 10.68 | 10.83 | 10.98 | 11.13 | 11.29 | 11.44 | 11.59 | 11.74 | 11.9 | 12.05
9.45 | 5.96 | 6.11 | 6.26 | 6.41 | 6.57 | 6.72 | 6.87 | 7.02 | 7.18 | 7.33 | 7.48 | 7.63 | 7.78 | 7.94 | 8.09 | 8.24 | 8.39 | 8.55 | 8.7 | 8.85 | 9 | 9.15 | 9.31 | 9.46 | 9.61 | 9.76 | 9.91 | 10.07 | 10.22 | 10.37 | 10.52 | 10.68 | 10.83 | 10.98 | 11.13 | 11.28 | 11.44 | 11.59 | 11.74 | 11.89 | 12.04
9.46 | 5.96 | 6.11 | 6.26 | 6.41 | 6.56 | 6.72 | 6.87 | 7.02 | 7.17 | 7.33 | 7.48 | 7.63 | 7.78 | 7.93 | 8.09 | 8.24 | 8.39 | 8.54 | 8.7 | 8.85 | 9 | 9.15 | 9.3 | 9.46 | 9.61 | 9.76 | 9.91 | 10.06 | 10.22 | 10.37 | 10.52 | 10.67 | 10.82 | 10.98 | 11.13 | 11.28 | 11.43 | 11.59 | 11.74 | 11.89 | 12.04
9.47 | 5.95 | 6.11 | 6.26 | 6.41 | 6.56 | 6.72 | 6.87 | 7.02 | 7.17 | 7.32 | 7.48 | 7.63 | 7.78 | 7.93 | 8.08 | 8.24 | 8.39 | 8.54 | 8.69 | 8.85 | 9 | 9.15 | 9.3 | 9.45 | 9.61 | 9.76 | 9.91 | 10.06 | 10.21 | 10.37 | 10.52 | 10.67 | 10.82 | 10.97 | 11.13 | 11.28 | 11.43 | 11.58 | 11.73 | 11.89 | 12.04
9.48 | 5.95 | 6.11 | 6.26 | 6.41 | 6.56 | 6.71 | 6.87 | 7.02 | 7.17 | 7.32 | 7.47 | 7.63 | 7.78 | 7.93 | 8.08 | 8.23 | 8.39 | 8.54 | 8.69 | 8.84 | 9 | 9.15 | 9.3 | 9.45 | 9.6 | 9.76 | 9.91 | 10.06 | 10.21 | 10.36 | 10.52 | 10.67 | 10.82 | 10.97 | 11.12 | 11.28 | 11.43 | 11.58 | 11.73 | 11.88 | 12.04
9.49 | 5.95 | 6.1 | 6.26 | 6.41 | 6.56 | 6.71 | 6.86 | 7.02 | 7.17 | 7.32 | 7.47 | 7.62 | 7.78 | 7.93 | 8.08 | 8.23 | 8.38 | 8.54 | 8.69 | 8.84 | 8.99 | 9.15 | 9.3 | 9.45 | 9.6 | 9.75 | 9.91 | 10.06 | 10.21 | 10.36 | 10.51 | 10.67 | 10.82 | 10.97 | 11.12 | 11.27 | 11.43 | 11.58 | 11.73 | 11.88 | 12.03
9.5 | 5.95 | 6.1 | 6.25 | 6.41 | 6.56 | 6.71 | 6.86 | 7.01 | 7.17 | 7.32 | 7.47 | 7.62 | 7.77 | 7.93 | 8.08 | 8.23 | 8.38 | 8.53 | 8.69 | 8.84 | 8.99 | 9.14 | 9.29 | 9.45 | 9.6 | 9.75 | 9.9 | 10.05 | 10.21 | 10.36 | 10.51 | 10.66 | 10.81 | 10.97 | 11.12 | 11.27 | 11.42 | 11.57 | 11.73 | 11.88 | 12.03
9.51 | 5.95 | 6.1 | 6.25 | 6.4 | 6.56 | 6.71 | 6.86 | 7.01 | 7.16 | 7.32 | 7.47 | 7.62 | 7.77 | 7.92 | 8.08 | 8.23 | 8.38 | 8.53 | 8.68 | 8.84 | 8.99 | 9.14 | 9.29 | 9.44 | 9.6 | 9.75 | 9.9 | 10.05 | 10.2 | 10.36 | 10.51 | 10.66 | 10.81 | 10.96 | 11.12 | 11.27 | 11.42 | 11.57 | 11.72 | 11.88 | 12.03
9.52 | 5.95 | 6.1 | 6.25 | 6.4 | 6.56 | 6.71 | 6.86 | 7.01 | 7.16 | 7.32 | 7.47 | 7.62 | 7.77 | 7.92 | 8.07 | 8.23 | 8.38 | 8.53 | 8.68 | 8.83 | 8.99 | 9.14 | 9.29 | 9.44 | 9.59 | 9.75 | 9.9 | 10.05 | 10.2 | 10.35 | 10.51 | 10.66 | 10.81 | 10.96 | 11.11 | 11.27 | 11.42 | 11.57 | 11.72 | 11.87 | 12.02
9.53 | 5.95 | 6.1 | 6.25 | 6.4 | 6.55 | 6.71 | 6.86 | 7.01 | 7.16 | 7.31 | 7.47 | 7.62 | 7.77 | 7.92 | 8.07 | 8.22 | 8.38 | 8.53 | 8.68 | 8.83 | 8.98 | 9.14 | 9.29 | 9.44 | 9.59 | 9.74 | 9.9 | 10.05 | 10.2 | 10.35 | 10.5 | 10.66 | 10.81 | 10.96 | 11.11 | 11.26 | 11.41 | 11.57 | 11.72 | 11.87 | 12.02
9.54 | 5.94 | 6.1 | 6.25 | 6.4 | 6.55 | 6.7 | 6.86 | 7.01 | 7.16 | 7.31 | 7.46 | 7.62 | 7.77 | 7.92 | 8.07 | 8.22 | 8.37 | 8.53 | 8.68 | 8.83 | 8.98 | 9.13 | 9.29 | 9.44 | 9.59 | 9.74 | 9.89 | 10.05 | 10.2 | 10.35 | 10.5 | 10.65 | 10.8 | 10.96 | 11.11 | 11.26 | 11.41 | 11.56 | 11.72 | 11.87 | 12.02
9.55 | 5.94 | 6.09 | 6.25 | 6.4 | 6.55 | 6.7 | 6.85 | 7.01 | 7.16 | 7.31 | 7.46 | 7.61 | 7.77 | 7.92 | 8.07 | 8.22 | 8.37 | 8.52 | 8.68 | 8.83 | 8.98 | 9.13 | 9.28 | 9.44 | 9.59 | 9.74 | 9.89 | 10.04 | 10.19 | 10.35 | 10.5 | 10.65 | 10.8 | 10.95 | 11.11 | 11.26 | 11.41 | 11.56 | 11.71 | 11.86 | 12.02
9.56 | 5.94 | 6.09 | 6.25 | 6.4 | 6.55 | 6.7 | 6.85 | 7 | 7.16 | 7.31 | 7.46 | 7.61 | 7.76 | 7.92 | 8.07 | 8.22 | 8.37 | 8.52 | 8.67 | 8.83 | 8.98 | 9.13 | 9.28 | 9.43 | 9.59 | 9.74 | 9.89 | 10.04 | 10.19 | 10.34 | 10.5 | 10.65 | 10.8 | 10.95 | 11.1 | 11.25 | 11.41 | 11.56 | 11.71 | 11.86 | 12.01
9.57 | 5.94 | 6.09 | 6.24 | 6.4 | 6.55 | 6.7 | 6.85 | 7 | 7.15 | 7.31 | 7.46 | 7.61 | 7.76 | 7.91 | 8.07 | 8.22 | 8.37 | 8.52 | 8.67 | 8.82 | 8.98 | 9.13 | 9.28 | 9.43 | 9.58 | 9.73 | 9.89 | 10.04 | 10.19 | 10.34 | 10.49 | 10.64 | 10.8 | 10.95 | 11.1 | 11.25 | 11.4 | 11.56 | 11.71 | 11.86 | 12.01
9.58 | 5.94 | 6.09 | 6.24 | 6.39 | 6.55 | 6.7 | 6.85 | 7 | 7.15 | 7.3 | 7.46 | 7.61 | 7.76 | 7.91 | 8.06 | 8.21 | 8.37 | 8.52 | 8.67 | 8.82 | 8.97 | 9.13 | 9.28 | 9.43 | 9.58 | 9.73 | 9.88 | 10.04 | 10.19 | 10.34 | 10.49 | 10.64 | 10.79 | 10.95 | 11.1 | 11.25 | 11.4 | 11.55 | 11.7 | 11.86 | 12.01
9.59 | 5.94 | 6.09 | 6.24 | 6.39 | 6.54 | 6.7 | 6.85 | 7 | 7.15 | 7.3 | 7.45 | 7.61 | 7.76 | 7.91 | 8.06 | 8.21 | 8.36 | 8.52 | 8.67 | 8.82 | 8.97 | 9.12 | 9.27 | 9.43 | 9.58 | 9.73 | 9.88 | 10.03 | 10.18 | 10.34 | 10.49 | 10.64 | 10.79 | 10.94 | 11.09 | 11.25 | 11.4 | 11.55 | 11.7 | 11.85 | 12
9.6 | 5.94 | 6.09 | 6.24 | 6.39 | 6.54 | 6.69 | 6.85 | 7 | 7.15 | 7.3 | 7.45 | 7.6 | 7.76 | 7.91 | 8.06 | 8.21 | 8.36 | 8.51 | 8.67 | 8.82 | 8.97 | 9.12 | 9.27 | 9.42 | 9.58 | 9.73 | 9.88 | 10.03 | 10.18 | 10.33 | 10.49 | 10.64 | 10.79 | 10.94 | 11.09 | 11.24 | 11.4 | 11.55 | 11.7 | 11.85 | 12
9.61 | 5.93 | 6.09 | 6.24 | 6.39 | 6.54 | 6.69 | 6.84 | 7 | 7.15 | 7.3 | 7.45 | 7.6 | 7.75 | 7.91 | 8.06 | 8.21 | 8.36 | 8.51 | 8.66 | 8.82 | 8.97 | 9.12 | 9.27 | 9.42 | 9.57 | 9.73 | 9.88 | 10.03 | 10.18 | 10.33 | 10.48 | 10.63 | 10.79 | 10.94 | 11.09 | 11.24 | 11.39 | 11.54 | 11.7 | 11.85 | 12
9.62 | 5.93 | 6.08 | 6.24 | 6.39 | 6.54 | 6.69 | 6.84 | 6.99 | 7.15 | 7.3 | 7.45 | 7.6 | 7.75 | 7.9 | 8.06 | 8.21 | 8.36 | 8.51 | 8.66 | 8.81 | 8.96 | 9.12 | 9.27 | 9.42 | 9.57 | 9.72 | 9.87 | 10.03 | 10.18 | 10.33 | 10.48 | 10.63 | 10.78 | 10.94 | 11.09 | 11.24 | 11.39 | 11.54 | 11.69 | 11.84 | 12
9.63 | 5.93 | 6.08 | 6.23 | 6.39 | 6.54 | 6.69 | 6.84 | 6.99 | 7.14 | 7.3 | 7.45 | 7.6 | 7.75 | 7.9 | 8.05 | 8.21 | 8.36 | 8.51 | 8.66 | 8.81 | 8.96 | 9.11 | 9.27 | 9.42 | 9.57 | 9.72 | 9.87 | 10.02 | 10.18 | 10.33 | 10.48 | 10.63 | 10.78 | 10.93 | 11.08 | 11.24 | 11.39 | 11.54 | 11.69 | 11.84 | 11.99
9.64 | 5.93 | 6.08 | 6.23 | 6.38 | 6.54 | 6.69 | 6.84 | 6.99 | 7.14 | 7.29 | 7.45 | 7.6 | 7.75 | 7.9 | 8.05 | 8.2 | 8.35 | 8.51 | 8.66 | 8.81 | 8.96 | 9.11 | 9.26 | 9.42 | 9.57 | 9.72 | 9.87 | 10.02 | 10.17 | 10.32 | 10.48 | 10.63 | 10.78 | 10.93 | 11.08 | 11.23 | 11.38 | 11.54 | 11.69 | 11.84 | 11.99
9.65 | 5.93 | 6.08 | 6.23 | 6.38 | 6.53 | 6.69 | 6.84 | 6.99 | 7.14 | 7.29 | 7.44 | 7.6 | 7.75 | 7.9 | 8.05 | 8.2 | 8.35 | 8.5 | 8.66 | 8.81 | 8.96 | 9.11 | 9.26 | 9.41 | 9.56 | 9.72 | 9.87 | 10.02 | 10.17 | 10.32 | 10.47 | 10.62 | 10.78 | 10.93 | 11.08 | 11.23 | 11.38 | 11.53 | 11.68 | 11.84 | 11.99
9.66 | 5.93 | 6.08 | 6.23 | 6.38 | 6.53 | 6.68 | 6.84 | 6.99 | 7.14 | 7.29 | 7.44 | 7.59 | 7.74 | 7.9 | 8.05 | 8.2 | 8.35 | 8.5 | 8.65 | 8.8 | 8.96 | 9.11 | 9.26 | 9.41 | 9.56 | 9.71 | 9.86 | 10.02 | 10.17 | 10.32 | 10.47 | 10.62 | 10.77 | 10.92 | 11.08 | 11.23 | 11.38 | 11.53 | 11.68 | 11.83 | 11.98
9.67 | 5.93 | 6.08 | 6.23 | 6.38 | 6.53 | 6.68 | 6.83 | 6.99 | 7.14 | 7.29 | 7.44 | 7.59 | 7.74 | 7.89 | 8.05 | 8.2 | 8.35 | 8.5 | 8.65 | 8.8 | 8.95 | 9.11 | 9.26 | 9.41 | 9.56 | 9.71 | 9.86 | 10.01 | 10.17 | 10.32 | 10.47 | 10.62 | 10.77 | 10.92 | 11.07 | 11.23 | 11.38 | 11.53 | 11.68 | 11.83 | 11.98
9.68 | 5.92 | 6.08 | 6.23 | 6.38 | 6.53 | 6.68 | 6.83 | 6.98 | 7.14 | 7.29 | 7.44 | 7.59 | 7.74 | 7.89 | 8.04 | 8.2 | 8.35 | 8.5 | 8.65 | 8.8 | 8.95 | 9.1 | 9.25 | 9.41 | 9.56 | 9.71 | 9.86 | 10.01 | 10.16 | 10.31 | 10.47 | 10.62 | 10.77 | 10.92 | 11.07 | 11.22 | 11.37 | 11.53 | 11.68 | 11.83 | 11.98
9.69 | 5.92 | 6.07 | 6.23 | 6.38 | 6.53 | 6.68 | 6.83 | 6.98 | 7.13 | 7.28 | 7.44 | 7.59 | 7.74 | 7.89 | 8.04 | 8.19 | 8.34 | 8.5 | 8.65 | 8.8 | 8.95 | 9.1 | 9.25 | 9.4 | 9.56 | 9.71 | 9.86 | 10.01 | 10.16 | 10.31 | 10.46 | 10.61 | 10.77 | 10.92 | 11.07 | 11.22 | 11.37 | 11.52 | 11.67 | 11.82 | 11.98
9.7 | 5.92 | 6.07 | 6.22 | 6.38 | 6.53 | 6.68 | 6.83 | 6.98 | 7.13 | 7.28 | 7.43 | 7.59 | 7.74 | 7.89 | 8.04 | 8.19 | 8.34 | 8.49 | 8.65 | 8.8 | 8.95 | 9.1 | 9.25 | 9.4 | 9.55 | 9.7 | 9.86 | 10.01 | 10.16 | 10.31 | 10.46 | 10.61 | 10.76 | 10.91 | 11.07 | 11.22 | 11.37 | 11.52 | 11.67 | 11.82 | 11.97
9.71 | 5.92 | 6.07 | 6.22 | 6.37 | 6.52 | 6.68 | 6.83 | 6.98 | 7.13 | 7.28 | 7.43 | 7.58 | 7.74 | 7.89 | 8.04 | 8.19 | 8.34 | 8.49 | 8.64 | 8.79 | 8.95 | 9.1 | 9.25 | 9.4 | 9.55 | 9.7 | 9.85 | 10 | 10.16 | 10.31 | 10.46 | 10.61 | 10.76 | 10.91 | 11.06 | 11.21 | 11.37 | 11.52 | 11.67 | 11.82 | 11.97
9.72 | 5.92 | 6.07 | 6.22 | 6.37 | 6.52 | 6.67 | 6.83 | 6.98 | 7.13 | 7.28 | 7.43 | 7.58 | 7.73 | 7.88 | 8.04 | 8.19 | 8.34 | 8.49 | 8.64 | 8.79 | 8.94 | 9.09 | 9.25 | 9.4 | 9.55 | 9.7 | 9.85 | 10 | 10.15 | 10.3 | 10.46 | 10.61 | 10.76 | 10.91 | 11.06 | 11.21 | 11.36 | 11.51 | 11.67 | 11.82 | 11.97
9.73 | 5.92 | 6.07 | 6.22 | 6.37 | 6.52 | 6.67 | 6.82 | 6.98 | 7.13 | 7.28 | 7.43 | 7.58 | 7.73 | 7.88 | 8.03 | 8.19 | 8.34 | 8.49 | 8.64 | 8.79 | 8.94 | 9.09 | 9.24 | 9.39 | 9.55 | 9.7 | 9.85 | 10 | 10.15 | 10.3 | 10.45 | 10.6 | 10.76 | 10.91 | 11.06 | 11.21 | 11.36 | 11.51 | 11.66 | 11.81 | 11.96
9.74 | 5.92 | 6.07 | 6.22 | 6.37 | 6.52 | 6.67 | 6.82 | 6.97 | 7.12 | 7.28 | 7.43 | 7.58 | 7.73 | 7.88 | 8.03 | 8.18 | 8.33 | 8.49 | 8.64 | 8.79 | 8.94 | 9.09 | 9.24 | 9.39 | 9.54 | 9.7 | 9.85 | 10 | 10.15 | 10.3 | 10.45 | 10.6 | 10.75 | 10.9 | 11.06 | 11.21 | 11.36 | 11.51 | 11.66 | 11.81 | 11.96
9.75 | 5.91 | 6.06 | 6.22 | 6.37 | 6.52 | 6.67 | 6.82 | 6.97 | 7.12 | 7.27 | 7.43 | 7.58 | 7.73 | 7.88 | 8.03 | 8.18 | 8.33 | 8.48 | 8.63 | 8.79 | 8.94 | 9.09 | 9.24 | 9.39 | 9.54 | 9.69 | 9.84 | 9.99 | 10.15 | 10.3 | 10.45 | 10.6 | 10.75 | 10.9 | 11.05 | 11.2 | 11.35 | 11.51 | 11.66 | 11.81 | 11.96
9.76 | 5.91 | 6.06 | 6.21 | 6.37 | 6.52 | 6.67 | 6.82 | 6.97 | 7.12 | 7.27 | 7.42 | 7.57 | 7.73 | 7.88 | 8.03 | 8.18 | 8.33 | 8.48 | 8.63 | 8.78 | 8.93 | 9.09 | 9.24 | 9.39 | 9.54 | 9.69 | 9.84 | 9.99 | 10.14 | 10.29 | 10.45 | 10.6 | 10.75 | 10.9 | 11.05 | 11.2 | 11.35 | 11.5 | 11.65 | 11.81 | 11.96
9.77 | 5.91 | 6.06 | 6.21 | 6.36 | 6.52 | 6.67 | 6.82 | 6.97 | 7.12 | 7.27 | 7.42 | 7.57 | 7.72 | 7.88 | 8.03 | 8.18 | 8.33 | 8.48 | 8.63 | 8.78 | 8.93 | 9.08 | 9.23 | 9.39 | 9.54 | 9.69 | 9.84 | 9.99 | 10.14 | 10.29 | 10.44 | 10.59 | 10.75 | 10.9 | 11.05 | 11.2 | 11.35 | 11.5 | 11.65 | 11.8 | 11.95
9.78 | 5.91 | 6.06 | 6.21 | 6.36 | 6.51 | 6.66 | 6.82 | 6.97 | 7.12 | 7.27 | 7.42 | 7.57 | 7.72 | 7.87 | 8.02 | 8.18 | 8.33 | 8.48 | 8.63 | 8.78 | 8.93 | 9.08 | 9.23 | 9.38 | 9.53 | 9.69 | 9.84 | 9.99 | 10.14 | 10.29 | 10.44 | 10.59 | 10.74 | 10.89 | 11.04 | 11.2 | 11.35 | 11.5 | 11.65 | 11.8 | 11.95
9.79 | 5.91 | 6.06 | 6.21 | 6.36 | 6.51 | 6.66 | 6.81 | 6.97 | 7.12 | 7.27 | 7.42 | 7.57 | 7.72 | 7.87 | 8.02 | 8.17 | 8.32 | 8.48 | 8.63 | 8.78 | 8.93 | 9.08 | 9.23 | 9.38 | 9.53 | 9.68 | 9.83 | 9.99 | 10.14 | 10.29 | 10.44 | 10.59 | 10.74 | 10.89 | 11.04 | 11.19 | 11.34 | 11.5 | 11.65 | 11.8 | 11.95
9.8 | 5.91 | 6.06 | 6.21 | 6.36 | 6.51 | 6.66 | 6.81 | 6.96 | 7.11 | 7.27 | 7.42 | 7.57 | 7.72 | 7.87 | 8.02 | 8.17 | 8.32 | 8.47 | 8.62 | 8.78 | 8.93 | 9.08 | 9.23 | 9.38 | 9.53 | 9.68 | 9.83 | 9.98 | 10.13 | 10.28 | 10.44 | 10.59 | 10.74 | 10.89 | 11.04 | 11.19 | 11.34 | 11.49 | 11.64 | 11.79 | 11.95
9.81 | 5.91 | 6.06 | 6.21 | 6.36 | 6.51 | 6.66 | 6.81 | 6.96 | 7.11 | 7.26 | 7.41 | 7.57 | 7.72 | 7.87 | 8.02 | 8.17 | 8.32 | 8.47 | 8.62 | 8.77 | 8.92 | 9.08 | 9.23 | 9.38 | 9.53 | 9.68 | 9.83 | 9.98 | 10.13 | 10.28 | 10.43 | 10.58 | 10.74 | 10.89 | 11.04 | 11.19 | 11.34 | 11.49 | 11.64 | 11.79 | 11.94
9.82 | 5.9 | 6.05 | 6.21 | 6.36 | 6.51 | 6.66 | 6.81 | 6.96 | 7.11 | 7.26 | 7.41 | 7.56 | 7.71 | 7.87 | 8.02 | 8.17 | 8.32 | 8.47 | 8.62 | 8.77 | 8.92 | 9.07 | 9.22 | 9.37 | 9.53 | 9.68 | 9.83 | 9.98 | 10.13 | 10.28 | 10.43 | 10.58 | 10.73 | 10.88 | 11.03 | 11.19 | 11.34 | 11.49 | 11.64 | 11.79 | 11.94
9.83 | 5.9 | 6.05 | 6.2 | 6.35 | 6.51 | 6.66 | 6.81 | 6.96 | 7.11 | 7.26 | 7.41 | 7.56 | 7.71 | 7.86 | 8.01 | 8.17 | 8.32 | 8.47 | 8.62 | 8.77 | 8.92 | 9.07 | 9.22 | 9.37 | 9.52 | 9.67 | 9.83 | 9.98 | 10.13 | 10.28 | 10.43 | 10.58 | 10.73 | 10.88 | 11.03 | 11.18 | 11.33 | 11.48 | 11.63 | 11.79 | 11.94
9.84 | 5.9 | 6.05 | 6.2 | 6.35 | 6.5 | 6.66 | 6.81 | 6.96 | 7.11 | 7.26 | 7.41 | 7.56 | 7.71 | 7.86 | 8.01 | 8.16 | 8.31 | 8.47 | 8.62 | 8.77 | 8.92 | 9.07 | 9.22 | 9.37 | 9.52 | 9.67 | 9.82 | 9.97 | 10.12 | 10.28 | 10.43 | 10.58 | 10.73 | 10.88 | 11.03 | 11.18 | 11.33 | 11.48 | 11.63 | 11.78 | 11.93
9.85 | 5.9 | 6.05 | 6.2 | 6.35 | 6.5 | 6.65 | 6.8 | 6.96 | 7.11 | 7.26 | 7.41 | 7.56 | 7.71 | 7.86 | 8.01 | 8.16 | 8.31 | 8.46 | 8.61 | 8.76 | 8.92 | 9.07 | 9.22 | 9.37 | 9.52 | 9.67 | 9.82 | 9.97 | 10.12 | 10.27 | 10.42 | 10.57 | 10.73 | 10.88 | 11.03 | 11.18 | 11.33 | 11.48 | 11.63 | 11.78 | 11.93
9.86 | 5.9 | 6.05 | 6.2 | 6.35 | 6.5 | 6.65 | 6.8 | 6.95 | 7.1 | 7.26 | 7.41 | 7.56 | 7.71 | 7.86 | 8.01 | 8.16 | 8.31 | 8.46 | 8.61 | 8.76 | 8.91 | 9.06 | 9.22 | 9.37 | 9.52 | 9.67 | 9.82 | 9.97 | 10.12 | 10.27 | 10.42 | 10.57 | 10.72 | 10.87 | 11.02 | 11.17 | 11.33 | 11.48 | 11.63 | 11.78 | 11.93
9.87 | 5.9 | 6.05 | 6.2 | 6.35 | 6.5 | 6.65 | 6.8 | 6.95 | 7.1 | 7.25 | 7.4 | 7.55 | 7.71 | 7.86 | 8.01 | 8.16 | 8.31 | 8.46 | 8.61 | 8.76 | 8.91 | 9.06 | 9.21 | 9.36 | 9.51 | 9.67 | 9.82 | 9.97 | 10.12 | 10.27 | 10.42 | 10.57 | 10.72 | 10.87 | 11.02 | 11.17 | 11.32 | 11.47 | 11.62 | 11.77 | 11.93
9.88 | 5.89 | 6.05 | 6.2 | 6.35 | 6.5 | 6.65 | 6.8 | 6.95 | 7.1 | 7.25 | 7.4 | 7.55 | 7.7 | 7.85 | 8.01 | 8.16 | 8.31 | 8.46 | 8.61 | 8.76 | 8.91 | 9.06 | 9.21 | 9.36 | 9.51 | 9.66 | 9.81 | 9.96 | 10.11 | 10.27 | 10.42 | 10.57 | 10.72 | 10.87 | 11.02 | 11.17 | 11.32 | 11.47 | 11.62 | 11.77 | 11.92
9.89 | 5.89 | 6.04 | 6.19 | 6.35 | 6.5 | 6.65 | 6.8 | 6.95 | 7.1 | 7.25 | 7.4 | 7.55 | 7.7 | 7.85 | 8 | 8.15 | 8.3 | 8.46 | 8.61 | 8.76 | 8.91 | 9.06 | 9.21 | 9.36 | 9.51 | 9.66 | 9.81 | 9.96 | 10.11 | 10.26 | 10.41 | 10.56 | 10.71 | 10.87 | 11.02 | 11.17 | 11.32 | 11.47 | 11.62 | 11.77 | 11.92
9.9 | 5.89 | 6.04 | 6.19 | 6.34 | 6.49 | 6.65 | 6.8 | 6.95 | 7.1 | 7.25 | 7.4 | 7.55 | 7.7 | 7.85 | 8 | 8.15 | 8.3 | 8.45 | 8.6 | 8.75 | 8.91 | 9.06 | 9.21 | 9.36 | 9.51 | 9.66 | 9.81 | 9.96 | 10.11 | 10.26 | 10.41 | 10.56 | 10.71 | 10.86 | 11.01 | 11.16 | 11.31 | 11.47 | 11.62 | 11.77 | 11.92
9.91 | 5.89 | 6.04 | 6.19 | 6.34 | 6.49 | 6.64 | 6.79 | 6.95 | 7.1 | 7.25 | 7.4 | 7.55 | 7.7 | 7.85 | 8 | 8.15 | 8.3 | 8.45 | 8.6 | 8.75 | 8.9 | 9.05 | 9.2 | 9.35 | 9.51 | 9.66 | 9.81 | 9.96 | 10.11 | 10.26 | 10.41 | 10.56 | 10.71 | 10.86 | 11.01 | 11.16 | 11.31 | 11.46 | 11.61 | 11.76 | 11.91
9.92 | 5.89 | 6.04 | 6.19 | 6.34 | 6.49 | 6.64 | 6.79 | 6.94 | 7.09 | 7.24 | 7.4 | 7.55 | 7.7 | 7.85 | 8 | 8.15 | 8.3 | 8.45 | 8.6 | 8.75 | 8.9 | 9.05 | 9.2 | 9.35 | 9.5 | 9.65 | 9.8 | 9.95 | 10.11 | 10.26 | 10.41 | 10.56 | 10.71 | 10.86 | 11.01 | 11.16 | 11.31 | 11.46 | 11.61 | 11.76 | 11.91
9.93 | 5.89 | 6.04 | 6.19 | 6.34 | 6.49 | 6.64 | 6.79 | 6.94 | 7.09 | 7.24 | 7.39 | 7.54 | 7.69 | 7.85 | 8 | 8.15 | 8.3 | 8.45 | 8.6 | 8.75 | 8.9 | 9.05 | 9.2 | 9.35 | 9.5 | 9.65 | 9.8 | 9.95 | 10.1 | 10.25 | 10.4 | 10.55 | 10.7 | 10.86 | 11.01 | 11.16 | 11.31 | 11.46 | 11.61 | 11.76 | 11.91
9.94 | 5.89 | 6.04 | 6.19 | 6.34 | 6.49 | 6.64 | 6.79 | 6.94 | 7.09 | 7.24 | 7.39 | 7.54 | 7.69 | 7.84 | 7.99 | 8.14 | 8.29 | 8.45 | 8.6 | 8.75 | 8.9 | 9.05 | 9.2 | 9.35 | 9.5 | 9.65 | 9.8 | 9.95 | 10.1 | 10.25 | 10.4 | 10.55 | 10.7 | 10.85 | 11 | 11.15 | 11.3 | 11.45 | 11.6 | 11.76 | 11.91
9.95 | 5.88 | 6.04 | 6.19 | 6.34 | 6.49 | 6.64 | 6.79 | 6.94 | 7.09 | 7.24 | 7.39 | 7.54 | 7.69 | 7.84 | 7.99 | 8.14 | 8.29 | 8.44 | 8.59 | 8.74 | 8.89 | 9.04 | 9.2 | 9.35 | 9.5 | 9.65 | 9.8 | 9.95 | 10.1 | 10.25 | 10.4 | 10.55 | 10.7 | 10.85 | 11 | 11.15 | 11.3 | 11.45 | 11.6 | 11.75 | 11.9
9.96 | 5.88 | 6.03 | 6.18 | 6.33 | 6.49 | 6.64 | 6.79 | 6.94 | 7.09 | 7.24 | 7.39 | 7.54 | 7.69 | 7.84 | 7.99 | 8.14 | 8.29 | 8.44 | 8.59 | 8.74 | 8.89 | 9.04 | 9.19 | 9.34 | 9.49 | 9.64 | 9.79 | 9.95 | 10.1 | 10.25 | 10.4 | 10.55 | 10.7 | 10.85 | 11 | 11.15 | 11.3 | 11.45 | 11.6 | 11.75 | 11.9
9.97 | 5.88 | 6.03 | 6.18 | 6.33 | 6.48 | 6.63 | 6.78 | 6.94 | 7.09 | 7.24 | 7.39 | 7.54 | 7.69 | 7.84 | 7.99 | 8.14 | 8.29 | 8.44 | 8.59 | 8.74 | 8.89 | 9.04 | 9.19 | 9.34 | 9.49 | 9.64 | 9.79 | 9.94 | 10.09 | 10.24 | 10.39 | 10.54 | 10.69 | 10.84 | 11 | 11.15 | 11.3 | 11.45 | 11.6 | 11.75 | 11.9
9.98 | 5.88 | 6.03 | 6.18 | 6.33 | 6.48 | 6.63 | 6.78 | 6.93 | 7.08 | 7.23 | 7.38 | 7.53 | 7.69 | 7.84 | 7.99 | 8.14 | 8.29 | 8.44 | 8.59 | 8.74 | 8.89 | 9.04 | 9.19 | 9.34 | 9.49 | 9.64 | 9.79 | 9.94 | 10.09 | 10.24 | 10.39 | 10.54 | 10.69 | 10.84 | 10.99 | 11.14 | 11.29 | 11.44 | 11.59 | 11.74 | 11.89
9.99 | 5.88 | 6.03 | 6.18 | 6.33 | 6.48 | 6.63 | 6.78 | 6.93 | 7.08 | 7.23 | 7.38 | 7.53 | 7.68 | 7.83 | 7.98 | 8.13 | 8.28 | 8.44 | 8.59 | 8.74 | 8.89 | 9.04 | 9.19 | 9.34 | 9.49 | 9.64 | 9.79 | 9.94 | 10.09 | 10.24 | 10.39 | 10.54 | 10.69 | 10.84 | 10.99 | 11.14 | 11.29 | 11.44 | 11.59 | 11.74 | 11.89
10 | 5.88 | 6.03 | 6.18 | 6.33 | 6.48 | 6.63 | 6.78 | 6.93 | 7.08 | 7.23 | 7.38 | 7.53 | 7.68 | 7.83 | 7.98 | 8.13 | 8.28 | 8.43 | 8.58 | 8.73 | 8.88 | 9.03 | 9.18 | 9.33 | 9.49 | 9.64 | 9.79 | 9.94 | 10.09 | 10.24 | 10.39 | 10.54 | 10.69 | 10.84 | 10.99 | 11.14 | 11.29 | 11.44 | 11.59 | 11.74 | 11.89
10.01 | 5.88 | 6.03 | 6.18 | 6.33 | 6.48 | 6.63 | 6.78 | 6.93 | 7.08 | 7.23 | 7.38 | 7.53 | 7.68 | 7.83 | 7.98 | 8.13 | 8.28 | 8.43 | 8.58 | 8.73 | 8.88 | 9.03 | 9.18 | 9.33 | 9.48 | 9.63 | 9.78 | 9.93 | 10.08 | 10.23 | 10.38 | 10.53 | 10.68 | 10.83 | 10.98 | 11.14 | 11.29 | 11.44 | 11.59 | 11.74 | 11.89
10.02 | 5.87 | 6.03 | 6.18 | 6.33 | 6.48 | 6.63 | 6.78 | 6.93 | 7.08 | 7.23 | 7.38 | 7.53 | 7.68 | 7.83 | 7.98 | 8.13 | 8.28 | 8.43 | 8.58 | 8.73 | 8.88 | 9.03 | 9.18 | 9.33 | 9.48 | 9.63 | 9.78 | 9.93 | 10.08 | 10.23 | 10.38 | 10.53 | 10.68 | 10.83 | 10.98 | 11.13 | 11.28 | 11.43 | 11.58 | 11.73 | 11.88
10.03 | 5.87 | 6.02 | 6.17 | 6.32 | 6.47 | 6.62 | 6.77 | 6.93 | 7.08 | 7.23 | 7.38 | 7.53 | 7.68 | 7.83 | 7.98 | 8.13 | 8.28 | 8.43 | 8.58 | 8.73 | 8.88 | 9.03 | 9.18 | 9.33 | 9.48 | 9.63 | 9.78 | 9.93 | 10.08 | 10.23 | 10.38 | 10.53 | 10.68 | 10.83 | 10.98 | 11.13 | 11.28 | 11.43 | 11.58 | 11.73 | 11.88
10.04 | 5.87 | 6.02 | 6.17 | 6.32 | 6.47 | 6.62 | 6.77 | 6.92 | 7.07 | 7.22 | 7.37 | 7.52 | 7.67 | 7.82 | 7.97 | 8.12 | 8.27 | 8.43 | 8.58 | 8.73 | 8.88 | 9.03 | 9.18 | 9.33 | 9.48 | 9.63 | 9.78 | 9.93 | 10.08 | 10.23 | 10.38 | 10.53 | 10.68 | 10.83 | 10.98 | 11.13 | 11.28 | 11.43 | 11.58 | 11.73 | 11.88
10.05 | 5.87 | 6.02 | 6.17 | 6.32 | 6.47 | 6.62 | 6.77 | 6.92 | 7.07 | 7.22 | 7.37 | 7.52 | 7.67 | 7.82 | 7.97 | 8.12 | 8.27 | 8.42 | 8.57 | 8.72 | 8.87 | 9.02 | 9.17 | 9.32 | 9.47 | 9.62 | 9.77 | 9.92 | 10.07 | 10.22 | 10.37 | 10.52 | 10.67 | 10.82 | 10.97 | 11.12 | 11.27 | 11.42 | 11.57 | 11.72 | 11.87
10.06 | 5.87 | 6.02 | 6.17 | 6.32 | 6.47 | 6.62 | 6.77 | 6.92 | 7.07 | 7.22 | 7.37 | 7.52 | 7.67 | 7.82 | 7.97 | 8.12 | 8.27 | 8.42 | 8.57 | 8.72 | 8.87 | 9.02 | 9.17 | 9.32 | 9.47 | 9.62 | 9.77 | 9.92 | 10.07 | 10.22 | 10.37 | 10.52 | 10.67 | 10.82 | 10.97 | 11.12 | 11.27 | 11.42 | 11.57 | 11.72 | 11.87
10.07 | 5.87 | 6.02 | 6.17 | 6.32 | 6.47 | 6.62 | 6.77 | 6.92 | 7.07 | 7.22 | 7.37 | 7.52 | 7.67 | 7.82 | 7.97 | 8.12 | 8.27 | 8.42 | 8.57 | 8.72 | 8.87 | 9.02 | 9.17 | 9.32 | 9.47 | 9.62 | 9.77 | 9.92 | 10.07 | 10.22 | 10.37 | 10.52 | 10.67 | 10.82 | 10.97 | 11.12 | 11.27 | 11.42 | 11.57 | 11.72 | 11.87
10.08 | 5.87 | 6.02 | 6.17 | 6.32 | 6.47 | 6.62 | 6.77 | 6.92 | 7.07 | 7.22 | 7.37 | 7.52 | 7.67 | 7.82 | 7.97 | 8.12 | 8.27 | 8.42 | 8.57 | 8.72 | 8.87 | 9.02 | 9.17 | 9.32 | 9.47 | 9.62 | 9.77 | 9.92 | 10.07 | 10.22 | 10.37 | 10.52 | 10.67 | 10.82 | 10.97 | 11.12 | 11.27 | 11.42 | 11.57 | 11.72 | 11.87
10.09 | 5.86 | 6.01 | 6.16 | 6.32 | 6.47 | 6.62 | 6.77 | 6.92 | 7.07 | 7.22 | 7.37 | 7.52 | 7.67 | 7.82 | 7.97 | 8.12 | 8.27 | 8.42 | 8.57 | 8.71 | 8.86 | 9.01 | 9.16 | 9.31 | 9.46 | 9.61 | 9.76 | 9.91 | 10.06 | 10.21 | 10.36 | 10.51 | 10.66 | 10.81 | 10.96 | 11.11 | 11.26 | 11.41 | 11.56 | 11.71 | 11.86
10.1 | 5.86 | 6.01 | 6.16 | 6.31 | 6.46 | 6.61 | 6.76 | 6.91 | 7.06 | 7.21 | 7.36 | 7.51 | 7.66 | 7.81 | 7.96 | 8.11 | 8.26 | 8.41 | 8.56 | 8.71 | 8.86 | 9.01 | 9.16 | 9.31 | 9.46 | 9.61 | 9.76 | 9.91 | 10.06 | 10.21 | 10.36 | 10.51 | 10.66 | 10.81 | 10.96 | 11.11 | 11.26 | 11.41 | 11.56 | 11.71 | 11.86
10.11 | 5.86 | 6.01 | 6.16 | 6.31 | 6.46 | 6.61 | 6.76 | 6.91 | 7.06 | 7.21 | 7.36 | 7.51 | 7.66 | 7.81 | 7.96 | 8.11 | 8.26 | 8.41 | 8.56 | 8.71 | 8.86 | 9.01 | 9.16 | 9.31 | 9.46 | 9.61 | 9.76 | 9.91 | 10.06 | 10.21 | 10.36 | 10.51 | 10.66 | 10.81 | 10.96 | 11.11 | 11.26 | 11.41 | 11.56 | 11.71 | 11.86
10.12 | 5.86 | 6.01 | 6.16 | 6.31 | 6.46 | 6.61 | 6.76 | 6.91 | 7.06 | 7.21 | 7.36 | 7.51 | 7.66 | 7.81 | 7.96 | 8.11 | 8.26 | 8.41 | 8.56 | 8.71 | 8.86 | 9.01 | 9.16 | 9.31 | 9.46 | 9.61 | 9.76 | 9.91 | 10.06 | 10.21 | 10.36 | 10.51 | 10.66 | 10.81 | 10.96 | 11.11 | 11.26 | 11.41 | 11.56 | 11.71 | 11.86
10.13 | 5.86 | 6.01 | 6.16 | 6.31 | 6.46 | 6.61 | 6.76 | 6.91 | 7.06 | 7.21 | 7.36 | 7.51 | 7.66 | 7.81 | 7.96 | 8.11 | 8.26 | 8.41 | 8.56 | 8.71 | 8.86 | 9.01 | 9.16 | 9.31 | 9.46 | 9.61 | 9.76 | 9.91 | 10.06 | 10.2 | 10.35 | 10.5 | 10.65 | 10.8 | 10.95 | 11.1 | 11.25 | 11.4 | 11.55 | 11.7 | 11.85
10.14 | 5.86 | 6.01 | 6.16 | 6.31 | 6.46 | 6.61 | 6.76 | 6.91 | 7.06 | 7.21 | 7.36 | 7.51 | 7.66 | 7.81 | 7.96 | 8.11 | 8.26 | 8.41 | 8.55 | 8.7 | 8.85 | 9 | 9.15 | 9.3 | 9.45 | 9.6 | 9.75 | 9.9 | 10.05 | 10.2 | 10.35 | 10.5 | 10.65 | 10.8 | 10.95 | 11.1 | 11.25 | 11.4 | 11.55 | 11.7 | 11.85
10.15 | 5.86 | 6.01 | 6.16 | 6.31 | 6.46 | 6.61 | 6.76 | 6.91 | 7.05 | 7.2 | 7.35 | 7.5 | 7.65 | 7.8 | 7.95 | 8.1 | 8.25 | 8.4 | 8.55 | 8.7 | 8.85 | 9 | 9.15 | 9.3 | 9.45 | 9.6 | 9.75 | 9.9 | 10.05 | 10.2 | 10.35 | 10.5 | 10.65 | 10.8 | 10.95 | 11.1 | 11.25 | 11.4 | 11.55 | 11.7 | 11.85
10.16 | 5.85 | 6 | 6.15 | 6.3 | 6.45 | 6.6 | 6.75 | 6.9 | 7.05 | 7.2 | 7.35 | 7.5 | 7.65 | 7.8 | 7.95 | 8.1 | 8.25 | 8.4 | 8.55 | 8.7 | 8.85 | 9 | 9.15 | 9.3 | 9.45 | 9.6 | 9.75 | 9.9 | 10.05 | 10.2 | 10.35 | 10.5 | 10.65 | 10.8 | 10.95 | 11.1 | 11.25 | 11.4 | 11.54 | 11.69 | 11.84
10.17 | 5.85 | 6 | 6.15 | 6.3 | 6.45 | 6.6 | 6.75 | 6.9 | 7.05 | 7.2 | 7.35 | 7.5 | 7.65 | 7.8 | 7.95 | 8.1 | 8.25 | 8.4 | 8.55 | 8.7 | 8.85 | 9 | 9.15 | 9.3 | 9.45 | 9.6 | 9.75 | 9.9 | 10.05 | 10.2 | 10.34 | 10.49 | 10.64 | 10.79 | 10.94 | 11.09 | 11.24 | 11.39 | 11.54 | 11.69 | 11.84
10.18 | 5.85 | 6 | 6.15 | 6.3 | 6.45 | 6.6 | 6.75 | 6.9 | 7.05 | 7.2 | 7.35 | 7.5 | 7.65 | 7.8 | 7.95 | 8.1 | 8.25 | 8.4 | 8.55 | 8.7 | 8.85 | 9 | 9.15 | 9.3 | 9.44 | 9.59 | 9.74 | 9.89 | 10.04 | 10.19 | 10.34 | 10.49 | 10.64 | 10.79 | 10.94 | 11.09 | 11.24 | 11.39 | 11.54 | 11.69 | 11.84
10.19 | 5.85 | 6 | 6.15 | 6.3 | 6.45 | 6.6 | 6.75 | 6.9 | 7.05 | 7.2 | 7.35 | 7.5 | 7.65 | 7.8 | 7.95 | 8.1 | 8.25 | 8.4 | 8.54 | 8.69 | 8.84 | 8.99 | 9.14 | 9.29 | 9.44 | 9.59 | 9.74 | 9.89 | 10.04 | 10.19 | 10.34 | 10.49 | 10.64 | 10.79 | 10.94 | 11.09 | 11.24 | 11.39 | 11.54 | 11.69 | 11.84
10.2 | 5.85 | 6 | 6.15 | 6.3 | 6.45 | 6.6 | 6.75 | 6.9 | 7.05 | 7.2 | 7.35 | 7.5 | 7.64 | 7.79 | 7.94 | 8.09 | 8.24 | 8.39 | 8.54 | 8.69 | 8.84 | 8.99 | 9.14 | 9.29 | 9.44 | 9.59 | 9.74 | 9.89 | 10.04 | 10.19 | 10.34 | 10.49 | 10.64 | 10.79 | 10.94 | 11.09 | 11.23 | 11.38 | 11.53 | 11.68 | 11.83
10.21 | 5.85 | 6 | 6.15 | 6.3 | 6.45 | 6.6 | 6.75 | 6.9 | 7.04 | 7.19 | 7.34 | 7.49 | 7.64 | 7.79 | 7.94 | 8.09 | 8.24 | 8.39 | 8.54 | 8.69 | 8.84 | 8.99 | 9.14 | 9.29 | 9.44 | 9.59 | 9.74 | 9.89 | 10.04 | 10.19 | 10.34 | 10.48 | 10.63 | 10.78 | 10.93 | 11.08 | 11.23 | 11.38 | 11.53 | 11.68 | 11.83
10.22 | 5.85 | 6 | 6.15 | 6.3 | 6.44 | 6.59 | 6.74 | 6.89 | 7.04 | 7.19 | 7.34 | 7.49 | 7.64 | 7.79 | 7.94 | 8.09 | 8.24 | 8.39 | 8.54 | 8.69 | 8.84 | 8.99 | 9.14 | 9.29 | 9.44 | 9.59 | 9.73 | 9.88 | 10.03 | 10.18 | 10.33 | 10.48 | 10.63 | 10.78 | 10.93 | 11.08 | 11.23 | 11.38 | 11.53 | 11.68 | 11.83
10.23 | 5.85 | 5.99 | 6.14 | 6.29 | 6.44 | 6.59 | 6.74 | 6.89 | 7.04 | 7.19 | 7.34 | 7.49 | 7.64 | 7.79 | 7.94 | 8.09 | 8.24 | 8.39 | 8.54 | 8.69 | 8.84 | 8.99 | 9.13 | 9.28 | 9.43 | 9.58 | 9.73 | 9.88 | 10.03 | 10.18 | 10.33 | 10.48 | 10.63 | 10.78 | 10.93 | 11.08 | 11.23 | 11.38 | 11.53 | 11.68 | 11.82
10.24 | 5.84 | 5.99 | 6.14 | 6.29 | 6.44 | 6.59 | 6.74 | 6.89 | 7.04 | 7.19 | 7.34 | 7.49 | 7.64 | 7.79 | 7.94 | 8.09 | 8.24 | 8.39 | 8.53 | 8.68 | 8.83 | 8.98 | 9.13 | 9.28 | 9.43 | 9.58 | 9.73 | 9.88 | 10.03 | 10.18 | 10.33 | 10.48 | 10.63 | 10.78 | 10.93 | 11.07 | 11.22 | 11.37 | 11.52 | 11.67 | 11.82
10.25 | 5.84 | 5.99 | 6.14 | 6.29 | 6.44 | 6.59 | 6.74 | 6.89 | 7.04 | 7.19 | 7.34 | 7.49 | 7.64 | 7.79 | 7.93 | 8.08 | 8.23 | 8.38 | 8.53 | 8.68 | 8.83 | 8.98 | 9.13 | 9.28 | 9.43 | 9.58 | 9.73 | 9.88 | 10.03 | 10.18 | 10.33 | 10.47 | 10.62 | 10.77 | 10.92 | 11.07 | 11.22 | 11.37 | 11.52 | 11.67 | 11.82
10.26 | 5.84 | 5.99 | 6.14 | 6.29 | 6.44 | 6.59 | 6.74 | 6.89 | 7.04 | 7.19 | 7.34 | 7.48 | 7.63 | 7.78 | 7.93 | 8.08 | 8.23 | 8.38 | 8.53 | 8.68 | 8.83 | 8.98 | 9.13 | 9.28 | 9.43 | 9.58 | 9.73 | 9.87 | 10.02 | 10.17 | 10.32 | 10.47 | 10.62 | 10.77 | 10.92 | 11.07 | 11.22 | 11.37 | 11.52 | 11.67 | 11.82
10.27 | 5.84 | 5.99 | 6.14 | 6.29 | 6.44 | 6.59 | 6.74 | 6.89 | 7.03 | 7.18 | 7.33 | 7.48 | 7.63 | 7.78 | 7.93 | 8.08 | 8.23 | 8.38 | 8.53 | 8.68 | 8.83 | 8.98 | 9.13 | 9.28 | 9.42 | 9.57 | 9.72 | 9.87 | 10.02 | 10.17 | 10.32 | 10.47 | 10.62 | 10.77 | 10.92 | 11.07 | 11.22 | 11.37 | 11.52 | 11.66 | 11.81
10.28 | 5.84 | 5.99 | 6.14 | 6.29 | 6.44 | 6.58 | 6.73 | 6.88 | 7.03 | 7.18 | 7.33 | 7.48 | 7.63 | 7.78 | 7.93 | 8.08 | 8.23 | 8.38 | 8.53 | 8.68 | 8.83 | 8.97 | 9.12 | 9.27 | 9.42 | 9.57 | 9.72 | 9.87 | 10.02 | 10.17 | 10.32 | 10.47 | 10.62 | 10.77 | 10.92 | 11.06 | 11.21 | 11.36 | 11.51 | 11.66 | 11.81
10.29 | 5.84 | 5.99 | 6.14 | 6.28 | 6.43 | 6.58 | 6.73 | 6.88 | 7.03 | 7.18 | 7.33 | 7.48 | 7.63 | 7.78 | 7.93 | 8.08 | 8.23 | 8.38 | 8.52 | 8.67 | 8.82 | 8.97 | 9.12 | 9.27 | 9.42 | 9.57 | 9.72 | 9.87 | 10.02 | 10.17 | 10.32 | 10.46 | 10.61 | 10.76 | 10.91 | 11.06 | 11.21 | 11.36 | 11.51 | 11.66 | 11.81
10.3 | 5.84 | 5.98 | 6.13 | 6.28 | 6.43 | 6.58 | 6.73 | 6.88 | 7.03 | 7.18 | 7.33 | 7.48 | 7.63 | 7.78 | 7.93 | 8.07 | 8.22 | 8.37 | 8.52 | 8.67 | 8.82 | 8.97 | 9.12 | 9.27 | 9.42 | 9.57 | 9.72 | 9.87 | 10.01 | 10.16 | 10.31 | 10.46 | 10.61 | 10.76 | 10.91 | 11.06 | 11.21 | 11.36 | 11.51 | 11.66 | 11.81
10.31 | 5.83 | 5.98 | 6.13 | 6.28 | 6.43 | 6.58 | 6.73 | 6.88 | 7.03 | 7.18 | 7.33 | 7.48 | 7.62 | 7.77 | 7.92 | 8.07 | 8.22 | 8.37 | 8.52 | 8.67 | 8.82 | 8.97 | 9.12 | 9.27 | 9.42 | 9.56 | 9.71 | 9.86 | 10.01 | 10.16 | 10.31 | 10.46 | 10.61 | 10.76 | 10.91 | 11.06 | 11.21 | 11.36 | 11.5 | 11.65 | 11.8
10.32 | 5.83 | 5.98 | 6.13 | 6.28 | 6.43 | 6.58 | 6.73 | 6.88 | 7.03 | 7.18 | 7.32 | 7.47 | 7.62 | 7.77 | 7.92 | 8.07 | 8.22 | 8.37 | 8.52 | 8.67 | 8.82 | 8.97 | 9.12 | 9.26 | 9.41 | 9.56 | 9.71 | 9.86 | 10.01 | 10.16 | 10.31 | 10.46 | 10.61 | 10.76 | 10.91 | 11.05 | 11.2 | 11.35 | 11.5 | 11.65 | 11.8
10.33 | 5.83 | 5.98 | 6.13 | 6.28 | 6.43 | 6.58 | 6.73 | 6.88 | 7.02 | 7.17 | 7.32 | 7.47 | 7.62 | 7.77 | 7.92 | 8.07 | 8.22 | 8.37 | 8.52 | 8.67 | 8.81 | 8.96 | 9.11 | 9.26 | 9.41 | 9.56 | 9.71 | 9.86 | 10.01 | 10.16 | 10.31 | 10.46 | 10.6 | 10.75 | 10.9 | 11.05 | 11.2 | 11.35 | 11.5 | 11.65 | 11.8
10.34 | 5.83 | 5.98 | 6.13 | 6.28 | 6.43 | 6.58 | 6.72 | 6.87 | 7.02 | 7.17 | 7.32 | 7.47 | 7.62 | 7.77 | 7.92 | 8.07 | 8.22 | 8.37 | 8.51 | 8.66 | 8.81 | 8.96 | 9.11 | 9.26 | 9.41 | 9.56 | 9.71 | 9.86 | 10.01 | 10.15 | 10.3 | 10.45 | 10.6 | 10.75 | 10.9 | 11.05 | 11.2 | 11.35 | 11.5 | 11.65 | 11.79
10.35 | 5.83 | 5.98 | 6.13 | 6.28 | 6.42 | 6.57 | 6.72 | 6.87 | 7.02 | 7.17 | 7.32 | 7.47 | 7.62 | 7.77 | 7.92 | 8.06 | 8.21 | 8.36 | 8.51 | 8.66 | 8.81 | 8.96 | 9.11 | 9.26 | 9.41 | 9.56 | 9.7 | 9.85 | 10 | 10.15 | 10.3 | 10.45 | 10.6 | 10.75 | 10.9 | 11.05 | 11.2 | 11.34 | 11.49 | 11.64 | 11.79
10.36 | 5.83 | 5.98 | 6.12 | 6.27 | 6.42 | 6.57 | 6.72 | 6.87 | 7.02 | 7.17 | 7.32 | 7.47 | 7.62 | 7.76 | 7.91 | 8.06 | 8.21 | 8.36 | 8.51 | 8.66 | 8.81 | 8.96 | 9.11 | 9.26 | 9.4 | 9.55 | 9.7 | 9.85 | 10 | 10.15 | 10.3 | 10.45 | 10.6 | 10.75 | 10.89 | 11.04 | 11.19 | 11.34 | 11.49 | 11.64 | 11.79
10.37 | 5.83 | 5.97 | 6.12 | 6.27 | 6.42 | 6.57 | 6.72 | 6.87 | 7.02 | 7.17 | 7.32 | 7.46 | 7.61 | 7.76 | 7.91 | 8.06 | 8.21 | 8.36 | 8.51 | 8.66 | 8.81 | 8.96 | 9.1 | 9.25 | 9.4 | 9.55 | 9.7 | 9.85 | 10 | 10.15 | 10.3 | 10.45 | 10.59 | 10.74 | 10.89 | 11.04 | 11.19 | 11.34 | 11.49 | 11.64 | 11.79
10.38 | 5.82 | 5.97 | 6.12 | 6.27 | 6.42 | 6.57 | 6.72 | 6.87 | 7.02 | 7.17 | 7.31 | 7.46 | 7.61 | 7.76 | 7.91 | 8.06 | 8.21 | 8.36 | 8.51 | 8.66 | 8.8 | 8.95 | 9.1 | 9.25 | 9.4 | 9.55 | 9.7 | 9.85 | 10 | 10.14 | 10.29 | 10.44 | 10.59 | 10.74 | 10.89 | 11.04 | 11.19 | 11.34 | 11.49 | 11.63 | 11.78
10.39 | 5.82 | 5.97 | 6.12 | 6.27 | 6.42 | 6.57 | 6.72 | 6.87 | 7.01 | 7.16 | 7.31 | 7.46 | 7.61 | 7.76 | 7.91 | 8.06 | 8.21 | 8.36 | 8.5 | 8.65 | 8.8 | 8.95 | 9.1 | 9.25 | 9.4 | 9.55 | 9.7 | 9.84 | 9.99 | 10.14 | 10.29 | 10.44 | 10.59 | 10.74 | 10.89 | 11.04 | 11.18 | 11.33 | 11.48 | 11.63 | 11.78
10.4 | 5.82 | 5.97 | 6.12 | 6.27 | 6.42 | 6.57 | 6.71 | 6.86 | 7.01 | 7.16 | 7.31 | 7.46 | 7.61 | 7.76 | 7.91 | 8.06 | 8.2 | 8.35 | 8.5 | 8.65 | 8.8 | 8.95 | 9.1 | 9.25 | 9.4 | 9.54 | 9.69 | 9.84 | 9.99 | 10.14 | 10.29 | 10.44 | 10.59 | 10.74 | 10.88 | 11.03 | 11.18 | 11.33 | 11.48 | 11.63 | 11.78
10.41 | 5.82 | 5.97 | 6.12 | 6.27 | 6.42 | 6.56 | 6.71 | 6.86 | 7.01 | 7.16 | 7.31 | 7.46 | 7.61 | 7.76 | 7.9 | 8.05 | 8.2 | 8.35 | 8.5 | 8.65 | 8.8 | 8.95 | 9.1 | 9.24 | 9.39 | 9.54 | 9.69 | 9.84 | 9.99 | 10.14 | 10.29 | 10.44 | 10.58 | 10.73 | 10.88 | 11.03 | 11.18 | 11.33 | 11.48 | 11.63 | 11.78
10.42 | 5.82 | 5.97 | 6.12 | 6.26 | 6.41 | 6.56 | 6.71 | 6.86 | 7.01 | 7.16 | 7.31 | 7.46 | 7.6 | 7.75 | 7.9 | 8.05 | 8.2 | 8.35 | 8.5 | 8.65 | 8.8 | 8.94 | 9.09 | 9.24 | 9.39 | 9.54 | 9.69 | 9.84 | 9.99 | 10.14 | 10.28 | 10.43 | 10.58 | 10.73 | 10.88 | 11.03 | 11.18 | 11.33 | 11.47 | 11.62 | 11.77
10.43 | 5.82 | 5.97 | 6.11 | 6.26 | 6.41 | 6.56 | 6.71 | 6.86 | 7.01 | 7.16 | 7.31 | 7.45 | 7.6 | 7.75 | 7.9 | 8.05 | 8.2 | 8.35 | 8.5 | 8.64 | 8.79 | 8.94 | 9.09 | 9.24 | 9.39 | 9.54 | 9.69 | 9.84 | 9.98 | 10.13 | 10.28 | 10.43 | 10.58 | 10.73 | 10.88 | 11.03 | 11.17 | 11.32 | 11.47 | 11.62 | 11.77
10.44 | 5.82 | 5.96 | 6.11 | 6.26 | 6.41 | 6.56 | 6.71 | 6.86 | 7.01 | 7.15 | 7.3 | 7.45 | 7.6 | 7.75 | 7.9 | 8.05 | 8.2 | 8.35 | 8.49 | 8.64 | 8.79 | 8.94 | 9.09 | 9.24 | 9.39 | 9.54 | 9.68 | 9.83 | 9.98 | 10.13 | 10.28 | 10.43 | 10.58 | 10.73 | 10.87 | 11.02 | 11.17 | 11.32 | 11.47 | 11.62 | 11.77
10.45 | 5.81 | 5.96 | 6.11 | 6.26 | 6.41 | 6.56 | 6.71 | 6.86 | 7 | 7.15 | 7.3 | 7.45 | 7.6 | 7.75 | 7.9 | 8.05 | 8.19 | 8.34 | 8.49 | 8.64 | 8.79 | 8.94 | 9.09 | 9.24 | 9.38 | 9.53 | 9.68 | 9.83 | 9.98 | 10.13 | 10.28 | 10.43 | 10.57 | 10.72 | 10.87 | 11.02 | 11.17 | 11.32 | 11.47 | 11.62 | 11.76
10.46 | 5.81 | 5.96 | 6.11 | 6.26 | 6.41 | 6.56 | 6.71 | 6.85 | 7 | 7.15 | 7.3 | 7.45 | 7.6 | 7.75 | 7.9 | 8.04 | 8.19 | 8.34 | 8.49 | 8.64 | 8.79 | 8.94 | 9.09 | 9.23 | 9.38 | 9.53 | 9.68 | 9.83 | 9.98 | 10.13 | 10.27 | 10.42 | 10.57 | 10.72 | 10.87 | 11.02 | 11.17 | 11.32 | 11.46 | 11.61 | 11.76
10.47 | 5.81 | 5.96 | 6.11 | 6.26 | 6.41 | 6.55 | 6.7 | 6.85 | 7 | 7.15 | 7.3 | 7.45 | 7.6 | 7.74 | 7.89 | 8.04 | 8.19 | 8.34 | 8.49 | 8.64 | 8.79 | 8.93 | 9.08 | 9.23 | 9.38 | 9.53 | 9.68 | 9.83 | 9.97 | 10.12 | 10.27 | 10.42 | 10.57 | 10.72 | 10.87 | 11.02 | 11.16 | 11.31 | 11.46 | 11.61 | 11.76
10.48 | 5.81 | 5.96 | 6.11 | 6.26 | 6.4 | 6.55 | 6.7 | 6.85 | 7 | 7.15 | 7.3 | 7.45 | 7.59 | 7.74 | 7.89 | 8.04 | 8.19 | 8.34 | 8.49 | 8.63 | 8.78 | 8.93 | 9.08 | 9.23 | 9.38 | 9.53 | 9.68 | 9.82 | 9.97 | 10.12 | 10.27 | 10.42 | 10.57 | 10.72 | 10.86 | 11.01 | 11.16 | 11.31 | 11.46 | 11.61 | 11.76
10.49 | 5.81 | 5.96 | 6.11 | 6.25 | 6.4 | 6.55 | 6.7 | 6.85 | 7 | 7.15 | 7.3 | 7.44 | 7.59 | 7.74 | 7.89 | 8.04 | 8.19 | 8.34 | 8.48 | 8.63 | 8.78 | 8.93 | 9.08 | 9.23 | 9.38 | 9.52 | 9.67 | 9.82 | 9.97 | 10.12 | 10.27 | 10.42 | 10.56 | 10.71 | 10.86 | 11.01 | 11.16 | 11.31 | 11.46 | 11.6 | 11.75
10.5 | 5.81 | 5.96 | 6.1 | 6.25 | 6.4 | 6.55 | 6.7 | 6.85 | 7 | 7.14 | 7.29 | 7.44 | 7.59 | 7.74 | 7.89 | 8.04 | 8.18 | 8.33 | 8.48 | 8.63 | 8.78 | 8.93 | 9.08 | 9.23 | 9.37 | 9.52 | 9.67 | 9.82 | 9.97 | 10.12 | 10.26 | 10.41 | 10.56 | 10.71 | 10.86 | 11.01 | 11.16 | 11.3 | 11.45 | 11.6 | 11.75
10.51 | 5.81 | 5.95 | 6.1 | 6.25 | 6.4 | 6.55 | 6.7 | 6.85 | 6.99 | 7.14 | 7.29 | 7.44 | 7.59 | 7.74 | 7.89 | 8.03 | 8.18 | 8.33 | 8.48 | 8.63 | 8.78 | 8.93 | 9.07 | 9.22 | 9.37 | 9.52 | 9.67 | 9.82 | 9.97 | 10.11 | 10.26 | 10.41 | 10.56 | 10.71 | 10.86 | 11.01 | 11.15 | 11.3 | 11.45 | 11.6 | 11.75
10.52 | 5.8 | 5.95 | 6.1 | 6.25 | 6.4 | 6.55 | 6.7 | 6.84 | 6.99 | 7.14 | 7.29 | 7.44 | 7.59 | 7.74 | 7.88 | 8.03 | 8.18 | 8.33 | 8.48 | 8.63 | 8.78 | 8.92 | 9.07 | 9.22 | 9.37 | 9.52 | 9.67 | 9.81 | 9.96 | 10.11 | 10.26 | 10.41 | 10.56 | 10.71 | 10.85 | 11 | 11.15 | 11.3 | 11.45 | 11.6 | 11.74
10.53 | 5.8 | 5.95 | 6.1 | 6.25 | 6.4 | 6.55 | 6.69 | 6.84 | 6.99 | 7.14 | 7.29 | 7.44 | 7.59 | 7.73 | 7.88 | 8.03 | 8.18 | 8.33 | 8.48 | 8.62 | 8.77 | 8.92 | 9.07 | 9.22 | 9.37 | 9.52 | 9.66 | 9.81 | 9.96 | 10.11 | 10.26 | 10.41 | 10.55 | 10.7 | 10.85 | 11 | 11.15 | 11.3 | 11.45 | 11.59 | 11.74
10.54 | 5.8 | 5.95 | 6.1 | 6.25 | 6.4 | 6.54 | 6.69 | 6.84 | 6.99 | 7.14 | 7.29 | 7.43 | 7.58 | 7.73 | 7.88 | 8.03 | 8.18 | 8.33 | 8.47 | 8.62 | 8.77 | 8.92 | 9.07 | 9.22 | 9.36 | 9.51 | 9.66 | 9.81 | 9.96 | 10.11 | 10.26 | 10.4 | 10.55 | 10.7 | 10.85 | 11 | 11.15 | 11.29 | 11.44 | 11.59 | 11.74
10.55 | 5.8 | 5.95 | 6.1 | 6.25 | 6.39 | 6.54 | 6.69 | 6.84 | 6.99 | 7.14 | 7.28 | 7.43 | 7.58 | 7.73 | 7.88 | 8.03 | 8.18 | 8.32 | 8.47 | 8.62 | 8.77 | 8.92 | 9.07 | 9.21 | 9.36 | 9.51 | 9.66 | 9.81 | 9.96 | 10.1 | 10.25 | 10.4 | 10.55 | 10.7 | 10.85 | 10.99 | 11.14 | 11.29 | 11.44 | 11.59 | 11.74
10.56 | 5.8 | 5.95 | 6.1 | 6.24 | 6.39 | 6.54 | 6.69 | 6.84 | 6.99 | 7.13 | 7.28 | 7.43 | 7.58 | 7.73 | 7.88 | 8.02 | 8.17 | 8.32 | 8.47 | 8.62 | 8.77 | 8.92 | 9.06 | 9.21 | 9.36 | 9.51 | 9.66 | 9.81 | 9.95 | 10.1 | 10.25 | 10.4 | 10.55 | 10.7 | 10.84 | 10.99 | 11.14 | 11.29 | 11.44 | 11.59 | 11.73
10.57 | 5.8 | 5.95 | 6.09 | 6.24 | 6.39 | 6.54 | 6.69 | 6.84 | 6.98 | 7.13 | 7.28 | 7.43 | 7.58 | 7.73 | 7.87 | 8.02 | 8.17 | 8.32 | 8.47 | 8.62 | 8.76 | 8.91 | 9.06 | 9.21 | 9.36 | 9.51 | 9.65 | 9.8 | 9.95 | 10.1 | 10.25 | 10.4 | 10.54 | 10.69 | 10.84 | 10.99 | 11.14 | 11.29 | 11.43 | 11.58 | 11.73
10.58 | 5.8 | 5.94 | 6.09 | 6.24 | 6.39 | 6.54 | 6.69 | 6.83 | 6.98 | 7.13 | 7.28 | 7.43 | 7.58 | 7.72 | 7.87 | 8.02 | 8.17 | 8.32 | 8.47 | 8.61 | 8.76 | 8.91 | 9.06 | 9.21 | 9.36 | 9.5 | 9.65 | 9.8 | 9.95 | 10.1 | 10.25 | 10.39 | 10.54 | 10.69 | 10.84 | 10.99 | 11.14 | 11.28 | 11.43 | 11.58 | 11.73
10.59 | 5.79 | 5.94 | 6.09 | 6.24 | 6.39 | 6.54 | 6.68 | 6.83 | 6.98 | 7.13 | 7.28 | 7.43 | 7.57 | 7.72 | 7.87 | 8.02 | 8.17 | 8.32 | 8.46 | 8.61 | 8.76 | 8.91 | 9.06 | 9.21 | 9.35 | 9.5 | 9.65 | 9.8 | 9.95 | 10.1 | 10.24 | 10.39 | 10.54 | 10.69 | 10.84 | 10.98 | 11.13 | 11.28 | 11.43 | 11.58 | 11.73
10.6 | 5.79 | 5.94 | 6.09 | 6.24 | 6.39 | 6.53 | 6.68 | 6.83 | 6.98 | 7.13 | 7.28 | 7.42 | 7.57 | 7.72 | 7.87 | 8.02 | 8.17 | 8.31 | 8.46 | 8.61 | 8.76 | 8.91 | 9.06 | 9.2 | 9.35 | 9.5 | 9.65 | 9.8 | 9.94 | 10.09 | 10.24 | 10.39 | 10.54 | 10.69 | 10.83 | 10.98 | 11.13 | 11.28 | 11.43 | 11.57 | 11.72
10.61 | 5.79 | 5.94 | 6.09 | 6.24 | 6.38 | 6.53 | 6.68 | 6.83 | 6.98 | 7.13 | 7.27 | 7.42 | 7.57 | 7.72 | 7.87 | 8.02 | 8.16 | 8.31 | 8.46 | 8.61 | 8.76 | 8.9 | 9.05 | 9.2 | 9.35 | 9.5 | 9.65 | 9.79 | 9.94 | 10.09 | 10.24 | 10.39 | 10.53 | 10.68 | 10.83 | 10.98 | 11.13 | 11.28 | 11.42 | 11.57 | 11.72
10.62 | 5.79 | 5.94 | 6.09 | 6.23 | 6.38 | 6.53 | 6.68 | 6.83 | 6.98 | 7.12 | 7.27 | 7.42 | 7.57 | 7.72 | 7.87 | 8.01 | 8.16 | 8.31 | 8.46 | 8.61 | 8.75 | 8.9 | 9.05 | 9.2 | 9.35 | 9.5 | 9.64 | 9.79 | 9.94 | 10.09 | 10.24 | 10.38 | 10.53 | 10.68 | 10.83 | 10.98 | 11.13 | 11.27 | 11.42 | 11.57 | 11.72
10.63 | 5.79 | 5.94 | 6.09 | 6.23 | 6.38 | 6.53 | 6.68 | 6.83 | 6.97 | 7.12 | 7.27 | 7.42 | 7.57 | 7.72 | 7.86 | 8.01 | 8.16 | 8.31 | 8.46 | 8.6 | 8.75 | 8.9 | 9.05 | 9.2 | 9.35 | 9.49 | 9.64 | 9.79 | 9.94 | 10.09 | 10.23 | 10.38 | 10.53 | 10.68 | 10.83 | 10.97 | 11.12 | 11.27 | 11.42 | 11.57 | 11.71
10.64 | 5.79 | 5.94 | 6.08 | 6.23 | 6.38 | 6.53 | 6.68 | 6.82 | 6.97 | 7.12 | 7.27 | 7.42 | 7.57 | 7.71 | 7.86 | 8.01 | 8.16 | 8.31 | 8.45 | 8.6 | 8.75 | 8.9 | 9.05 | 9.19 | 9.34 | 9.49 | 9.64 | 9.79 | 9.94 | 10.08 | 10.23 | 10.38 | 10.53 | 10.68 | 10.82 | 10.97 | 11.12 | 11.27 | 11.42 | 11.56 | 11.71
10.65 | 5.79 | 5.93 | 6.08 | 6.23 | 6.38 | 6.53 | 6.67 | 6.82 | 6.97 | 7.12 | 7.27 | 7.42 | 7.56 | 7.71 | 7.86 | 8.01 | 8.16 | 8.3 | 8.45 | 8.6 | 8.75 | 8.9 | 9.04 | 9.19 | 9.34 | 9.49 | 9.64 | 9.78 | 9.93 | 10.08 | 10.23 | 10.38 | 10.53 | 10.67 | 10.82 | 10.97 | 11.12 | 11.27 | 11.41 | 11.56 | 11.71
10.66 | 5.78 | 5.93 | 6.08 | 6.23 | 6.38 | 6.53 | 6.67 | 6.82 | 6.97 | 7.12 | 7.27 | 7.41 | 7.56 | 7.71 | 7.86 | 8.01 | 8.15 | 8.3 | 8.45 | 8.6 | 8.75 | 8.89 | 9.04 | 9.19 | 9.34 | 9.49 | 9.63 | 9.78 | 9.93 | 10.08 | 10.23 | 10.37 | 10.52 | 10.67 | 10.82 | 10.97 | 11.11 | 11.26 | 11.41 | 11.56 | 11.71
10.67 | 5.78 | 5.93 | 6.08 | 6.23 | 6.38 | 6.52 | 6.67 | 6.82 | 6.97 | 7.12 | 7.26 | 7.41 | 7.56 | 7.71 | 7.86 | 8 | 8.15 | 8.3 | 8.45 | 8.6 | 8.74 | 8.89 | 9.04 | 9.19 | 9.34 | 9.48 | 9.63 | 9.78 | 9.93 | 10.08 | 10.22 | 10.37 | 10.52 | 10.67 | 10.82 | 10.96 | 11.11 | 11.26 | 11.41 | 11.56 | 11.7
10.68 | 5.78 | 5.93 | 6.08 | 6.23 | 6.37 | 6.52 | 6.67 | 6.82 | 6.97 | 7.11 | 7.26 | 7.41 | 7.56 | 7.71 | 7.85 | 8 | 8.15 | 8.3 | 8.45 | 8.59 | 8.74 | 8.89 | 9.04 | 9.19 | 9.33 | 9.48 | 9.63 | 9.78 | 9.93 | 10.07 | 10.22 | 10.37 | 10.52 | 10.67 | 10.81 | 10.96 | 11.11 | 11.26 | 11.41 | 11.55 | 11.7
10.69 | 5.78 | 5.93 | 6.08 | 6.22 | 6.37 | 6.52 | 6.67 | 6.82 | 6.96 | 7.11 | 7.26 | 7.41 | 7.56 | 7.7 | 7.85 | 8 | 8.15 | 8.3 | 8.44 | 8.59 | 8.74 | 8.89 | 9.04 | 9.18 | 9.33 | 9.48 | 9.63 | 9.78 | 9.92 | 10.07 | 10.22 | 10.37 | 10.52 | 10.66 | 10.81 | 10.96 | 11.11 | 11.25 | 11.4 | 11.55 | 11.7
10.7 | 5.78 | 5.93 | 6.07 | 6.22 | 6.37 | 6.52 | 6.67 | 6.81 | 6.96 | 7.11 | 7.26 | 7.41 | 7.55 | 7.7 | 7.85 | 8 | 8.15 | 8.29 | 8.44 | 8.59 | 8.74 | 8.89 | 9.03 | 9.18 | 9.33 | 9.48 | 9.63 | 9.77 | 9.92 | 10.07 | 10.22 | 10.36 | 10.51 | 10.66 | 10.81 | 10.96 | 11.1 | 11.25 | 11.4 | 11.55 | 11.7
10.71 | 5.78 | 5.93 | 6.07 | 6.22 | 6.37 | 6.52 | 6.67 | 6.81 | 6.96 | 7.11 | 7.26 | 7.4 | 7.55 | 7.7 | 7.85 | 8 | 8.14 | 8.29 | 8.44 | 8.59 | 8.74 | 8.88 | 9.03 | 9.18 | 9.33 | 9.48 | 9.62 | 9.77 | 9.92 | 10.07 | 10.21 | 10.36 | 10.51 | 10.66 | 10.81 | 10.95 | 11.1 | 11.25 | 11.4 | 11.55 | 11.69
10.72 | 5.78 | 5.92 | 6.07 | 6.22 | 6.37 | 6.52 | 6.66 | 6.81 | 6.96 | 7.11 | 7.26 | 7.4 | 7.55 | 7.7 | 7.85 | 7.99 | 8.14 | 8.29 | 8.44 | 8.59 | 8.73 | 8.88 | 9.03 | 9.18 | 9.33 | 9.47 | 9.62 | 9.77 | 9.92 | 10.06 | 10.21 | 10.36 | 10.51 | 10.66 | 10.8 | 10.95 | 11.1 | 11.25 | 11.39 | 11.54 | 11.69
10.73 | 5.77 | 5.92 | 6.07 | 6.22 | 6.37 | 6.51 | 6.66 | 6.81 | 6.96 | 7.11 | 7.25 | 7.4 | 7.55 | 7.7 | 7.84 | 7.99 | 8.14 | 8.29 | 8.44 | 8.58 | 8.73 | 8.88 | 9.03 | 9.18 | 9.32 | 9.47 | 9.62 | 9.77 | 9.91 | 10.06 | 10.21 | 10.36 | 10.51 | 10.65 | 10.8 | 10.95 | 11.1 | 11.24 | 11.39 | 11.54 | 11.69
10.74 | 5.77 | 5.92 | 6.07 | 6.22 | 6.36 | 6.51 | 6.66 | 6.81 | 6.96 | 7.1 | 7.25 | 7.4 | 7.55 | 7.7 | 7.84 | 7.99 | 8.14 | 8.29 | 8.43 | 8.58 | 8.73 | 8.88 | 9.03 | 9.17 | 9.32 | 9.47 | 9.62 | 9.76 | 9.91 | 10.06 | 10.21 | 10.36 | 10.5 | 10.65 | 10.8 | 10.95 | 11.09 | 11.24 | 11.39 | 11.54 | 11.68
10.75 | 5.77 | 5.92 | 6.07 | 6.22 | 6.36 | 6.51 | 6.66 | 6.81 | 6.95 | 7.1 | 7.25 | 7.4 | 7.55 | 7.69 | 7.84 | 7.99 | 8.14 | 8.28 | 8.43 | 8.58 | 8.73 | 8.88 | 9.02 | 9.17 | 9.32 | 9.47 | 9.61 | 9.76 | 9.91 | 10.06 | 10.21 | 10.35 | 10.5 | 10.65 | 10.8 | 10.94 | 11.09 | 11.24 | 11.39 | 11.53 | 11.68
10.76 | 5.77 | 5.92 | 6.07 | 6.21 | 6.36 | 6.51 | 6.66 | 6.81 | 6.95 | 7.1 | 7.25 | 7.4 | 7.54 | 7.69 | 7.84 | 7.99 | 8.13 | 8.28 | 8.43 | 8.58 | 8.73 | 8.87 | 9.02 | 9.17 | 9.32 | 9.46 | 9.61 | 9.76 | 9.91 | 10.06 | 10.2 | 10.35 | 10.5 | 10.65 | 10.79 | 10.94 | 11.09 | 11.24 | 11.38 | 11.53 | 11.68
10.77 | 5.77 | 5.92 | 6.06 | 6.21 | 6.36 | 6.51 | 6.66 | 6.8 | 6.95 | 7.1 | 7.25 | 7.39 | 7.54 | 7.69 | 7.84 | 7.99 | 8.13 | 8.28 | 8.43 | 8.58 | 8.72 | 8.87 | 9.02 | 9.17 | 9.31 | 9.46 | 9.61 | 9.76 | 9.91 | 10.05 | 10.2 | 10.35 | 10.5 | 10.64 | 10.79 | 10.94 | 11.09 | 11.23 | 11.38 | 11.53 | 11.68
10.78 | 5.77 | 5.92 | 6.06 | 6.21 | 6.36 | 6.51 | 6.65 | 6.8 | 6.95 | 7.1 | 7.24 | 7.39 | 7.54 | 7.69 | 7.84 | 7.98 | 8.13 | 8.28 | 8.43 | 8.57 | 8.72 | 8.87 | 9.02 | 9.16 | 9.31 | 9.46 | 9.61 | 9.76 | 9.9 | 10.05 | 10.2 | 10.35 | 10.49 | 10.64 | 10.79 | 10.94 | 11.08 | 11.23 | 11.38 | 11.53 | 11.67
10.79 | 5.77 | 5.91 | 6.06 | 6.21 | 6.36 | 6.5 | 6.65 | 6.8 | 6.95 | 7.1 | 7.24 | 7.39 | 7.54 | 7.69 | 7.83 | 7.98 | 8.13 | 8.28 | 8.42 | 8.57 | 8.72 | 8.87 | 9.01 | 9.16 | 9.31 | 9.46 | 9.61 | 9.75 | 9.9 | 10.05 | 10.2 | 10.34 | 10.49 | 10.64 | 10.79 | 10.93 | 11.08 | 11.23 | 11.38 | 11.52 | 11.67
10.8 | 5.77 | 5.91 | 6.06 | 6.21 | 6.36 | 6.5 | 6.65 | 6.8 | 6.95 | 7.09 | 7.24 | 7.39 | 7.54 | 7.68 | 7.83 | 7.98 | 8.13 | 8.27 | 8.42 | 8.57 | 8.72 | 8.87 | 9.01 | 9.16 | 9.31 | 9.46 | 9.6 | 9.75 | 9.9 | 10.05 | 10.19 | 10.34 | 10.49 | 10.64 | 10.78 | 10.93 | 11.08 | 11.23 | 11.37 | 11.52 | 11.67
10.81 | 5.76 | 5.91 | 6.06 | 6.21 | 6.35 | 6.5 | 6.65 | 6.8 | 6.94 | 7.09 | 7.24 | 7.39 | 7.53 | 7.68 | 7.83 | 7.98 | 8.13 | 8.27 | 8.42 | 8.57 | 8.72 | 8.86 | 9.01 | 9.16 | 9.31 | 9.45 | 9.6 | 9.75 | 9.9 | 10.04 | 10.19 | 10.34 | 10.49 | 10.63 | 10.78 | 10.93 | 11.08 | 11.22 | 11.37 | 11.52 | 11.67
10.82 | 5.76 | 5.91 | 6.06 | 6.21 | 6.35 | 6.5 | 6.65 | 6.8 | 6.94 | 7.09 | 7.24 | 7.39 | 7.53 | 7.68 | 7.83 | 7.98 | 8.12 | 8.27 | 8.42 | 8.57 | 8.71 | 8.86 | 9.01 | 9.16 | 9.3 | 9.45 | 9.6 | 9.75 | 9.89 | 10.04 | 10.19 | 10.34 | 10.48 | 10.63 | 10.78 | 10.93 | 11.07 | 11.22 | 11.37 | 11.52 | 11.66
10.83 | 5.76 | 5.91 | 6.06 | 6.2 | 6.35 | 6.5 | 6.65 | 6.79 | 6.94 | 7.09 | 7.24 | 7.38 | 7.53 | 7.68 | 7.83 | 7.97 | 8.12 | 8.27 | 8.42 | 8.56 | 8.71 | 8.86 | 9.01 | 9.15 | 9.3 | 9.45 | 9.6 | 9.74 | 9.89 | 10.04 | 10.19 | 10.33 | 10.48 | 10.63 | 10.78 | 10.92 | 11.07 | 11.22 | 11.37 | 11.51 | 11.66
10.84 | 5.76 | 5.91 | 6.05 | 6.2 | 6.35 | 6.5 | 6.64 | 6.79 | 6.94 | 7.09 | 7.23 | 7.38 | 7.53 | 7.68 | 7.82 | 7.97 | 8.12 | 8.27 | 8.41 | 8.56 | 8.71 | 8.86 | 9 | 9.15 | 9.3 | 9.45 | 9.59 | 9.74 | 9.89 | 10.04 | 10.18 | 10.33 | 10.48 | 10.63 | 10.77 | 10.92 | 11.07 | 11.22 | 11.36 | 11.51 | 11.66
10.85 | 5.76 | 5.91 | 6.05 | 6.2 | 6.35 | 6.5 | 6.64 | 6.79 | 6.94 | 7.09 | 7.23 | 7.38 | 7.53 | 7.68 | 7.82 | 7.97 | 8.12 | 8.26 | 8.41 | 8.56 | 8.71 | 8.85 | 9 | 9.15 | 9.3 | 9.44 | 9.59 | 9.74 | 9.89 | 10.03 | 10.18 | 10.33 | 10.48 | 10.62 | 10.77 | 10.92 | 11.07 | 11.21 | 11.36 | 11.51 | 11.66
10.86 | 5.76 | 5.9 | 6.05 | 6.2 | 6.35 | 6.49 | 6.64 | 6.79 | 6.94 | 7.08 | 7.23 | 7.38 | 7.53 | 7.67 | 7.82 | 7.97 | 8.12 | 8.26 | 8.41 | 8.56 | 8.71 | 8.85 | 9 | 9.15 | 9.29 | 9.44 | 9.59 | 9.74 | 9.88 | 10.03 | 10.18 | 10.33 | 10.47 | 10.62 | 10.77 | 10.92 | 11.06 | 11.21 | 11.36 | 11.51 | 11.65
10.87 | 5.76 | 5.9 | 6.05 | 6.2 | 6.35 | 6.49 | 6.64 | 6.79 | 6.93 | 7.08 | 7.23 | 7.38 | 7.52 | 7.67 | 7.82 | 7.97 | 8.11 | 8.26 | 8.41 | 8.56 | 8.7 | 8.85 | 9 | 9.15 | 9.29 | 9.44 | 9.59 | 9.73 | 9.88 | 10.03 | 10.18 | 10.32 | 10.47 | 10.62 | 10.77 | 10.91 | 11.06 | 11.21 | 11.36 | 11.5 | 11.65
10.88 | 5.75 | 5.9 | 6.05 | 6.2 | 6.34 | 6.49 | 6.64 | 6.79 | 6.93 | 7.08 | 7.23 | 7.38 | 7.52 | 7.67 | 7.82 | 7.96 | 8.11 | 8.26 | 8.41 | 8.55 | 8.7 | 8.85 | 9 | 9.14 | 9.29 | 9.44 | 9.59 | 9.73 | 9.88 | 10.03 | 10.17 | 10.32 | 10.47 | 10.62 | 10.76 | 10.91 | 11.06 | 11.21 | 11.35 | 11.5 | 11.65
10.89 | 5.75 | 5.9 | 6.05 | 6.19 | 6.34 | 6.49 | 6.64 | 6.78 | 6.93 | 7.08 | 7.23 | 7.37 | 7.52 | 7.67 | 7.82 | 7.96 | 8.11 | 8.26 | 8.4 | 8.55 | 8.7 | 8.85 | 8.99 | 9.14 | 9.29 | 9.44 | 9.58 | 9.73 | 9.88 | 10.02 | 10.17 | 10.32 | 10.47 | 10.61 | 10.76 | 10.91 | 11.06 | 11.2 | 11.35 | 11.5 | 11.64
10.9 | 5.75 | 5.9 | 6.05 | 6.19 | 6.34 | 6.49 | 6.64 | 6.78 | 6.93 | 7.08 | 7.22 | 7.37 | 7.52 | 7.67 | 7.81 | 7.96 | 8.11 | 8.26 | 8.4 | 8.55 | 8.7 | 8.84 | 8.99 | 9.14 | 9.29 | 9.43 | 9.58 | 9.73 | 9.88 | 10.02 | 10.17 | 10.32 | 10.46 | 10.61 | 10.76 | 10.91 | 11.05 | 11.2 | 11.35 | 11.49 | 11.64
10.91 | 5.75 | 5.9 | 6.04 | 6.19 | 6.34 | 6.49 | 6.63 | 6.78 | 6.93 | 7.08 | 7.22 | 7.37 | 7.52 | 7.66 | 7.81 | 7.96 | 8.11 | 8.25 | 8.4 | 8.55 | 8.7 | 8.84 | 8.99 | 9.14 | 9.28 | 9.43 | 9.58 | 9.73 | 9.87 | 10.02 | 10.17 | 10.31 | 10.46 | 10.61 | 10.76 | 10.9 | 11.05 | 11.2 | 11.34 | 11.49 | 11.64
10.92 | 5.75 | 5.9 | 6.04 | 6.19 | 6.34 | 6.48 | 6.63 | 6.78 | 6.93 | 7.07 | 7.22 | 7.37 | 7.52 | 7.66 | 7.81 | 7.96 | 8.1 | 8.25 | 8.4 | 8.55 | 8.69 | 8.84 | 8.99 | 9.13 | 9.28 | 9.43 | 9.58 | 9.72 | 9.87 | 10.02 | 10.16 | 10.31 | 10.46 | 10.61 | 10.75 | 10.9 | 11.05 | 11.19 | 11.34 | 11.49 | 11.64
10.93 | 5.75 | 5.89 | 6.04 | 6.19 | 6.34 | 6.48 | 6.63 | 6.78 | 6.92 | 7.07 | 7.22 | 7.37 | 7.51 | 7.66 | 7.81 | 7.96 | 8.1 | 8.25 | 8.4 | 8.54 | 8.69 | 8.84 | 8.99 | 9.13 | 9.28 | 9.43 | 9.57 | 9.72 | 9.87 | 10.02 | 10.16 | 10.31 | 10.46 | 10.6 | 10.75 | 10.9 | 11.05 | 11.19 | 11.34 | 11.49 | 11.63
10.94 | 5.75 | 5.89 | 6.04 | 6.19 | 6.33 | 6.48 | 6.63 | 6.78 | 6.92 | 7.07 | 7.22 | 7.36 | 7.51 | 7.66 | 7.81 | 7.95 | 8.1 | 8.25 | 8.39 | 8.54 | 8.69 | 8.84 | 8.98 | 9.13 | 9.28 | 9.42 | 9.57 | 9.72 | 9.87 | 10.01 | 10.16 | 10.31 | 10.45 | 10.6 | 10.75 | 10.9 | 11.04 | 11.19 | 11.34 | 11.48 | 11.63
10.95 | 5.74 | 5.89 | 6.04 | 6.19 | 6.33 | 6.48 | 6.63 | 6.77 | 6.92 | 7.07 | 7.22 | 7.36 | 7.51 | 7.66 | 7.8 | 7.95 | 8.1 | 8.25 | 8.39 | 8.54 | 8.69 | 8.83 | 8.98 | 9.13 | 9.28 | 9.42 | 9.57 | 9.72 | 9.86 | 10.01 | 10.16 | 10.3 | 10.45 | 10.6 | 10.75 | 10.89 | 11.04 | 11.19 | 11.33 | 11.48 | 11.63
10.96 | 5.74 | 5.89 | 6.04 | 6.18 | 6.33 | 6.48 | 6.63 | 6.77 | 6.92 | 7.07 | 7.21 | 7.36 | 7.51 | 7.66 | 7.8 | 7.95 | 8.1 | 8.24 | 8.39 | 8.54 | 8.68 | 8.83 | 8.98 | 9.13 | 9.27 | 9.42 | 9.57 | 9.71 | 9.86 | 10.01 | 10.16 | 10.3 | 10.45 | 10.6 | 10.74 | 10.89 | 11.04 | 11.18 | 11.33 | 11.48 | 11.63
10.97 | 5.74 | 5.89 | 6.04 | 6.18 | 6.33 | 6.48 | 6.62 | 6.77 | 6.92 | 7.07 | 7.21 | 7.36 | 7.51 | 7.65 | 7.8 | 7.95 | 8.09 | 8.24 | 8.39 | 8.54 | 8.68 | 8.83 | 8.98 | 9.12 | 9.27 | 9.42 | 9.56 | 9.71 | 9.86 | 10.01 | 10.15 | 10.3 | 10.45 | 10.59 | 10.74 | 10.89 | 11.03 | 11.18 | 11.33 | 11.48 | 11.62
10.98 | 5.74 | 5.89 | 6.03 | 6.18 | 6.33 | 6.48 | 6.62 | 6.77 | 6.92 | 7.06 | 7.21 | 7.36 | 7.5 | 7.65 | 7.8 | 7.95 | 8.09 | 8.24 | 8.39 | 8.53 | 8.68 | 8.83 | 8.97 | 9.12 | 9.27 | 9.42 | 9.56 | 9.71 | 9.86 | 10 | 10.15 | 10.3 | 10.44 | 10.59 | 10.74 | 10.89 | 11.03 | 11.18 | 11.33 | 11.47 | 11.62
10.99 | 5.74 | 5.89 | 6.03 | 6.18 | 6.33 | 6.47 | 6.62 | 6.77 | 6.91 | 7.06 | 7.21 | 7.36 | 7.5 | 7.65 | 7.8 | 7.94 | 8.09 | 8.24 | 8.38 | 8.53 | 8.68 | 8.83 | 8.97 | 9.12 | 9.27 | 9.41 | 9.56 | 9.71 | 9.85 | 10 | 10.15 | 10.3 | 10.44 | 10.59 | 10.74 | 10.88 | 11.03 | 11.18 | 11.32 | 11.47 | 11.62
11 | 5.74 | 5.88 | 6.03 | 6.18 | 6.33 | 6.47 | 6.62 | 6.77 | 6.91 | 7.06 | 7.21 | 7.35 | 7.5 | 7.65 | 7.8 | 7.94 | 8.09 | 8.24 | 8.38 | 8.53 | 8.68 | 8.82 | 8.97 | 9.12 | 9.26 | 9.41 | 9.56 | 9.71 | 9.85 | 10 | 10.15 | 10.29 | 10.44 | 10.59 | 10.73 | 10.88 | 11.03 | 11.17 | 11.32 | 11.47 | 11.61
11.01 | 5.74 | 5.88 | 6.03 | 6.18 | 6.32 | 6.47 | 6.62 | 6.76 | 6.91 | 7.06 | 7.21 | 7.35 | 7.5 | 7.65 | 7.79 | 7.94 | 8.09 | 8.23 | 8.38 | 8.53 | 8.67 | 8.82 | 8.97 | 9.12 | 9.26 | 9.41 | 9.56 | 9.7 | 9.85 | 10 | 10.14 | 10.29 | 10.44 | 10.58 | 10.73 | 10.88 | 11.02 | 11.17 | 11.32 | 11.47 | 11.61
11.02 | 5.73 | 5.88 | 6.03 | 6.18 | 6.32 | 6.47 | 6.62 | 6.76 | 6.91 | 7.06 | 7.2 | 7.35 | 7.5 | 7.64 | 7.79 | 7.94 | 8.09 | 8.23 | 8.38 | 8.53 | 8.67 | 8.82 | 8.97 | 9.11 | 9.26 | 9.41 | 9.55 | 9.7 | 9.85 | 9.99 | 10.14 | 10.29 | 10.43 | 10.58 | 10.73 | 10.88 | 11.02 | 11.17 | 11.32 | 11.46 | 11.61
11.03 | 5.73 | 5.88 | 6.03 | 6.17 | 6.32 | 6.47 | 6.61 | 6.76 | 6.91 | 7.06 | 7.2 | 7.35 | 7.5 | 7.64 | 7.79 | 7.94 | 8.08 | 8.23 | 8.38 | 8.52 | 8.67 | 8.82 | 8.96 | 9.11 | 9.26 | 9.4 | 9.55 | 9.7 | 9.85 | 9.99 | 10.14 | 10.29 | 10.43 | 10.58 | 10.73 | 10.87 | 11.02 | 11.17 | 11.31 | 11.46 | 11.61
11.04 | 5.73 | 5.88 | 6.03 | 6.17 | 6.32 | 6.47 | 6.61 | 6.76 | 6.91 | 7.05 | 7.2 | 7.35 | 7.49 | 7.64 | 7.79 | 7.93 | 8.08 | 8.23 | 8.38 | 8.52 | 8.67 | 8.82 | 8.96 | 9.11 | 9.26 | 9.4 | 9.55 | 9.7 | 9.84 | 9.99 | 10.14 | 10.28 | 10.43 | 10.58 | 10.72 | 10.87 | 11.02 | 11.16 | 11.31 | 11.46 | 11.6
11.05 | 5.73 | 5.88 | 6.02 | 6.17 | 6.32 | 6.46 | 6.61 | 6.76 | 6.91 | 7.05 | 7.2 | 7.35 | 7.49 | 7.64 | 7.79 | 7.93 | 8.08 | 8.23 | 8.37 | 8.52 | 8.67 | 8.81 | 8.96 | 9.11 | 9.25 | 9.4 | 9.55 | 9.69 | 9.84 | 9.99 | 10.13 | 10.28 | 10.43 | 10.57 | 10.72 | 10.87 | 11.01 | 11.16 | 11.31 | 11.45 | 11.6
11.06 | 5.73 | 5.88 | 6.02 | 6.17 | 6.32 | 6.46 | 6.61 | 6.76 | 6.9 | 7.05 | 7.2 | 7.34 | 7.49 | 7.64 | 7.78 | 7.93 | 8.08 | 8.22 | 8.37 | 8.52 | 8.66 | 8.81 | 8.96 | 9.1 | 9.25 | 9.4 | 9.54 | 9.69 | 9.84 | 9.99 | 10.13 | 10.28 | 10.43 | 10.57 | 10.72 | 10.87 | 11.01 | 11.16 | 11.31 | 11.45 | 11.6
11.07 | 5.73 | 5.87 | 6.02 | 6.17 | 6.31 | 6.46 | 6.61 | 6.76 | 6.9 | 7.05 | 7.2 | 7.34 | 7.49 | 7.64 | 7.78 | 7.93 | 8.08 | 8.22 | 8.37 | 8.52 | 8.66 | 8.81 | 8.96 | 9.1 | 9.25 | 9.4 | 9.54 | 9.69 | 9.84 | 9.98 | 10.13 | 10.28 | 10.42 | 10.57 | 10.72 | 10.86 | 11.01 | 11.16 | 11.3 | 11.45 | 11.6
11.08 | 5.73 | 5.87 | 6.02 | 6.17 | 6.31 | 6.46 | 6.61 | 6.75 | 6.9 | 7.05 | 7.19 | 7.34 | 7.49 | 7.63 | 7.78 | 7.93 | 8.07 | 8.22 | 8.37 | 8.51 | 8.66 | 8.81 | 8.95 | 9.1 | 9.25 | 9.39 | 9.54 | 9.69 | 9.83 | 9.98 | 10.13 | 10.27 | 10.42 | 10.57 | 10.71 | 10.86 | 11.01 | 11.15 | 11.3 | 11.45 | 11.59
11.09 | 5.73 | 5.87 | 6.02 | 6.17 | 6.31 | 6.46 | 6.61 | 6.75 | 6.9 | 7.05 | 7.19 | 7.34 | 7.49 | 7.63 | 7.78 | 7.93 | 8.07 | 8.22 | 8.37 | 8.51 | 8.66 | 8.81 | 8.95 | 9.1 | 9.25 | 9.39 | 9.54 | 9.68 | 9.83 | 9.98 | 10.12 | 10.27 | 10.42 | 10.56 | 10.71 | 10.86 | 11 | 11.15 | 11.3 | 11.44 | 11.59
11.1 | 5.72 | 5.87 | 6.02 | 6.16 | 6.31 | 6.46 | 6.6 | 6.75 | 6.9 | 7.04 | 7.19 | 7.34 | 7.48 | 7.63 | 7.78 | 7.92 | 8.07 | 8.22 | 8.36 | 8.51 | 8.66 | 8.8 | 8.95 | 9.1 | 9.24 | 9.39 | 9.54 | 9.68 | 9.83 | 9.98 | 10.12 | 10.27 | 10.42 | 10.56 | 10.71 | 10.86 | 11 | 11.15 | 11.29 | 11.44 | 11.59
11.11 | 5.72 | 5.87 | 6.02 | 6.16 | 6.31 | 6.46 | 6.6 | 6.75 | 6.9 | 7.04 | 7.19 | 7.34 | 7.48 | 7.63 | 7.77 | 7.92 | 8.07 | 8.21 | 8.36 | 8.51 | 8.65 | 8.8 | 8.95 | 9.09 | 9.24 | 9.39 | 9.53 | 9.68 | 9.83 | 9.97 | 10.12 | 10.27 | 10.41 | 10.56 | 10.71 | 10.85 | 11 | 11.15 | 11.29 | 11.44 | 11.59
11.12 | 5.72 | 5.87 | 6.01 | 6.16 | 6.31 | 6.45 | 6.6 | 6.75 | 6.89 | 7.04 | 7.19 | 7.33 | 7.48 | 7.63 | 7.77 | 7.92 | 8.07 | 8.21 | 8.36 | 8.51 | 8.65 | 8.8 | 8.95 | 9.09 | 9.24 | 9.39 | 9.53 | 9.68 | 9.82 | 9.97 | 10.12 | 10.26 | 10.41 | 10.56 | 10.7 | 10.85 | 11 | 11.14 | 11.29 | 11.44 | 11.58
11.13 | 5.72 | 5.87 | 6.01 | 6.16 | 6.31 | 6.45 | 6.6 | 6.75 | 6.89 | 7.04 | 7.19 | 7.33 | 7.48 | 7.62 | 7.77 | 7.92 | 8.06 | 8.21 | 8.36 | 8.5 | 8.65 | 8.8 | 8.94 | 9.09 | 9.24 | 9.38 | 9.53 | 9.68 | 9.82 | 9.97 | 10.12 | 10.26 | 10.41 | 10.55 | 10.7 | 10.85 | 10.99 | 11.14 | 11.29 | 11.43 | 11.58
11.14 | 5.72 | 5.86 | 6.01 | 6.16 | 6.3 | 6.45 | 6.6 | 6.74 | 6.89 | 7.04 | 7.18 | 7.33 | 7.48 | 7.62 | 7.77 | 7.92 | 8.06 | 8.21 | 8.36 | 8.5 | 8.65 | 8.79 | 8.94 | 9.09 | 9.23 | 9.38 | 9.53 | 9.67 | 9.82 | 9.97 | 10.11 | 10.26 | 10.41 | 10.55 | 10.7 | 10.85 | 10.99 | 11.14 | 11.28 | 11.43 | 11.58
11.15 | 5.72 | 5.86 | 6.01 | 6.16 | 6.3 | 6.45 | 6.6 | 6.74 | 6.89 | 7.04 | 7.18 | 7.33 | 7.47 | 7.62 | 7.77 | 7.91 | 8.06 | 8.21 | 8.35 | 8.5 | 8.65 | 8.79 | 8.94 | 9.09 | 9.23 | 9.38 | 9.52 | 9.67 | 9.82 | 9.96 | 10.11 | 10.26 | 10.4 | 10.55 | 10.7 | 10.84 | 10.99 | 11.14 | 11.28 | 11.43 | 11.57
11.16 | 5.72 | 5.86 | 6.01 | 6.15 | 6.3 | 6.45 | 6.59 | 6.74 | 6.89 | 7.03 | 7.18 | 7.33 | 7.47 | 7.62 | 7.77 | 7.91 | 8.06 | 8.21 | 8.35 | 8.5 | 8.64 | 8.79 | 8.94 | 9.08 | 9.23 | 9.38 | 9.52 | 9.67 | 9.82 | 9.96 | 10.11 | 10.25 | 10.4 | 10.55 | 10.69 | 10.84 | 10.99 | 11.13 | 11.28 | 11.43 | 11.57
11.17 | 5.71 | 5.86 | 6.01 | 6.15 | 6.3 | 6.45 | 6.59 | 6.74 | 6.89 | 7.03 | 7.18 | 7.32 | 7.47 | 7.62 | 7.76 | 7.91 | 8.06 | 8.2 | 8.35 | 8.5 | 8.64 | 8.79 | 8.94 | 9.08 | 9.23 | 9.37 | 9.52 | 9.67 | 9.81 | 9.96 | 10.11 | 10.25 | 10.4 | 10.54 | 10.69 | 10.84 | 10.98 | 11.13 | 11.28 | 11.42 | 11.57
11.18 | 5.71 | 5.86 | 6.01 | 6.15 | 6.3 | 6.44 | 6.59 | 6.74 | 6.88 | 7.03 | 7.18 | 7.32 | 7.47 | 7.62 | 7.76 | 7.91 | 8.05 | 8.2 | 8.35 | 8.49 | 8.64 | 8.79 | 8.93 | 9.08 | 9.23 | 9.37 | 9.52 | 9.66 | 9.81 | 9.96 | 10.1 | 10.25 | 10.4 | 10.54 | 10.69 | 10.84 | 10.98 | 11.13 | 11.27 | 11.42 | 11.57
11.19 | 5.71 | 5.86 | 6 | 6.15 | 6.3 | 6.44 | 6.59 | 6.74 | 6.88 | 7.03 | 7.18 | 7.32 | 7.47 | 7.61 | 7.76 | 7.91 | 8.05 | 8.2 | 8.35 | 8.49 | 8.64 | 8.78 | 8.93 | 9.08 | 9.22 | 9.37 | 9.52 | 9.66 | 9.81 | 9.95 | 10.1 | 10.25 | 10.39 | 10.54 | 10.69 | 10.83 | 10.98 | 11.13 | 11.27 | 11.42 | 11.56
11.2 | 5.71 | 5.86 | 6 | 6.15 | 6.3 | 6.44 | 6.59 | 6.73 | 6.88 | 7.03 | 7.17 | 7.32 | 7.47 | 7.61 | 7.76 | 7.9 | 8.05 | 8.2 | 8.34 | 8.49 | 8.64 | 8.78 | 8.93 | 9.08 | 9.22 | 9.37 | 9.51 | 9.66 | 9.81 | 9.95 | 10.1 | 10.25 | 10.39 | 10.54 | 10.68 | 10.83 | 10.98 | 11.12 | 11.27 | 11.42 | 11.56
11.21 | 5.71 | 5.86 | 6 | 6.15 | 6.29 | 6.44 | 6.59 | 6.73 | 6.88 | 7.03 | 7.17 | 7.32 | 7.46 | 7.61 | 7.76 | 7.9 | 8.05 | 8.2 | 8.34 | 8.49 | 8.63 | 8.78 | 8.93 | 9.07 | 9.22 | 9.37 | 9.51 | 9.66 | 9.8 | 9.95 | 10.1 | 10.24 | 10.39 | 10.54 | 10.68 | 10.83 | 10.97 | 11.12 | 11.27 | 11.41 | 11.56
11.22 | 5.71 | 5.85 | 6 | 6.15 | 6.29 | 6.44 | 6.58 | 6.73 | 6.88 | 7.02 | 7.17 | 7.32 | 7.46 | 7.61 | 7.75 | 7.9 | 8.05 | 8.19 | 8.34 | 8.49 | 8.63 | 8.78 | 8.92 | 9.07 | 9.22 | 9.36 | 9.51 | 9.66 | 9.8 | 9.95 | 10.09 | 10.24 | 10.39 | 10.53 | 10.68 | 10.83 | 10.97 | 11.12 | 11.26 | 11.41 | 11.56
11.23 | 5.71 | 5.85 | 6 | 6.14 | 6.29 | 6.44 | 6.58 | 6.73 | 6.88 | 7.02 | 7.17 | 7.31 | 7.46 | 7.61 | 7.75 | 7.9 | 8.05 | 8.19 | 8.34 | 8.48 | 8.63 | 8.78 | 8.92 | 9.07 | 9.21 | 9.36 | 9.51 | 9.65 | 9.8 | 9.95 | 10.09 | 10.24 | 10.38 | 10.53 | 10.68 | 10.82 | 10.97 | 11.11 | 11.26 | 11.41 | 11.55
11.24 | 5.7 | 5.85 | 6 | 6.14 | 6.29 | 6.44 | 6.58 | 6.73 | 6.87 | 7.02 | 7.17 | 7.31 | 7.46 | 7.61 | 7.75 | 7.9 | 8.04 | 8.19 | 8.34 | 8.48 | 8.63 | 8.77 | 8.92 | 9.07 | 9.21 | 9.36 | 9.51 | 9.65 | 9.8 | 9.94 | 10.09 | 10.24 | 10.38 | 10.53 | 10.67 | 10.82 | 10.97 | 11.11 | 11.26 | 11.4 | 11.55
11.25 | 5.7 | 5.85 | 6 | 6.14 | 6.29 | 6.43 | 6.58 | 6.73 | 6.87 | 7.02 | 7.16 | 7.31 | 7.46 | 7.6 | 7.75 | 7.9 | 8.04 | 8.19 | 8.33 | 8.48 | 8.63 | 8.77 | 8.92 | 9.06 | 9.21 | 9.36 | 9.5 | 9.65 | 9.8 | 9.94 | 10.09 | 10.23 | 10.38 | 10.53 | 10.67 | 10.82 | 10.96 | 11.11 | 11.26 | 11.4 | 11.55
11.26 | 5.7 | 5.85 | 5.99 | 6.14 | 6.29 | 6.43 | 6.58 | 6.72 | 6.87 | 7.02 | 7.16 | 7.31 | 7.46 | 7.6 | 7.75 | 7.89 | 8.04 | 8.19 | 8.33 | 8.48 | 8.62 | 8.77 | 8.92 | 9.06 | 9.21 | 9.35 | 9.5 | 9.65 | 9.79 | 9.94 | 10.08 | 10.23 | 10.38 | 10.52 | 10.67 | 10.82 | 10.96 | 11.11 | 11.25 | 11.4 | 11.55
11.27 | 5.7 | 5.85 | 5.99 | 6.14 | 6.28 | 6.43 | 6.58 | 6.72 | 6.87 | 7.02 | 7.16 | 7.31 | 7.45 | 7.6 | 7.75 | 7.89 | 8.04 | 8.18 | 8.33 | 8.48 | 8.62 | 8.77 | 8.91 | 9.06 | 9.21 | 9.35 | 9.5 | 9.64 | 9.79 | 9.94 | 10.08 | 10.23 | 10.37 | 10.52 | 10.67 | 10.81 | 10.96 | 11.1 | 11.25 | 11.4 | 11.54
11.28 | 5.7 | 5.85 | 5.99 | 6.14 | 6.28 | 6.43 | 6.58 | 6.72 | 6.87 | 7.01 | 7.16 | 7.31 | 7.45 | 7.6 | 7.74 | 7.89 | 8.04 | 8.18 | 8.33 | 8.47 | 8.62 | 8.77 | 8.91 | 9.06 | 9.2 | 9.35 | 9.5 | 9.64 | 9.79 | 9.93 | 10.08 | 10.23 | 10.37 | 10.52 | 10.66 | 10.81 | 10.96 | 11.1 | 11.25 | 11.39 | 11.54
11.29 | 5.7 | 5.84 | 5.99 | 6.14 | 6.28 | 6.43 | 6.57 | 6.72 | 6.87 | 7.01 | 7.16 | 7.3 | 7.45 | 7.6 | 7.74 | 7.89 | 8.03 | 8.18 | 8.33 | 8.47 | 8.62 | 8.76 | 8.91 | 9.06 | 9.2 | 9.35 | 9.49 | 9.64 | 9.79 | 9.93 | 10.08 | 10.22 | 10.37 | 10.52 | 10.66 | 10.81 | 10.95 | 11.1 | 11.25 | 11.39 | 11.54
11.3 | 5.7 | 5.84 | 5.99 | 6.13 | 6.28 | 6.43 | 6.57 | 6.72 | 6.86 | 7.01 | 7.16 | 7.3 | 7.45 | 7.59 | 7.74 | 7.89 | 8.03 | 8.18 | 8.32 | 8.47 | 8.62 | 8.76 | 8.91 | 9.05 | 9.2 | 9.35 | 9.49 | 9.64 | 9.78 | 9.93 | 10.08 | 10.22 | 10.37 | 10.51 | 10.66 | 10.81 | 10.95 | 11.1 | 11.24 | 11.39 | 11.53
11.31 | 5.7 | 5.84 | 5.99 | 6.13 | 6.28 | 6.43 | 6.57 | 6.72 | 6.86 | 7.01 | 7.15 | 7.3 | 7.45 | 7.59 | 7.74 | 7.88 | 8.03 | 8.18 | 8.32 | 8.47 | 8.61 | 8.76 | 8.91 | 9.05 | 9.2 | 9.34 | 9.49 | 9.64 | 9.78 | 9.93 | 10.07 | 10.22 | 10.37 | 10.51 | 10.66 | 10.8 | 10.95 | 11.09 | 11.24 | 11.39 | 11.53
11.32 | 5.69 | 5.84 | 5.99 | 6.13 | 6.28 | 6.42 | 6.57 | 6.72 | 6.86 | 7.01 | 7.15 | 7.3 | 7.44 | 7.59 | 7.74 | 7.88 | 8.03 | 8.17 | 8.32 | 8.47 | 8.61 | 8.76 | 8.9 | 9.05 | 9.2 | 9.34 | 9.49 | 9.63 | 9.78 | 9.93 | 10.07 | 10.22 | 10.36 | 10.51 | 10.65 | 10.8 | 10.95 | 11.09 | 11.24 | 11.38 | 11.53
11.33 | 5.69 | 5.84 | 5.98 | 6.13 | 6.28 | 6.42 | 6.57 | 6.71 | 6.86 | 7.01 | 7.15 | 7.3 | 7.44 | 7.59 | 7.73 | 7.88 | 8.03 | 8.17 | 8.32 | 8.46 | 8.61 | 8.76 | 8.9 | 9.05 | 9.19 | 9.34 | 9.49 | 9.63 | 9.78 | 9.92 | 10.07 | 10.21 | 10.36 | 10.51 | 10.65 | 10.8 | 10.94 | 11.09 | 11.24 | 11.38 | 11.53
11.34 | 5.69 | 5.84 | 5.98 | 6.13 | 6.27 | 6.42 | 6.57 | 6.71 | 6.86 | 7 | 7.15 | 7.3 | 7.44 | 7.59 | 7.73 | 7.88 | 8.02 | 8.17 | 8.32 | 8.46 | 8.61 | 8.75 | 8.9 | 9.05 | 9.19 | 9.34 | 9.48 | 9.63 | 9.77 | 9.92 | 10.07 | 10.21 | 10.36 | 10.5 | 10.65 | 10.8 | 10.94 | 11.09 | 11.23 | 11.38 | 11.52
11.35 | 5.69 | 5.84 | 5.98 | 6.13 | 6.27 | 6.42 | 6.56 | 6.71 | 6.86 | 7 | 7.15 | 7.29 | 7.44 | 7.59 | 7.73 | 7.88 | 8.02 | 8.17 | 8.31 | 8.46 | 8.61 | 8.75 | 8.9 | 9.04 | 9.19 | 9.34 | 9.48 | 9.63 | 9.77 | 9.92 | 10.06 | 10.21 | 10.36 | 10.5 | 10.65 | 10.79 | 10.94 | 11.08 | 11.23 | 11.38 | 11.52
11.36 | 5.69 | 5.83 | 5.98 | 6.13 | 6.27 | 6.42 | 6.56 | 6.71 | 6.85 | 7 | 7.15 | 7.29 | 7.44 | 7.58 | 7.73 | 7.88 | 8.02 | 8.17 | 8.31 | 8.46 | 8.6 | 8.75 | 8.9 | 9.04 | 9.19 | 9.33 | 9.48 | 9.62 | 9.77 | 9.92 | 10.06 | 10.21 | 10.35 | 10.5 | 10.64 | 10.79 | 10.94 | 11.08 | 11.23 | 11.37 | 11.52
11.37 | 5.69 | 5.83 | 5.98 | 6.12 | 6.27 | 6.42 | 6.56 | 6.71 | 6.85 | 7 | 7.14 | 7.29 | 7.44 | 7.58 | 7.73 | 7.87 | 8.02 | 8.16 | 8.31 | 8.46 | 8.6 | 8.75 | 8.89 | 9.04 | 9.19 | 9.33 | 9.48 | 9.62 | 9.77 | 9.91 | 10.06 | 10.2 | 10.35 | 10.5 | 10.64 | 10.79 | 10.93 | 11.08 | 11.22 | 11.37 | 11.52
11.38 | 5.69 | 5.83 | 5.98 | 6.12 | 6.27 | 6.41 | 6.56 | 6.71 | 6.85 | 7 | 7.14 | 7.29 | 7.43 | 7.58 | 7.73 | 7.87 | 8.02 | 8.16 | 8.31 | 8.45 | 8.6 | 8.75 | 8.89 | 9.04 | 9.18 | 9.33 | 9.47 | 9.62 | 9.77 | 9.91 | 10.06 | 10.2 | 10.35 | 10.49 | 10.64 | 10.79 | 10.93 | 11.08 | 11.22 | 11.37 | 11.51
11.39 | 5.68 | 5.83 | 5.98 | 6.12 | 6.27 | 6.41 | 6.56 | 6.7 | 6.85 | 7 | 7.14 | 7.29 | 7.43 | 7.58 | 7.72 | 7.87 | 8.02 | 8.16 | 8.31 | 8.45 | 8.6 | 8.74 | 8.89 | 9.04 | 9.18 | 9.33 | 9.47 | 9.62 | 9.76 | 9.91 | 10.05 | 10.2 | 10.35 | 10.49 | 10.64 | 10.78 | 10.93 | 11.07 | 11.22 | 11.37 | 11.51
11.4 | 5.68 | 5.83 | 5.97 | 6.12 | 6.27 | 6.41 | 6.56 | 6.7 | 6.85 | 6.99 | 7.14 | 7.29 | 7.43 | 7.58 | 7.72 | 7.87 | 8.01 | 8.16 | 8.3 | 8.45 | 8.6 | 8.74 | 8.89 | 9.03 | 9.18 | 9.32 | 9.47 | 9.62 | 9.76 | 9.91 | 10.05 | 10.2 | 10.34 | 10.49 | 10.63 | 10.78 | 10.93 | 11.07 | 11.22 | 11.36 | 11.51
11.41 | 5.68 | 5.83 | 5.97 | 6.12 | 6.26 | 6.41 | 6.56 | 6.7 | 6.85 | 6.99 | 7.14 | 7.28 | 7.43 | 7.57 | 7.72 | 7.87 | 8.01 | 8.16 | 8.3 | 8.45 | 8.59 | 8.74 | 8.89 | 9.03 | 9.18 | 9.32 | 9.47 | 9.61 | 9.76 | 9.9 | 10.05 | 10.2 | 10.34 | 10.49 | 10.63 | 10.78 | 10.92 | 11.07 | 11.21 | 11.36 | 11.51
11.42 | 5.68 | 5.83 | 5.97 | 6.12 | 6.26 | 6.41 | 6.55 | 6.7 | 6.85 | 6.99 | 7.14 | 7.28 | 7.43 | 7.57 | 7.72 | 7.86 | 8.01 | 8.16 | 8.3 | 8.45 | 8.59 | 8.74 | 8.88 | 9.03 | 9.17 | 9.32 | 9.47 | 9.61 | 9.76 | 9.9 | 10.05 | 10.19 | 10.34 | 10.48 | 10.63 | 10.78 | 10.92 | 11.07 | 11.21 | 11.36 | 11.5
11.43 | 5.68 | 5.82 | 5.97 | 6.12 | 6.26 | 6.41 | 6.55 | 6.7 | 6.84 | 6.99 | 7.13 | 7.28 | 7.43 | 7.57 | 7.72 | 7.86 | 8.01 | 8.15 | 8.3 | 8.44 | 8.59 | 8.74 | 8.88 | 9.03 | 9.17 | 9.32 | 9.46 | 9.61 | 9.75 | 9.9 | 10.05 | 10.19 | 10.34 | 10.48 | 10.63 | 10.77 | 10.92 | 11.06 | 11.21 | 11.35 | 11.5
11.44 | 5.68 | 5.82 | 5.97 | 6.11 | 6.26 | 6.41 | 6.55 | 6.7 | 6.84 | 6.99 | 7.13 | 7.28 | 7.42 | 7.57 | 7.72 | 7.86 | 8.01 | 8.15 | 8.3 | 8.44 | 8.59 | 8.73 | 8.88 | 9.02 | 9.17 | 9.32 | 9.46 | 9.61 | 9.75 | 9.9 | 10.04 | 10.19 | 10.33 | 10.48 | 10.62 | 10.77 | 10.92 | 11.06 | 11.21 | 11.35 | 11.5
11.45 | 5.68 | 5.82 | 5.97 | 6.11 | 6.26 | 6.4 | 6.55 | 6.69 | 6.84 | 6.99 | 7.13 | 7.28 | 7.42 | 7.57 | 7.71 | 7.86 | 8 | 8.15 | 8.3 | 8.44 | 8.59 | 8.73 | 8.88 | 9.02 | 9.17 | 9.31 | 9.46 | 9.6 | 9.75 | 9.9 | 10.04 | 10.19 | 10.33 | 10.48 | 10.62 | 10.77 | 10.91 | 11.06 | 11.2 | 11.35 | 11.49
11.46 | 5.67 | 5.82 | 5.97 | 6.11 | 6.26 | 6.4 | 6.55 | 6.69 | 6.84 | 6.98 | 7.13 | 7.28 | 7.42 | 7.57 | 7.71 | 7.86 | 8 | 8.15 | 8.29 | 8.44 | 8.58 | 8.73 | 8.88 | 9.02 | 9.17 | 9.31 | 9.46 | 9.6 | 9.75 | 9.89 | 10.04 | 10.18 | 10.33 | 10.47 | 10.62 | 10.77 | 10.91 | 11.06 | 11.2 | 11.35 | 11.49
11.47 | 5.67 | 5.82 | 5.96 | 6.11 | 6.26 | 6.4 | 6.55 | 6.69 | 6.84 | 6.98 | 7.13 | 7.27 | 7.42 | 7.56 | 7.71 | 7.86 | 8 | 8.15 | 8.29 | 8.44 | 8.58 | 8.73 | 8.87 | 9.02 | 9.16 | 9.31 | 9.45 | 9.6 | 9.75 | 9.89 | 10.04 | 10.18 | 10.33 | 10.47 | 10.62 | 10.76 | 10.91 | 11.05 | 11.2 | 11.34 | 11.49
11.48 | 5.67 | 5.82 | 5.96 | 6.11 | 6.25 | 6.4 | 6.54 | 6.69 | 6.84 | 6.98 | 7.13 | 7.27 | 7.42 | 7.56 | 7.71 | 7.85 | 8 | 8.14 | 8.29 | 8.43 | 8.58 | 8.73 | 8.87 | 9.02 | 9.16 | 9.31 | 9.45 | 9.6 | 9.74 | 9.89 | 10.03 | 10.18 | 10.32 | 10.47 | 10.62 | 10.76 | 10.91 | 11.05 | 11.2 | 11.34 | 11.49
11.49 | 5.67 | 5.82 | 5.96 | 6.11 | 6.25 | 6.4 | 6.54 | 6.69 | 6.83 | 6.98 | 7.12 | 7.27 | 7.42 | 7.56 | 7.71 | 7.85 | 8 | 8.14 | 8.29 | 8.43 | 8.58 | 8.72 | 8.87 | 9.01 | 9.16 | 9.3 | 9.45 | 9.6 | 9.74 | 9.89 | 10.03 | 10.18 | 10.32 | 10.47 | 10.61 | 10.76 | 10.9 | 11.05 | 11.19 | 11.34 | 11.48
11.5 | 5.67 | 5.81 | 5.96 | 6.11 | 6.25 | 6.4 | 6.54 | 6.69 | 6.83 | 6.98 | 7.12 | 7.27 | 7.41 | 7.56 | 7.7 | 7.85 | 7.99 | 8.14 | 8.29 | 8.43 | 8.58 | 8.72 | 8.87 | 9.01 | 9.16 | 9.3 | 9.45 | 9.59 | 9.74 | 9.88 | 10.03 | 10.17 | 10.32 | 10.47 | 10.61 | 10.76 | 10.9 | 11.05 | 11.19 | 11.34 | 11.48
11.51 | 5.67 | 5.81 | 5.96 | 6.1 | 6.25 | 6.39 | 6.54 | 6.69 | 6.83 | 6.98 | 7.12 | 7.27 | 7.41 | 7.56 | 7.7 | 7.85 | 7.99 | 8.14 | 8.28 | 8.43 | 8.57 | 8.72 | 8.86 | 9.01 | 9.16 | 9.3 | 9.45 | 9.59 | 9.74 | 9.88 | 10.03 | 10.17 | 10.32 | 10.46 | 10.61 | 10.75 | 10.9 | 11.04 | 11.19 | 11.33 | 11.48
11.52 | 5.67 | 5.81 | 5.96 | 6.1 | 6.25 | 6.39 | 6.54 | 6.68 | 6.83 | 6.97 | 7.12 | 7.26 | 7.41 | 7.56 | 7.7 | 7.85 | 7.99 | 8.14 | 8.28 | 8.43 | 8.57 | 8.72 | 8.86 | 9.01 | 9.15 | 9.3 | 9.44 | 9.59 | 9.73 | 9.88 | 10.02 | 10.17 | 10.31 | 10.46 | 10.61 | 10.75 | 10.9 | 11.04 | 11.19 | 11.33 | 11.48
11.53 | 5.67 | 5.81 | 5.96 | 6.1 | 6.25 | 6.39 | 6.54 | 6.68 | 6.83 | 6.97 | 7.12 | 7.26 | 7.41 | 7.55 | 7.7 | 7.84 | 7.99 | 8.13 | 8.28 | 8.43 | 8.57 | 8.72 | 8.86 | 9.01 | 9.15 | 9.3 | 9.44 | 9.59 | 9.73 | 9.88 | 10.02 | 10.17 | 10.31 | 10.46 | 10.6 | 10.75 | 10.89 | 11.04 | 11.18 | 11.33 | 11.47
11.54 | 5.66 | 5.81 | 5.95 | 6.1 | 6.24 | 6.39 | 6.54 | 6.68 | 6.83 | 6.97 | 7.12 | 7.26 | 7.41 | 7.55 | 7.7 | 7.84 | 7.99 | 8.13 | 8.28 | 8.42 | 8.57 | 8.71 | 8.86 | 9 | 9.15 | 9.29 | 9.44 | 9.58 | 9.73 | 9.87 | 10.02 | 10.17 | 10.31 | 10.46 | 10.6 | 10.75 | 10.89 | 11.04 | 11.18 | 11.33 | 11.47
11.55 | 5.66 | 5.81 | 5.95 | 6.1 | 6.24 | 6.39 | 6.53 | 6.68 | 6.82 | 6.97 | 7.11 | 7.26 | 7.4 | 7.55 | 7.7 | 7.84 | 7.99 | 8.13 | 8.28 | 8.42 | 8.57 | 8.71 | 8.86 | 9 | 9.15 | 9.29 | 9.44 | 9.58 | 9.73 | 9.87 | 10.02 | 10.16 | 10.31 | 10.45 | 10.6 | 10.74 | 10.89 | 11.03 | 11.18 | 11.32 | 11.47
11.56 | 5.66 | 5.81 | 5.95 | 6.1 | 6.24 | 6.39 | 6.53 | 6.68 | 6.82 | 6.97 | 7.11 | 7.26 | 7.4 | 7.55 | 7.69 | 7.84 | 7.98 | 8.13 | 8.27 | 8.42 | 8.56 | 8.71 | 8.85 | 9 | 9.14 | 9.29 | 9.43 | 9.58 | 9.73 | 9.87 | 10.02 | 10.16 | 10.31 | 10.45 | 10.6 | 10.74 | 10.89 | 11.03 | 11.18 | 11.32 | 11.47
11.57 | 5.66 | 5.81 | 5.95 | 6.1 | 6.24 | 6.39 | 6.53 | 6.68 | 6.82 | 6.97 | 7.11 | 7.26 | 7.4 | 7.55 | 7.69 | 7.84 | 7.98 | 8.13 | 8.27 | 8.42 | 8.56 | 8.71 | 8.85 | 9 | 9.14 | 9.29 | 9.43 | 9.58 | 9.72 | 9.87 | 10.01 | 10.16 | 10.3 | 10.45 | 10.59 | 10.74 | 10.88 | 11.03 | 11.17 | 11.32 | 11.46
11.58 | 5.66 | 5.8 | 5.95 | 6.09 | 6.24 | 6.38 | 6.53 | 6.67 | 6.82 | 6.96 | 7.11 | 7.25 | 7.4 | 7.54 | 7.69 | 7.84 | 7.98 | 8.13 | 8.27 | 8.42 | 8.56 | 8.71 | 8.85 | 9 | 9.14 | 9.29 | 9.43 | 9.58 | 9.72 | 9.87 | 10.01 | 10.16 | 10.3 | 10.45 | 10.59 | 10.74 | 10.88 | 11.03 | 11.17 | 11.32 | 11.46
11.59 | 5.66 | 5.8 | 5.95 | 6.09 | 6.24 | 6.38 | 6.53 | 6.67 | 6.82 | 6.96 | 7.11 | 7.25 | 7.4 | 7.54 | 7.69 | 7.83 | 7.98 | 8.12 | 8.27 | 8.41 | 8.56 | 8.7 | 8.85 | 8.99 | 9.14 | 9.28 | 9.43 | 9.57 | 9.72 | 9.86 | 10.01 | 10.15 | 10.3 | 10.44 | 10.59 | 10.73 | 10.88 | 11.02 | 11.17 | 11.31 | 11.46
11.6 | 5.66 | 5.8 | 5.95 | 6.09 | 6.24 | 6.38 | 6.53 | 6.67 | 6.82 | 6.96 | 7.11 | 7.25 | 7.4 | 7.54 | 7.69 | 7.83 | 7.98 | 8.12 | 8.27 | 8.41 | 8.56 | 8.7 | 8.85 | 8.99 | 9.14 | 9.28 | 9.43 | 9.57 | 9.72 | 9.86 | 10.01 | 10.15 | 10.3 | 10.44 | 10.59 | 10.73 | 10.88 | 11.02 | 11.17 | 11.31 | 11.46
11.61 | 5.65 | 5.8 | 5.94 | 6.09 | 6.23 | 6.38 | 6.52 | 6.67 | 6.81 | 6.96 | 7.1 | 7.25 | 7.39 | 7.54 | 7.68 | 7.83 | 7.97 | 8.12 | 8.26 | 8.41 | 8.55 | 8.7 | 8.84 | 8.99 | 9.13 | 9.28 | 9.42 | 9.57 | 9.71 | 9.86 | 10 | 10.15 | 10.29 | 10.44 | 10.58 | 10.73 | 10.87 | 11.02 | 11.16 | 11.31 | 11.45
11.62 | 5.65 | 5.8 | 5.94 | 6.09 | 6.23 | 6.38 | 6.52 | 6.67 | 6.81 | 6.96 | 7.1 | 7.25 | 7.39 | 7.54 | 7.68 | 7.83 | 7.97 | 8.12 | 8.26 | 8.41 | 8.55 | 8.7 | 8.84 | 8.99 | 9.13 | 9.28 | 9.42 | 9.57 | 9.71 | 9.86 | 10 | 10.15 | 10.29 | 10.44 | 10.58 | 10.73 | 10.87 | 11.02 | 11.16 | 11.31 | 11.45
11.63 | 5.65 | 5.8 | 5.94 | 6.09 | 6.23 | 6.38 | 6.52 | 6.67 | 6.81 | 6.96 | 7.1 | 7.25 | 7.39 | 7.54 | 7.68 | 7.83 | 7.97 | 8.12 | 8.26 | 8.41 | 8.55 | 8.7 | 8.84 | 8.99 | 9.13 | 9.27 | 9.42 | 9.56 | 9.71 | 9.85 | 10 | 10.14 | 10.29 | 10.43 | 10.58 | 10.72 | 10.87 | 11.01 | 11.16 | 11.3 | 11.45
11.64 | 5.65 | 5.8 | 5.94 | 6.09 | 6.23 | 6.38 | 6.52 | 6.66 | 6.81 | 6.95 | 7.1 | 7.24 | 7.39 | 7.53 | 7.68 | 7.82 | 7.97 | 8.11 | 8.26 | 8.4 | 8.55 | 8.69 | 8.84 | 8.98 | 9.13 | 9.27 | 9.42 | 9.56 | 9.71 | 9.85 | 10 | 10.14 | 10.29 | 10.43 | 10.58 | 10.72 | 10.87 | 11.01 | 11.16 | 11.3 | 11.45
11.65 | 5.65 | 5.79 | 5.94 | 6.08 | 6.23 | 6.37 | 6.52 | 6.66 | 6.81 | 6.95 | 7.1 | 7.24 | 7.39 | 7.53 | 7.68 | 7.82 | 7.97 | 8.11 | 8.26 | 8.4 | 8.55 | 8.69 | 8.84 | 8.98 | 9.13 | 9.27 | 9.42 | 9.56 | 9.71 | 9.85 | 9.99 | 10.14 | 10.28 | 10.43 | 10.57 | 10.72 | 10.86 | 11.01 | 11.15 | 11.3 | 11.44
11.66 | 5.65 | 5.79 | 5.94 | 6.08 | 6.23 | 6.37 | 6.52 | 6.66 | 6.81 | 6.95 | 7.1 | 7.24 | 7.39 | 7.53 | 7.68 | 7.82 | 7.97 | 8.11 | 8.25 | 8.4 | 8.54 | 8.69 | 8.83 | 8.98 | 9.12 | 9.27 | 9.41 | 9.56 | 9.7 | 9.85 | 9.99 | 10.14 | 10.28 | 10.43 | 10.57 | 10.72 | 10.86 | 11.01 | 11.15 | 11.3 | 11.44
11.67 | 5.65 | 5.79 | 5.94 | 6.08 | 6.23 | 6.37 | 6.52 | 6.66 | 6.81 | 6.95 | 7.09 | 7.24 | 7.38 | 7.53 | 7.67 | 7.82 | 7.96 | 8.11 | 8.25 | 8.4 | 8.54 | 8.69 | 8.83 | 8.98 | 9.12 | 9.27 | 9.41 | 9.56 | 9.7 | 9.85 | 9.99 | 10.13 | 10.28 | 10.42 | 10.57 | 10.71 | 10.86 | 11 | 11.15 | 11.29 | 11.44
11.68 | 5.65 | 5.79 | 5.93 | 6.08 | 6.22 | 6.37 | 6.51 | 6.66 | 6.8 | 6.95 | 7.09 | 7.24 | 7.38 | 7.53 | 7.67 | 7.82 | 7.96 | 8.11 | 8.25 | 8.4 | 8.54 | 8.69 | 8.83 | 8.97 | 9.12 | 9.26 | 9.41 | 9.55 | 9.7 | 9.84 | 9.99 | 10.13 | 10.28 | 10.42 | 10.57 | 10.71 | 10.86 | 11 | 11.15 | 11.29 | 11.43
11.69 | 5.64 | 5.79 | 5.93 | 6.08 | 6.22 | 6.37 | 6.51 | 6.66 | 6.8 | 6.95 | 7.09 | 7.24 | 7.38 | 7.53 | 7.67 | 7.82 | 7.96 | 8.1 | 8.25 | 8.39 | 8.54 | 8.68 | 8.83 | 8.97 | 9.12 | 9.26 | 9.41 | 9.55 | 9.7 | 9.84 | 9.99 | 10.13 | 10.27 | 10.42 | 10.56 | 10.71 | 10.85 | 11 | 11.14 | 11.29 | 11.43
11.7 | 5.64 | 5.79 | 5.93 | 6.08 | 6.22 | 6.37 | 6.51 | 6.66 | 6.8 | 6.95 | 7.09 | 7.23 | 7.38 | 7.52 | 7.67 | 7.81 | 7.96 | 8.1 | 8.25 | 8.39 | 8.54 | 8.68 | 8.83 | 8.97 | 9.12 | 9.26 | 9.4 | 9.55 | 9.69 | 9.84 | 9.98 | 10.13 | 10.27 | 10.42 | 10.56 | 10.71 | 10.85 | 11 | 11.14 | 11.28 | 11.43
11.71 | 5.64 | 5.79 | 5.93 | 6.08 | 6.22 | 6.36 | 6.51 | 6.65 | 6.8 | 6.94 | 7.09 | 7.23 | 7.38 | 7.52 | 7.67 | 7.81 | 7.96 | 8.1 | 8.25 | 8.39 | 8.53 | 8.68 | 8.82 | 8.97 | 9.11 | 9.26 | 9.4 | 9.55 | 9.69 | 9.84 | 9.98 | 10.13 | 10.27 | 10.41 | 10.56 | 10.7 | 10.85 | 10.99 | 11.14 | 11.28 | 11.43
11.72 | 5.64 | 5.78 | 5.93 | 6.07 | 6.22 | 6.36 | 6.51 | 6.65 | 6.8 | 6.94 | 7.09 | 7.23 | 7.38 | 7.52 | 7.66 | 7.81 | 7.95 | 8.1 | 8.24 | 8.39 | 8.53 | 8.68 | 8.82 | 8.97 | 9.11 | 9.26 | 9.4 | 9.54 | 9.69 | 9.83 | 9.98 | 10.12 | 10.27 | 10.41 | 10.56 | 10.7 | 10.85 | 10.99 | 11.14 | 11.28 | 11.42
11.73 | 5.64 | 5.78 | 5.93 | 6.07 | 6.22 | 6.36 | 6.51 | 6.65 | 6.8 | 6.94 | 7.08 | 7.23 | 7.37 | 7.52 | 7.66 | 7.81 | 7.95 | 8.1 | 8.24 | 8.39 | 8.53 | 8.68 | 8.82 | 8.96 | 9.11 | 9.25 | 9.4 | 9.54 | 9.69 | 9.83 | 9.98 | 10.12 | 10.27 | 10.41 | 10.55 | 10.7 | 10.84 | 10.99 | 11.13 | 11.28 | 11.42
11.74 | 5.64 | 5.78 | 5.93 | 6.07 | 6.22 | 6.36 | 6.5 | 6.65 | 6.79 | 6.94 | 7.08 | 7.23 | 7.37 | 7.52 | 7.66 | 7.81 | 7.95 | 8.1 | 8.24 | 8.38 | 8.53 | 8.67 | 8.82 | 8.96 | 9.11 | 9.25 | 9.4 | 9.54 | 9.68 | 9.83 | 9.97 | 10.12 | 10.26 | 10.41 | 10.55 | 10.7 | 10.84 | 10.99 | 11.13 | 11.27 | 11.42
11.75 | 5.64 | 5.78 | 5.92 | 6.07 | 6.21 | 6.36 | 6.5 | 6.65 | 6.79 | 6.94 | 7.08 | 7.23 | 7.37 | 7.52 | 7.66 | 7.8 | 7.95 | 8.09 | 8.24 | 8.38 | 8.53 | 8.67 | 8.82 | 8.96 | 9.1 | 9.25 | 9.39 | 9.54 | 9.68 | 9.83 | 9.97 | 10.12 | 10.26 | 10.41 | 10.55 | 10.69 | 10.84 | 10.98 | 11.13 | 11.27 | 11.42
11.76 | 5.63 | 5.78 | 5.92 | 6.07 | 6.21 | 6.36 | 6.5 | 6.65 | 6.79 | 6.94 | 7.08 | 7.22 | 7.37 | 7.51 | 7.66 | 7.8 | 7.95 | 8.09 | 8.24 | 8.38 | 8.52 | 8.67 | 8.81 | 8.96 | 9.1 | 9.25 | 9.39 | 9.54 | 9.68 | 9.82 | 9.97 | 10.11 | 10.26 | 10.4 | 10.55 | 10.69 | 10.84 | 10.98 | 11.12 | 11.27 | 11.41
11.77 | 5.63 | 5.78 | 5.92 | 6.07 | 6.21 | 6.36 | 6.5 | 6.64 | 6.79 | 6.93 | 7.08 | 7.22 | 7.37 | 7.51 | 7.66 | 7.8 | 7.94 | 8.09 | 8.23 | 8.38 | 8.52 | 8.67 | 8.81 | 8.96 | 9.1 | 9.25 | 9.39 | 9.53 | 9.68 | 9.82 | 9.97 | 10.11 | 10.26 | 10.4 | 10.54 | 10.69 | 10.83 | 10.98 | 11.12 | 11.27 | 11.41
11.78 | 5.63 | 5.78 | 5.92 | 6.07 | 6.21 | 6.35 | 6.5 | 6.64 | 6.79 | 6.93 | 7.08 | 7.22 | 7.37 | 7.51 | 7.65 | 7.8 | 7.94 | 8.09 | 8.23 | 8.38 | 8.52 | 8.67 | 8.81 | 8.95 | 9.1 | 9.24 | 9.39 | 9.53 | 9.68 | 9.82 | 9.96 | 10.11 | 10.25 | 10.4 | 10.54 | 10.69 | 10.83 | 10.98 | 11.12 | 11.26 | 11.41
11.79 | 5.63 | 5.77 | 5.92 | 6.06 | 6.21 | 6.35 | 6.5 | 6.64 | 6.79 | 6.93 | 7.07 | 7.22 | 7.36 | 7.51 | 7.65 | 7.8 | 7.94 | 8.09 | 8.23 | 8.37 | 8.52 | 8.66 | 8.81 | 8.95 | 9.1 | 9.24 | 9.39 | 9.53 | 9.67 | 9.82 | 9.96 | 10.11 | 10.25 | 10.4 | 10.54 | 10.68 | 10.83 | 10.97 | 11.12 | 11.26 | 11.41
11.8 | 5.63 | 5.77 | 5.92 | 6.06 | 6.21 | 6.35 | 6.5 | 6.64 | 6.78 | 6.93 | 7.07 | 7.22 | 7.36 | 7.51 | 7.65 | 7.8 | 7.94 | 8.08 | 8.23 | 8.37 | 8.52 | 8.66 | 8.81 | 8.95 | 9.09 | 9.24 | 9.38 | 9.53 | 9.67 | 9.82 | 9.96 | 10.1 | 10.25 | 10.39 | 10.54 | 10.68 | 10.83 | 10.97 | 11.11 | 11.26 | 11.4
11.81 | 5.63 | 5.77 | 5.92 | 6.06 | 6.21 | 6.35 | 6.49 | 6.64 | 6.78 | 6.93 | 7.07 | 7.22 | 7.36 | 7.5 | 7.65 | 7.79 | 7.94 | 8.08 | 8.23 | 8.37 | 8.51 | 8.66 | 8.8 | 8.95 | 9.09 | 9.24 | 9.38 | 9.53 | 9.67 | 9.81 | 9.96 | 10.1 | 10.25 | 10.39 | 10.54 | 10.68 | 10.82 | 10.97 | 11.11 | 11.26 | 11.4
11.82 | 5.63 | 5.77 | 5.92 | 6.06 | 6.2 | 6.35 | 6.49 | 6.64 | 6.78 | 6.93 | 7.07 | 7.21 | 7.36 | 7.5 | 7.65 | 7.79 | 7.94 | 8.08 | 8.22 | 8.37 | 8.51 | 8.66 | 8.8 | 8.95 | 9.09 | 9.23 | 9.38 | 9.52 | 9.67 | 9.81 | 9.96 | 10.1 | 10.24 | 10.39 | 10.53 | 10.68 | 10.82 | 10.97 | 11.11 | 11.25 | 11.4
11.83 | 5.63 | 5.77 | 5.91 | 6.06 | 6.2 | 6.35 | 6.49 | 6.64 | 6.78 | 6.92 | 7.07 | 7.21 | 7.36 | 7.5 | 7.65 | 7.79 | 7.93 | 8.08 | 8.22 | 8.37 | 8.51 | 8.66 | 8.8 | 8.94 | 9.09 | 9.23 | 9.38 | 9.52 | 9.66 | 9.81 | 9.95 | 10.1 | 10.24 | 10.39 | 10.53 | 10.67 | 10.82 | 10.96 | 11.11 | 11.25 | 11.4
11.84 | 5.62 | 5.77 | 5.91 | 6.06 | 6.2 | 6.35 | 6.49 | 6.63 | 6.78 | 6.92 | 7.07 | 7.21 | 7.36 | 7.5 | 7.64 | 7.79 | 7.93 | 8.08 | 8.22 | 8.36 | 8.51 | 8.65 | 8.8 | 8.94 | 9.09 | 9.23 | 9.37 | 9.52 | 9.66 | 9.81 | 9.95 | 10.1 | 10.24 | 10.38 | 10.53 | 10.67 | 10.82 | 10.96 | 11.1 | 11.25 | 11.39
11.85 | 5.62 | 5.77 | 5.91 | 6.06 | 6.2 | 6.34 | 6.49 | 6.63 | 6.78 | 6.92 | 7.06 | 7.21 | 7.35 | 7.5 | 7.64 | 7.79 | 7.93 | 8.07 | 8.22 | 8.36 | 8.51 | 8.65 | 8.8 | 8.94 | 9.08 | 9.23 | 9.37 | 9.52 | 9.66 | 9.8 | 9.95 | 10.09 | 10.24 | 10.38 | 10.53 | 10.67 | 10.81 | 10.96 | 11.1 | 11.25 | 11.39
11.86 | 5.62 | 5.77 | 5.91 | 6.05 | 6.2 | 6.34 | 6.49 | 6.63 | 6.77 | 6.92 | 7.06 | 7.21 | 7.35 | 7.5 | 7.64 | 7.78 | 7.93 | 8.07 | 8.22 | 8.36 | 8.51 | 8.65 | 8.79 | 8.94 | 9.08 | 9.23 | 9.37 | 9.51 | 9.66 | 9.8 | 9.95 | 10.09 | 10.23 | 10.38 | 10.52 | 10.67 | 10.81 | 10.96 | 11.1 | 11.24 | 11.39
11.87 | 5.62 | 5.76 | 5.91 | 6.05 | 6.2 | 6.34 | 6.48 | 6.63 | 6.77 | 6.92 | 7.06 | 7.21 | 7.35 | 7.49 | 7.64 | 7.78 | 7.93 | 8.07 | 8.21 | 8.36 | 8.5 | 8.65 | 8.79 | 8.94 | 9.08 | 9.22 | 9.37 | 9.51 | 9.66 | 9.8 | 9.94 | 10.09 | 10.23 | 10.38 | 10.52 | 10.66 | 10.81 | 10.95 | 11.1 | 11.24 | 11.39
11.88 | 5.62 | 5.76 | 5.91 | 6.05 | 6.2 | 6.34 | 6.48 | 6.63 | 6.77 | 6.92 | 7.06 | 7.2 | 7.35 | 7.49 | 7.64 | 7.78 | 7.92 | 8.07 | 8.21 | 8.36 | 8.5 | 8.65 | 8.79 | 8.93 | 9.08 | 9.22 | 9.37 | 9.51 | 9.65 | 9.8 | 9.94 | 10.09 | 10.23 | 10.37 | 10.52 | 10.66 | 10.81 | 10.95 | 11.09 | 11.24 | 11.38
11.89 | 5.62 | 5.76 | 5.91 | 6.05 | 6.19 | 6.34 | 6.48 | 6.63 | 6.77 | 6.91 | 7.06 | 7.2 | 7.35 | 7.49 | 7.63 | 7.78 | 7.92 | 8.07 | 8.21 | 8.36 | 8.5 | 8.64 | 8.79 | 8.93 | 9.08 | 9.22 | 9.36 | 9.51 | 9.65 | 9.8 | 9.94 | 10.08 | 10.23 | 10.37 | 10.52 | 10.66 | 10.8 | 10.95 | 11.09 | 11.24 | 11.38
11.9 | 5.62 | 5.76 | 5.9 | 6.05 | 6.19 | 6.34 | 6.48 | 6.62 | 6.77 | 6.91 | 7.06 | 7.2 | 7.34 | 7.49 | 7.63 | 7.78 | 7.92 | 8.07 | 8.21 | 8.35 | 8.5 | 8.64 | 8.79 | 8.93 | 9.07 | 9.22 | 9.36 | 9.51 | 9.65 | 9.79 | 9.94 | 10.08 | 10.23 | 10.37 | 10.51 | 10.66 | 10.8 | 10.95 | 11.09 | 11.23 | 11.38
11.91 | 5.61 | 5.76 | 5.9 | 6.05 | 6.19 | 6.33 | 6.48 | 6.62 | 6.77 | 6.91 | 7.06 | 7.2 | 7.34 | 7.49 | 7.63 | 7.78 | 7.92 | 8.06 | 8.21 | 8.35 | 8.5 | 8.64 | 8.78 | 8.93 | 9.07 | 9.22 | 9.36 | 9.5 | 9.65 | 9.79 | 9.94 | 10.08 | 10.22 | 10.37 | 10.51 | 10.66 | 10.8 | 10.94 | 11.09 | 11.23 | 11.37
11.92 | 5.61 | 5.76 | 5.9 | 6.05 | 6.19 | 6.33 | 6.48 | 6.62 | 6.77 | 6.91 | 7.05 | 7.2 | 7.34 | 7.49 | 7.63 | 7.77 | 7.92 | 8.06 | 8.21 | 8.35 | 8.49 | 8.64 | 8.78 | 8.93 | 9.07 | 9.21 | 9.36 | 9.5 | 9.65 | 9.79 | 9.93 | 10.08 | 10.22 | 10.36 | 10.51 | 10.65 | 10.8 | 10.94 | 11.08 | 11.23 | 11.37
11.93 | 5.61 | 5.76 | 5.9 | 6.04 | 6.19 | 6.33 | 6.48 | 6.62 | 6.76 | 6.91 | 7.05 | 7.2 | 7.34 | 7.48 | 7.63 | 7.77 | 7.92 | 8.06 | 8.2 | 8.35 | 8.49 | 8.64 | 8.78 | 8.92 | 9.07 | 9.21 | 9.36 | 9.5 | 9.64 | 9.79 | 9.93 | 10.07 | 10.22 | 10.36 | 10.51 | 10.65 | 10.79 | 10.94 | 11.08 | 11.23 | 11.37
11.94 | 5.61 | 5.75 | 5.9 | 6.04 | 6.19 | 6.33 | 6.47 | 6.62 | 6.76 | 6.91 | 7.05 | 7.19 | 7.34 | 7.48 | 7.63 | 7.77 | 7.91 | 8.06 | 8.2 | 8.35 | 8.49 | 8.63 | 8.78 | 8.92 | 9.07 | 9.21 | 9.35 | 9.5 | 9.64 | 9.78 | 9.93 | 10.07 | 10.22 | 10.36 | 10.5 | 10.65 | 10.79 | 10.94 | 11.08 | 11.22 | 11.37
11.95 | 5.61 | 5.75 | 5.9 | 6.04 | 6.18 | 6.33 | 6.47 | 6.62 | 6.76 | 6.9 | 7.05 | 7.19 | 7.34 | 7.48 | 7.62 | 7.77 | 7.91 | 8.06 | 8.2 | 8.34 | 8.49 | 8.63 | 8.78 | 8.92 | 9.06 | 9.21 | 9.35 | 9.49 | 9.64 | 9.78 | 9.93 | 10.07 | 10.21 | 10.36 | 10.5 | 10.65 | 10.79 | 10.93 | 11.08 | 11.22 | 11.36
11.96 | 5.61 | 5.75 | 5.9 | 6.04 | 6.18 | 6.33 | 6.47 | 6.62 | 6.76 | 6.9 | 7.05 | 7.19 | 7.33 | 7.48 | 7.62 | 7.77 | 7.91 | 8.05 | 8.2 | 8.34 | 8.49 | 8.63 | 8.77 | 8.92 | 9.06 | 9.2 | 9.35 | 9.49 | 9.64 | 9.78 | 9.92 | 10.07 | 10.21 | 10.36 | 10.5 | 10.64 | 10.79 | 10.93 | 11.07 | 11.22 | 11.36
11.97 | 5.61 | 5.75 | 5.89 | 6.04 | 6.18 | 6.33 | 6.47 | 6.61 | 6.76 | 6.9 | 7.05 | 7.19 | 7.33 | 7.48 | 7.62 | 7.76 | 7.91 | 8.05 | 8.2 | 8.34 | 8.48 | 8.63 | 8.77 | 8.92 | 9.06 | 9.2 | 9.35 | 9.49 | 9.63 | 9.78 | 9.92 | 10.07 | 10.21 | 10.35 | 10.5 | 10.64 | 10.78 | 10.93 | 11.07 | 11.22 | 11.36
11.98 | 5.61 | 5.75 | 5.89 | 6.04 | 6.18 | 6.32 | 6.47 | 6.61 | 6.76 | 6.9 | 7.04 | 7.19 | 7.33 | 7.48 | 7.62 | 7.76 | 7.91 | 8.05 | 8.19 | 8.34 | 8.48 | 8.63 | 8.77 | 8.91 | 9.06 | 9.2 | 9.34 | 9.49 | 9.63 | 9.78 | 9.92 | 10.06 | 10.21 | 10.35 | 10.49 | 10.64 | 10.78 | 10.93 | 11.07 | 11.21 | 11.36
11.99 | 5.6 | 5.75 | 5.89 | 6.04 | 6.18 | 6.32 | 6.47 | 6.61 | 6.75 | 6.9 | 7.04 | 7.19 | 7.33 | 7.47 | 7.62 | 7.76 | 7.9 | 8.05 | 8.19 | 8.34 | 8.48 | 8.62 | 8.77 | 8.91 | 9.05 | 9.2 | 9.34 | 9.49 | 9.63 | 9.77 | 9.92 | 10.06 | 10.2 | 10.35 | 10.49 | 10.64 | 10.78 | 10.92 | 11.07 | 11.21 | 11.35
12 | 5.6 | 5.75 | 5.89 | 6.03 | 6.18 | 6.32 | 6.47 | 6.61 | 6.75 | 6.9 | 7.04 | 7.18 | 7.33 | 7.47 | 7.62 | 7.76 | 7.9 | 8.05 | 8.19 | 8.33 | 8.48 | 8.62 | 8.77 | 8.91 | 9.05 | 9.2 | 9.34 | 9.48 | 9.63 | 9.77 | 9.91 | 10.06 | 10.2 | 10.35 | 10.49 | 10.63 | 10.78 | 10.92 | 11.06 | 11.21 | 11.35
12.01 | 5.6 | 5.75 | 5.89 | 6.03 | 6.18 | 6.32 | 6.46 | 6.61 | 6.75 | 6.89 | 7.04 | 7.18 | 7.33 | 7.47 | 7.61 | 7.76 | 7.9 | 8.04 | 8.19 | 8.33 | 8.48 | 8.62 | 8.76 | 8.91 | 9.05 | 9.19 | 9.34 | 9.48 | 9.63 | 9.77 | 9.91 | 10.06 | 10.2 | 10.34 | 10.49 | 10.63 | 10.77 | 10.92 | 11.06 | 11.21 | 11.35
12.02 | 5.6 | 5.74 | 5.89 | 6.03 | 6.17 | 6.32 | 6.46 | 6.61 | 6.75 | 6.89 | 7.04 | 7.18 | 7.32 | 7.47 | 7.61 | 7.76 | 7.9 | 8.04 | 8.19 | 8.33 | 8.47 | 8.62 | 8.76 | 8.9 | 9.05 | 9.19 | 9.34 | 9.48 | 9.62 | 9.77 | 9.91 | 10.05 | 10.2 | 10.34 | 10.48 | 10.63 | 10.77 | 10.92 | 11.06 | 11.2 | 11.35
12.03 | 5.6 | 5.74 | 5.89 | 6.03 | 6.17 | 6.32 | 6.46 | 6.6 | 6.75 | 6.89 | 7.04 | 7.18 | 7.32 | 7.47 | 7.61 | 7.75 | 7.9 | 8.04 | 8.18 | 8.33 | 8.47 | 8.62 | 8.76 | 8.9 | 9.05 | 9.19 | 9.33 | 9.48 | 9.62 | 9.76 | 9.91 | 10.05 | 10.2 | 10.34 | 10.48 | 10.63 | 10.77 | 10.91 | 11.06 | 11.2 | 11.34
12.04 | 5.6 | 5.74 | 5.88 | 6.03 | 6.17 | 6.32 | 6.46 | 6.6 | 6.75 | 6.89 | 7.03 | 7.18 | 7.32 | 7.46 | 7.61 | 7.75 | 7.9 | 8.04 | 8.18 | 8.33 | 8.47 | 8.61 | 8.76 | 8.9 | 9.04 | 9.19 | 9.33 | 9.48 | 9.62 | 9.76 | 9.91 | 10.05 | 10.19 | 10.34 | 10.48 | 10.62 | 10.77 | 10.91 | 11.05 | 11.2 | 11.34
12.05 | 5.6 | 5.74 | 5.88 | 6.03 | 6.17 | 6.31 | 6.46 | 6.6 | 6.74 | 6.89 | 7.03 | 7.18 | 7.32 | 7.46 | 7.61 | 7.75 | 7.89 | 8.04 | 8.18 | 8.32 | 8.47 | 8.61 | 8.76 | 8.9 | 9.04 | 9.19 | 9.33 | 9.47 | 9.62 | 9.76 | 9.9 | 10.05 | 10.19 | 10.33 | 10.48 | 10.62 | 10.76 | 10.91 | 11.05 | 11.2 | 11.34
12.06 | 5.59 | 5.74 | 5.88 | 6.03 | 6.17 | 6.31 | 6.46 | 6.6 | 6.74 | 6.89 | 7.03 | 7.17 | 7.32 | 7.46 | 7.6 | 7.75 | 7.89 | 8.04 | 8.18 | 8.32 | 8.47 | 8.61 | 8.75 | 8.9 | 9.04 | 9.18 | 9.33 | 9.47 | 9.61 | 9.76 | 9.9 | 10.04 | 10.19 | 10.33 | 10.48 | 10.62 | 10.76 | 10.91 | 11.05 | 11.19 | 11.34
12.07 | 5.59 | 5.74 | 5.88 | 6.02 | 6.17 | 6.31 | 6.45 | 6.6 | 6.74 | 6.89 | 7.03 | 7.17 | 7.32 | 7.46 | 7.6 | 7.75 | 7.89 | 8.03 | 8.18 | 8.32 | 8.46 | 8.61 | 8.75 | 8.89 | 9.04 | 9.18 | 9.33 | 9.47 | 9.61 | 9.76 | 9.9 | 10.04 | 10.19 | 10.33 | 10.47 | 10.62 | 10.76 | 10.9 | 11.05 | 11.19 | 11.33
12.08 | 5.59 | 5.74 | 5.88 | 6.02 | 6.17 | 6.31 | 6.45 | 6.6 | 6.74 | 6.88 | 7.03 | 7.17 | 7.31 | 7.46 | 7.6 | 7.74 | 7.89 | 8.03 | 8.18 | 8.32 | 8.46 | 8.61 | 8.75 | 8.89 | 9.04 | 9.18 | 9.32 | 9.47 | 9.61 | 9.75 | 9.9 | 10.04 | 10.18 | 10.33 | 10.47 | 10.61 | 10.76 | 10.9 | 11.04 | 11.19 | 11.33
12.09 | 5.59 | 5.73 | 5.88 | 6.02 | 6.16 | 6.31 | 6.45 | 6.6 | 6.74 | 6.88 | 7.03 | 7.17 | 7.31 | 7.46 | 7.6 | 7.74 | 7.89 | 8.03 | 8.17 | 8.32 | 8.46 | 8.6 | 8.75 | 8.89 | 9.03 | 9.18 | 9.32 | 9.46 | 9.61 | 9.75 | 9.89 | 10.04 | 10.18 | 10.32 | 10.47 | 10.61 | 10.75 | 10.9 | 11.04 | 11.19 | 11.33
12.1 | 5.59 | 5.73 | 5.88 | 6.02 | 6.16 | 6.31 | 6.45 | 6.59 | 6.74 | 6.88 | 7.02 | 7.17 | 7.31 | 7.45 | 7.6 | 7.74 | 7.88 | 8.03 | 8.17 | 8.31 | 8.46 | 8.6 | 8.75 | 8.89 | 9.03 | 9.18 | 9.32 | 9.46 | 9.61 | 9.75 | 9.89 | 10.04 | 10.18 | 10.32 | 10.47 | 10.61 | 10.75 | 10.9 | 11.04 | 11.18 | 11.33
12.11 | 5.59 | 5.73 | 5.87 | 6.02 | 6.16 | 6.31 | 6.45 | 6.59 | 6.74 | 6.88 | 7.02 | 7.17 | 7.31 | 7.45 | 7.6 | 7.74 | 7.88 | 8.03 | 8.17 | 8.31 | 8.46 | 8.6 | 8.74 | 8.89 | 9.03 | 9.17 | 9.32 | 9.46 | 9.6 | 9.75 | 9.89 | 10.03 | 10.18 | 10.32 | 10.46 | 10.61 | 10.75 | 10.89 | 11.04 | 11.18 | 11.32
12.12 | 5.59 | 5.73 | 5.87 | 6.02 | 6.16 | 6.3 | 6.45 | 6.59 | 6.73 | 6.88 | 7.02 | 7.16 | 7.31 | 7.45 | 7.59 | 7.74 | 7.88 | 8.02 | 8.17 | 8.31 | 8.45 | 8.6 | 8.74 | 8.88 | 9.03 | 9.17 | 9.31 | 9.46 | 9.6 | 9.74 | 9.89 | 10.03 | 10.17 | 10.32 | 10.46 | 10.6 | 10.75 | 10.89 | 11.03 | 11.18 | 11.32
12.13 | 5.59 | 5.73 | 5.87 | 6.02 | 6.16 | 6.3 | 6.45 | 6.59 | 6.73 | 6.88 | 7.02 | 7.16 | 7.31 | 7.45 | 7.59 | 7.74 | 7.88 | 8.02 | 8.17 | 8.31 | 8.45 | 8.6 | 8.74 | 8.88 | 9.03 | 9.17 | 9.31 | 9.46 | 9.6 | 9.74 | 9.89 | 10.03 | 10.17 | 10.32 | 10.46 | 10.6 | 10.75 | 10.89 | 11.03 | 11.17 | 11.32
12.14 | 5.58 | 5.73 | 5.87 | 6.01 | 6.16 | 6.3 | 6.44 | 6.59 | 6.73 | 6.87 | 7.02 | 7.16 | 7.3 | 7.45 | 7.59 | 7.73 | 7.88 | 8.02 | 8.16 | 8.31 | 8.45 | 8.59 | 8.74 | 8.88 | 9.02 | 9.17 | 9.31 | 9.45 | 9.6 | 9.74 | 9.88 | 10.03 | 10.17 | 10.31 | 10.46 | 10.6 | 10.74 | 10.89 | 11.03 | 11.17 | 11.32
12.15 | 5.58 | 5.73 | 5.87 | 6.01 | 6.16 | 6.3 | 6.44 | 6.59 | 6.73 | 6.87 | 7.02 | 7.16 | 7.3 | 7.45 | 7.59 | 7.73 | 7.88 | 8.02 | 8.16 | 8.31 | 8.45 | 8.59 | 8.74 | 8.88 | 9.02 | 9.16 | 9.31 | 9.45 | 9.59 | 9.74 | 9.88 | 10.02 | 10.17 | 10.31 | 10.45 | 10.6 | 10.74 | 10.88 | 11.03 | 11.17 | 11.31
12.16 | 5.58 | 5.72 | 5.87 | 6.01 | 6.15 | 6.3 | 6.44 | 6.58 | 6.73 | 6.87 | 7.01 | 7.16 | 7.3 | 7.44 | 7.59 | 7.73 | 7.87 | 8.02 | 8.16 | 8.3 | 8.45 | 8.59 | 8.73 | 8.88 | 9.02 | 9.16 | 9.31 | 9.45 | 9.59 | 9.74 | 9.88 | 10.02 | 10.17 | 10.31 | 10.45 | 10.59 | 10.74 | 10.88 | 11.02 | 11.17 | 11.31
12.17 | 5.58 | 5.72 | 5.87 | 6.01 | 6.15 | 6.3 | 6.44 | 6.58 | 6.73 | 6.87 | 7.01 | 7.16 | 7.3 | 7.44 | 7.59 | 7.73 | 7.87 | 8.02 | 8.16 | 8.3 | 8.44 | 8.59 | 8.73 | 8.87 | 9.02 | 9.16 | 9.3 | 9.45 | 9.59 | 9.73 | 9.88 | 10.02 | 10.16 | 10.31 | 10.45 | 10.59 | 10.74 | 10.88 | 11.02 | 11.16 | 11.31
12.18 | 5.58 | 5.72 | 5.87 | 6.01 | 6.15 | 6.29 | 6.44 | 6.58 | 6.72 | 6.87 | 7.01 | 7.15 | 7.3 | 7.44 | 7.58 | 7.73 | 7.87 | 8.01 | 8.16 | 8.3 | 8.44 | 8.59 | 8.73 | 8.87 | 9.02 | 9.16 | 9.3 | 9.44 | 9.59 | 9.73 | 9.87 | 10.02 | 10.16 | 10.3 | 10.45 | 10.59 | 10.73 | 10.88 | 11.02 | 11.16 | 11.31
12.19 | 5.58 | 5.72 | 5.86 | 6.01 | 6.15 | 6.29 | 6.44 | 6.58 | 6.72 | 6.87 | 7.01 | 7.15 | 7.3 | 7.44 | 7.58 | 7.73 | 7.87 | 8.01 | 8.15 | 8.3 | 8.44 | 8.58 | 8.73 | 8.87 | 9.01 | 9.16 | 9.3 | 9.44 | 9.59 | 9.73 | 9.87 | 10.02 | 10.16 | 10.3 | 10.44 | 10.59 | 10.73 | 10.87 | 11.02 | 11.16 | 11.3
12.2 | 5.58 | 5.72 | 5.86 | 6.01 | 6.15 | 6.29 | 6.44 | 6.58 | 6.72 | 6.86 | 7.01 | 7.15 | 7.29 | 7.44 | 7.58 | 7.72 | 7.87 | 8.01 | 8.15 | 8.3 | 8.44 | 8.58 | 8.73 | 8.87 | 9.01 | 9.15 | 9.3 | 9.44 | 9.58 | 9.73 | 9.87 | 10.01 | 10.16 | 10.3 | 10.44 | 10.59 | 10.73 | 10.87 | 11.01 | 11.16 | 11.3
12.21 | 5.57 | 5.72 | 5.86 | 6 | 6.15 | 6.29 | 6.43 | 6.58 | 6.72 | 6.86 | 7.01 | 7.15 | 7.29 | 7.44 | 7.58 | 7.72 | 7.86 | 8.01 | 8.15 | 8.29 | 8.44 | 8.58 | 8.72 | 8.87 | 9.01 | 9.15 | 9.3 | 9.44 | 9.58 | 9.72 | 9.87 | 10.01 | 10.15 | 10.3 | 10.44 | 10.58 | 10.73 | 10.87 | 11.01 | 11.15 | 11.3
12.22 | 5.57 | 5.72 | 5.86 | 6 | 6.15 | 6.29 | 6.43 | 6.58 | 6.72 | 6.86 | 7 | 7.15 | 7.29 | 7.43 | 7.58 | 7.72 | 7.86 | 8.01 | 8.15 | 8.29 | 8.44 | 8.58 | 8.72 | 8.86 | 9.01 | 9.15 | 9.29 | 9.44 | 9.58 | 9.72 | 9.87 | 10.01 | 10.15 | 10.29 | 10.44 | 10.58 | 10.72 | 10.87 | 11.01 | 11.15 | 11.3
12.23 | 5.57 | 5.72 | 5.86 | 6 | 6.14 | 6.29 | 6.43 | 6.57 | 6.72 | 6.86 | 7 | 7.15 | 7.29 | 7.43 | 7.57 | 7.72 | 7.86 | 8 | 8.15 | 8.29 | 8.43 | 8.58 | 8.72 | 8.86 | 9.01 | 9.15 | 9.29 | 9.43 | 9.58 | 9.72 | 9.86 | 10.01 | 10.15 | 10.29 | 10.43 | 10.58 | 10.72 | 10.86 | 11.01 | 11.15 | 11.29
12.24 | 5.57 | 5.71 | 5.86 | 6 | 6.14 | 6.29 | 6.43 | 6.57 | 6.72 | 6.86 | 7 | 7.14 | 7.29 | 7.43 | 7.57 | 7.72 | 7.86 | 8 | 8.15 | 8.29 | 8.43 | 8.57 | 8.72 | 8.86 | 9 | 9.15 | 9.29 | 9.43 | 9.57 | 9.72 | 9.86 | 10 | 10.15 | 10.29 | 10.43 | 10.58 | 10.72 | 10.86 | 11 | 11.15 | 11.29
12.25 | 5.57 | 5.71 | 5.86 | 6 | 6.14 | 6.28 | 6.43 | 6.57 | 6.71 | 6.86 | 7 | 7.14 | 7.29 | 7.43 | 7.57 | 7.71 | 7.86 | 8 | 8.14 | 8.29 | 8.43 | 8.57 | 8.72 | 8.86 | 9 | 9.14 | 9.29 | 9.43 | 9.57 | 9.72 | 9.86 | 10 | 10.14 | 10.29 | 10.43 | 10.57 | 10.72 | 10.86 | 11 | 11.14 | 11.29
12.26 | 5.57 | 5.71 | 5.85 | 6 | 6.14 | 6.28 | 6.43 | 6.57 | 6.71 | 6.86 | 7 | 7.14 | 7.28 | 7.43 | 7.57 | 7.71 | 7.86 | 8 | 8.14 | 8.28 | 8.43 | 8.57 | 8.71 | 8.86 | 9 | 9.14 | 9.28 | 9.43 | 9.57 | 9.71 | 9.86 | 10 | 10.14 | 10.28 | 10.43 | 10.57 | 10.71 | 10.86 | 11 | 11.14 | 11.28
12.27 | 5.57 | 5.71 | 5.85 | 6 | 6.14 | 6.28 | 6.42 | 6.57 | 6.71 | 6.85 | 7 | 7.14 | 7.28 | 7.43 | 7.57 | 7.71 | 7.85 | 8 | 8.14 | 8.28 | 8.43 | 8.57 | 8.71 | 8.85 | 9 | 9.14 | 9.28 | 9.43 | 9.57 | 9.71 | 9.85 | 10 | 10.14 | 10.28 | 10.43 | 10.57 | 10.71 | 10.85 | 11 | 11.14 | 11.28
12.28 | 5.57 | 5.71 | 5.85 | 5.99 | 6.14 | 6.28 | 6.42 | 6.57 | 6.71 | 6.85 | 6.99 | 7.14 | 7.28 | 7.42 | 7.57 | 7.71 | 7.85 | 7.99 | 8.14 | 8.28 | 8.42 | 8.57 | 8.71 | 8.85 | 8.99 | 9.14 | 9.28 | 9.42 | 9.57 | 9.71 | 9.85 | 9.99 | 10.14 | 10.28 | 10.42 | 10.57 | 10.71 | 10.85 | 10.99 | 11.14 | 11.28
12.29 | 5.56 | 5.71 | 5.85 | 5.99 | 6.14 | 6.28 | 6.42 | 6.56 | 6.71 | 6.85 | 6.99 | 7.14 | 7.28 | 7.42 | 7.56 | 7.71 | 7.85 | 7.99 | 8.14 | 8.28 | 8.42 | 8.56 | 8.71 | 8.85 | 8.99 | 9.14 | 9.28 | 9.42 | 9.56 | 9.71 | 9.85 | 9.99 | 10.14 | 10.28 | 10.42 | 10.56 | 10.71 | 10.85 | 10.99 | 11.13 | 11.28
12.3 | 5.56 | 5.71 | 5.85 | 5.99 | 6.13 | 6.28 | 6.42 | 6.56 | 6.71 | 6.85 | 6.99 | 7.13 | 7.28 | 7.42 | 7.56 | 7.71 | 7.85 | 7.99 | 8.13 | 8.28 | 8.42 | 8.56 | 8.71 | 8.85 | 8.99 | 9.13 | 9.28 | 9.42 | 9.56 | 9.7 | 9.85 | 9.99 | 10.13 | 10.28 | 10.42 | 10.56 | 10.7 | 10.85 | 10.99 | 11.13 | 11.27
12.31 | 5.56 | 5.7 | 5.85 | 5.99 | 6.13 | 6.28 | 6.42 | 6.56 | 6.7 | 6.85 | 6.99 | 7.13 | 7.28 | 7.42 | 7.56 | 7.7 | 7.85 | 7.99 | 8.13 | 8.27 | 8.42 | 8.56 | 8.7 | 8.85 | 8.99 | 9.13 | 9.27 | 9.42 | 9.56 | 9.7 | 9.85 | 9.99 | 10.13 | 10.27 | 10.42 | 10.56 | 10.7 | 10.84 | 10.99 | 11.13 | 11.27
12.32 | 5.56 | 5.7 | 5.85 | 5.99 | 6.13 | 6.27 | 6.42 | 6.56 | 6.7 | 6.85 | 6.99 | 7.13 | 7.27 | 7.42 | 7.56 | 7.7 | 7.84 | 7.99 | 8.13 | 8.27 | 8.42 | 8.56 | 8.7 | 8.84 | 8.99 | 9.13 | 9.27 | 9.41 | 9.56 | 9.7 | 9.84 | 9.99 | 10.13 | 10.27 | 10.41 | 10.56 | 10.7 | 10.84 | 10.98 | 11.13 | 11.27
12.33 | 5.56 | 5.7 | 5.84 | 5.99 | 6.13 | 6.27 | 6.42 | 6.56 | 6.7 | 6.84 | 6.99 | 7.13 | 7.27 | 7.41 | 7.56 | 7.7 | 7.84 | 7.99 | 8.13 | 8.27 | 8.41 | 8.56 | 8.7 | 8.84 | 8.98 | 9.13 | 9.27 | 9.41 | 9.56 | 9.7 | 9.84 | 9.98 | 10.13 | 10.27 | 10.41 | 10.55 | 10.7 | 10.84 | 10.98 | 11.12 | 11.27
12.34 | 5.56 | 5.7 | 5.84 | 5.99 | 6.13 | 6.27 | 6.41 | 6.56 | 6.7 | 6.84 | 6.99 | 7.13 | 7.27 | 7.41 | 7.56 | 7.7 | 7.84 | 7.98 | 8.13 | 8.27 | 8.41 | 8.55 | 8.7 | 8.84 | 8.98 | 9.13 | 9.27 | 9.41 | 9.55 | 9.7 | 9.84 | 9.98 | 10.12 | 10.27 | 10.41 | 10.55 | 10.69 | 10.84 | 10.98 | 11.12 | 11.26
12.35 | 5.56 | 5.7 | 5.84 | 5.98 | 6.13 | 6.27 | 6.41 | 6.56 | 6.7 | 6.84 | 6.98 | 7.13 | 7.27 | 7.41 | 7.55 | 7.7 | 7.84 | 7.98 | 8.12 | 8.27 | 8.41 | 8.55 | 8.7 | 8.84 | 8.98 | 9.12 | 9.27 | 9.41 | 9.55 | 9.69 | 9.84 | 9.98 | 10.12 | 10.26 | 10.41 | 10.55 | 10.69 | 10.83 | 10.98 | 11.12 | 11.26
12.36 | 5.56 | 5.7 | 5.84 | 5.98 | 6.13 | 6.27 | 6.41 | 6.55 | 6.7 | 6.84 | 6.98 | 7.12 | 7.27 | 7.41 | 7.55 | 7.69 | 7.84 | 7.98 | 8.12 | 8.27 | 8.41 | 8.55 | 8.69 | 8.84 | 8.98 | 9.12 | 9.26 | 9.41 | 9.55 | 9.69 | 9.83 | 9.98 | 10.12 | 10.26 | 10.4 | 10.55 | 10.69 | 10.83 | 10.97 | 11.12 | 11.26
12.37 | 5.55 | 5.7 | 5.84 | 5.98 | 6.12 | 6.27 | 6.41 | 6.55 | 6.69 | 6.84 | 6.98 | 7.12 | 7.27 | 7.41 | 7.55 | 7.69 | 7.84 | 7.98 | 8.12 | 8.26 | 8.41 | 8.55 | 8.69 | 8.83 | 8.98 | 9.12 | 9.26 | 9.4 | 9.55 | 9.69 | 9.83 | 9.97 | 10.12 | 10.26 | 10.4 | 10.54 | 10.69 | 10.83 | 10.97 | 11.11 | 11.26
12.38 | 5.55 | 5.7 | 5.84 | 5.98 | 6.12 | 6.27 | 6.41 | 6.55 | 6.69 | 6.84 | 6.98 | 7.12 | 7.26 | 7.41 | 7.55 | 7.69 | 7.83 | 7.98 | 8.12 | 8.26 | 8.4 | 8.55 | 8.69 | 8.83 | 8.97 | 9.12 | 9.26 | 9.4 | 9.54 | 9.69 | 9.83 | 9.97 | 10.11 | 10.26 | 10.4 | 10.54 | 10.68 | 10.83 | 10.97 | 11.11 | 11.25
12.39 | 5.55 | 5.69 | 5.84 | 5.98 | 6.12 | 6.26 | 6.41 | 6.55 | 6.69 | 6.83 | 6.98 | 7.12 | 7.26 | 7.4 | 7.55 | 7.69 | 7.83 | 7.97 | 8.12 | 8.26 | 8.4 | 8.54 | 8.69 | 8.83 | 8.97 | 9.11 | 9.26 | 9.4 | 9.54 | 9.68 | 9.83 | 9.97 | 10.11 | 10.25 | 10.4 | 10.54 | 10.68 | 10.82 | 10.97 | 11.11 | 11.25
12.4 | 5.55 | 5.69 | 5.84 | 5.98 | 6.12 | 6.26 | 6.41 | 6.55 | 6.69 | 6.83 | 6.98 | 7.12 | 7.26 | 7.4 | 7.55 | 7.69 | 7.83 | 7.97 | 8.12 | 8.26 | 8.4 | 8.54 | 8.69 | 8.83 | 8.97 | 9.11 | 9.26 | 9.4 | 9.54 | 9.68 | 9.83 | 9.97 | 10.11 | 10.25 | 10.39 | 10.54 | 10.68 | 10.82 | 10.96 | 11.11 | 11.25
12.41 | 5.55 | 5.69 | 5.83 | 5.98 | 6.12 | 6.26 | 6.4 | 6.55 | 6.69 | 6.83 | 6.97 | 7.12 | 7.26 | 7.4 | 7.54 | 7.69 | 7.83 | 7.97 | 8.11 | 8.26 | 8.4 | 8.54 | 8.68 | 8.83 | 8.97 | 9.11 | 9.25 | 9.4 | 9.54 | 9.68 | 9.82 | 9.97 | 10.11 | 10.25 | 10.39 | 10.53 | 10.68 | 10.82 | 10.96 | 11.1 | 11.25
12.42 | 5.55 | 5.69 | 5.83 | 5.97 | 6.12 | 6.26 | 6.4 | 6.54 | 6.69 | 6.83 | 6.97 | 7.11 | 7.26 | 7.4 | 7.54 | 7.68 | 7.83 | 7.97 | 8.11 | 8.25 | 8.4 | 8.54 | 8.68 | 8.82 | 8.97 | 9.11 | 9.25 | 9.39 | 9.54 | 9.68 | 9.82 | 9.96 | 10.11 | 10.25 | 10.39 | 10.53 | 10.67 | 10.82 | 10.96 | 11.1 | 11.24
12.43 | 5.55 | 5.69 | 5.83 | 5.97 | 6.12 | 6.26 | 6.4 | 6.54 | 6.69 | 6.83 | 6.97 | 7.11 | 7.26 | 7.4 | 7.54 | 7.68 | 7.83 | 7.97 | 8.11 | 8.25 | 8.39 | 8.54 | 8.68 | 8.82 | 8.96 | 9.11 | 9.25 | 9.39 | 9.53 | 9.68 | 9.82 | 9.96 | 10.1 | 10.25 | 10.39 | 10.53 | 10.67 | 10.81 | 10.96 | 11.1 | 11.24
12.44 | 5.54 | 5.69 | 5.83 | 5.97 | 6.11 | 6.26 | 6.4 | 6.54 | 6.68 | 6.83 | 6.97 | 7.11 | 7.25 | 7.4 | 7.54 | 7.68 | 7.82 | 7.97 | 8.11 | 8.25 | 8.39 | 8.54 | 8.68 | 8.82 | 8.96 | 9.1 | 9.25 | 9.39 | 9.53 | 9.67 | 9.82 | 9.96 | 10.1 | 10.24 | 10.39 | 10.53 | 10.67 | 10.81 | 10.95 | 11.1 | 11.24
12.45 | 5.54 | 5.69 | 5.83 | 5.97 | 6.11 | 6.26 | 6.4 | 6.54 | 6.68 | 6.82 | 6.97 | 7.11 | 7.25 | 7.39 | 7.54 | 7.68 | 7.82 | 7.96 | 8.11 | 8.25 | 8.39 | 8.53 | 8.68 | 8.82 | 8.96 | 9.1 | 9.24 | 9.39 | 9.53 | 9.67 | 9.81 | 9.96 | 10.1 | 10.24 | 10.38 | 10.53 | 10.67 | 10.81 | 10.95 | 11.09 | 11.24
12.46 | 5.54 | 5.68 | 5.83 | 5.97 | 6.11 | 6.25 | 6.4 | 6.54 | 6.68 | 6.82 | 6.97 | 7.11 | 7.25 | 7.39 | 7.54 | 7.68 | 7.82 | 7.96 | 8.1 | 8.25 | 8.39 | 8.53 | 8.67 | 8.82 | 8.96 | 9.1 | 9.24 | 9.38 | 9.53 | 9.67 | 9.81 | 9.95 | 10.1 | 10.24 | 10.38 | 10.52 | 10.67 | 10.81 | 10.95 | 11.09 | 11.23
12.47 | 5.54 | 5.68 | 5.83 | 5.97 | 6.11 | 6.25 | 6.39 | 6.54 | 6.68 | 6.82 | 6.96 | 7.11 | 7.25 | 7.39 | 7.53 | 7.68 | 7.82 | 7.96 | 8.1 | 8.24 | 8.39 | 8.53 | 8.67 | 8.81 | 8.96 | 9.1 | 9.24 | 9.38 | 9.52 | 9.67 | 9.81 | 9.95 | 10.09 | 10.24 | 10.38 | 10.52 | 10.66 | 10.81 | 10.95 | 11.09 | 11.23
12.48 | 5.54 | 5.68 | 5.82 | 5.97 | 6.11 | 6.25 | 6.39 | 6.54 | 6.68 | 6.82 | 6.96 | 7.1 | 7.25 | 7.39 | 7.53 | 7.67 | 7.82 | 7.96 | 8.1 | 8.24 | 8.39 | 8.53 | 8.67 | 8.81 | 8.95 | 9.1 | 9.24 | 9.38 | 9.52 | 9.67 | 9.81 | 9.95 | 10.09 | 10.23 | 10.38 | 10.52 | 10.66 | 10.8 | 10.94 | 11.09 | 11.23
12.49 | 5.54 | 5.68 | 5.82 | 5.97 | 6.11 | 6.25 | 6.39 | 6.53 | 6.68 | 6.82 | 6.96 | 7.1 | 7.25 | 7.39 | 7.53 | 7.67 | 7.81 | 7.96 | 8.1 | 8.24 | 8.38 | 8.53 | 8.67 | 8.81 | 8.95 | 9.09 | 9.24 | 9.38 | 9.52 | 9.66 | 9.81 | 9.95 | 10.09 | 10.23 | 10.37 | 10.52 | 10.66 | 10.8 | 10.94 | 11.08 | 11.23
12.5 | 5.54 | 5.68 | 5.82 | 5.96 | 6.11 | 6.25 | 6.39 | 6.53 | 6.67 | 6.82 | 6.96 | 7.1 | 7.24 | 7.39 | 7.53 | 7.67 | 7.81 | 7.95 | 8.1 | 8.24 | 8.38 | 8.52 | 8.67 | 8.81 | 8.95 | 9.09 | 9.23 | 9.38 | 9.52 | 9.66 | 9.8 | 9.94 | 10.09 | 10.23 | 10.37 | 10.51 | 10.66 | 10.8 | 10.94 | 11.08 | 11.22
12.51 | 5.54 | 5.68 | 5.82 | 5.96 | 6.1 | 6.25 | 6.39 | 6.53 | 6.67 | 6.82 | 6.96 | 7.1 | 7.24 | 7.38 | 7.53 | 7.67 | 7.81 | 7.95 | 8.09 | 8.24 | 8.38 | 8.52 | 8.66 | 8.81 | 8.95 | 9.09 | 9.23 | 9.37 | 9.52 | 9.66 | 9.8 | 9.94 | 10.08 | 10.23 | 10.37 | 10.51 | 10.65 | 10.8 | 10.94 | 11.08 | 11.22
12.52 | 5.53 | 5.68 | 5.82 | 5.96 | 6.1 | 6.25 | 6.39 | 6.53 | 6.67 | 6.81 | 6.96 | 7.1 | 7.24 | 7.38 | 7.52 | 7.67 | 7.81 | 7.95 | 8.09 | 8.24 | 8.38 | 8.52 | 8.66 | 8.8 | 8.95 | 9.09 | 9.23 | 9.37 | 9.51 | 9.66 | 9.8 | 9.94 | 10.08 | 10.22 | 10.37 | 10.51 | 10.65 | 10.79 | 10.93 | 11.08 | 11.22
12.53 | 5.53 | 5.68 | 5.82 | 5.96 | 6.1 | 6.24 | 6.39 | 6.53 | 6.67 | 6.81 | 6.95 | 7.1 | 7.24 | 7.38 | 7.52 | 7.67 | 7.81 | 7.95 | 8.09 | 8.23 | 8.38 | 8.52 | 8.66 | 8.8 | 8.94 | 9.09 | 9.23 | 9.37 | 9.51 | 9.65 | 9.8 | 9.94 | 10.08 | 10.22 | 10.36 | 10.51 | 10.65 | 10.79 | 10.93 | 11.07 | 11.22
12.54 | 5.53 | 5.67 | 5.82 | 5.96 | 6.1 | 6.24 | 6.38 | 6.53 | 6.67 | 6.81 | 6.95 | 7.09 | 7.24 | 7.38 | 7.52 | 7.66 | 7.81 | 7.95 | 8.09 | 8.23 | 8.37 | 8.52 | 8.66 | 8.8 | 8.94 | 9.08 | 9.23 | 9.37 | 9.51 | 9.65 | 9.79 | 9.94 | 10.08 | 10.22 | 10.36 | 10.5 | 10.65 | 10.79 | 10.93 | 11.07 | 11.21
12.55 | 5.53 | 5.67 | 5.81 | 5.96 | 6.1 | 6.24 | 6.38 | 6.53 | 6.67 | 6.81 | 6.95 | 7.09 | 7.24 | 7.38 | 7.52 | 7.66 | 7.8 | 7.95 | 8.09 | 8.23 | 8.37 | 8.51 | 8.66 | 8.8 | 8.94 | 9.08 | 9.22 | 9.37 | 9.51 | 9.65 | 9.79 | 9.93 | 10.08 | 10.22 | 10.36 | 10.5 | 10.64 | 10.79 | 10.93 | 11.07 | 11.21
12.56 | 5.53 | 5.67 | 5.81 | 5.96 | 6.1 | 6.24 | 6.38 | 6.52 | 6.67 | 6.81 | 6.95 | 7.09 | 7.23 | 7.38 | 7.52 | 7.66 | 7.8 | 7.94 | 8.09 | 8.23 | 8.37 | 8.51 | 8.65 | 8.8 | 8.94 | 9.08 | 9.22 | 9.36 | 9.51 | 9.65 | 9.79 | 9.93 | 10.07 | 10.22 | 10.36 | 10.5 | 10.64 | 10.78 | 10.93 | 11.07 | 11.21
12.57 | 5.53 | 5.67 | 5.81 | 5.95 | 6.1 | 6.24 | 6.38 | 6.52 | 6.66 | 6.81 | 6.95 | 7.09 | 7.23 | 7.37 | 7.52 | 7.66 | 7.8 | 7.94 | 8.08 | 8.23 | 8.37 | 8.51 | 8.65 | 8.79 | 8.94 | 9.08 | 9.22 | 9.36 | 9.5 | 9.65 | 9.79 | 9.93 | 10.07 | 10.21 | 10.35 | 10.5 | 10.64 | 10.78 | 10.92 | 11.06 | 11.21
12.58 | 5.53 | 5.67 | 5.81 | 5.95 | 6.09 | 6.24 | 6.38 | 6.52 | 6.66 | 6.8 | 6.95 | 7.09 | 7.23 | 7.37 | 7.51 | 7.66 | 7.8 | 7.94 | 8.08 | 8.22 | 8.37 | 8.51 | 8.65 | 8.79 | 8.93 | 9.08 | 9.22 | 9.36 | 9.5 | 9.64 | 9.79 | 9.93 | 10.07 | 10.21 | 10.35 | 10.49 | 10.64 | 10.78 | 10.92 | 11.06 | 11.2
12.59 | 5.53 | 5.67 | 5.81 | 5.95 | 6.09 | 6.24 | 6.38 | 6.52 | 6.66 | 6.8 | 6.94 | 7.09 | 7.23 | 7.37 | 7.51 | 7.65 | 7.8 | 7.94 | 8.08 | 8.22 | 8.36 | 8.51 | 8.65 | 8.79 | 8.93 | 9.07 | 9.22 | 9.36 | 9.5 | 9.64 | 9.78 | 9.92 | 10.07 | 10.21 | 10.35 | 10.49 | 10.63 | 10.78 | 10.92 | 11.06 | 11.2
12.6 | 5.52 | 5.67 | 5.81 | 5.95 | 6.09 | 6.23 | 6.38 | 6.52 | 6.66 | 6.8 | 6.94 | 7.09 | 7.23 | 7.37 | 7.51 | 7.65 | 7.79 | 7.94 | 8.08 | 8.22 | 8.36 | 8.5 | 8.65 | 8.79 | 8.93 | 9.07 | 9.21 | 9.36 | 9.5 | 9.64 | 9.78 | 9.92 | 10.06 | 10.21 | 10.35 | 10.49 | 10.63 | 10.77 | 10.92 | 11.06 | 11.2
12.61 | 5.52 | 5.66 | 5.81 | 5.95 | 6.09 | 6.23 | 6.37 | 6.52 | 6.66 | 6.8 | 6.94 | 7.08 | 7.23 | 7.37 | 7.51 | 7.65 | 7.79 | 7.93 | 8.08 | 8.22 | 8.36 | 8.5 | 8.64 | 8.79 | 8.93 | 9.07 | 9.21 | 9.35 | 9.49 | 9.64 | 9.78 | 9.92 | 10.06 | 10.2 | 10.35 | 10.49 | 10.63 | 10.77 | 10.91 | 11.05 | 11.2
12.62 | 5.52 | 5.66 | 5.81 | 5.95 | 6.09 | 6.23 | 6.37 | 6.51 | 6.66 | 6.8 | 6.94 | 7.08 | 7.22 | 7.37 | 7.51 | 7.65 | 7.79 | 7.93 | 8.07 | 8.22 | 8.36 | 8.5 | 8.64 | 8.78 | 8.93 | 9.07 | 9.21 | 9.35 | 9.49 | 9.63 | 9.78 | 9.92 | 10.06 | 10.2 | 10.34 | 10.49 | 10.63 | 10.77 | 10.91 | 11.05 | 11.19
12.63 | 5.52 | 5.66 | 5.8 | 5.95 | 6.09 | 6.23 | 6.37 | 6.51 | 6.65 | 6.8 | 6.94 | 7.08 | 7.22 | 7.36 | 7.51 | 7.65 | 7.79 | 7.93 | 8.07 | 8.21 | 8.36 | 8.5 | 8.64 | 8.78 | 8.92 | 9.07 | 9.21 | 9.35 | 9.49 | 9.63 | 9.77 | 9.92 | 10.06 | 10.2 | 10.34 | 10.48 | 10.62 | 10.77 | 10.91 | 11.05 | 11.19
12.64 | 5.52 | 5.66 | 5.8 | 5.94 | 6.09 | 6.23 | 6.37 | 6.51 | 6.65 | 6.8 | 6.94 | 7.08 | 7.22 | 7.36 | 7.5 | 7.65 | 7.79 | 7.93 | 8.07 | 8.21 | 8.35 | 8.5 | 8.64 | 8.78 | 8.92 | 9.06 | 9.2 | 9.35 | 9.49 | 9.63 | 9.77 | 9.91 | 10.06 | 10.2 | 10.34 | 10.48 | 10.62 | 10.76 | 10.91 | 11.05 | 11.19
12.65 | 5.52 | 5.66 | 5.8 | 5.94 | 6.08 | 6.23 | 6.37 | 6.51 | 6.65 | 6.79 | 6.94 | 7.08 | 7.22 | 7.36 | 7.5 | 7.64 | 7.79 | 7.93 | 8.07 | 8.21 | 8.35 | 8.49 | 8.64 | 8.78 | 8.92 | 9.06 | 9.2 | 9.34 | 9.49 | 9.63 | 9.77 | 9.91 | 10.05 | 10.19 | 10.34 | 10.48 | 10.62 | 10.76 | 10.9 | 11.04 | 11.19
12.66 | 5.52 | 5.66 | 5.8 | 5.94 | 6.08 | 6.23 | 6.37 | 6.51 | 6.65 | 6.79 | 6.93 | 7.08 | 7.22 | 7.36 | 7.5 | 7.64 | 7.78 | 7.93 | 8.07 | 8.21 | 8.35 | 8.49 | 8.63 | 8.78 | 8.92 | 9.06 | 9.2 | 9.34 | 9.48 | 9.63 | 9.77 | 9.91 | 10.05 | 10.19 | 10.33 | 10.48 | 10.62 | 10.76 | 10.9 | 11.04 | 11.18
12.67 | 5.52 | 5.66 | 5.8 | 5.94 | 6.08 | 6.22 | 6.37 | 6.51 | 6.65 | 6.79 | 6.93 | 7.07 | 7.22 | 7.36 | 7.5 | 7.64 | 7.78 | 7.92 | 8.07 | 8.21 | 8.35 | 8.49 | 8.63 | 8.77 | 8.92 | 9.06 | 9.2 | 9.34 | 9.48 | 9.62 | 9.77 | 9.91 | 10.05 | 10.19 | 10.33 | 10.47 | 10.61 | 10.76 | 10.9 | 11.04 | 11.18
12.68 | 5.51 | 5.66 | 5.8 | 5.94 | 6.08 | 6.22 | 6.36 | 6.51 | 6.65 | 6.79 | 6.93 | 7.07 | 7.21 | 7.36 | 7.5 | 7.64 | 7.78 | 7.92 | 8.06 | 8.21 | 8.35 | 8.49 | 8.63 | 8.77 | 8.91 | 9.05 | 9.2 | 9.34 | 9.48 | 9.62 | 9.76 | 9.9 | 10.05 | 10.19 | 10.33 | 10.47 | 10.61 | 10.75 | 10.9 | 11.04 | 11.18
12.69 | 5.51 | 5.65 | 5.8 | 5.94 | 6.08 | 6.22 | 6.36 | 6.5 | 6.65 | 6.79 | 6.93 | 7.07 | 7.21 | 7.35 | 7.5 | 7.64 | 7.78 | 7.92 | 8.06 | 8.2 | 8.35 | 8.49 | 8.63 | 8.77 | 8.91 | 9.05 | 9.19 | 9.34 | 9.48 | 9.62 | 9.76 | 9.9 | 10.04 | 10.19 | 10.33 | 10.47 | 10.61 | 10.75 | 10.89 | 11.03 | 11.18
12.7 | 5.51 | 5.65 | 5.79 | 5.94 | 6.08 | 6.22 | 6.36 | 6.5 | 6.64 | 6.79 | 6.93 | 7.07 | 7.21 | 7.35 | 7.49 | 7.64 | 7.78 | 7.92 | 8.06 | 8.2 | 8.34 | 8.48 | 8.63 | 8.77 | 8.91 | 9.05 | 9.19 | 9.33 | 9.48 | 9.62 | 9.76 | 9.9 | 10.04 | 10.18 | 10.32 | 10.47 | 10.61 | 10.75 | 10.89 | 11.03 | 11.17
12.71 | 5.51 | 5.65 | 5.79 | 5.93 | 6.08 | 6.22 | 6.36 | 6.5 | 6.64 | 6.78 | 6.93 | 7.07 | 7.21 | 7.35 | 7.49 | 7.63 | 7.78 | 7.92 | 8.06 | 8.2 | 8.34 | 8.48 | 8.62 | 8.77 | 8.91 | 9.05 | 9.19 | 9.33 | 9.47 | 9.61 | 9.76 | 9.9 | 10.04 | 10.18 | 10.32 | 10.46 | 10.61 | 10.75 | 10.89 | 11.03 | 11.17
12.72 | 5.51 | 5.65 | 5.79 | 5.93 | 6.07 | 6.22 | 6.36 | 6.5 | 6.64 | 6.78 | 6.92 | 7.07 | 7.21 | 7.35 | 7.49 | 7.63 | 7.77 | 7.91 | 8.06 | 8.2 | 8.34 | 8.48 | 8.62 | 8.76 | 8.91 | 9.05 | 9.19 | 9.33 | 9.47 | 9.61 | 9.75 | 9.9 | 10.04 | 10.18 | 10.32 | 10.46 | 10.6 | 10.74 | 10.89 | 11.03 | 11.17
12.73 | 5.51 | 5.65 | 5.79 | 5.93 | 6.07 | 6.22 | 6.36 | 6.5 | 6.64 | 6.78 | 6.92 | 7.06 | 7.21 | 7.35 | 7.49 | 7.63 | 7.77 | 7.91 | 8.05 | 8.2 | 8.34 | 8.48 | 8.62 | 8.76 | 8.9 | 9.04 | 9.19 | 9.33 | 9.47 | 9.61 | 9.75 | 9.89 | 10.03 | 10.18 | 10.32 | 10.46 | 10.6 | 10.74 | 10.88 | 11.02 | 11.17
12.74 | 5.51 | 5.65 | 5.79 | 5.93 | 6.07 | 6.21 | 6.36 | 6.5 | 6.64 | 6.78 | 6.92 | 7.06 | 7.2 | 7.35 | 7.49 | 7.63 | 7.77 | 7.91 | 8.05 | 8.19 | 8.34 | 8.48 | 8.62 | 8.76 | 8.9 | 9.04 | 9.18 | 9.33 | 9.47 | 9.61 | 9.75 | 9.89 | 10.03 | 10.17 | 10.32 | 10.46 | 10.6 | 10.74 | 10.88 | 11.02 | 11.16
12.75 | 5.5 | 5.65 | 5.79 | 5.93 | 6.07 | 6.21 | 6.35 | 6.5 | 6.64 | 6.78 | 6.92 | 7.06 | 7.2 | 7.34 | 7.49 | 7.63 | 7.77 | 7.91 | 8.05 | 8.19 | 8.33 | 8.48 | 8.62 | 8.76 | 8.9 | 9.04 | 9.18 | 9.32 | 9.46 | 9.61 | 9.75 | 9.89 | 10.03 | 10.17 | 10.31 | 10.45 | 10.6 | 10.74 | 10.88 | 11.02 | 11.16
12.76 | 5.5 | 5.64 | 5.79 | 5.93 | 6.07 | 6.21 | 6.35 | 6.49 | 6.63 | 6.78 | 6.92 | 7.06 | 7.2 | 7.34 | 7.48 | 7.62 | 7.77 | 7.91 | 8.05 | 8.19 | 8.33 | 8.47 | 8.61 | 8.76 | 8.9 | 9.04 | 9.18 | 9.32 | 9.46 | 9.6 | 9.75 | 9.89 | 10.03 | 10.17 | 10.31 | 10.45 | 10.59 | 10.73 | 10.88 | 11.02 | 11.16
12.77 | 5.5 | 5.64 | 5.79 | 5.93 | 6.07 | 6.21 | 6.35 | 6.49 | 6.63 | 6.77 | 6.92 | 7.06 | 7.2 | 7.34 | 7.48 | 7.62 | 7.76 | 7.91 | 8.05 | 8.19 | 8.33 | 8.47 | 8.61 | 8.75 | 8.9 | 9.04 | 9.18 | 9.32 | 9.46 | 9.6 | 9.74 | 9.88 | 10.03 | 10.17 | 10.31 | 10.45 | 10.59 | 10.73 | 10.87 | 11.01 | 11.16
12.78 | 5.5 | 5.64 | 5.78 | 5.93 | 6.07 | 6.21 | 6.35 | 6.49 | 6.63 | 6.77 | 6.91 | 7.06 | 7.2 | 7.34 | 7.48 | 7.62 | 7.76 | 7.9 | 8.05 | 8.19 | 8.33 | 8.47 | 8.61 | 8.75 | 8.89 | 9.03 | 9.18 | 9.32 | 9.46 | 9.6 | 9.74 | 9.88 | 10.02 | 10.16 | 10.31 | 10.45 | 10.59 | 10.73 | 10.87 | 11.01 | 11.15
12.79 | 5.5 | 5.64 | 5.78 | 5.92 | 6.07 | 6.21 | 6.35 | 6.49 | 6.63 | 6.77 | 6.91 | 7.05 | 7.2 | 7.34 | 7.48 | 7.62 | 7.76 | 7.9 | 8.04 | 8.18 | 8.33 | 8.47 | 8.61 | 8.75 | 8.89 | 9.03 | 9.17 | 9.32 | 9.46 | 9.6 | 9.74 | 9.88 | 10.02 | 10.16 | 10.3 | 10.45 | 10.59 | 10.73 | 10.87 | 11.01 | 11.15
12.8 | 5.5 | 5.64 | 5.78 | 5.92 | 6.06 | 6.2 | 6.35 | 6.49 | 6.63 | 6.77 | 6.91 | 7.05 | 7.19 | 7.34 | 7.48 | 7.62 | 7.76 | 7.9 | 8.04 | 8.18 | 8.32 | 8.47 | 8.61 | 8.75 | 8.89 | 9.03 | 9.17 | 9.31 | 9.45 | 9.6 | 9.74 | 9.88 | 10.02 | 10.16 | 10.3 | 10.44 | 10.58 | 10.73 | 10.87 | 11.01 | 11.15
12.81 | 5.5 | 5.64 | 5.78 | 5.92 | 6.06 | 6.2 | 6.34 | 6.49 | 6.63 | 6.77 | 6.91 | 7.05 | 7.19 | 7.33 | 7.47 | 7.62 | 7.76 | 7.9 | 8.04 | 8.18 | 8.32 | 8.46 | 8.6 | 8.75 | 8.89 | 9.03 | 9.17 | 9.31 | 9.45 | 9.59 | 9.73 | 9.88 | 10.02 | 10.16 | 10.3 | 10.44 | 10.58 | 10.72 | 10.86 | 11.01 | 11.15
12.82 | 5.5 | 5.64 | 5.78 | 5.92 | 6.06 | 6.2 | 6.34 | 6.48 | 6.63 | 6.77 | 6.91 | 7.05 | 7.19 | 7.33 | 7.47 | 7.61 | 7.76 | 7.9 | 8.04 | 8.18 | 8.32 | 8.46 | 8.6 | 8.74 | 8.89 | 9.03 | 9.17 | 9.31 | 9.45 | 9.59 | 9.73 | 9.87 | 10.01 | 10.16 | 10.3 | 10.44 | 10.58 | 10.72 | 10.86 | 11 | 11.14
12.83 | 5.49 | 5.64 | 5.78 | 5.92 | 6.06 | 6.2 | 6.34 | 6.48 | 6.62 | 6.77 | 6.91 | 7.05 | 7.19 | 7.33 | 7.47 | 7.61 | 7.75 | 7.9 | 8.04 | 8.18 | 8.32 | 8.46 | 8.6 | 8.74 | 8.88 | 9.02 | 9.17 | 9.31 | 9.45 | 9.59 | 9.73 | 9.87 | 10.01 | 10.15 | 10.29 | 10.44 | 10.58 | 10.72 | 10.86 | 11 | 11.14
12.84 | 5.49 | 5.63 | 5.78 | 5.92 | 6.06 | 6.2 | 6.34 | 6.48 | 6.62 | 6.76 | 6.91 | 7.05 | 7.19 | 7.33 | 7.47 | 7.61 | 7.75 | 7.89 | 8.03 | 8.18 | 8.32 | 8.46 | 8.6 | 8.74 | 8.88 | 9.02 | 9.16 | 9.3 | 9.45 | 9.59 | 9.73 | 9.87 | 10.01 | 10.15 | 10.29 | 10.43 | 10.57 | 10.72 | 10.86 | 11 | 11.14
12.85 | 5.49 | 5.63 | 5.77 | 5.92 | 6.06 | 6.2 | 6.34 | 6.48 | 6.62 | 6.76 | 6.9 | 7.04 | 7.19 | 7.33 | 7.47 | 7.61 | 7.75 | 7.89 | 8.03 | 8.17 | 8.31 | 8.46 | 8.6 | 8.74 | 8.88 | 9.02 | 9.16 | 9.3 | 9.44 | 9.58 | 9.73 | 9.87 | 10.01 | 10.15 | 10.29 | 10.43 | 10.57 | 10.71 | 10.85 | 11 | 11.14
12.86 | 5.49 | 5.63 | 5.77 | 5.91 | 6.06 | 6.2 | 6.34 | 6.48 | 6.62 | 6.76 | 6.9 | 7.04 | 7.18 | 7.33 | 7.47 | 7.61 | 7.75 | 7.89 | 8.03 | 8.17 | 8.31 | 8.45 | 8.6 | 8.74 | 8.88 | 9.02 | 9.16 | 9.3 | 9.44 | 9.58 | 9.72 | 9.86 | 10.01 | 10.15 | 10.29 | 10.43 | 10.57 | 10.71 | 10.85 | 10.99 | 11.13
12.87 | 5.49 | 5.63 | 5.77 | 5.91 | 6.05 | 6.19 | 6.34 | 6.48 | 6.62 | 6.76 | 6.9 | 7.04 | 7.18 | 7.32 | 7.46 | 7.61 | 7.75 | 7.89 | 8.03 | 8.17 | 8.31 | 8.45 | 8.59 | 8.73 | 8.88 | 9.02 | 9.16 | 9.3 | 9.44 | 9.58 | 9.72 | 9.86 | 10 | 10.14 | 10.29 | 10.43 | 10.57 | 10.71 | 10.85 | 10.99 | 11.13
12.88 | 5.49 | 5.63 | 5.77 | 5.91 | 6.05 | 6.19 | 6.33 | 6.48 | 6.62 | 6.76 | 6.9 | 7.04 | 7.18 | 7.32 | 7.46 | 7.6 | 7.74 | 7.89 | 8.03 | 8.17 | 8.31 | 8.45 | 8.59 | 8.73 | 8.87 | 9.01 | 9.16 | 9.3 | 9.44 | 9.58 | 9.72 | 9.86 | 10 | 10.14 | 10.28 | 10.42 | 10.57 | 10.71 | 10.85 | 10.99 | 11.13
12.89 | 5.49 | 5.63 | 5.77 | 5.91 | 6.05 | 6.19 | 6.33 | 6.47 | 6.62 | 6.76 | 6.9 | 7.04 | 7.18 | 7.32 | 7.46 | 7.6 | 7.74 | 7.88 | 8.03 | 8.17 | 8.31 | 8.45 | 8.59 | 8.73 | 8.87 | 9.01 | 9.15 | 9.29 | 9.44 | 9.58 | 9.72 | 9.86 | 10 | 10.14 | 10.28 | 10.42 | 10.56 | 10.7 | 10.84 | 10.99 | 11.13
12.9 | 5.49 | 5.63 | 5.77 | 5.91 | 6.05 | 6.19 | 6.33 | 6.47 | 6.61 | 6.75 | 6.9 | 7.04 | 7.18 | 7.32 | 7.46 | 7.6 | 7.74 | 7.88 | 8.02 | 8.16 | 8.31 | 8.45 | 8.59 | 8.73 | 8.87 | 9.01 | 9.15 | 9.29 | 9.43 | 9.57 | 9.71 | 9.86 | 10 | 10.14 | 10.28 | 10.42 | 10.56 | 10.7 | 10.84 | 10.98 | 11.12
12.91 | 5.48 | 5.63 | 5.77 | 5.91 | 6.05 | 6.19 | 6.33 | 6.47 | 6.61 | 6.75 | 6.89 | 7.03 | 7.18 | 7.32 | 7.46 | 7.6 | 7.74 | 7.88 | 8.02 | 8.16 | 8.3 | 8.44 | 8.59 | 8.73 | 8.87 | 9.01 | 9.15 | 9.29 | 9.43 | 9.57 | 9.71 | 9.85 | 9.99 | 10.14 | 10.28 | 10.42 | 10.56 | 10.7 | 10.84 | 10.98 | 11.12
12.92 | 5.48 | 5.62 | 5.76 | 5.91 | 6.05 | 6.19 | 6.33 | 6.47 | 6.61 | 6.75 | 6.89 | 7.03 | 7.17 | 7.32 | 7.46 | 7.6 | 7.74 | 7.88 | 8.02 | 8.16 | 8.3 | 8.44 | 8.58 | 8.72 | 8.87 | 9.01 | 9.15 | 9.29 | 9.43 | 9.57 | 9.71 | 9.85 | 9.99 | 10.13 | 10.27 | 10.41 | 10.56 | 10.7 | 10.84 | 10.98 | 11.12
12.93 | 5.48 | 5.62 | 5.76 | 5.9 | 6.05 | 6.19 | 6.33 | 6.47 | 6.61 | 6.75 | 6.89 | 7.03 | 7.17 | 7.31 | 7.45 | 7.6 | 7.74 | 7.88 | 8.02 | 8.16 | 8.3 | 8.44 | 8.58 | 8.72 | 8.86 | 9 | 9.14 | 9.29 | 9.43 | 9.57 | 9.71 | 9.85 | 9.99 | 10.13 | 10.27 | 10.41 | 10.55 | 10.69 | 10.83 | 10.98 | 11.12
12.94 | 5.48 | 5.62 | 5.76 | 5.9 | 6.04 | 6.18 | 6.33 | 6.47 | 6.61 | 6.75 | 6.89 | 7.03 | 7.17 | 7.31 | 7.45 | 7.59 | 7.73 | 7.88 | 8.02 | 8.16 | 8.3 | 8.44 | 8.58 | 8.72 | 8.86 | 9 | 9.14 | 9.28 | 9.42 | 9.57 | 9.71 | 9.85 | 9.99 | 10.13 | 10.27 | 10.41 | 10.55 | 10.69 | 10.83 | 10.97 | 11.11
12.95 | 5.48 | 5.62 | 5.76 | 5.9 | 6.04 | 6.18 | 6.32 | 6.47 | 6.61 | 6.75 | 6.89 | 7.03 | 7.17 | 7.31 | 7.45 | 7.59 | 7.73 | 7.87 | 8.01 | 8.16 | 8.3 | 8.44 | 8.58 | 8.72 | 8.86 | 9 | 9.14 | 9.28 | 9.42 | 9.56 | 9.7 | 9.84 | 9.99 | 10.13 | 10.27 | 10.41 | 10.55 | 10.69 | 10.83 | 10.97 | 11.11
12.96 | 5.48 | 5.62 | 5.76 | 5.9 | 6.04 | 6.18 | 6.32 | 6.46 | 6.6 | 6.75 | 6.89 | 7.03 | 7.17 | 7.31 | 7.45 | 7.59 | 7.73 | 7.87 | 8.01 | 8.15 | 8.29 | 8.43 | 8.58 | 8.72 | 8.86 | 9 | 9.14 | 9.28 | 9.42 | 9.56 | 9.7 | 9.84 | 9.98 | 10.12 | 10.26 | 10.41 | 10.55 | 10.69 | 10.83 | 10.97 | 11.11
12.97 | 5.48 | 5.62 | 5.76 | 5.9 | 6.04 | 6.18 | 6.32 | 6.46 | 6.6 | 6.74 | 6.88 | 7.03 | 7.17 | 7.31 | 7.45 | 7.59 | 7.73 | 7.87 | 8.01 | 8.15 | 8.29 | 8.43 | 8.57 | 8.71 | 8.86 | 9 | 9.14 | 9.28 | 9.42 | 9.56 | 9.7 | 9.84 | 9.98 | 10.12 | 10.26 | 10.4 | 10.54 | 10.68 | 10.83 | 10.97 | 11.11
12.98 | 5.48 | 5.62 | 5.76 | 5.9 | 6.04 | 6.18 | 6.32 | 6.46 | 6.6 | 6.74 | 6.88 | 7.02 | 7.16 | 7.31 | 7.45 | 7.59 | 7.73 | 7.87 | 8.01 | 8.15 | 8.29 | 8.43 | 8.57 | 8.71 | 8.85 | 8.99 | 9.13 | 9.28 | 9.42 | 9.56 | 9.7 | 9.84 | 9.98 | 10.12 | 10.26 | 10.4 | 10.54 | 10.68 | 10.82 | 10.96 | 11.1
12.99 | 5.47 | 5.61 | 5.76 | 5.9 | 6.04 | 6.18 | 6.32 | 6.46 | 6.6 | 6.74 | 6.88 | 7.02 | 7.16 | 7.3 | 7.44 | 7.58 | 7.73 | 7.87 | 8.01 | 8.15 | 8.29 | 8.43 | 8.57 | 8.71 | 8.85 | 8.99 | 9.13 | 9.27 | 9.41 | 9.55 | 9.7 | 9.84 | 9.98 | 10.12 | 10.26 | 10.4 | 10.54 | 10.68 | 10.82 | 10.96 | 11.1
13 | 5.47 | 5.61 | 5.75 | 5.89 | 6.04 | 6.18 | 6.32 | 6.46 | 6.6 | 6.74 | 6.88 | 7.02 | 7.16 | 7.3 | 7.44 | 7.58 | 7.72 | 7.86 | 8.01 | 8.15 | 8.29 | 8.43 | 8.57 | 8.71 | 8.85 | 8.99 | 9.13 | 9.27 | 9.41 | 9.55 | 9.69 | 9.83 | 9.97 | 10.11 | 10.26 | 10.4 | 10.54 | 10.68 | 10.82 | 10.96 | 11.1
13.01 | 5.47 | 5.61 | 5.75 | 5.89 | 6.03 | 6.17 | 6.32 | 6.46 | 6.6 | 6.74 | 6.88 | 7.02 | 7.16 | 7.3 | 7.44 | 7.58 | 7.72 | 7.86 | 8 | 8.14 | 8.28 | 8.43 | 8.57 | 8.71 | 8.85 | 8.99 | 9.13 | 9.27 | 9.41 | 9.55 | 9.69 | 9.83 | 9.97 | 10.11 | 10.25 | 10.39 | 10.53 | 10.67 | 10.82 | 10.96 | 11.1
13.02 | 5.47 | 5.61 | 5.75 | 5.89 | 6.03 | 6.17 | 6.31 | 6.45 | 6.6 | 6.74 | 6.88 | 7.02 | 7.16 | 7.3 | 7.44 | 7.58 | 7.72 | 7.86 | 8 | 8.14 | 8.28 | 8.42 | 8.56 | 8.7 | 8.85 | 8.99 | 9.13 | 9.27 | 9.41 | 9.55 | 9.69 | 9.83 | 9.97 | 10.11 | 10.25 | 10.39 | 10.53 | 10.67 | 10.81 | 10.95 | 11.09
13.03 | 5.47 | 5.61 | 5.75 | 5.89 | 6.03 | 6.17 | 6.31 | 6.45 | 6.59 | 6.73 | 6.88 | 7.02 | 7.16 | 7.3 | 7.44 | 7.58 | 7.72 | 7.86 | 8 | 8.14 | 8.28 | 8.42 | 8.56 | 8.7 | 8.84 | 8.98 | 9.12 | 9.26 | 9.41 | 9.55 | 9.69 | 9.83 | 9.97 | 10.11 | 10.25 | 10.39 | 10.53 | 10.67 | 10.81 | 10.95 | 11.09
13.04 | 5.47 | 5.61 | 5.75 | 5.89 | 6.03 | 6.17 | 6.31 | 6.45 | 6.59 | 6.73 | 6.87 | 7.01 | 7.15 | 7.3 | 7.44 | 7.58 | 7.72 | 7.86 | 8 | 8.14 | 8.28 | 8.42 | 8.56 | 8.7 | 8.84 | 8.98 | 9.12 | 9.26 | 9.4 | 9.54 | 9.68 | 9.82 | 9.97 | 10.11 | 10.25 | 10.39 | 10.53 | 10.67 | 10.81 | 10.95 | 11.09
13.05 | 5.47 | 5.61 | 5.75 | 5.89 | 6.03 | 6.17 | 6.31 | 6.45 | 6.59 | 6.73 | 6.87 | 7.01 | 7.15 | 7.29 | 7.43 | 7.57 | 7.72 | 7.86 | 8 | 8.14 | 8.28 | 8.42 | 8.56 | 8.7 | 8.84 | 8.98 | 9.12 | 9.26 | 9.4 | 9.54 | 9.68 | 9.82 | 9.96 | 10.1 | 10.24 | 10.38 | 10.52 | 10.67 | 10.81 | 10.95 | 11.09
13.06 | 5.47 | 5.61 | 5.75 | 5.89 | 6.03 | 6.17 | 6.31 | 6.45 | 6.59 | 6.73 | 6.87 | 7.01 | 7.15 | 7.29 | 7.43 | 7.57 | 7.71 | 7.85 | 7.99 | 8.13 | 8.28 | 8.42 | 8.56 | 8.7 | 8.84 | 8.98 | 9.12 | 9.26 | 9.4 | 9.54 | 9.68 | 9.82 | 9.96 | 10.1 | 10.24 | 10.38 | 10.52 | 10.66 | 10.8 | 10.94 | 11.08
13.07 | 5.46 | 5.6 | 5.74 | 5.89 | 6.03 | 6.17 | 6.31 | 6.45 | 6.59 | 6.73 | 6.87 | 7.01 | 7.15 | 7.29 | 7.43 | 7.57 | 7.71 | 7.85 | 7.99 | 8.13 | 8.27 | 8.41 | 8.55 | 8.69 | 8.84 | 8.98 | 9.12 | 9.26 | 9.4 | 9.54 | 9.68 | 9.82 | 9.96 | 10.1 | 10.24 | 10.38 | 10.52 | 10.66 | 10.8 | 10.94 | 11.08
13.08 | 5.46 | 5.6 | 5.74 | 5.88 | 6.02 | 6.16 | 6.31 | 6.45 | 6.59 | 6.73 | 6.87 | 7.01 | 7.15 | 7.29 | 7.43 | 7.57 | 7.71 | 7.85 | 7.99 | 8.13 | 8.27 | 8.41 | 8.55 | 8.69 | 8.83 | 8.97 | 9.11 | 9.25 | 9.39 | 9.54 | 9.68 | 9.82 | 9.96 | 10.1 | 10.24 | 10.38 | 10.52 | 10.66 | 10.8 | 10.94 | 11.08
13.09 | 5.46 | 5.6 | 5.74 | 5.88 | 6.02 | 6.16 | 6.3 | 6.44 | 6.58 | 6.73 | 6.87 | 7.01 | 7.15 | 7.29 | 7.43 | 7.57 | 7.71 | 7.85 | 7.99 | 8.13 | 8.27 | 8.41 | 8.55 | 8.69 | 8.83 | 8.97 | 9.11 | 9.25 | 9.39 | 9.53 | 9.67 | 9.81 | 9.95 | 10.09 | 10.23 | 10.38 | 10.52 | 10.66 | 10.8 | 10.94 | 11.08
13.1 | 5.46 | 5.6 | 5.74 | 5.88 | 6.02 | 6.16 | 6.3 | 6.44 | 6.58 | 6.72 | 6.86 | 7 | 7.14 | 7.29 | 7.43 | 7.57 | 7.71 | 7.85 | 7.99 | 8.13 | 8.27 | 8.41 | 8.55 | 8.69 | 8.83 | 8.97 | 9.11 | 9.25 | 9.39 | 9.53 | 9.67 | 9.81 | 9.95 | 10.09 | 10.23 | 10.37 | 10.51 | 10.65 | 10.79 | 10.93 | 11.07
13.11 | 5.46 | 5.6 | 5.74 | 5.88 | 6.02 | 6.16 | 6.3 | 6.44 | 6.58 | 6.72 | 6.86 | 7 | 7.14 | 7.28 | 7.42 | 7.56 | 7.7 | 7.84 | 7.99 | 8.13 | 8.27 | 8.41 | 8.55 | 8.69 | 8.83 | 8.97 | 9.11 | 9.25 | 9.39 | 9.53 | 9.67 | 9.81 | 9.95 | 10.09 | 10.23 | 10.37 | 10.51 | 10.65 | 10.79 | 10.93 | 11.07
13.12 | 5.46 | 5.6 | 5.74 | 5.88 | 6.02 | 6.16 | 6.3 | 6.44 | 6.58 | 6.72 | 6.86 | 7 | 7.14 | 7.28 | 7.42 | 7.56 | 7.7 | 7.84 | 7.98 | 8.12 | 8.26 | 8.4 | 8.54 | 8.68 | 8.83 | 8.97 | 9.11 | 9.25 | 9.39 | 9.53 | 9.67 | 9.81 | 9.95 | 10.09 | 10.23 | 10.37 | 10.51 | 10.65 | 10.79 | 10.93 | 11.07
13.13 | 5.46 | 5.6 | 5.74 | 5.88 | 6.02 | 6.16 | 6.3 | 6.44 | 6.58 | 6.72 | 6.86 | 7 | 7.14 | 7.28 | 7.42 | 7.56 | 7.7 | 7.84 | 7.98 | 8.12 | 8.26 | 8.4 | 8.54 | 8.68 | 8.82 | 8.96 | 9.1 | 9.24 | 9.38 | 9.52 | 9.66 | 9.81 | 9.95 | 10.09 | 10.23 | 10.37 | 10.51 | 10.65 | 10.79 | 10.93 | 11.07
13.14 | 5.46 | 5.6 | 5.74 | 5.88 | 6.02 | 6.16 | 6.3 | 6.44 | 6.58 | 6.72 | 6.86 | 7 | 7.14 | 7.28 | 7.42 | 7.56 | 7.7 | 7.84 | 7.98 | 8.12 | 8.26 | 8.4 | 8.54 | 8.68 | 8.82 | 8.96 | 9.1 | 9.24 | 9.38 | 9.52 | 9.66 | 9.8 | 9.94 | 10.08 | 10.22 | 10.36 | 10.5 | 10.64 | 10.78 | 10.92 | 11.06
13.15 | 5.45 | 5.59 | 5.73 | 5.87 | 6.01 | 6.16 | 6.3 | 6.44 | 6.58 | 6.72 | 6.86 | 7 | 7.14 | 7.28 | 7.42 | 7.56 | 7.7 | 7.84 | 7.98 | 8.12 | 8.26 | 8.4 | 8.54 | 8.68 | 8.82 | 8.96 | 9.1 | 9.24 | 9.38 | 9.52 | 9.66 | 9.8 | 9.94 | 10.08 | 10.22 | 10.36 | 10.5 | 10.64 | 10.78 | 10.92 | 11.06
13.16 | 5.45 | 5.59 | 5.73 | 5.87 | 6.01 | 6.15 | 6.29 | 6.43 | 6.57 | 6.71 | 6.85 | 6.99 | 7.14 | 7.28 | 7.42 | 7.56 | 7.7 | 7.84 | 7.98 | 8.12 | 8.26 | 8.4 | 8.54 | 8.68 | 8.82 | 8.96 | 9.1 | 9.24 | 9.38 | 9.52 | 9.66 | 9.8 | 9.94 | 10.08 | 10.22 | 10.36 | 10.5 | 10.64 | 10.78 | 10.92 | 11.06
13.17 | 5.45 | 5.59 | 5.73 | 5.87 | 6.01 | 6.15 | 6.29 | 6.43 | 6.57 | 6.71 | 6.85 | 6.99 | 7.13 | 7.27 | 7.41 | 7.55 | 7.69 | 7.83 | 7.97 | 8.11 | 8.25 | 8.39 | 8.54 | 8.68 | 8.82 | 8.96 | 9.1 | 9.24 | 9.38 | 9.52 | 9.66 | 9.8 | 9.94 | 10.08 | 10.22 | 10.36 | 10.5 | 10.64 | 10.78 | 10.92 | 11.06
13.18 | 5.45 | 5.59 | 5.73 | 5.87 | 6.01 | 6.15 | 6.29 | 6.43 | 6.57 | 6.71 | 6.85 | 6.99 | 7.13 | 7.27 | 7.41 | 7.55 | 7.69 | 7.83 | 7.97 | 8.11 | 8.25 | 8.39 | 8.53 | 8.67 | 8.81 | 8.95 | 9.09 | 9.23 | 9.37 | 9.51 | 9.65 | 9.79 | 9.93 | 10.07 | 10.21 | 10.35 | 10.49 | 10.63 | 10.77 | 10.91 | 11.05
13.19 | 5.45 | 5.59 | 5.73 | 5.87 | 6.01 | 6.15 | 6.29 | 6.43 | 6.57 | 6.71 | 6.85 | 6.99 | 7.13 | 7.27 | 7.41 | 7.55 | 7.69 | 7.83 | 7.97 | 8.11 | 8.25 | 8.39 | 8.53 | 8.67 | 8.81 | 8.95 | 9.09 | 9.23 | 9.37 | 9.51 | 9.65 | 9.79 | 9.93 | 10.07 | 10.21 | 10.35 | 10.49 | 10.63 | 10.77 | 10.91 | 11.05
13.2 | 5.45 | 5.59 | 5.73 | 5.87 | 6.01 | 6.15 | 6.29 | 6.43 | 6.57 | 6.71 | 6.85 | 6.99 | 7.13 | 7.27 | 7.41 | 7.55 | 7.69 | 7.83 | 7.97 | 8.11 | 8.25 | 8.39 | 8.53 | 8.67 | 8.81 | 8.95 | 9.09 | 9.23 | 9.37 | 9.51 | 9.65 | 9.79 | 9.93 | 10.07 | 10.21 | 10.35 | 10.49 | 10.63 | 10.77 | 10.91 | 11.05
13.21 | 5.45 | 5.59 | 5.73 | 5.87 | 6.01 | 6.15 | 6.29 | 6.43 | 6.57 | 6.71 | 6.85 | 6.99 | 7.13 | 7.27 | 7.41 | 7.55 | 7.69 | 7.83 | 7.97 | 8.11 | 8.25 | 8.39 | 8.53 | 8.67 | 8.81 | 8.95 | 9.09 | 9.23 | 9.37 | 9.51 | 9.65 | 9.79 | 9.93 | 10.07 | 10.21 | 10.35 | 10.49 | 10.63 | 10.77 | 10.91 | 11.05
13.22 | 5.44 | 5.58 | 5.73 | 5.87 | 6.01 | 6.15 | 6.29 | 6.43 | 6.57 | 6.71 | 6.85 | 6.99 | 7.13 | 7.27 | 7.41 | 7.55 | 7.69 | 7.83 | 7.97 | 8.11 | 8.25 | 8.39 | 8.53 | 8.67 | 8.81 | 8.95 | 9.09 | 9.23 | 9.37 | 9.51 | 9.65 | 9.79 | 9.93 | 10.07 | 10.21 | 10.35 | 10.49 | 10.63 | 10.76 | 10.9 | 11.04
13.23 | 5.44 | 5.58 | 5.72 | 5.86 | 6 | 6.14 | 6.28 | 6.42 | 6.56 | 6.7 | 6.84 | 6.98 | 7.12 | 7.26 | 7.4 | 7.54 | 7.68 | 7.82 | 7.96 | 8.1 | 8.24 | 8.38 | 8.52 | 8.66 | 8.8 | 8.94 | 9.08 | 9.22 | 9.36 | 9.5 | 9.64 | 9.78 | 9.92 | 10.06 | 10.2 | 10.34 | 10.48 | 10.62 | 10.76 | 10.9 | 11.04
13.24 | 5.44 | 5.58 | 5.72 | 5.86 | 6 | 6.14 | 6.28 | 6.42 | 6.56 | 6.7 | 6.84 | 6.98 | 7.12 | 7.26 | 7.4 | 7.54 | 7.68 | 7.82 | 7.96 | 8.1 | 8.24 | 8.38 | 8.52 | 8.66 | 8.8 | 8.94 | 9.08 | 9.22 | 9.36 | 9.5 | 9.64 | 9.78 | 9.92 | 10.06 | 10.2 | 10.34 | 10.48 | 10.62 | 10.76 | 10.9 | 11.04
13.25 | 5.44 | 5.58 | 5.72 | 5.86 | 6 | 6.14 | 6.28 | 6.42 | 6.56 | 6.7 | 6.84 | 6.98 | 7.12 | 7.26 | 7.4 | 7.54 | 7.68 | 7.82 | 7.96 | 8.1 | 8.24 | 8.38 | 8.52 | 8.66 | 8.8 | 8.94 | 9.08 | 9.22 | 9.36 | 9.5 | 9.64 | 9.78 | 9.92 | 10.06 | 10.2 | 10.34 | 10.48 | 10.62 | 10.76 | 10.9 | 11.04
13.26 | 5.44 | 5.58 | 5.72 | 5.86 | 6 | 6.14 | 6.28 | 6.42 | 6.56 | 6.7 | 6.84 | 6.98 | 7.12 | 7.26 | 7.4 | 7.54 | 7.68 | 7.82 | 7.96 | 8.1 | 8.24 | 8.38 | 8.52 | 8.66 | 8.8 | 8.94 | 9.08 | 9.22 | 9.36 | 9.5 | 9.64 | 9.78 | 9.92 | 10.06 | 10.2 | 10.34 | 10.48 | 10.62 | 10.76 | 10.9 | 11.04
13.27 | 5.44 | 5.58 | 5.72 | 5.86 | 6 | 6.14 | 6.28 | 6.42 | 6.56 | 6.7 | 6.84 | 6.98 | 7.12 | 7.26 | 7.4 | 7.54 | 7.68 | 7.82 | 7.96 | 8.1 | 8.24 | 8.38 | 8.52 | 8.66 | 8.8 | 8.94 | 9.08 | 9.22 | 9.35 | 9.49 | 9.63 | 9.77 | 9.91 | 10.05 | 10.19 | 10.33 | 10.47 | 10.61 | 10.75 | 10.89 | 11.03
13.28 | 5.44 | 5.58 | 5.72 | 5.86 | 6 | 6.14 | 6.28 | 6.42 | 6.56 | 6.7 | 6.84 | 6.98 | 7.12 | 7.26 | 7.4 | 7.54 | 7.68 | 7.81 | 7.95 | 8.09 | 8.23 | 8.37 | 8.51 | 8.65 | 8.79 | 8.93 | 9.07 | 9.21 | 9.35 | 9.49 | 9.63 | 9.77 | 9.91 | 10.05 | 10.19 | 10.33 | 10.47 | 10.61 | 10.75 | 10.89 | 11.03
13.29 | 5.44 | 5.58 | 5.72 | 5.86 | 6 | 6.14 | 6.28 | 6.41 | 6.55 | 6.69 | 6.83 | 6.97 | 7.11 | 7.25 | 7.39 | 7.53 | 7.67 | 7.81 | 7.95 | 8.09 | 8.23 | 8.37 | 8.51 | 8.65 | 8.79 | 8.93 | 9.07 | 9.21 | 9.35 | 9.49 | 9.63 | 9.77 | 9.91 | 10.05 | 10.19 | 10.33 | 10.47 | 10.61 | 10.75 | 10.89 | 11.03
13.3 | 5.43 | 5.57 | 5.71 | 5.85 | 5.99 | 6.13 | 6.27 | 6.41 | 6.55 | 6.69 | 6.83 | 6.97 | 7.11 | 7.25 | 7.39 | 7.53 | 7.67 | 7.81 | 7.95 | 8.09 | 8.23 | 8.37 | 8.51 | 8.65 | 8.79 | 8.93 | 9.07 | 9.21 | 9.35 | 9.49 | 9.63 | 9.77 | 9.91 | 10.05 | 10.19 | 10.33 | 10.47 | 10.61 | 10.75 | 10.89 | 11.03
13.31 | 5.43 | 5.57 | 5.71 | 5.85 | 5.99 | 6.13 | 6.27 | 6.41 | 6.55 | 6.69 | 6.83 | 6.97 | 7.11 | 7.25 | 7.39 | 7.53 | 7.67 | 7.81 | 7.95 | 8.09 | 8.23 | 8.37 | 8.51 | 8.65 | 8.79 | 8.93 | 9.07 | 9.21 | 9.35 | 9.49 | 9.63 | 9.77 | 9.91 | 10.05 | 10.18 | 10.32 | 10.46 | 10.6 | 10.74 | 10.88 | 11.02
13.32 | 5.43 | 5.57 | 5.71 | 5.85 | 5.99 | 6.13 | 6.27 | 6.41 | 6.55 | 6.69 | 6.83 | 6.97 | 7.11 | 7.25 | 7.39 | 7.53 | 7.67 | 7.81 | 7.95 | 8.09 | 8.23 | 8.37 | 8.51 | 8.65 | 8.79 | 8.93 | 9.07 | 9.2 | 9.34 | 9.48 | 9.62 | 9.76 | 9.9 | 10.04 | 10.18 | 10.32 | 10.46 | 10.6 | 10.74 | 10.88 | 11.02
13.33 | 5.43 | 5.57 | 5.71 | 5.85 | 5.99 | 6.13 | 6.27 | 6.41 | 6.55 | 6.69 | 6.83 | 6.97 | 7.11 | 7.25 | 7.39 | 7.53 | 7.67 | 7.81 | 7.95 | 8.09 | 8.23 | 8.36 | 8.5 | 8.64 | 8.78 | 8.92 | 9.06 | 9.2 | 9.34 | 9.48 | 9.62 | 9.76 | 9.9 | 10.04 | 10.18 | 10.32 | 10.46 | 10.6 | 10.74 | 10.88 | 11.02
13.34 | 5.43 | 5.57 | 5.71 | 5.85 | 5.99 | 6.13 | 6.27 | 6.41 | 6.55 | 6.69 | 6.83 | 6.97 | 7.11 | 7.25 | 7.39 | 7.52 | 7.66 | 7.8 | 7.94 | 8.08 | 8.22 | 8.36 | 8.5 | 8.64 | 8.78 | 8.92 | 9.06 | 9.2 | 9.34 | 9.48 | 9.62 | 9.76 | 9.9 | 10.04 | 10.18 | 10.32 | 10.46 | 10.6 | 10.74 | 10.88 | 11.02
13.35 | 5.43 | 5.57 | 5.71 | 5.85 | 5.99 | 6.13 | 6.27 | 6.41 | 6.55 | 6.69 | 6.83 | 6.96 | 7.1 | 7.24 | 7.38 | 7.52 | 7.66 | 7.8 | 7.94 | 8.08 | 8.22 | 8.36 | 8.5 | 8.64 | 8.78 | 8.92 | 9.06 | 9.2 | 9.34 | 9.48 | 9.62 | 9.76 | 9.9 | 10.04 | 10.18 | 10.32 | 10.45 | 10.59 | 10.73 | 10.87 | 11.01
13.36 | 5.43 | 5.57 | 5.71 | 5.85 | 5.99 | 6.13 | 6.27 | 6.4 | 6.54 | 6.68 | 6.82 | 6.96 | 7.1 | 7.24 | 7.38 | 7.52 | 7.66 | 7.8 | 7.94 | 8.08 | 8.22 | 8.36 | 8.5 | 8.64 | 8.78 | 8.92 | 9.06 | 9.2 | 9.34 | 9.48 | 9.62 | 9.75 | 9.89 | 10.03 | 10.17 | 10.31 | 10.45 | 10.59 | 10.73 | 10.87 | 11.01
13.37 | 5.43 | 5.57 | 5.71 | 5.84 | 5.98 | 6.12 | 6.26 | 6.4 | 6.54 | 6.68 | 6.82 | 6.96 | 7.1 | 7.24 | 7.38 | 7.52 | 7.66 | 7.8 | 7.94 | 8.08 | 8.22 | 8.36 | 8.5 | 8.64 | 8.78 | 8.92 | 9.05 | 9.19 | 9.33 | 9.47 | 9.61 | 9.75 | 9.89 | 10.03 | 10.17 | 10.31 | 10.45 | 10.59 | 10.73 | 10.87 | 11.01
13.38 | 5.42 | 5.56 | 5.7 | 5.84 | 5.98 | 6.12 | 6.26 | 6.4 | 6.54 | 6.68 | 6.82 | 6.96 | 7.1 | 7.24 | 7.38 | 7.52 | 7.66 | 7.8 | 7.94 | 8.08 | 8.22 | 8.36 | 8.49 | 8.63 | 8.77 | 8.91 | 9.05 | 9.19 | 9.33 | 9.47 | 9.61 | 9.75 | 9.89 | 10.03 | 10.17 | 10.31 | 10.45 | 10.59 | 10.73 | 10.87 | 11.01
13.39 | 5.42 | 5.56 | 5.7 | 5.84 | 5.98 | 6.12 | 6.26 | 6.4 | 6.54 | 6.68 | 6.82 | 6.96 | 7.1 | 7.24 | 7.38 | 7.52 | 7.66 | 7.8 | 7.94 | 8.07 | 8.21 | 8.35 | 8.49 | 8.63 | 8.77 | 8.91 | 9.05 | 9.19 | 9.33 | 9.47 | 9.61 | 9.75 | 9.89 | 10.03 | 10.17 | 10.31 | 10.45 | 10.58 | 10.72 | 10.86 | 11
13.4 | 5.42 | 5.56 | 5.7 | 5.84 | 5.98 | 6.12 | 6.26 | 6.4 | 6.54 | 6.68 | 6.82 | 6.96 | 7.1 | 7.24 | 7.38 | 7.51 | 7.65 | 7.79 | 7.93 | 8.07 | 8.21 | 8.35 | 8.49 | 8.63 | 8.77 | 8.91 | 9.05 | 9.19 | 9.33 | 9.47 | 9.61 | 9.75 | 9.89 | 10.02 | 10.16 | 10.3 | 10.44 | 10.58 | 10.72 | 10.86 | 11
13.41 | 5.42 | 5.56 | 5.7 | 5.84 | 5.98 | 6.12 | 6.26 | 6.4 | 6.54 | 6.68 | 6.82 | 6.96 | 7.09 | 7.23 | 7.37 | 7.51 | 7.65 | 7.79 | 7.93 | 8.07 | 8.21 | 8.35 | 8.49 | 8.63 | 8.77 | 8.91 | 9.05 | 9.19 | 9.33 | 9.47 | 9.6 | 9.74 | 9.88 | 10.02 | 10.16 | 10.3 | 10.44 | 10.58 | 10.72 | 10.86 | 11
13.42 | 5.42 | 5.56 | 5.7 | 5.84 | 5.98 | 6.12 | 6.26 | 6.4 | 6.54 | 6.67 | 6.81 | 6.95 | 7.09 | 7.23 | 7.37 | 7.51 | 7.65 | 7.79 | 7.93 | 8.07 | 8.21 | 8.35 | 8.49 | 8.63 | 8.77 | 8.91 | 9.04 | 9.18 | 9.32 | 9.46 | 9.6 | 9.74 | 9.88 | 10.02 | 10.16 | 10.3 | 10.44 | 10.58 | 10.72 | 10.86 | 11
13.43 | 5.42 | 5.56 | 5.7 | 5.84 | 5.98 | 6.12 | 6.26 | 6.39 | 6.53 | 6.67 | 6.81 | 6.95 | 7.09 | 7.23 | 7.37 | 7.51 | 7.65 | 7.79 | 7.93 | 8.07 | 8.21 | 8.35 | 8.49 | 8.62 | 8.76 | 8.9 | 9.04 | 9.18 | 9.32 | 9.46 | 9.6 | 9.74 | 9.88 | 10.02 | 10.16 | 10.3 | 10.44 | 10.58 | 10.71 | 10.85 | 10.99
13.44 | 5.42 | 5.56 | 5.7 | 5.84 | 5.97 | 6.11 | 6.25 | 6.39 | 6.53 | 6.67 | 6.81 | 6.95 | 7.09 | 7.23 | 7.37 | 7.51 | 7.65 | 7.79 | 7.93 | 8.07 | 8.2 | 8.34 | 8.48 | 8.62 | 8.76 | 8.9 | 9.04 | 9.18 | 9.32 | 9.46 | 9.6 | 9.74 | 9.88 | 10.02 | 10.16 | 10.29 | 10.43 | 10.57 | 10.71 | 10.85 | 10.99
13.45 | 5.42 | 5.56 | 5.69 | 5.83 | 5.97 | 6.11 | 6.25 | 6.39 | 6.53 | 6.67 | 6.81 | 6.95 | 7.09 | 7.23 | 7.37 | 7.51 | 7.65 | 7.78 | 7.92 | 8.06 | 8.2 | 8.34 | 8.48 | 8.62 | 8.76 | 8.9 | 9.04 | 9.18 | 9.32 | 9.46 | 9.6 | 9.74 | 9.87 | 10.01 | 10.15 | 10.29 | 10.43 | 10.57 | 10.71 | 10.85 | 10.99
13.46 | 5.41 | 5.55 | 5.69 | 5.83 | 5.97 | 6.11 | 6.25 | 6.39 | 6.53 | 6.67 | 6.81 | 6.95 | 7.09 | 7.23 | 7.37 | 7.5 | 7.64 | 7.78 | 7.92 | 8.06 | 8.2 | 8.34 | 8.48 | 8.62 | 8.76 | 8.9 | 9.04 | 9.18 | 9.32 | 9.45 | 9.59 | 9.73 | 9.87 | 10.01 | 10.15 | 10.29 | 10.43 | 10.57 | 10.71 | 10.85 | 10.99
13.47 | 5.41 | 5.55 | 5.69 | 5.83 | 5.97 | 6.11 | 6.25 | 6.39 | 6.53 | 6.67 | 6.81 | 6.95 | 7.09 | 7.22 | 7.36 | 7.5 | 7.64 | 7.78 | 7.92 | 8.06 | 8.2 | 8.34 | 8.48 | 8.62 | 8.76 | 8.9 | 9.03 | 9.17 | 9.31 | 9.45 | 9.59 | 9.73 | 9.87 | 10.01 | 10.15 | 10.29 | 10.43 | 10.57 | 10.71 | 10.84 | 10.98
13.48 | 5.41 | 5.55 | 5.69 | 5.83 | 5.97 | 6.11 | 6.25 | 6.39 | 6.53 | 6.67 | 6.8 | 6.94 | 7.08 | 7.22 | 7.36 | 7.5 | 7.64 | 7.78 | 7.92 | 8.06 | 8.2 | 8.34 | 8.48 | 8.62 | 8.75 | 8.89 | 9.03 | 9.17 | 9.31 | 9.45 | 9.59 | 9.73 | 9.87 | 10.01 | 10.15 | 10.29 | 10.42 | 10.56 | 10.7 | 10.84 | 10.98
13.49 | 5.41 | 5.55 | 5.69 | 5.83 | 5.97 | 6.11 | 6.25 | 6.39 | 6.52 | 6.66 | 6.8 | 6.94 | 7.08 | 7.22 | 7.36 | 7.5 | 7.64 | 7.78 | 7.92 | 8.06 | 8.2 | 8.33 | 8.47 | 8.61 | 8.75 | 8.89 | 9.03 | 9.17 | 9.31 | 9.45 | 9.59 | 9.73 | 9.87 | 10 | 10.14 | 10.28 | 10.42 | 10.56 | 10.7 | 10.84 | 10.98
13.5 | 5.41 | 5.55 | 5.69 | 5.83 | 5.97 | 6.11 | 6.25 | 6.38 | 6.52 | 6.66 | 6.8 | 6.94 | 7.08 | 7.22 | 7.36 | 7.5 | 7.64 | 7.78 | 7.92 | 8.05 | 8.19 | 8.33 | 8.47 | 8.61 | 8.75 | 8.89 | 9.03 | 9.17 | 9.31 | 9.45 | 9.59 | 9.72 | 9.86 | 10 | 10.14 | 10.28 | 10.42 | 10.56 | 10.7 | 10.84 | 10.98
13.51 | 5.41 | 5.55 | 5.69 | 5.83 | 5.97 | 6.1 | 6.24 | 6.38 | 6.52 | 6.66 | 6.8 | 6.94 | 7.08 | 7.22 | 7.36 | 7.5 | 7.64 | 7.77 | 7.91 | 8.05 | 8.19 | 8.33 | 8.47 | 8.61 | 8.75 | 8.89 | 9.03 | 9.17 | 9.3 | 9.44 | 9.58 | 9.72 | 9.86 | 10 | 10.14 | 10.28 | 10.42 | 10.56 | 10.7 | 10.84 | 10.97
13.52 | 5.41 | 5.55 | 5.69 | 5.82 | 5.96 | 6.1 | 6.24 | 6.38 | 6.52 | 6.66 | 6.8 | 6.94 | 7.08 | 7.22 | 7.36 | 7.49 | 7.63 | 7.77 | 7.91 | 8.05 | 8.19 | 8.33 | 8.47 | 8.61 | 8.75 | 8.89 | 9.02 | 9.16 | 9.3 | 9.44 | 9.58 | 9.72 | 9.86 | 10 | 10.14 | 10.28 | 10.42 | 10.55 | 10.69 | 10.83 | 10.97
13.53 | 5.41 | 5.55 | 5.68 | 5.82 | 5.96 | 6.1 | 6.24 | 6.38 | 6.52 | 6.66 | 6.8 | 6.94 | 7.08 | 7.21 | 7.35 | 7.49 | 7.63 | 7.77 | 7.91 | 8.05 | 8.19 | 8.33 | 8.47 | 8.61 | 8.74 | 8.88 | 9.02 | 9.16 | 9.3 | 9.44 | 9.58 | 9.72 | 9.86 | 10 | 10.14 | 10.27 | 10.41 | 10.55 | 10.69 | 10.83 | 10.97
13.54 | 5.4 | 5.54 | 5.68 | 5.82 | 5.96 | 6.1 | 6.24 | 6.38 | 6.52 | 6.66 | 6.8 | 6.93 | 7.07 | 7.21 | 7.35 | 7.49 | 7.63 | 7.77 | 7.91 | 8.05 | 8.19 | 8.33 | 8.46 | 8.6 | 8.74 | 8.88 | 9.02 | 9.16 | 9.3 | 9.44 | 9.58 | 9.72 | 9.85 | 9.99 | 10.13 | 10.27 | 10.41 | 10.55 | 10.69 | 10.83 | 10.97
13.55 | 5.4 | 5.54 | 5.68 | 5.82 | 5.96 | 6.1 | 6.24 | 6.38 | 6.52 | 6.66 | 6.79 | 6.93 | 7.07 | 7.21 | 7.35 | 7.49 | 7.63 | 7.77 | 7.91 | 8.05 | 8.18 | 8.32 | 8.46 | 8.6 | 8.74 | 8.88 | 9.02 | 9.16 | 9.3 | 9.44 | 9.57 | 9.71 | 9.85 | 9.99 | 10.13 | 10.27 | 10.41 | 10.55 | 10.69 | 10.83 | 10.96
13.56 | 5.4 | 5.54 | 5.68 | 5.82 | 5.96 | 6.1 | 6.24 | 6.38 | 6.51 | 6.65 | 6.79 | 6.93 | 7.07 | 7.21 | 7.35 | 7.49 | 7.63 | 7.77 | 7.9 | 8.04 | 8.18 | 8.32 | 8.46 | 8.6 | 8.74 | 8.88 | 9.02 | 9.16 | 9.29 | 9.43 | 9.57 | 9.71 | 9.85 | 9.99 | 10.13 | 10.27 | 10.41 | 10.55 | 10.68 | 10.82 | 10.96
13.57 | 5.4 | 5.54 | 5.68 | 5.82 | 5.96 | 6.1 | 6.24 | 6.37 | 6.51 | 6.65 | 6.79 | 6.93 | 7.07 | 7.21 | 7.35 | 7.49 | 7.63 | 7.76 | 7.9 | 8.04 | 8.18 | 8.32 | 8.46 | 8.6 | 8.74 | 8.88 | 9.01 | 9.15 | 9.29 | 9.43 | 9.57 | 9.71 | 9.85 | 9.99 | 10.13 | 10.27 | 10.4 | 10.54 | 10.68 | 10.82 | 10.96
13.58 | 5.4 | 5.54 | 5.68 | 5.82 | 5.96 | 6.09 | 6.23 | 6.37 | 6.51 | 6.65 | 6.79 | 6.93 | 7.07 | 7.21 | 7.35 | 7.48 | 7.62 | 7.76 | 7.9 | 8.04 | 8.18 | 8.32 | 8.46 | 8.6 | 8.73 | 8.87 | 9.01 | 9.15 | 9.29 | 9.43 | 9.57 | 9.71 | 9.85 | 9.98 | 10.12 | 10.26 | 10.4 | 10.54 | 10.68 | 10.82 | 10.96
13.59 | 5.4 | 5.54 | 5.68 | 5.82 | 5.95 | 6.09 | 6.23 | 6.37 | 6.51 | 6.65 | 6.79 | 6.93 | 7.07 | 7.2 | 7.34 | 7.48 | 7.62 | 7.76 | 7.9 | 8.04 | 8.18 | 8.32 | 8.46 | 8.59 | 8.73 | 8.87 | 9.01 | 9.15 | 9.29 | 9.43 | 9.57 | 9.71 | 9.84 | 9.98 | 10.12 | 10.26 | 10.4 | 10.54 | 10.68 | 10.82 | 10.95
13.6 | 5.4 | 5.54 | 5.68 | 5.81 | 5.95 | 6.09 | 6.23 | 6.37 | 6.51 | 6.65 | 6.79 | 6.93 | 7.06 | 7.2 | 7.34 | 7.48 | 7.62 | 7.76 | 7.9 | 8.04 | 8.18 | 8.31 | 8.45 | 8.59 | 8.73 | 8.87 | 9.01 | 9.15 | 9.29 | 9.43 | 9.56 | 9.7 | 9.84 | 9.98 | 10.12 | 10.26 | 10.4 | 10.54 | 10.67 | 10.81 | 10.95
13.61 | 5.4 | 5.53 | 5.67 | 5.81 | 5.95 | 6.09 | 6.23 | 6.37 | 6.51 | 6.65 | 6.78 | 6.92 | 7.06 | 7.2 | 7.34 | 7.48 | 7.62 | 7.76 | 7.9 | 8.03 | 8.17 | 8.31 | 8.45 | 8.59 | 8.73 | 8.87 | 9.01 | 9.15 | 9.28 | 9.42 | 9.56 | 9.7 | 9.84 | 9.98 | 10.12 | 10.26 | 10.39 | 10.53 | 10.67 | 10.81 | 10.95
13.62 | 5.39 | 5.53 | 5.67 | 5.81 | 5.95 | 6.09 | 6.23 | 6.37 | 6.51 | 6.64 | 6.78 | 6.92 | 7.06 | 7.2 | 7.34 | 7.48 | 7.62 | 7.76 | 7.89 | 8.03 | 8.17 | 8.31 | 8.45 | 8.59 | 8.73 | 8.87 | 9 | 9.14 | 9.28 | 9.42 | 9.56 | 9.7 | 9.84 | 9.98 | 10.11 | 10.25 | 10.39 | 10.53 | 10.67 | 10.81 | 10.95
13.63 | 5.39 | 5.53 | 5.67 | 5.81 | 5.95 | 6.09 | 6.23 | 6.37 | 6.5 | 6.64 | 6.78 | 6.92 | 7.06 | 7.2 | 7.34 | 7.48 | 7.61 | 7.75 | 7.89 | 8.03 | 8.17 | 8.31 | 8.45 | 8.59 | 8.73 | 8.86 | 9 | 9.14 | 9.28 | 9.42 | 9.56 | 9.7 | 9.84 | 9.97 | 10.11 | 10.25 | 10.39 | 10.53 | 10.67 | 10.81 | 10.95
13.64 | 5.39 | 5.53 | 5.67 | 5.81 | 5.95 | 6.09 | 6.23 | 6.36 | 6.5 | 6.64 | 6.78 | 6.92 | 7.06 | 7.2 | 7.34 | 7.47 | 7.61 | 7.75 | 7.89 | 8.03 | 8.17 | 8.31 | 8.45 | 8.58 | 8.72 | 8.86 | 9 | 9.14 | 9.28 | 9.42 | 9.56 | 9.69 | 9.83 | 9.97 | 10.11 | 10.25 | 10.39 | 10.53 | 10.67 | 10.8 | 10.94
13.65 | 5.39 | 5.53 | 5.67 | 5.81 | 5.95 | 6.08 | 6.22 | 6.36 | 6.5 | 6.64 | 6.78 | 6.92 | 7.06 | 7.2 | 7.33 | 7.47 | 7.61 | 7.75 | 7.89 | 8.03 | 8.17 | 8.3 | 8.44 | 8.58 | 8.72 | 8.86 | 9 | 9.14 | 9.28 | 9.41 | 9.55 | 9.69 | 9.83 | 9.97 | 10.11 | 10.25 | 10.39 | 10.52 | 10.66 | 10.8 | 10.94
13.66 | 5.39 | 5.53 | 5.67 | 5.81 | 5.94 | 6.08 | 6.22 | 6.36 | 6.5 | 6.64 | 6.78 | 6.92 | 7.05 | 7.19 | 7.33 | 7.47 | 7.61 | 7.75 | 7.89 | 8.03 | 8.16 | 8.3 | 8.44 | 8.58 | 8.72 | 8.86 | 9 | 9.14 | 9.27 | 9.41 | 9.55 | 9.69 | 9.83 | 9.97 | 10.11 | 10.24 | 10.38 | 10.52 | 10.66 | 10.8 | 10.94
13.67 | 5.39 | 5.53 | 5.67 | 5.8 | 5.94 | 6.08 | 6.22 | 6.36 | 6.5 | 6.64 | 6.78 | 6.91 | 7.05 | 7.19 | 7.33 | 7.47 | 7.61 | 7.75 | 7.89 | 8.02 | 8.16 | 8.3 | 8.44 | 8.58 | 8.72 | 8.86 | 8.99 | 9.13 | 9.27 | 9.41 | 9.55 | 9.69 | 9.83 | 9.97 | 10.1 | 10.24 | 10.38 | 10.52 | 10.66 | 10.8 | 10.94
13.68 | 5.39 | 5.53 | 5.66 | 5.8 | 5.94 | 6.08 | 6.22 | 6.36 | 6.5 | 6.64 | 6.77 | 6.91 | 7.05 | 7.19 | 7.33 | 7.47 | 7.61 | 7.74 | 7.88 | 8.02 | 8.16 | 8.3 | 8.44 | 8.58 | 8.72 | 8.85 | 8.99 | 9.13 | 9.27 | 9.41 | 9.55 | 9.69 | 9.82 | 9.96 | 10.1 | 10.24 | 10.38 | 10.52 | 10.66 | 10.79 | 10.93
13.69 | 5.39 | 5.52 | 5.66 | 5.8 | 5.94 | 6.08 | 6.22 | 6.36 | 6.5 | 6.63 | 6.77 | 6.91 | 7.05 | 7.19 | 7.33 | 7.47 | 7.6 | 7.74 | 7.88 | 8.02 | 8.16 | 8.3 | 8.44 | 8.57 | 8.71 | 8.85 | 8.99 | 9.13 | 9.27 | 9.41 | 9.54 | 9.68 | 9.82 | 9.96 | 10.1 | 10.24 | 10.38 | 10.51 | 10.65 | 10.79 | 10.93
13.7 | 5.38 | 5.52 | 5.66 | 5.8 | 5.94 | 6.08 | 6.22 | 6.36 | 6.49 | 6.63 | 6.77 | 6.91 | 7.05 | 7.19 | 7.33 | 7.46 | 7.6 | 7.74 | 7.88 | 8.02 | 8.16 | 8.3 | 8.43 | 8.57 | 8.71 | 8.85 | 8.99 | 9.13 | 9.27 | 9.4 | 9.54 | 9.68 | 9.82 | 9.96 | 10.1 | 10.24 | 10.37 | 10.51 | 10.65 | 10.79 | 10.93
13.71 | 5.38 | 5.52 | 5.66 | 5.8 | 5.94 | 6.08 | 6.22 | 6.35 | 6.49 | 6.63 | 6.77 | 6.91 | 7.05 | 7.19 | 7.32 | 7.46 | 7.6 | 7.74 | 7.88 | 8.02 | 8.16 | 8.29 | 8.43 | 8.57 | 8.71 | 8.85 | 8.99 | 9.13 | 9.26 | 9.4 | 9.54 | 9.68 | 9.82 | 9.96 | 10.09 | 10.23 | 10.37 | 10.51 | 10.65 | 10.79 | 10.93
13.72 | 5.38 | 5.52 | 5.66 | 5.8 | 5.94 | 6.08 | 6.21 | 6.35 | 6.49 | 6.63 | 6.77 | 6.91 | 7.05 | 7.18 | 7.32 | 7.46 | 7.6 | 7.74 | 7.88 | 8.01 | 8.15 | 8.29 | 8.43 | 8.57 | 8.71 | 8.85 | 8.98 | 9.12 | 9.26 | 9.4 | 9.54 | 9.68 | 9.82 | 9.95 | 10.09 | 10.23 | 10.37 | 10.51 | 10.65 | 10.78 | 10.92
13.73 | 5.38 | 5.52 | 5.66 | 5.8 | 5.94 | 6.07 | 6.21 | 6.35 | 6.49 | 6.63 | 6.77 | 6.91 | 7.04 | 7.18 | 7.32 | 7.46 | 7.6 | 7.74 | 7.87 | 8.01 | 8.15 | 8.29 | 8.43 | 8.57 | 8.71 | 8.84 | 8.98 | 9.12 | 9.26 | 9.4 | 9.54 | 9.67 | 9.81 | 9.95 | 10.09 | 10.23 | 10.37 | 10.51 | 10.64 | 10.78 | 10.92
13.74 | 5.38 | 5.52 | 5.66 | 5.8 | 5.93 | 6.07 | 6.21 | 6.35 | 6.49 | 6.63 | 6.77 | 6.9 | 7.04 | 7.18 | 7.32 | 7.46 | 7.6 | 7.73 | 7.87 | 8.01 | 8.15 | 8.29 | 8.43 | 8.57 | 8.7 | 8.84 | 8.98 | 9.12 | 9.26 | 9.4 | 9.53 | 9.67 | 9.81 | 9.95 | 10.09 | 10.23 | 10.36 | 10.5 | 10.64 | 10.78 | 10.92
13.75 | 5.38 | 5.52 | 5.66 | 5.79 | 5.93 | 6.07 | 6.21 | 6.35 | 6.49 | 6.63 | 6.76 | 6.9 | 7.04 | 7.18 | 7.32 | 7.46 | 7.59 | 7.73 | 7.87 | 8.01 | 8.15 | 8.29 | 8.42 | 8.56 | 8.7 | 8.84 | 8.98 | 9.12 | 9.26 | 9.39 | 9.53 | 9.67 | 9.81 | 9.95 | 10.09 | 10.22 | 10.36 | 10.5 | 10.64 | 10.78 | 10.92
13.76 | 5.38 | 5.52 | 5.65 | 5.79 | 5.93 | 6.07 | 6.21 | 6.35 | 6.49 | 6.62 | 6.76 | 6.9 | 7.04 | 7.18 | 7.32 | 7.45 | 7.59 | 7.73 | 7.87 | 8.01 | 8.15 | 8.28 | 8.42 | 8.56 | 8.7 | 8.84 | 8.98 | 9.11 | 9.25 | 9.39 | 9.53 | 9.67 | 9.81 | 9.95 | 10.08 | 10.22 | 10.36 | 10.5 | 10.64 | 10.78 | 10.91
13.77 | 5.38 | 5.51 | 5.65 | 5.79 | 5.93 | 6.07 | 6.21 | 6.35 | 6.48 | 6.62 | 6.76 | 6.9 | 7.04 | 7.18 | 7.31 | 7.45 | 7.59 | 7.73 | 7.87 | 8.01 | 8.14 | 8.28 | 8.42 | 8.56 | 8.7 | 8.84 | 8.97 | 9.11 | 9.25 | 9.39 | 9.53 | 9.67 | 9.8 | 9.94 | 10.08 | 10.22 | 10.36 | 10.5 | 10.63 | 10.77 | 10.91
13.78 | 5.37 | 5.51 | 5.65 | 5.79 | 5.93 | 6.07 | 6.21 | 6.34 | 6.48 | 6.62 | 6.76 | 6.9 | 7.04 | 7.17 | 7.31 | 7.45 | 7.59 | 7.73 | 7.87 | 8 | 8.14 | 8.28 | 8.42 | 8.56 | 8.7 | 8.83 | 8.97 | 9.11 | 9.25 | 9.39 | 9.53 | 9.66 | 9.8 | 9.94 | 10.08 | 10.22 | 10.36 | 10.49 | 10.63 | 10.77 | 10.91
13.79 | 5.37 | 5.51 | 5.65 | 5.79 | 5.93 | 6.07 | 6.2 | 6.34 | 6.48 | 6.62 | 6.76 | 6.9 | 7.03 | 7.17 | 7.31 | 7.45 | 7.59 | 7.73 | 7.86 | 8 | 8.14 | 8.28 | 8.42 | 8.56 | 8.69 | 8.83 | 8.97 | 9.11 | 9.25 | 9.39 | 9.52 | 9.66 | 9.8 | 9.94 | 10.08 | 10.22 | 10.35 | 10.49 | 10.63 | 10.77 | 10.91
13.8 | 5.37 | 5.51 | 5.65 | 5.79 | 5.93 | 6.06 | 6.2 | 6.34 | 6.48 | 6.62 | 6.76 | 6.89 | 7.03 | 7.17 | 7.31 | 7.45 | 7.59 | 7.72 | 7.86 | 8 | 8.14 | 8.28 | 8.42 | 8.55 | 8.69 | 8.83 | 8.97 | 9.11 | 9.25 | 9.38 | 9.52 | 9.66 | 9.8 | 9.94 | 10.07 | 10.21 | 10.35 | 10.49 | 10.63 | 10.77 | 10.9
13.81 | 5.37 | 5.51 | 5.65 | 5.79 | 5.92 | 6.06 | 6.2 | 6.34 | 6.48 | 6.62 | 6.75 | 6.89 | 7.03 | 7.17 | 7.31 | 7.45 | 7.58 | 7.72 | 7.86 | 8 | 8.14 | 8.28 | 8.41 | 8.55 | 8.69 | 8.83 | 8.97 | 9.1 | 9.24 | 9.38 | 9.52 | 9.66 | 9.8 | 9.93 | 10.07 | 10.21 | 10.35 | 10.49 | 10.63 | 10.76 | 10.9
13.82 | 5.37 | 5.51 | 5.65 | 5.78 | 5.92 | 6.06 | 6.2 | 6.34 | 6.48 | 6.61 | 6.75 | 6.89 | 7.03 | 7.17 | 7.31 | 7.44 | 7.58 | 7.72 | 7.86 | 8 | 8.14 | 8.27 | 8.41 | 8.55 | 8.69 | 8.83 | 8.96 | 9.1 | 9.24 | 9.38 | 9.52 | 9.66 | 9.79 | 9.93 | 10.07 | 10.21 | 10.35 | 10.48 | 10.62 | 10.76 | 10.9
13.83 | 5.37 | 5.51 | 5.65 | 5.78 | 5.92 | 6.06 | 6.2 | 6.34 | 6.47 | 6.61 | 6.75 | 6.89 | 7.03 | 7.17 | 7.3 | 7.44 | 7.58 | 7.72 | 7.86 | 8 | 8.13 | 8.27 | 8.41 | 8.55 | 8.69 | 8.82 | 8.96 | 9.1 | 9.24 | 9.38 | 9.52 | 9.65 | 9.79 | 9.93 | 10.07 | 10.21 | 10.34 | 10.48 | 10.62 | 10.76 | 10.9
13.84 | 5.37 | 5.51 | 5.64 | 5.78 | 5.92 | 6.06 | 6.2 | 6.34 | 6.47 | 6.61 | 6.75 | 6.89 | 7.03 | 7.16 | 7.3 | 7.44 | 7.58 | 7.72 | 7.86 | 7.99 | 8.13 | 8.27 | 8.41 | 8.55 | 8.68 | 8.82 | 8.96 | 9.1 | 9.24 | 9.38 | 9.51 | 9.65 | 9.79 | 9.93 | 10.07 | 10.2 | 10.34 | 10.48 | 10.62 | 10.76 | 10.89
13.85 | 5.37 | 5.5 | 5.64 | 5.78 | 5.92 | 6.06 | 6.2 | 6.33 | 6.47 | 6.61 | 6.75 | 6.89 | 7.02 | 7.16 | 7.3 | 7.44 | 7.58 | 7.72 | 7.85 | 7.99 | 8.13 | 8.27 | 8.41 | 8.54 | 8.68 | 8.82 | 8.96 | 9.1 | 9.23 | 9.37 | 9.51 | 9.65 | 9.79 | 9.93 | 10.06 | 10.2 | 10.34 | 10.48 | 10.62 | 10.75 | 10.89
13.86 | 5.37 | 5.5 | 5.64 | 5.78 | 5.92 | 6.06 | 6.19 | 6.33 | 6.47 | 6.61 | 6.75 | 6.88 | 7.02 | 7.16 | 7.3 | 7.44 | 7.58 | 7.71 | 7.85 | 7.99 | 8.13 | 8.27 | 8.4 | 8.54 | 8.68 | 8.82 | 8.96 | 9.09 | 9.23 | 9.37 | 9.51 | 9.65 | 9.79 | 9.92 | 10.06 | 10.2 | 10.34 | 10.48 | 10.61 | 10.75 | 10.89
13.87 | 5.36 | 5.5 | 5.64 | 5.78 | 5.92 | 6.05 | 6.19 | 6.33 | 6.47 | 6.61 | 6.75 | 6.88 | 7.02 | 7.16 | 7.3 | 7.44 | 7.57 | 7.71 | 7.85 | 7.99 | 8.13 | 8.26 | 8.4 | 8.54 | 8.68 | 8.82 | 8.95 | 9.09 | 9.23 | 9.37 | 9.51 | 9.65 | 9.78 | 9.92 | 10.06 | 10.2 | 10.34 | 10.47 | 10.61 | 10.75 | 10.89
13.88 | 5.36 | 5.5 | 5.64 | 5.78 | 5.92 | 6.05 | 6.19 | 6.33 | 6.47 | 6.61 | 6.74 | 6.88 | 7.02 | 7.16 | 7.3 | 7.43 | 7.57 | 7.71 | 7.85 | 7.99 | 8.12 | 8.26 | 8.4 | 8.54 | 8.68 | 8.81 | 8.95 | 9.09 | 9.23 | 9.37 | 9.5 | 9.64 | 9.78 | 9.92 | 10.06 | 10.2 | 10.33 | 10.47 | 10.61 | 10.75 | 10.89
13.89 | 5.36 | 5.5 | 5.64 | 5.78 | 5.91 | 6.05 | 6.19 | 6.33 | 6.47 | 6.6 | 6.74 | 6.88 | 7.02 | 7.16 | 7.29 | 7.43 | 7.57 | 7.71 | 7.85 | 7.98 | 8.12 | 8.26 | 8.4 | 8.54 | 8.67 | 8.81 | 8.95 | 9.09 | 9.23 | 9.36 | 9.5 | 9.64 | 9.78 | 9.92 | 10.05 | 10.19 | 10.33 | 10.47 | 10.61 | 10.74 | 10.88
13.9 | 5.36 | 5.5 | 5.64 | 5.77 | 5.91 | 6.05 | 6.19 | 6.33 | 6.46 | 6.6 | 6.74 | 6.88 | 7.02 | 7.15 | 7.29 | 7.43 | 7.57 | 7.71 | 7.84 | 7.98 | 8.12 | 8.26 | 8.4 | 8.53 | 8.67 | 8.81 | 8.95 | 9.09 | 9.22 | 9.36 | 9.5 | 9.64 | 9.78 | 9.91 | 10.05 | 10.19 | 10.33 | 10.47 | 10.6 | 10.74 | 10.88
13.91 | 5.36 | 5.5 | 5.63 | 5.77 | 5.91 | 6.05 | 6.19 | 6.33 | 6.46 | 6.6 | 6.74 | 6.88 | 7.02 | 7.15 | 7.29 | 7.43 | 7.57 | 7.71 | 7.84 | 7.98 | 8.12 | 8.26 | 8.39 | 8.53 | 8.67 | 8.81 | 8.95 | 9.08 | 9.22 | 9.36 | 9.5 | 9.64 | 9.77 | 9.91 | 10.05 | 10.19 | 10.33 | 10.46 | 10.6 | 10.74 | 10.88
13.92 | 5.36 | 5.5 | 5.63 | 5.77 | 5.91 | 6.05 | 6.19 | 6.32 | 6.46 | 6.6 | 6.74 | 6.88 | 7.01 | 7.15 | 7.29 | 7.43 | 7.57 | 7.7 | 7.84 | 7.98 | 8.12 | 8.26 | 8.39 | 8.53 | 8.67 | 8.81 | 8.94 | 9.08 | 9.22 | 9.36 | 9.5 | 9.63 | 9.77 | 9.91 | 10.05 | 10.19 | 10.32 | 10.46 | 10.6 | 10.74 | 10.88
13.93 | 5.36 | 5.49 | 5.63 | 5.77 | 5.91 | 6.05 | 6.18 | 6.32 | 6.46 | 6.6 | 6.74 | 6.87 | 7.01 | 7.15 | 7.29 | 7.43 | 7.56 | 7.7 | 7.84 | 7.98 | 8.12 | 8.25 | 8.39 | 8.53 | 8.67 | 8.8 | 8.94 | 9.08 | 9.22 | 9.36 | 9.49 | 9.63 | 9.77 | 9.91 | 10.05 | 10.18 | 10.32 | 10.46 | 10.6 | 10.74 | 10.87
13.94 | 5.36 | 5.49 | 5.63 | 5.77 | 5.91 | 6.04 | 6.18 | 6.32 | 6.46 | 6.6 | 6.73 | 6.87 | 7.01 | 7.15 | 7.29 | 7.42 | 7.56 | 7.7 | 7.84 | 7.98 | 8.11 | 8.25 | 8.39 | 8.53 | 8.67 | 8.8 | 8.94 | 9.08 | 9.22 | 9.35 | 9.49 | 9.63 | 9.77 | 9.91 | 10.04 | 10.18 | 10.32 | 10.46 | 10.6 | 10.73 | 10.87
13.95 | 5.35 | 5.49 | 5.63 | 5.77 | 5.91 | 6.04 | 6.18 | 6.32 | 6.46 | 6.6 | 6.73 | 6.87 | 7.01 | 7.15 | 7.28 | 7.42 | 7.56 | 7.7 | 7.84 | 7.97 | 8.11 | 8.25 | 8.39 | 8.53 | 8.66 | 8.8 | 8.94 | 9.08 | 9.21 | 9.35 | 9.49 | 9.63 | 9.77 | 9.9 | 10.04 | 10.18 | 10.32 | 10.45 | 10.59 | 10.73 | 10.87
13.96 | 5.35 | 5.49 | 5.63 | 5.77 | 5.9 | 6.04 | 6.18 | 6.32 | 6.46 | 6.59 | 6.73 | 6.87 | 7.01 | 7.15 | 7.28 | 7.42 | 7.56 | 7.7 | 7.83 | 7.97 | 8.11 | 8.25 | 8.39 | 8.52 | 8.66 | 8.8 | 8.94 | 9.07 | 9.21 | 9.35 | 9.49 | 9.63 | 9.76 | 9.9 | 10.04 | 10.18 | 10.31 | 10.45 | 10.59 | 10.73 | 10.87
13.97 | 5.35 | 5.49 | 5.63 | 5.77 | 5.9 | 6.04 | 6.18 | 6.32 | 6.45 | 6.59 | 6.73 | 6.87 | 7.01 | 7.14 | 7.28 | 7.42 | 7.56 | 7.69 | 7.83 | 7.97 | 8.11 | 8.25 | 8.38 | 8.52 | 8.66 | 8.8 | 8.93 | 9.07 | 9.21 | 9.35 | 9.49 | 9.62 | 9.76 | 9.9 | 10.04 | 10.17 | 10.31 | 10.45 | 10.59 | 10.73 | 10.86
13.98 | 5.35 | 5.49 | 5.63 | 5.76 | 5.9 | 6.04 | 6.18 | 6.32 | 6.45 | 6.59 | 6.73 | 6.87 | 7 | 7.14 | 7.28 | 7.42 | 7.56 | 7.69 | 7.83 | 7.97 | 8.11 | 8.24 | 8.38 | 8.52 | 8.66 | 8.8 | 8.93 | 9.07 | 9.21 | 9.35 | 9.48 | 9.62 | 9.76 | 9.9 | 10.03 | 10.17 | 10.31 | 10.45 | 10.59 | 10.72 | 10.86
13.99 | 5.35 | 5.49 | 5.62 | 5.76 | 5.9 | 6.04 | 6.18 | 6.31 | 6.45 | 6.59 | 6.73 | 6.86 | 7 | 7.14 | 7.28 | 7.42 | 7.55 | 7.69 | 7.83 | 7.97 | 8.1 | 8.24 | 8.38 | 8.52 | 8.66 | 8.79 | 8.93 | 9.07 | 9.21 | 9.34 | 9.48 | 9.62 | 9.76 | 9.89 | 10.03 | 10.17 | 10.31 | 10.45 | 10.58 | 10.72 | 10.86
14 | 5.35 | 5.49 | 5.62 | 5.76 | 5.9 | 6.04 | 6.17 | 6.31 | 6.45 | 6.59 | 6.73 | 6.86 | 7 | 7.14 | 7.28 | 7.41 | 7.55 | 7.69 | 7.83 | 7.97 | 8.1 | 8.24 | 8.38 | 8.52 | 8.65 | 8.79 | 8.93 | 9.07 | 9.2 | 9.34 | 9.48 | 9.62 | 9.76 | 9.89 | 10.03 | 10.17 | 10.31 | 10.44 | 10.58 | 10.72 | 10.86
14.01 | 5.35 | 5.48 | 5.62 | 5.76 | 5.9 | 6.04 | 6.17 | 6.31 | 6.45 | 6.59 | 6.72 | 6.86 | 7 | 7.14 | 7.27 | 7.41 | 7.55 | 7.69 | 7.83 | 7.96 | 8.1 | 8.24 | 8.38 | 8.51 | 8.65 | 8.79 | 8.93 | 9.06 | 9.2 | 9.34 | 9.48 | 9.62 | 9.75 | 9.89 | 10.03 | 10.17 | 10.3 | 10.44 | 10.58 | 10.72 | 10.85
14.02 | 5.35 | 5.48 | 5.62 | 5.76 | 5.9 | 6.03 | 6.17 | 6.31 | 6.45 | 6.58 | 6.72 | 6.86 | 7 | 7.14 | 7.27 | 7.41 | 7.55 | 7.69 | 7.82 | 7.96 | 8.1 | 8.24 | 8.37 | 8.51 | 8.65 | 8.79 | 8.93 | 9.06 | 9.2 | 9.34 | 9.48 | 9.61 | 9.75 | 9.89 | 10.03 | 10.16 | 10.3 | 10.44 | 10.58 | 10.71 | 10.85
14.03 | 5.34 | 5.48 | 5.62 | 5.76 | 5.89 | 6.03 | 6.17 | 6.31 | 6.45 | 6.58 | 6.72 | 6.86 | 7 | 7.13 | 7.27 | 7.41 | 7.55 | 7.68 | 7.82 | 7.96 | 8.1 | 8.23 | 8.37 | 8.51 | 8.65 | 8.79 | 8.92 | 9.06 | 9.2 | 9.34 | 9.47 | 9.61 | 9.75 | 9.89 | 10.02 | 10.16 | 10.3 | 10.44 | 10.57 | 10.71 | 10.85
14.04 | 5.34 | 5.48 | 5.62 | 5.76 | 5.89 | 6.03 | 6.17 | 6.31 | 6.44 | 6.58 | 6.72 | 6.86 | 6.99 | 7.13 | 7.27 | 7.41 | 7.55 | 7.68 | 7.82 | 7.96 | 8.1 | 8.23 | 8.37 | 8.51 | 8.65 | 8.78 | 8.92 | 9.06 | 9.2 | 9.33 | 9.47 | 9.61 | 9.75 | 9.88 | 10.02 | 10.16 | 10.3 | 10.43 | 10.57 | 10.71 | 10.85
14.05 | 5.34 | 5.48 | 5.62 | 5.75 | 5.89 | 6.03 | 6.17 | 6.31 | 6.44 | 6.58 | 6.72 | 6.86 | 6.99 | 7.13 | 7.27 | 7.41 | 7.54 | 7.68 | 7.82 | 7.96 | 8.09 | 8.23 | 8.37 | 8.51 | 8.64 | 8.78 | 8.92 | 9.06 | 9.19 | 9.33 | 9.47 | 9.61 | 9.74 | 9.88 | 10.02 | 10.16 | 10.29 | 10.43 | 10.57 | 10.71 | 10.84
14.06 | 5.34 | 5.48 | 5.62 | 5.75 | 5.89 | 6.03 | 6.17 | 6.3 | 6.44 | 6.58 | 6.72 | 6.85 | 6.99 | 7.13 | 7.27 | 7.4 | 7.54 | 7.68 | 7.82 | 7.95 | 8.09 | 8.23 | 8.37 | 8.5 | 8.64 | 8.78 | 8.92 | 9.05 | 9.19 | 9.33 | 9.47 | 9.6 | 9.74 | 9.88 | 10.02 | 10.15 | 10.29 | 10.43 | 10.57 | 10.7 | 10.84
14.07 | 5.34 | 5.48 | 5.61 | 5.75 | 5.89 | 6.03 | 6.16 | 6.3 | 6.44 | 6.58 | 6.71 | 6.85 | 6.99 | 7.13 | 7.27 | 7.4 | 7.54 | 7.68 | 7.82 | 7.95 | 8.09 | 8.23 | 8.37 | 8.5 | 8.64 | 8.78 | 8.92 | 9.05 | 9.19 | 9.33 | 9.47 | 9.6 | 9.74 | 9.88 | 10.02 | 10.15 | 10.29 | 10.43 | 10.56 | 10.7 | 10.84
14.08 | 5.34 | 5.48 | 5.61 | 5.75 | 5.89 | 6.03 | 6.16 | 6.3 | 6.44 | 6.58 | 6.71 | 6.85 | 6.99 | 7.13 | 7.26 | 7.4 | 7.54 | 7.68 | 7.81 | 7.95 | 8.09 | 8.23 | 8.36 | 8.5 | 8.64 | 8.78 | 8.91 | 9.05 | 9.19 | 9.33 | 9.46 | 9.6 | 9.74 | 9.88 | 10.01 | 10.15 | 10.29 | 10.43 | 10.56 | 10.7 | 10.84
14.09 | 5.34 | 5.47 | 5.61 | 5.75 | 5.89 | 6.02 | 6.16 | 6.3 | 6.44 | 6.57 | 6.71 | 6.85 | 6.99 | 7.12 | 7.26 | 7.4 | 7.54 | 7.67 | 7.81 | 7.95 | 8.09 | 8.22 | 8.36 | 8.5 | 8.64 | 8.77 | 8.91 | 9.05 | 9.19 | 9.32 | 9.46 | 9.6 | 9.74 | 9.87 | 10.01 | 10.15 | 10.29 | 10.42 | 10.56 | 10.7 | 10.84
14.1 | 5.34 | 5.47 | 5.61 | 5.75 | 5.89 | 6.02 | 6.16 | 6.3 | 6.44 | 6.57 | 6.71 | 6.85 | 6.99 | 7.12 | 7.26 | 7.4 | 7.54 | 7.67 | 7.81 | 7.95 | 8.08 | 8.22 | 8.36 | 8.5 | 8.63 | 8.77 | 8.91 | 9.05 | 9.18 | 9.32 | 9.46 | 9.6 | 9.73 | 9.87 | 10.01 | 10.15 | 10.28 | 10.42 | 10.56 | 10.7 | 10.83
14.11 | 5.33 | 5.47 | 5.61 | 5.75 | 5.88 | 6.02 | 6.16 | 6.3 | 6.43 | 6.57 | 6.71 | 6.85 | 6.98 | 7.12 | 7.26 | 7.4 | 7.53 | 7.67 | 7.81 | 7.95 | 8.08 | 8.22 | 8.36 | 8.5 | 8.63 | 8.77 | 8.91 | 9.04 | 9.18 | 9.32 | 9.46 | 9.59 | 9.73 | 9.87 | 10.01 | 10.14 | 10.28 | 10.42 | 10.56 | 10.69 | 10.83
14.12 | 5.33 | 5.47 | 5.61 | 5.75 | 5.88 | 6.02 | 6.16 | 6.3 | 6.43 | 6.57 | 6.71 | 6.84 | 6.98 | 7.12 | 7.26 | 7.39 | 7.53 | 7.67 | 7.81 | 7.94 | 8.08 | 8.22 | 8.36 | 8.49 | 8.63 | 8.77 | 8.91 | 9.04 | 9.18 | 9.32 | 9.45 | 9.59 | 9.73 | 9.87 | 10 | 10.14 | 10.28 | 10.42 | 10.55 | 10.69 | 10.83
14.13 | 5.33 | 5.47 | 5.61 | 5.74 | 5.88 | 6.02 | 6.16 | 6.29 | 6.43 | 6.57 | 6.71 | 6.84 | 6.98 | 7.12 | 7.26 | 7.39 | 7.53 | 7.67 | 7.8 | 7.94 | 8.08 | 8.22 | 8.35 | 8.49 | 8.63 | 8.77 | 8.9 | 9.04 | 9.18 | 9.32 | 9.45 | 9.59 | 9.73 | 9.86 | 10 | 10.14 | 10.28 | 10.41 | 10.55 | 10.69 | 10.83
14.14 | 5.33 | 5.47 | 5.61 | 5.74 | 5.88 | 6.02 | 6.15 | 6.29 | 6.43 | 6.57 | 6.7 | 6.84 | 6.98 | 7.12 | 7.25 | 7.39 | 7.53 | 7.67 | 7.8 | 7.94 | 8.08 | 8.21 | 8.35 | 8.49 | 8.63 | 8.76 | 8.9 | 9.04 | 9.18 | 9.31 | 9.45 | 9.59 | 9.73 | 9.86 | 10 | 10.14 | 10.27 | 10.41 | 10.55 | 10.69 | 10.82
14.15 | 5.33 | 5.47 | 5.6 | 5.74 | 5.88 | 6.02 | 6.15 | 6.29 | 6.43 | 6.57 | 6.7 | 6.84 | 6.98 | 7.11 | 7.25 | 7.39 | 7.53 | 7.66 | 7.8 | 7.94 | 8.08 | 8.21 | 8.35 | 8.49 | 8.62 | 8.76 | 8.9 | 9.04 | 9.17 | 9.31 | 9.45 | 9.59 | 9.72 | 9.86 | 10 | 10.13 | 10.27 | 10.41 | 10.55 | 10.68 | 10.82
14.16 | 5.33 | 5.47 | 5.6 | 5.74 | 5.88 | 6.01 | 6.15 | 6.29 | 6.43 | 6.56 | 6.7 | 6.84 | 6.98 | 7.11 | 7.25 | 7.39 | 7.52 | 7.66 | 7.8 | 7.94 | 8.07 | 8.21 | 8.35 | 8.49 | 8.62 | 8.76 | 8.9 | 9.03 | 9.17 | 9.31 | 9.45 | 9.58 | 9.72 | 9.86 | 10 | 10.13 | 10.27 | 10.41 | 10.54 | 10.68 | 10.82
14.17 | 5.33 | 5.46 | 5.6 | 5.74 | 5.88 | 6.01 | 6.15 | 6.29 | 6.43 | 6.56 | 6.7 | 6.84 | 6.97 | 7.11 | 7.25 | 7.39 | 7.52 | 7.66 | 7.8 | 7.93 | 8.07 | 8.21 | 8.35 | 8.48 | 8.62 | 8.76 | 8.9 | 9.03 | 9.17 | 9.31 | 9.44 | 9.58 | 9.72 | 9.86 | 9.99 | 10.13 | 10.27 | 10.4 | 10.54 | 10.68 | 10.82
14.18 | 5.33 | 5.46 | 5.6 | 5.74 | 5.87 | 6.01 | 6.15 | 6.29 | 6.42 | 6.56 | 6.7 | 6.84 | 6.97 | 7.11 | 7.25 | 7.38 | 7.52 | 7.66 | 7.8 | 7.93 | 8.07 | 8.21 | 8.34 | 8.48 | 8.62 | 8.76 | 8.89 | 9.03 | 9.17 | 9.31 | 9.44 | 9.58 | 9.72 | 9.85 | 9.99 | 10.13 | 10.27 | 10.4 | 10.54 | 10.68 | 10.81
14.19 | 5.32 | 5.46 | 5.6 | 5.74 | 5.87 | 6.01 | 6.15 | 6.29 | 6.42 | 6.56 | 6.7 | 6.83 | 6.97 | 7.11 | 7.25 | 7.38 | 7.52 | 7.66 | 7.79 | 7.93 | 8.07 | 8.21 | 8.34 | 8.48 | 8.62 | 8.75 | 8.89 | 9.03 | 9.17 | 9.3 | 9.44 | 9.58 | 9.71 | 9.85 | 9.99 | 10.13 | 10.26 | 10.4 | 10.54 | 10.67 | 10.81
14.2 | 5.32 | 5.46 | 5.6 | 5.74 | 5.87 | 6.01 | 6.15 | 6.28 | 6.42 | 6.56 | 6.7 | 6.83 | 6.97 | 7.11 | 7.24 | 7.38 | 7.52 | 7.66 | 7.79 | 7.93 | 8.07 | 8.2 | 8.34 | 8.48 | 8.62 | 8.75 | 8.89 | 9.03 | 9.16 | 9.3 | 9.44 | 9.58 | 9.71 | 9.85 | 9.99 | 10.12 | 10.26 | 10.4 | 10.53 | 10.67 | 10.81
14.21 | 5.32 | 5.46 | 5.6 | 5.73 | 5.87 | 6.01 | 6.15 | 6.28 | 6.42 | 6.56 | 6.69 | 6.83 | 6.97 | 7.11 | 7.24 | 7.38 | 7.52 | 7.65 | 7.79 | 7.93 | 8.07 | 8.2 | 8.34 | 8.48 | 8.61 | 8.75 | 8.89 | 9.02 | 9.16 | 9.3 | 9.44 | 9.57 | 9.71 | 9.85 | 9.98 | 10.12 | 10.26 | 10.4 | 10.53 | 10.67 | 10.81
14.22 | 5.32 | 5.46 | 5.6 | 5.73 | 5.87 | 6.01 | 6.14 | 6.28 | 6.42 | 6.56 | 6.69 | 6.83 | 6.97 | 7.1 | 7.24 | 7.38 | 7.51 | 7.65 | 7.79 | 7.93 | 8.06 | 8.2 | 8.34 | 8.47 | 8.61 | 8.75 | 8.89 | 9.02 | 9.16 | 9.3 | 9.43 | 9.57 | 9.71 | 9.85 | 9.98 | 10.12 | 10.26 | 10.39 | 10.53 | 10.67 | 10.8
14.23 | 5.32 | 5.46 | 5.59 | 5.73 | 5.87 | 6.01 | 6.14 | 6.28 | 6.42 | 6.55 | 6.69 | 6.83 | 6.96 | 7.1 | 7.24 | 7.38 | 7.51 | 7.65 | 7.79 | 7.92 | 8.06 | 8.2 | 8.34 | 8.47 | 8.61 | 8.75 | 8.88 | 9.02 | 9.16 | 9.29 | 9.43 | 9.57 | 9.71 | 9.84 | 9.98 | 10.12 | 10.25 | 10.39 | 10.53 | 10.67 | 10.8
14.24 | 5.32 | 5.46 | 5.59 | 5.73 | 5.87 | 6 | 6.14 | 6.28 | 6.42 | 6.55 | 6.69 | 6.83 | 6.96 | 7.1 | 7.24 | 7.37 | 7.51 | 7.65 | 7.79 | 7.92 | 8.06 | 8.2 | 8.33 | 8.47 | 8.61 | 8.74 | 8.88 | 9.02 | 9.16 | 9.29 | 9.43 | 9.57 | 9.7 | 9.84 | 9.98 | 10.11 | 10.25 | 10.39 | 10.53 | 10.66 | 10.8
14.25 | 5.32 | 5.45 | 5.59 | 5.73 | 5.87 | 6 | 6.14 | 6.28 | 6.41 | 6.55 | 6.69 | 6.82 | 6.96 | 7.1 | 7.24 | 7.37 | 7.51 | 7.65 | 7.78 | 7.92 | 8.06 | 8.19 | 8.33 | 8.47 | 8.61 | 8.74 | 8.88 | 9.02 | 9.15 | 9.29 | 9.43 | 9.56 | 9.7 | 9.84 | 9.98 | 10.11 | 10.25 | 10.39 | 10.52 | 10.66 | 10.8
14.26 | 5.32 | 5.45 | 5.59 | 5.73 | 5.86 | 6 | 6.14 | 6.28 | 6.41 | 6.55 | 6.69 | 6.82 | 6.96 | 7.1 | 7.23 | 7.37 | 7.51 | 7.65 | 7.78 | 7.92 | 8.06 | 8.19 | 8.33 | 8.47 | 8.6 | 8.74 | 8.88 | 9.01 | 9.15 | 9.29 | 9.43 | 9.56 | 9.7 | 9.84 | 9.97 | 10.11 | 10.25 | 10.38 | 10.52 | 10.66 | 10.79
14.27 | 5.31 | 5.45 | 5.59 | 5.73 | 5.86 | 6 | 6.14 | 6.27 | 6.41 | 6.55 | 6.68 | 6.82 | 6.96 | 7.1 | 7.23 | 7.37 | 7.51 | 7.64 | 7.78 | 7.92 | 8.05 | 8.19 | 8.33 | 8.47 | 8.6 | 8.74 | 8.88 | 9.01 | 9.15 | 9.29 | 9.42 | 9.56 | 9.7 | 9.83 | 9.97 | 10.11 | 10.25 | 10.38 | 10.52 | 10.66 | 10.79
14.28 | 5.31 | 5.45 | 5.59 | 5.72 | 5.86 | 6 | 6.14 | 6.27 | 6.41 | 6.55 | 6.68 | 6.82 | 6.96 | 7.09 | 7.23 | 7.37 | 7.5 | 7.64 | 7.78 | 7.92 | 8.05 | 8.19 | 8.33 | 8.46 | 8.6 | 8.74 | 8.87 | 9.01 | 9.15 | 9.28 | 9.42 | 9.56 | 9.7 | 9.83 | 9.97 | 10.11 | 10.24 | 10.38 | 10.52 | 10.65 | 10.79
14.29 | 5.31 | 5.45 | 5.59 | 5.72 | 5.86 | 6 | 6.13 | 6.27 | 6.41 | 6.54 | 6.68 | 6.82 | 6.96 | 7.09 | 7.23 | 7.37 | 7.5 | 7.64 | 7.78 | 7.91 | 8.05 | 8.19 | 8.32 | 8.46 | 8.6 | 8.74 | 8.87 | 9.01 | 9.15 | 9.28 | 9.42 | 9.56 | 9.69 | 9.83 | 9.97 | 10.1 | 10.24 | 10.38 | 10.51 | 10.65 | 10.79
14.3 | 5.31 | 5.45 | 5.59 | 5.72 | 5.86 | 6 | 6.13 | 6.27 | 6.41 | 6.54 | 6.68 | 6.82 | 6.95 | 7.09 | 7.23 | 7.36 | 7.5 | 7.64 | 7.78 | 7.91 | 8.05 | 8.19 | 8.32 | 8.46 | 8.6 | 8.73 | 8.87 | 9.01 | 9.14 | 9.28 | 9.42 | 9.55 | 9.69 | 9.83 | 9.96 | 10.1 | 10.24 | 10.38 | 10.51 | 10.65 | 10.79
14.31 | 5.31 | 5.45 | 5.58 | 5.72 | 5.86 | 5.99 | 6.13 | 6.27 | 6.41 | 6.54 | 6.68 | 6.82 | 6.95 | 7.09 | 7.23 | 7.36 | 7.5 | 7.64 | 7.77 | 7.91 | 8.05 | 8.18 | 8.32 | 8.46 | 8.59 | 8.73 | 8.87 | 9 | 9.14 | 9.28 | 9.42 | 9.55 | 9.69 | 9.83 | 9.96 | 10.1 | 10.24 | 10.37 | 10.51 | 10.65 | 10.78
14.32 | 5.31 | 5.45 | 5.58 | 5.72 | 5.86 | 5.99 | 6.13 | 6.27 | 6.4 | 6.54 | 6.68 | 6.81 | 6.95 | 7.09 | 7.22 | 7.36 | 7.5 | 7.63 | 7.77 | 7.91 | 8.05 | 8.18 | 8.32 | 8.46 | 8.59 | 8.73 | 8.87 | 9 | 9.14 | 9.28 | 9.41 | 9.55 | 9.69 | 9.82 | 9.96 | 10.1 | 10.23 | 10.37 | 10.51 | 10.64 | 10.78
14.33 | 5.31 | 5.44 | 5.58 | 5.72 | 5.85 | 5.99 | 6.13 | 6.27 | 6.4 | 6.54 | 6.68 | 6.81 | 6.95 | 7.09 | 7.22 | 7.36 | 7.5 | 7.63 | 7.77 | 7.91 | 8.04 | 8.18 | 8.32 | 8.45 | 8.59 | 8.73 | 8.86 | 9 | 9.14 | 9.27 | 9.41 | 9.55 | 9.68 | 9.82 | 9.96 | 10.09 | 10.23 | 10.37 | 10.51 | 10.64 | 10.78
14.34 | 5.31 | 5.44 | 5.58 | 5.72 | 5.85 | 5.99 | 6.13 | 6.26 | 6.4 | 6.54 | 6.67 | 6.81 | 6.95 | 7.08 | 7.22 | 7.36 | 7.49 | 7.63 | 7.77 | 7.91 | 8.04 | 8.18 | 8.32 | 8.45 | 8.59 | 8.73 | 8.86 | 9 | 9.14 | 9.27 | 9.41 | 9.55 | 9.68 | 9.82 | 9.96 | 10.09 | 10.23 | 10.37 | 10.5 | 10.64 | 10.78
14.35 | 5.31 | 5.44 | 5.58 | 5.72 | 5.85 | 5.99 | 6.13 | 6.26 | 6.4 | 6.54 | 6.67 | 6.81 | 6.95 | 7.08 | 7.22 | 7.36 | 7.49 | 7.63 | 7.77 | 7.9 | 8.04 | 8.18 | 8.31 | 8.45 | 8.59 | 8.72 | 8.86 | 9 | 9.13 | 9.27 | 9.41 | 9.54 | 9.68 | 9.82 | 9.95 | 10.09 | 10.23 | 10.36 | 10.5 | 10.64 | 10.77
14.36 | 5.3 | 5.44 | 5.58 | 5.71 | 5.85 | 5.99 | 6.12 | 6.26 | 6.4 | 6.53 | 6.67 | 6.81 | 6.94 | 7.08 | 7.22 | 7.35 | 7.49 | 7.63 | 7.76 | 7.9 | 8.04 | 8.17 | 8.31 | 8.45 | 8.59 | 8.72 | 8.86 | 9 | 9.13 | 9.27 | 9.41 | 9.54 | 9.68 | 9.81 | 9.95 | 10.09 | 10.22 | 10.36 | 10.5 | 10.63 | 10.77
14.37 | 5.3 | 5.44 | 5.58 | 5.71 | 5.85 | 5.99 | 6.12 | 6.26 | 6.4 | 6.53 | 6.67 | 6.81 | 6.94 | 7.08 | 7.22 | 7.35 | 7.49 | 7.63 | 7.76 | 7.9 | 8.04 | 8.17 | 8.31 | 8.45 | 8.58 | 8.72 | 8.86 | 8.99 | 9.13 | 9.27 | 9.4 | 9.54 | 9.68 | 9.81 | 9.95 | 10.09 | 10.22 | 10.36 | 10.5 | 10.63 | 10.77
14.38 | 5.3 | 5.44 | 5.57 | 5.71 | 5.85 | 5.98 | 6.12 | 6.26 | 6.39 | 6.53 | 6.67 | 6.8 | 6.94 | 7.08 | 7.21 | 7.35 | 7.49 | 7.62 | 7.76 | 7.9 | 8.03 | 8.17 | 8.31 | 8.44 | 8.58 | 8.72 | 8.85 | 8.99 | 9.13 | 9.26 | 9.4 | 9.54 | 9.67 | 9.81 | 9.95 | 10.08 | 10.22 | 10.36 | 10.49 | 10.63 | 10.77
14.39 | 5.3 | 5.44 | 5.57 | 5.71 | 5.85 | 5.98 | 6.12 | 6.26 | 6.39 | 6.53 | 6.67 | 6.8 | 6.94 | 7.08 | 7.21 | 7.35 | 7.49 | 7.62 | 7.76 | 7.9 | 8.03 | 8.17 | 8.31 | 8.44 | 8.58 | 8.72 | 8.85 | 8.99 | 9.13 | 9.26 | 9.4 | 9.54 | 9.67 | 9.81 | 9.95 | 10.08 | 10.22 | 10.35 | 10.49 | 10.63 | 10.76
14.4 | 5.3 | 5.44 | 5.57 | 5.71 | 5.85 | 5.98 | 6.12 | 6.26 | 6.39 | 6.53 | 6.67 | 6.8 | 6.94 | 7.08 | 7.21 | 7.35 | 7.48 | 7.62 | 7.76 | 7.89 | 8.03 | 8.17 | 8.3 | 8.44 | 8.58 | 8.71 | 8.85 | 8.99 | 9.12 | 9.26 | 9.4 | 9.53 | 9.67 | 9.81 | 9.94 | 10.08 | 10.22 | 10.35 | 10.49 | 10.63 | 10.76
14.41 | 5.3 | 5.43 | 5.57 | 5.71 | 5.84 | 5.98 | 6.12 | 6.25 | 6.39 | 6.53 | 6.66 | 6.8 | 6.94 | 7.07 | 7.21 | 7.35 | 7.48 | 7.62 | 7.76 | 7.89 | 8.03 | 8.17 | 8.3 | 8.44 | 8.58 | 8.71 | 8.85 | 8.99 | 9.12 | 9.26 | 9.39 | 9.53 | 9.67 | 9.8 | 9.94 | 10.08 | 10.21 | 10.35 | 10.49 | 10.62 | 10.76
14.42 | 5.3 | 5.43 | 5.57 | 5.71 | 5.84 | 5.98 | 6.12 | 6.25 | 6.39 | 6.53 | 6.66 | 6.8 | 6.94 | 7.07 | 7.21 | 7.35 | 7.48 | 7.62 | 7.75 | 7.89 | 8.03 | 8.16 | 8.3 | 8.44 | 8.57 | 8.71 | 8.85 | 8.98 | 9.12 | 9.26 | 9.39 | 9.53 | 9.67 | 9.8 | 9.94 | 10.08 | 10.21 | 10.35 | 10.48 | 10.62 | 10.76
14.43 | 5.3 | 5.43 | 5.57 | 5.71 | 5.84 | 5.98 | 6.11 | 6.25 | 6.39 | 6.52 | 6.66 | 6.8 | 6.93 | 7.07 | 7.21 | 7.34 | 7.48 | 7.62 | 7.75 | 7.89 | 8.03 | 8.16 | 8.3 | 8.44 | 8.57 | 8.71 | 8.84 | 8.98 | 9.12 | 9.25 | 9.39 | 9.53 | 9.66 | 9.8 | 9.94 | 10.07 | 10.21 | 10.35 | 10.48 | 10.62 | 10.76
14.44 | 5.29 | 5.43 | 5.57 | 5.7 | 5.84 | 5.98 | 6.11 | 6.25 | 6.39 | 6.52 | 6.66 | 6.8 | 6.93 | 7.07 | 7.21 | 7.34 | 7.48 | 7.61 | 7.75 | 7.89 | 8.02 | 8.16 | 8.3 | 8.43 | 8.57 | 8.71 | 8.84 | 8.98 | 9.12 | 9.25 | 9.39 | 9.52 | 9.66 | 9.8 | 9.93 | 10.07 | 10.21 | 10.34 | 10.48 | 10.62 | 10.75
14.45 | 5.29 | 5.43 | 5.57 | 5.7 | 5.84 | 5.98 | 6.11 | 6.25 | 6.38 | 6.52 | 6.66 | 6.79 | 6.93 | 7.07 | 7.2 | 7.34 | 7.48 | 7.61 | 7.75 | 7.89 | 8.02 | 8.16 | 8.3 | 8.43 | 8.57 | 8.7 | 8.84 | 8.98 | 9.11 | 9.25 | 9.39 | 9.52 | 9.66 | 9.8 | 9.93 | 10.07 | 10.2 | 10.34 | 10.48 | 10.61 | 10.75
14.46 | 5.29 | 5.43 | 5.56 | 5.7 | 5.84 | 5.97 | 6.11 | 6.25 | 6.38 | 6.52 | 6.66 | 6.79 | 6.93 | 7.07 | 7.2 | 7.34 | 7.47 | 7.61 | 7.75 | 7.88 | 8.02 | 8.16 | 8.29 | 8.43 | 8.57 | 8.7 | 8.84 | 8.98 | 9.11 | 9.25 | 9.38 | 9.52 | 9.66 | 9.79 | 9.93 | 10.07 | 10.2 | 10.34 | 10.48 | 10.61 | 10.75
14.47 | 5.29 | 5.43 | 5.56 | 5.7 | 5.84 | 5.97 | 6.11 | 6.25 | 6.38 | 6.52 | 6.65 | 6.79 | 6.93 | 7.06 | 7.2 | 7.34 | 7.47 | 7.61 | 7.75 | 7.88 | 8.02 | 8.16 | 8.29 | 8.43 | 8.56 | 8.7 | 8.84 | 8.97 | 9.11 | 9.25 | 9.38 | 9.52 | 9.66 | 9.79 | 9.93 | 10.06 | 10.2 | 10.34 | 10.47 | 10.61 | 10.75
14.48 | 5.29 | 5.43 | 5.56 | 5.7 | 5.84 | 5.97 | 6.11 | 6.24 | 6.38 | 6.52 | 6.65 | 6.79 | 6.93 | 7.06 | 7.2 | 7.34 | 7.47 | 7.61 | 7.74 | 7.88 | 8.02 | 8.15 | 8.29 | 8.43 | 8.56 | 8.7 | 8.84 | 8.97 | 9.11 | 9.24 | 9.38 | 9.52 | 9.65 | 9.79 | 9.93 | 10.06 | 10.2 | 10.33 | 10.47 | 10.61 | 10.74
14.49 | 5.29 | 5.42 | 5.56 | 5.7 | 5.83 | 5.97 | 6.11 | 6.24 | 6.38 | 6.52 | 6.65 | 6.79 | 6.92 | 7.06 | 7.2 | 7.33 | 7.47 | 7.61 | 7.74 | 7.88 | 8.02 | 8.15 | 8.29 | 8.42 | 8.56 | 8.7 | 8.83 | 8.97 | 9.11 | 9.24 | 9.38 | 9.51 | 9.65 | 9.79 | 9.92 | 10.06 | 10.2 | 10.33 | 10.47 | 10.6 | 10.74
14.5 | 5.29 | 5.42 | 5.56 | 5.7 | 5.83 | 5.97 | 6.11 | 6.24 | 6.38 | 6.51 | 6.65 | 6.79 | 6.92 | 7.06 | 7.2 | 7.33 | 7.47 | 7.6 | 7.74 | 7.88 | 8.01 | 8.15 | 8.29 | 8.42 | 8.56 | 8.69 | 8.83 | 8.97 | 9.1 | 9.24 | 9.38 | 9.51 | 9.65 | 9.78 | 9.92 | 10.06 | 10.19 | 10.33 | 10.47 | 10.6 | 10.74
14.51 | 5.29 | 5.42 | 5.56 | 5.69 | 5.83 | 5.97 | 6.1 | 6.24 | 6.38 | 6.51 | 6.65 | 6.79 | 6.92 | 7.06 | 7.19 | 7.33 | 7.47 | 7.6 | 7.74 | 7.88 | 8.01 | 8.15 | 8.28 | 8.42 | 8.56 | 8.69 | 8.83 | 8.97 | 9.1 | 9.24 | 9.37 | 9.51 | 9.65 | 9.78 | 9.92 | 10.06 | 10.19 | 10.33 | 10.46 | 10.6 | 10.74
14.52 | 5.28 | 5.42 | 5.56 | 5.69 | 5.83 | 5.97 | 6.1 | 6.24 | 6.37 | 6.51 | 6.65 | 6.78 | 6.92 | 7.06 | 7.19 | 7.33 | 7.46 | 7.6 | 7.74 | 7.87 | 8.01 | 8.15 | 8.28 | 8.42 | 8.55 | 8.69 | 8.83 | 8.96 | 9.1 | 9.24 | 9.37 | 9.51 | 9.64 | 9.78 | 9.92 | 10.05 | 10.19 | 10.33 | 10.46 | 10.6 | 10.73
14.53 | 5.28 | 5.42 | 5.56 | 5.69 | 5.83 | 5.96 | 6.1 | 6.24 | 6.37 | 6.51 | 6.65 | 6.78 | 6.92 | 7.05 | 7.19 | 7.33 | 7.46 | 7.6 | 7.74 | 7.87 | 8.01 | 8.14 | 8.28 | 8.42 | 8.55 | 8.69 | 8.83 | 8.96 | 9.1 | 9.23 | 9.37 | 9.51 | 9.64 | 9.78 | 9.91 | 10.05 | 10.19 | 10.32 | 10.46 | 10.6 | 10.73
14.54 | 5.28 | 5.42 | 5.55 | 5.69 | 5.83 | 5.96 | 6.1 | 6.24 | 6.37 | 6.51 | 6.64 | 6.78 | 6.92 | 7.05 | 7.19 | 7.33 | 7.46 | 7.6 | 7.73 | 7.87 | 8.01 | 8.14 | 8.28 | 8.41 | 8.55 | 8.69 | 8.82 | 8.96 | 9.1 | 9.23 | 9.37 | 9.5 | 9.64 | 9.78 | 9.91 | 10.05 | 10.18 | 10.32 | 10.46 | 10.59 | 10.73
14.55 | 5.28 | 5.42 | 5.55 | 5.69 | 5.83 | 5.96 | 6.1 | 6.23 | 6.37 | 6.51 | 6.64 | 6.78 | 6.92 | 7.05 | 7.19 | 7.32 | 7.46 | 7.6 | 7.73 | 7.87 | 8 | 8.14 | 8.28 | 8.41 | 8.55 | 8.69 | 8.82 | 8.96 | 9.09 | 9.23 | 9.37 | 9.5 | 9.64 | 9.77 | 9.91 | 10.05 | 10.18 | 10.32 | 10.45 | 10.59 | 10.73
14.56 | 5.28 | 5.42 | 5.55 | 5.69 | 5.82 | 5.96 | 6.1 | 6.23 | 6.37 | 6.51 | 6.64 | 6.78 | 6.91 | 7.05 | 7.19 | 7.32 | 7.46 | 7.59 | 7.73 | 7.87 | 8 | 8.14 | 8.28 | 8.41 | 8.55 | 8.68 | 8.82 | 8.96 | 9.09 | 9.23 | 9.36 | 9.5 | 9.64 | 9.77 | 9.91 | 10.04 | 10.18 | 10.32 | 10.45 | 10.59 | 10.72
14.57 | 5.28 | 5.41 | 5.55 | 5.69 | 5.82 | 5.96 | 6.1 | 6.23 | 6.37 | 6.5 | 6.64 | 6.78 | 6.91 | 7.05 | 7.18 | 7.32 | 7.46 | 7.59 | 7.73 | 7.87 | 8 | 8.14 | 8.27 | 8.41 | 8.55 | 8.68 | 8.82 | 8.95 | 9.09 | 9.23 | 9.36 | 9.5 | 9.63 | 9.77 | 9.91 | 10.04 | 10.18 | 10.31 | 10.45 | 10.59 | 10.72
14.58 | 5.28 | 5.41 | 5.55 | 5.69 | 5.82 | 5.96 | 6.09 | 6.23 | 6.37 | 6.5 | 6.64 | 6.77 | 6.91 | 7.05 | 7.18 | 7.32 | 7.46 | 7.59 | 7.73 | 7.86 | 8 | 8.14 | 8.27 | 8.41 | 8.54 | 8.68 | 8.82 | 8.95 | 9.09 | 9.22 | 9.36 | 9.5 | 9.63 | 9.77 | 9.9 | 10.04 | 10.18 | 10.31 | 10.45 | 10.58 | 10.72
14.59 | 5.28 | 5.41 | 5.55 | 5.68 | 5.82 | 5.96 | 6.09 | 6.23 | 6.36 | 6.5 | 6.64 | 6.77 | 6.91 | 7.05 | 7.18 | 7.32 | 7.45 | 7.59 | 7.73 | 7.86 | 8 | 8.13 | 8.27 | 8.41 | 8.54 | 8.68 | 8.81 | 8.95 | 9.09 | 9.22 | 9.36 | 9.49 | 9.63 | 9.77 | 9.9 | 10.04 | 10.17 | 10.31 | 10.45 | 10.58 | 10.72
14.6 | 5.28 | 5.41 | 5.55 | 5.68 | 5.82 | 5.96 | 6.09 | 6.23 | 6.36 | 6.5 | 6.64 | 6.77 | 6.91 | 7.04 | 7.18 | 7.32 | 7.45 | 7.59 | 7.72 | 7.86 | 8 | 8.13 | 8.27 | 8.4 | 8.54 | 8.68 | 8.81 | 8.95 | 9.08 | 9.22 | 9.36 | 9.49 | 9.63 | 9.76 | 9.9 | 10.04 | 10.17 | 10.31 | 10.44 | 10.58 | 10.72
14.61 | 5.27 | 5.41 | 5.55 | 5.68 | 5.82 | 5.95 | 6.09 | 6.23 | 6.36 | 6.5 | 6.63 | 6.77 | 6.91 | 7.04 | 7.18 | 7.31 | 7.45 | 7.59 | 7.72 | 7.86 | 7.99 | 8.13 | 8.27 | 8.4 | 8.54 | 8.67 | 8.81 | 8.95 | 9.08 | 9.22 | 9.35 | 9.49 | 9.63 | 9.76 | 9.9 | 10.03 | 10.17 | 10.31 | 10.44 | 10.58 | 10.71
14.62 | 5.27 | 5.41 | 5.54 | 5.68 | 5.82 | 5.95 | 6.09 | 6.22 | 6.36 | 6.5 | 6.63 | 6.77 | 6.9 | 7.04 | 7.18 | 7.31 | 7.45 | 7.58 | 7.72 | 7.86 | 7.99 | 8.13 | 8.26 | 8.4 | 8.54 | 8.67 | 8.81 | 8.94 | 9.08 | 9.22 | 9.35 | 9.49 | 9.62 | 9.76 | 9.9 | 10.03 | 10.17 | 10.3 | 10.44 | 10.57 | 10.71
14.63 | 5.27 | 5.41 | 5.54 | 5.68 | 5.82 | 5.95 | 6.09 | 6.22 | 6.36 | 6.5 | 6.63 | 6.77 | 6.9 | 7.04 | 7.17 | 7.31 | 7.45 | 7.58 | 7.72 | 7.85 | 7.99 | 8.13 | 8.26 | 8.4 | 8.53 | 8.67 | 8.81 | 8.94 | 9.08 | 9.21 | 9.35 | 9.49 | 9.62 | 9.76 | 9.89 | 10.03 | 10.16 | 10.3 | 10.44 | 10.57 | 10.71
14.64 | 5.27 | 5.41 | 5.54 | 5.68 | 5.81 | 5.95 | 6.09 | 6.22 | 6.36 | 6.49 | 6.63 | 6.77 | 6.9 | 7.04 | 7.17 | 7.31 | 7.45 | 7.58 | 7.72 | 7.85 | 7.99 | 8.12 | 8.26 | 8.4 | 8.53 | 8.67 | 8.8 | 8.94 | 9.08 | 9.21 | 9.35 | 9.48 | 9.62 | 9.76 | 9.89 | 10.03 | 10.16 | 10.3 | 10.43 | 10.57 | 10.71
14.65 | 5.27 | 5.41 | 5.54 | 5.68 | 5.81 | 5.95 | 6.08 | 6.22 | 6.36 | 6.49 | 6.63 | 6.76 | 6.9 | 7.04 | 7.17 | 7.31 | 7.44 | 7.58 | 7.72 | 7.85 | 7.99 | 8.12 | 8.26 | 8.39 | 8.53 | 8.67 | 8.8 | 8.94 | 9.07 | 9.21 | 9.35 | 9.48 | 9.62 | 9.75 | 9.89 | 10.02 | 10.16 | 10.3 | 10.43 | 10.57 | 10.7
14.66 | 5.27 | 5.4 | 5.54 | 5.68 | 5.81 | 5.95 | 6.08 | 6.22 | 6.35 | 6.49 | 6.63 | 6.76 | 6.9 | 7.03 | 7.17 | 7.31 | 7.44 | 7.58 | 7.71 | 7.85 | 7.99 | 8.12 | 8.26 | 8.39 | 8.53 | 8.66 | 8.8 | 8.94 | 9.07 | 9.21 | 9.34 | 9.48 | 9.62 | 9.75 | 9.89 | 10.02 | 10.16 | 10.29 | 10.43 | 10.57 | 10.7
14.67 | 5.27 | 5.4 | 5.54 | 5.67 | 5.81 | 5.95 | 6.08 | 6.22 | 6.35 | 6.49 | 6.63 | 6.76 | 6.9 | 7.03 | 7.17 | 7.3 | 7.44 | 7.58 | 7.71 | 7.85 | 7.98 | 8.12 | 8.26 | 8.39 | 8.53 | 8.66 | 8.8 | 8.93 | 9.07 | 9.21 | 9.34 | 9.48 | 9.61 | 9.75 | 9.88 | 10.02 | 10.16 | 10.29 | 10.43 | 10.56 | 10.7
14.68 | 5.27 | 5.4 | 5.54 | 5.67 | 5.81 | 5.94 | 6.08 | 6.22 | 6.35 | 6.49 | 6.62 | 6.76 | 6.9 | 7.03 | 7.17 | 7.3 | 7.44 | 7.57 | 7.71 | 7.85 | 7.98 | 8.12 | 8.25 | 8.39 | 8.52 | 8.66 | 8.8 | 8.93 | 9.07 | 9.2 | 9.34 | 9.48 | 9.61 | 9.75 | 9.88 | 10.02 | 10.15 | 10.29 | 10.43 | 10.56 | 10.7
14.69 | 5.26 | 5.4 | 5.54 | 5.67 | 5.81 | 5.94 | 6.08 | 6.21 | 6.35 | 6.49 | 6.62 | 6.76 | 6.89 | 7.03 | 7.17 | 7.3 | 7.44 | 7.57 | 7.71 | 7.84 | 7.98 | 8.12 | 8.25 | 8.39 | 8.52 | 8.66 | 8.79 | 8.93 | 9.07 | 9.2 | 9.34 | 9.47 | 9.61 | 9.74 | 9.88 | 10.02 | 10.15 | 10.29 | 10.42 | 10.56 | 10.69
14.7 | 5.26 | 5.4 | 5.53 | 5.67 | 5.81 | 5.94 | 6.08 | 6.21 | 6.35 | 6.49 | 6.62 | 6.76 | 6.89 | 7.03 | 7.16 | 7.3 | 7.44 | 7.57 | 7.71 | 7.84 | 7.98 | 8.11 | 8.25 | 8.39 | 8.52 | 8.66 | 8.79 | 8.93 | 9.06 | 9.2 | 9.34 | 9.47 | 9.61 | 9.74 | 9.88 | 10.01 | 10.15 | 10.29 | 10.42 | 10.56 | 10.69
14.71 | 5.26 | 5.4 | 5.53 | 5.67 | 5.8 | 5.94 | 6.08 | 6.21 | 6.35 | 6.48 | 6.62 | 6.76 | 6.89 | 7.03 | 7.16 | 7.3 | 7.43 | 7.57 | 7.71 | 7.84 | 7.98 | 8.11 | 8.25 | 8.38 | 8.52 | 8.65 | 8.79 | 8.93 | 9.06 | 9.2 | 9.33 | 9.47 | 9.6 | 9.74 | 9.88 | 10.01 | 10.15 | 10.28 | 10.42 | 10.55 | 10.69
14.72 | 5.26 | 5.4 | 5.53 | 5.67 | 5.8 | 5.94 | 6.08 | 6.21 | 6.35 | 6.48 | 6.62 | 6.75 | 6.89 | 7.02 | 7.16 | 7.3 | 7.43 | 7.57 | 7.7 | 7.84 | 7.97 | 8.11 | 8.25 | 8.38 | 8.52 | 8.65 | 8.79 | 8.92 | 9.06 | 9.2 | 9.33 | 9.47 | 9.6 | 9.74 | 9.87 | 10.01 | 10.15 | 10.28 | 10.42 | 10.55 | 10.69
14.73 | 5.26 | 5.4 | 5.53 | 5.67 | 5.8 | 5.94 | 6.07 | 6.21 | 6.35 | 6.48 | 6.62 | 6.75 | 6.89 | 7.02 | 7.16 | 7.29 | 7.43 | 7.57 | 7.7 | 7.84 | 7.97 | 8.11 | 8.24 | 8.38 | 8.52 | 8.65 | 8.79 | 8.92 | 9.06 | 9.19 | 9.33 | 9.46 | 9.6 | 9.74 | 9.87 | 10.01 | 10.14 | 10.28 | 10.41 | 10.55 | 10.69
14.74 | 5.26 | 5.39 | 5.53 | 5.67 | 5.8 | 5.94 | 6.07 | 6.21 | 6.34 | 6.48 | 6.61 | 6.75 | 6.89 | 7.02 | 7.16 | 7.29 | 7.43 | 7.56 | 7.7 | 7.84 | 7.97 | 8.11 | 8.24 | 8.38 | 8.51 | 8.65 | 8.78 | 8.92 | 9.06 | 9.19 | 9.33 | 9.46 | 9.6 | 9.73 | 9.87 | 10.01 | 10.14 | 10.28 | 10.41 | 10.55 | 10.68
14.75 | 5.26 | 5.39 | 5.53 | 5.66 | 5.8 | 5.94 | 6.07 | 6.21 | 6.34 | 6.48 | 6.61 | 6.75 | 6.88 | 7.02 | 7.16 | 7.29 | 7.43 | 7.56 | 7.7 | 7.83 | 7.97 | 8.11 | 8.24 | 8.38 | 8.51 | 8.65 | 8.78 | 8.92 | 9.05 | 9.19 | 9.33 | 9.46 | 9.6 | 9.73 | 9.87 | 10 | 10.14 | 10.27 | 10.41 | 10.55 | 10.68
14.76 | 5.26 | 5.39 | 5.53 | 5.66 | 5.8 | 5.93 | 6.07 | 6.21 | 6.34 | 6.48 | 6.61 | 6.75 | 6.88 | 7.02 | 7.15 | 7.29 | 7.43 | 7.56 | 7.7 | 7.83 | 7.97 | 8.1 | 8.24 | 8.37 | 8.51 | 8.65 | 8.78 | 8.92 | 9.05 | 9.19 | 9.32 | 9.46 | 9.59 | 9.73 | 9.87 | 10 | 10.14 | 10.27 | 10.41 | 10.54 | 10.68
14.77 | 5.25 | 5.39 | 5.53 | 5.66 | 5.8 | 5.93 | 6.07 | 6.2 | 6.34 | 6.47 | 6.61 | 6.75 | 6.88 | 7.02 | 7.15 | 7.29 | 7.42 | 7.56 | 7.69 | 7.83 | 7.97 | 8.1 | 8.24 | 8.37 | 8.51 | 8.64 | 8.78 | 8.91 | 9.05 | 9.19 | 9.32 | 9.46 | 9.59 | 9.73 | 9.86 | 10 | 10.13 | 10.27 | 10.41 | 10.54 | 10.68
14.78 | 5.25 | 5.39 | 5.52 | 5.66 | 5.8 | 5.93 | 6.07 | 6.2 | 6.34 | 6.47 | 6.61 | 6.74 | 6.88 | 7.02 | 7.15 | 7.29 | 7.42 | 7.56 | 7.69 | 7.83 | 7.96 | 8.1 | 8.24 | 8.37 | 8.51 | 8.64 | 8.78 | 8.91 | 9.05 | 9.18 | 9.32 | 9.45 | 9.59 | 9.73 | 9.86 | 10 | 10.13 | 10.27 | 10.4 | 10.54 | 10.67
14.79 | 5.25 | 5.39 | 5.52 | 5.66 | 5.79 | 5.93 | 6.07 | 6.2 | 6.34 | 6.47 | 6.61 | 6.74 | 6.88 | 7.01 | 7.15 | 7.29 | 7.42 | 7.56 | 7.69 | 7.83 | 7.96 | 8.1 | 8.23 | 8.37 | 8.5 | 8.64 | 8.78 | 8.91 | 9.05 | 9.18 | 9.32 | 9.45 | 9.59 | 9.72 | 9.86 | 9.99 | 10.13 | 10.27 | 10.4 | 10.54 | 10.67
14.8 | 5.25 | 5.39 | 5.52 | 5.66 | 5.79 | 5.93 | 6.06 | 6.2 | 6.34 | 6.47 | 6.61 | 6.74 | 6.88 | 7.01 | 7.15 | 7.28 | 7.42 | 7.55 | 7.69 | 7.83 | 7.96 | 8.1 | 8.23 | 8.37 | 8.5 | 8.64 | 8.77 | 8.91 | 9.04 | 9.18 | 9.32 | 9.45 | 9.59 | 9.72 | 9.86 | 9.99 | 10.13 | 10.26 | 10.4 | 10.53 | 10.67
14.81 | 5.25 | 5.39 | 5.52 | 5.66 | 5.79 | 5.93 | 6.06 | 6.2 | 6.33 | 6.47 | 6.6 | 6.74 | 6.88 | 7.01 | 7.15 | 7.28 | 7.42 | 7.55 | 7.69 | 7.82 | 7.96 | 8.09 | 8.23 | 8.37 | 8.5 | 8.64 | 8.77 | 8.91 | 9.04 | 9.18 | 9.31 | 9.45 | 9.58 | 9.72 | 9.85 | 9.99 | 10.13 | 10.26 | 10.4 | 10.53 | 10.67
14.82 | 5.25 | 5.38 | 5.52 | 5.66 | 5.79 | 5.93 | 6.06 | 6.2 | 6.33 | 6.47 | 6.6 | 6.74 | 6.87 | 7.01 | 7.14 | 7.28 | 7.42 | 7.55 | 7.69 | 7.82 | 7.96 | 8.09 | 8.23 | 8.36 | 8.5 | 8.63 | 8.77 | 8.9 | 9.04 | 9.18 | 9.31 | 9.45 | 9.58 | 9.72 | 9.85 | 9.99 | 10.12 | 10.26 | 10.39 | 10.53 | 10.66
14.83 | 5.25 | 5.38 | 5.52 | 5.65 | 5.79 | 5.92 | 6.06 | 6.2 | 6.33 | 6.47 | 6.6 | 6.74 | 6.87 | 7.01 | 7.14 | 7.28 | 7.41 | 7.55 | 7.68 | 7.82 | 7.96 | 8.09 | 8.23 | 8.36 | 8.5 | 8.63 | 8.77 | 8.9 | 9.04 | 9.17 | 9.31 | 9.44 | 9.58 | 9.72 | 9.85 | 9.99 | 10.12 | 10.26 | 10.39 | 10.53 | 10.66
14.84 | 5.25 | 5.38 | 5.52 | 5.65 | 5.79 | 5.92 | 6.06 | 6.19 | 6.33 | 6.46 | 6.6 | 6.74 | 6.87 | 7.01 | 7.14 | 7.28 | 7.41 | 7.55 | 7.68 | 7.82 | 7.95 | 8.09 | 8.22 | 8.36 | 8.5 | 8.63 | 8.77 | 8.9 | 9.04 | 9.17 | 9.31 | 9.44 | 9.58 | 9.71 | 9.85 | 9.98 | 10.12 | 10.25 | 10.39 | 10.52 | 10.66
14.85 | 5.25 | 5.38 | 5.52 | 5.65 | 5.79 | 5.92 | 6.06 | 6.19 | 6.33 | 6.46 | 6.6 | 6.73 | 6.87 | 7 | 7.14 | 7.28 | 7.41 | 7.55 | 7.68 | 7.82 | 7.95 | 8.09 | 8.22 | 8.36 | 8.49 | 8.63 | 8.76 | 8.9 | 9.03 | 9.17 | 9.3 | 9.44 | 9.58 | 9.71 | 9.85 | 9.98 | 10.12 | 10.25 | 10.39 | 10.52 | 10.66
14.86 | 5.24 | 5.38 | 5.51 | 5.65 | 5.79 | 5.92 | 6.06 | 6.19 | 6.33 | 6.46 | 6.6 | 6.73 | 6.87 | 7 | 7.14 | 7.27 | 7.41 | 7.54 | 7.68 | 7.81 | 7.95 | 8.09 | 8.22 | 8.36 | 8.49 | 8.63 | 8.76 | 8.9 | 9.03 | 9.17 | 9.3 | 9.44 | 9.57 | 9.71 | 9.84 | 9.98 | 10.11 | 10.25 | 10.38 | 10.52 | 10.66
14.87 | 5.24 | 5.38 | 5.51 | 5.65 | 5.78 | 5.92 | 6.05 | 6.19 | 6.33 | 6.46 | 6.6 | 6.73 | 6.87 | 7 | 7.14 | 7.27 | 7.41 | 7.54 | 7.68 | 7.81 | 7.95 | 8.08 | 8.22 | 8.35 | 8.49 | 8.62 | 8.76 | 8.9 | 9.03 | 9.17 | 9.3 | 9.44 | 9.57 | 9.71 | 9.84 | 9.98 | 10.11 | 10.25 | 10.38 | 10.52 | 10.65
14.88 | 5.24 | 5.38 | 5.51 | 5.65 | 5.78 | 5.92 | 6.05 | 6.19 | 6.32 | 6.46 | 6.59 | 6.73 | 6.86 | 7 | 7.14 | 7.27 | 7.41 | 7.54 | 7.68 | 7.81 | 7.95 | 8.08 | 8.22 | 8.35 | 8.49 | 8.62 | 8.76 | 8.89 | 9.03 | 9.16 | 9.3 | 9.43 | 9.57 | 9.7 | 9.84 | 9.97 | 10.11 | 10.25 | 10.38 | 10.52 | 10.65
14.89 | 5.24 | 5.38 | 5.51 | 5.65 | 5.78 | 5.92 | 6.05 | 6.19 | 6.32 | 6.46 | 6.59 | 6.73 | 6.86 | 7 | 7.13 | 7.27 | 7.4 | 7.54 | 7.67 | 7.81 | 7.95 | 8.08 | 8.22 | 8.35 | 8.49 | 8.62 | 8.76 | 8.89 | 9.03 | 9.16 | 9.3 | 9.43 | 9.57 | 9.7 | 9.84 | 9.97 | 10.11 | 10.24 | 10.38 | 10.51 | 10.65
14.9 | 5.24 | 5.37 | 5.51 | 5.64 | 5.78 | 5.92 | 6.05 | 6.19 | 6.32 | 6.46 | 6.59 | 6.73 | 6.86 | 7 | 7.13 | 7.27 | 7.4 | 7.54 | 7.67 | 7.81 | 7.94 | 8.08 | 8.21 | 8.35 | 8.48 | 8.62 | 8.75 | 8.89 | 9.02 | 9.16 | 9.29 | 9.43 | 9.57 | 9.7 | 9.84 | 9.97 | 10.11 | 10.24 | 10.38 | 10.51 | 10.65
14.91 | 5.24 | 5.37 | 5.51 | 5.64 | 5.78 | 5.91 | 6.05 | 6.18 | 6.32 | 6.45 | 6.59 | 6.73 | 6.86 | 7 | 7.13 | 7.27 | 7.4 | 7.54 | 7.67 | 7.81 | 7.94 | 8.08 | 8.21 | 8.35 | 8.48 | 8.62 | 8.75 | 8.89 | 9.02 | 9.16 | 9.29 | 9.43 | 9.56 | 9.7 | 9.83 | 9.97 | 10.1 | 10.24 | 10.37 | 10.51 | 10.64
14.92 | 5.24 | 5.37 | 5.51 | 5.64 | 5.78 | 5.91 | 6.05 | 6.18 | 6.32 | 6.45 | 6.59 | 6.72 | 6.86 | 6.99 | 7.13 | 7.26 | 7.4 | 7.53 | 7.67 | 7.8 | 7.94 | 8.07 | 8.21 | 8.35 | 8.48 | 8.62 | 8.75 | 8.89 | 9.02 | 9.16 | 9.29 | 9.43 | 9.56 | 9.7 | 9.83 | 9.97 | 10.1 | 10.24 | 10.37 | 10.51 | 10.64
14.93 | 5.24 | 5.37 | 5.51 | 5.64 | 5.78 | 5.91 | 6.05 | 6.18 | 6.32 | 6.45 | 6.59 | 6.72 | 6.86 | 6.99 | 7.13 | 7.26 | 7.4 | 7.53 | 7.67 | 7.8 | 7.94 | 8.07 | 8.21 | 8.34 | 8.48 | 8.61 | 8.75 | 8.88 | 9.02 | 9.15 | 9.29 | 9.42 | 9.56 | 9.69 | 9.83 | 9.96 | 10.1 | 10.23 | 10.37 | 10.5 | 10.64
14.94 | 5.23 | 5.37 | 5.5 | 5.64 | 5.78 | 5.91 | 6.05 | 6.18 | 6.32 | 6.45 | 6.59 | 6.72 | 6.86 | 6.99 | 7.13 | 7.26 | 7.4 | 7.53 | 7.67 | 7.8 | 7.94 | 8.07 | 8.21 | 8.34 | 8.48 | 8.61 | 8.75 | 8.88 | 9.02 | 9.15 | 9.29 | 9.42 | 9.56 | 9.69 | 9.83 | 9.96 | 10.1 | 10.23 | 10.37 | 10.5 | 10.64
14.95 | 5.23 | 5.37 | 5.5 | 5.64 | 5.77 | 5.91 | 6.04 | 6.18 | 6.31 | 6.45 | 6.58 | 6.72 | 6.85 | 6.99 | 7.12 | 7.26 | 7.39 | 7.53 | 7.66 | 7.8 | 7.93 | 8.07 | 8.2 | 8.34 | 8.47 | 8.61 | 8.74 | 8.88 | 9.01 | 9.15 | 9.28 | 9.42 | 9.55 | 9.69 | 9.82 | 9.96 | 10.09 | 10.23 | 10.36 | 10.5 | 10.63
14.96 | 5.23 | 5.37 | 5.5 | 5.64 | 5.77 | 5.91 | 6.04 | 6.18 | 6.31 | 6.45 | 6.58 | 6.72 | 6.85 | 6.99 | 7.12 | 7.26 | 7.39 | 7.53 | 7.66 | 7.8 | 7.93 | 8.07 | 8.2 | 8.34 | 8.47 | 8.61 | 8.74 | 8.88 | 9.01 | 9.15 | 9.28 | 9.42 | 9.55 | 9.69 | 9.82 | 9.96 | 10.09 | 10.23 | 10.36 | 10.5 | 10.63
14.97 | 5.23 | 5.37 | 5.5 | 5.64 | 5.77 | 5.91 | 6.04 | 6.18 | 6.31 | 6.45 | 6.58 | 6.72 | 6.85 | 6.99 | 7.12 | 7.26 | 7.39 | 7.53 | 7.66 | 7.8 | 7.93 | 8.07 | 8.2 | 8.34 | 8.47 | 8.61 | 8.74 | 8.88 | 9.01 | 9.15 | 9.28 | 9.42 | 9.55 | 9.69 | 9.82 | 9.96 | 10.09 | 10.23 | 10.36 | 10.5 | 10.63
14.98 | 5.23 | 5.36 | 5.5 | 5.63 | 5.77 | 5.9 | 6.04 | 6.17 | 6.31 | 6.44 | 6.58 | 6.71 | 6.85 | 6.98 | 7.12 | 7.25 | 7.39 | 7.52 | 7.66 | 7.79 | 7.93 | 8.06 | 8.2 | 8.33 | 8.47 | 8.6 | 8.74 | 8.87 | 9.01 | 9.14 | 9.28 | 9.41 | 9.55 | 9.68 | 9.82 | 9.95 | 10.09 | 10.22 | 10.36 | 10.49 | 10.63
14.99 | 5.23 | 5.36 | 5.5 | 5.63 | 5.77 | 5.9 | 6.04 | 6.17 | 6.31 | 6.44 | 6.58 | 6.71 | 6.85 | 6.98 | 7.12 | 7.25 | 7.39 | 7.52 | 7.66 | 7.79 | 7.93 | 8.06 | 8.2 | 8.33 | 8.47 | 8.6 | 8.74 | 8.87 | 9.01 | 9.14 | 9.28 | 9.41 | 9.55 | 9.68 | 9.82 | 9.95 | 10.09 | 10.22 | 10.36 | 10.49 | 10.63
15 | 5.23 | 5.36 | 5.5 | 5.63 | 5.77 | 5.9 | 6.04 | 6.17 | 6.31 | 6.44 | 6.58 | 6.71 | 6.85 | 6.98 | 7.12 | 7.25 | 7.39 | 7.52 | 7.66 | 7.79 | 7.93 | 8.06 | 8.2 | 8.33 | 8.47 | 8.6 | 8.74 | 8.87 | 9 | 9.14 | 9.27 | 9.41 | 9.54 | 9.68 | 9.81 | 9.95 | 10.08 | 10.22 | 10.35 | 10.49 | 10.62
15.01 | 5.23 | 5.36 | 5.5 | 5.63 | 5.77 | 5.9 | 6.04 | 6.17 | 6.31 | 6.44 | 6.58 | 6.71 | 6.85 | 6.98 | 7.11 | 7.25 | 7.38 | 7.52 | 7.65 | 7.79 | 7.92 | 8.06 | 8.19 | 8.33 | 8.46 | 8.6 | 8.73 | 8.87 | 9 | 9.14 | 9.27 | 9.41 | 9.54 | 9.68 | 9.81 | 9.95 | 10.08 | 10.22 | 10.35 | 10.49 | 10.62
15.02 | 5.23 | 5.36 | 5.49 | 5.63 | 5.76 | 5.9 | 6.03 | 6.17 | 6.3 | 6.44 | 6.57 | 6.71 | 6.84 | 6.98 | 7.11 | 7.25 | 7.38 | 7.52 | 7.65 | 7.79 | 7.92 | 8.06 | 8.19 | 8.33 | 8.46 | 8.6 | 8.73 | 8.87 | 9 | 9.14 | 9.27 | 9.41 | 9.54 | 9.68 | 9.81 | 9.94 | 10.08 | 10.21 | 10.35 | 10.48 | 10.62
15.03 | 5.22 | 5.36 | 5.49 | 5.63 | 5.76 | 5.9 | 6.03 | 6.17 | 6.3 | 6.44 | 6.57 | 6.71 | 6.84 | 6.98 | 7.11 | 7.25 | 7.38 | 7.52 | 7.65 | 7.79 | 7.92 | 8.06 | 8.19 | 8.33 | 8.46 | 8.59 | 8.73 | 8.86 | 9 | 9.13 | 9.27 | 9.4 | 9.54 | 9.67 | 9.81 | 9.94 | 10.08 | 10.21 | 10.35 | 10.48 | 10.62
15.04 | 5.22 | 5.36 | 5.49 | 5.63 | 5.76 | 5.9 | 6.03 | 6.17 | 6.3 | 6.44 | 6.57 | 6.71 | 6.84 | 6.98 | 7.11 | 7.25 | 7.38 | 7.51 | 7.65 | 7.78 | 7.92 | 8.05 | 8.19 | 8.32 | 8.46 | 8.59 | 8.73 | 8.86 | 9 | 9.13 | 9.27 | 9.4 | 9.54 | 9.67 | 9.81 | 9.94 | 10.08 | 10.21 | 10.34 | 10.48 | 10.61
15.05 | 5.22 | 5.36 | 5.49 | 5.63 | 5.76 | 5.9 | 6.03 | 6.17 | 6.3 | 6.43 | 6.57 | 6.7 | 6.84 | 6.97 | 7.11 | 7.24 | 7.38 | 7.51 | 7.65 | 7.78 | 7.92 | 8.05 | 8.19 | 8.32 | 8.46 | 8.59 | 8.73 | 8.86 | 9 | 9.13 | 9.26 | 9.4 | 9.53 | 9.67 | 9.8 | 9.94 | 10.07 | 10.21 | 10.34 | 10.48 | 10.61
15.06 | 5.22 | 5.36 | 5.49 | 5.62 | 5.76 | 5.89 | 6.03 | 6.16 | 6.3 | 6.43 | 6.57 | 6.7 | 6.84 | 6.97 | 7.11 | 7.24 | 7.38 | 7.51 | 7.65 | 7.78 | 7.92 | 8.05 | 8.18 | 8.32 | 8.45 | 8.59 | 8.72 | 8.86 | 8.99 | 9.13 | 9.26 | 9.4 | 9.53 | 9.67 | 9.8 | 9.94 | 10.07 | 10.21 | 10.34 | 10.47 | 10.61
15.07 | 5.22 | 5.35 | 5.49 | 5.62 | 5.76 | 5.89 | 6.03 | 6.16 | 6.3 | 6.43 | 6.57 | 6.7 | 6.84 | 6.97 | 7.11 | 7.24 | 7.37 | 7.51 | 7.64 | 7.78 | 7.91 | 8.05 | 8.18 | 8.32 | 8.45 | 8.59 | 8.72 | 8.86 | 8.99 | 9.13 | 9.26 | 9.4 | 9.53 | 9.66 | 9.8 | 9.93 | 10.07 | 10.2 | 10.34 | 10.47 | 10.61
15.08 | 5.22 | 5.35 | 5.49 | 5.62 | 5.76 | 5.89 | 6.03 | 6.16 | 6.3 | 6.43 | 6.57 | 6.7 | 6.83 | 6.97 | 7.1 | 7.24 | 7.37 | 7.51 | 7.64 | 7.78 | 7.91 | 8.05 | 8.18 | 8.32 | 8.45 | 8.59 | 8.72 | 8.85 | 8.99 | 9.12 | 9.26 | 9.39 | 9.53 | 9.66 | 9.8 | 9.93 | 10.07 | 10.2 | 10.34 | 10.47 | 10.6
15.09 | 5.22 | 5.35 | 5.49 | 5.62 | 5.76 | 5.89 | 6.02 | 6.16 | 6.29 | 6.43 | 6.56 | 6.7 | 6.83 | 6.97 | 7.1 | 7.24 | 7.37 | 7.51 | 7.64 | 7.78 | 7.91 | 8.04 | 8.18 | 8.31 | 8.45 | 8.58 | 8.72 | 8.85 | 8.99 | 9.12 | 9.26 | 9.39 | 9.53 | 9.66 | 9.8 | 9.93 | 10.06 | 10.2 | 10.33 | 10.47 | 10.6
15.1 | 5.22 | 5.35 | 5.49 | 5.62 | 5.75 | 5.89 | 6.02 | 6.16 | 6.29 | 6.43 | 6.56 | 6.7 | 6.83 | 6.97 | 7.1 | 7.24 | 7.37 | 7.5 | 7.64 | 7.77 | 7.91 | 8.04 | 8.18 | 8.31 | 8.45 | 8.58 | 8.72 | 8.85 | 8.99 | 9.12 | 9.25 | 9.39 | 9.52 | 9.66 | 9.79 | 9.93 | 10.06 | 10.2 | 10.33 | 10.47 | 10.6
15.11 | 5.21 | 5.35 | 5.48 | 5.62 | 5.75 | 5.89 | 6.02 | 6.16 | 6.29 | 6.43 | 6.56 | 6.7 | 6.83 | 6.96 | 7.1 | 7.23 | 7.37 | 7.5 | 7.64 | 7.77 | 7.91 | 8.04 | 8.18 | 8.31 | 8.45 | 8.58 | 8.71 | 8.85 | 8.98 | 9.12 | 9.25 | 9.39 | 9.52 | 9.66 | 9.79 | 9.93 | 10.06 | 10.19 | 10.33 | 10.46 | 10.6
15.12 | 5.21 | 5.35 | 5.48 | 5.62 | 5.75 | 5.89 | 6.02 | 6.16 | 6.29 | 6.42 | 6.56 | 6.69 | 6.83 | 6.96 | 7.1 | 7.23 | 7.37 | 7.5 | 7.64 | 7.77 | 7.91 | 8.04 | 8.17 | 8.31 | 8.44 | 8.58 | 8.71 | 8.85 | 8.98 | 9.12 | 9.25 | 9.39 | 9.52 | 9.65 | 9.79 | 9.92 | 10.06 | 10.19 | 10.33 | 10.46 | 10.6
15.13 | 5.21 | 5.35 | 5.48 | 5.62 | 5.75 | 5.89 | 6.02 | 6.15 | 6.29 | 6.42 | 6.56 | 6.69 | 6.83 | 6.96 | 7.1 | 7.23 | 7.37 | 7.5 | 7.63 | 7.77 | 7.9 | 8.04 | 8.17 | 8.31 | 8.44 | 8.58 | 8.71 | 8.85 | 8.98 | 9.11 | 9.25 | 9.38 | 9.52 | 9.65 | 9.79 | 9.92 | 10.06 | 10.19 | 10.32 | 10.46 | 10.59
15.14 | 5.21 | 5.35 | 5.48 | 5.61 | 5.75 | 5.88 | 6.02 | 6.15 | 6.29 | 6.42 | 6.56 | 6.69 | 6.83 | 6.96 | 7.09 | 7.23 | 7.36 | 7.5 | 7.63 | 7.77 | 7.9 | 8.04 | 8.17 | 8.31 | 8.44 | 8.57 | 8.71 | 8.84 | 8.98 | 9.11 | 9.25 | 9.38 | 9.52 | 9.65 | 9.78 | 9.92 | 10.05 | 10.19 | 10.32 | 10.46 | 10.59
15.15 | 5.21 | 5.34 | 5.48 | 5.61 | 5.75 | 5.88 | 6.02 | 6.15 | 6.29 | 6.42 | 6.56 | 6.69 | 6.82 | 6.96 | 7.09 | 7.23 | 7.36 | 7.5 | 7.63 | 7.77 | 7.9 | 8.03 | 8.17 | 8.3 | 8.44 | 8.57 | 8.71 | 8.84 | 8.98 | 9.11 | 9.24 | 9.38 | 9.51 | 9.65 | 9.78 | 9.92 | 10.05 | 10.19 | 10.32 | 10.45 | 10.59
15.16 | 5.21 | 5.34 | 5.48 | 5.61 | 5.75 | 5.88 | 6.02 | 6.15 | 6.28 | 6.42 | 6.55 | 6.69 | 6.82 | 6.96 | 7.09 | 7.23 | 7.36 | 7.49 | 7.63 | 7.76 | 7.9 | 8.03 | 8.17 | 8.3 | 8.44 | 8.57 | 8.7 | 8.84 | 8.97 | 9.11 | 9.24 | 9.38 | 9.51 | 9.65 | 9.78 | 9.91 | 10.05 | 10.18 | 10.32 | 10.45 | 10.59
15.17 | 5.21 | 5.34 | 5.48 | 5.61 | 5.75 | 5.88 | 6.01 | 6.15 | 6.28 | 6.42 | 6.55 | 6.69 | 6.82 | 6.96 | 7.09 | 7.22 | 7.36 | 7.49 | 7.63 | 7.76 | 7.9 | 8.03 | 8.17 | 8.3 | 8.43 | 8.57 | 8.7 | 8.84 | 8.97 | 9.11 | 9.24 | 9.38 | 9.51 | 9.64 | 9.78 | 9.91 | 10.05 | 10.18 | 10.32 | 10.45 | 10.58
15.18 | 5.21 | 5.34 | 5.48 | 5.61 | 5.74 | 5.88 | 6.01 | 6.15 | 6.28 | 6.42 | 6.55 | 6.69 | 6.82 | 6.95 | 7.09 | 7.22 | 7.36 | 7.49 | 7.63 | 7.76 | 7.89 | 8.03 | 8.16 | 8.3 | 8.43 | 8.57 | 8.7 | 8.84 | 8.97 | 9.1 | 9.24 | 9.37 | 9.51 | 9.64 | 9.78 | 9.91 | 10.04 | 10.18 | 10.31 | 10.45 | 10.58
15.19 | 5.21 | 5.34 | 5.47 | 5.61 | 5.74 | 5.88 | 6.01 | 6.15 | 6.28 | 6.41 | 6.55 | 6.68 | 6.82 | 6.95 | 7.09 | 7.22 | 7.36 | 7.49 | 7.62 | 7.76 | 7.89 | 8.03 | 8.16 | 8.3 | 8.43 | 8.56 | 8.7 | 8.83 | 8.97 | 9.1 | 9.24 | 9.37 | 9.51 | 9.64 | 9.77 | 9.91 | 10.04 | 10.18 | 10.31 | 10.45 | 10.58
15.2 | 5.2 | 5.34 | 5.47 | 5.61 | 5.74 | 5.88 | 6.01 | 6.14 | 6.28 | 6.41 | 6.55 | 6.68 | 6.82 | 6.95 | 7.09 | 7.22 | 7.35 | 7.49 | 7.62 | 7.76 | 7.89 | 8.03 | 8.16 | 8.29 | 8.43 | 8.56 | 8.7 | 8.83 | 8.97 | 9.1 | 9.23 | 9.37 | 9.5 | 9.64 | 9.77 | 9.91 | 10.04 | 10.17 | 10.31 | 10.44 | 10.58
15.21 | 5.2 | 5.34 | 5.47 | 5.61 | 5.74 | 5.87 | 6.01 | 6.14 | 6.28 | 6.41 | 6.55 | 6.68 | 6.81 | 6.95 | 7.08 | 7.22 | 7.35 | 7.49 | 7.62 | 7.76 | 7.89 | 8.02 | 8.16 | 8.29 | 8.43 | 8.56 | 8.7 | 8.83 | 8.96 | 9.1 | 9.23 | 9.37 | 9.5 | 9.64 | 9.77 | 9.9 | 10.04 | 10.17 | 10.31 | 10.44 | 10.58
15.22 | 5.2 | 5.34 | 5.47 | 5.6 | 5.74 | 5.87 | 6.01 | 6.14 | 6.28 | 6.41 | 6.54 | 6.68 | 6.81 | 6.95 | 7.08 | 7.22 | 7.35 | 7.49 | 7.62 | 7.75 | 7.89 | 8.02 | 8.16 | 8.29 | 8.43 | 8.56 | 8.69 | 8.83 | 8.96 | 9.1 | 9.23 | 9.36 | 9.5 | 9.63 | 9.77 | 9.9 | 10.04 | 10.17 | 10.3 | 10.44 | 10.57
15.23 | 5.2 | 5.33 | 5.47 | 5.6 | 5.74 | 5.87 | 6.01 | 6.14 | 6.27 | 6.41 | 6.54 | 6.68 | 6.81 | 6.95 | 7.08 | 7.21 | 7.35 | 7.48 | 7.62 | 7.75 | 7.89 | 8.02 | 8.15 | 8.29 | 8.42 | 8.56 | 8.69 | 8.83 | 8.96 | 9.09 | 9.23 | 9.36 | 9.5 | 9.63 | 9.77 | 9.9 | 10.03 | 10.17 | 10.3 | 10.44 | 10.57
15.24 | 5.2 | 5.33 | 5.47 | 5.6 | 5.74 | 5.87 | 6 | 6.14 | 6.27 | 6.41 | 6.54 | 6.68 | 6.81 | 6.94 | 7.08 | 7.21 | 7.35 | 7.48 | 7.62 | 7.75 | 7.88 | 8.02 | 8.15 | 8.29 | 8.42 | 8.56 | 8.69 | 8.82 | 8.96 | 9.09 | 9.23 | 9.36 | 9.5 | 9.63 | 9.76 | 9.9 | 10.03 | 10.17 | 10.3 | 10.43 | 10.57
15.25 | 5.2 | 5.33 | 5.47 | 5.6 | 5.74 | 5.87 | 6 | 6.14 | 6.27 | 6.41 | 6.54 | 6.67 | 6.81 | 6.94 | 7.08 | 7.21 | 7.35 | 7.48 | 7.61 | 7.75 | 7.88 | 8.02 | 8.15 | 8.29 | 8.42 | 8.55 | 8.69 | 8.82 | 8.96 | 9.09 | 9.22 | 9.36 | 9.49 | 9.63 | 9.76 | 9.9 | 10.03 | 10.16 | 10.3 | 10.43 | 10.57
15.26 | 5.2 | 5.33 | 5.47 | 5.6 | 5.73 | 5.87 | 6 | 6.14 | 6.27 | 6.4 | 6.54 | 6.67 | 6.81 | 6.94 | 7.08 | 7.21 | 7.34 | 7.48 | 7.61 | 7.75 | 7.88 | 8.02 | 8.15 | 8.28 | 8.42 | 8.55 | 8.69 | 8.82 | 8.95 | 9.09 | 9.22 | 9.36 | 9.49 | 9.63 | 9.76 | 9.89 | 10.03 | 10.16 | 10.3 | 10.43 | 10.56
15.27 | 5.2 | 5.33 | 5.46 | 5.6 | 5.73 | 5.87 | 6 | 6.14 | 6.27 | 6.4 | 6.54 | 6.67 | 6.81 | 6.94 | 7.07 | 7.21 | 7.34 | 7.48 | 7.61 | 7.75 | 7.88 | 8.01 | 8.15 | 8.28 | 8.42 | 8.55 | 8.68 | 8.82 | 8.95 | 9.09 | 9.22 | 9.35 | 9.49 | 9.62 | 9.76 | 9.89 | 10.03 | 10.16 | 10.29 | 10.43 | 10.56
15.28 | 5.19 | 5.33 | 5.46 | 5.6 | 5.73 | 5.87 | 6 | 6.13 | 6.27 | 6.4 | 6.54 | 6.67 | 6.8 | 6.94 | 7.07 | 7.21 | 7.34 | 7.48 | 7.61 | 7.74 | 7.88 | 8.01 | 8.15 | 8.28 | 8.41 | 8.55 | 8.68 | 8.82 | 8.95 | 9.08 | 9.22 | 9.35 | 9.49 | 9.62 | 9.76 | 9.89 | 10.02 | 10.16 | 10.29 | 10.43 | 10.56
15.29 | 5.19 | 5.33 | 5.46 | 5.6 | 5.73 | 5.86 | 6 | 6.13 | 6.27 | 6.4 | 6.53 | 6.67 | 6.8 | 6.94 | 7.07 | 7.21 | 7.34 | 7.47 | 7.61 | 7.74 | 7.88 | 8.01 | 8.14 | 8.28 | 8.41 | 8.55 | 8.68 | 8.81 | 8.95 | 9.08 | 9.22 | 9.35 | 9.48 | 9.62 | 9.75 | 9.89 | 10.02 | 10.16 | 10.29 | 10.42 | 10.56
15.3 | 5.19 | 5.33 | 5.46 | 5.59 | 5.73 | 5.86 | 6 | 6.13 | 6.27 | 6.4 | 6.53 | 6.67 | 6.8 | 6.94 | 7.07 | 7.2 | 7.34 | 7.47 | 7.61 | 7.74 | 7.87 | 8.01 | 8.14 | 8.28 | 8.41 | 8.54 | 8.68 | 8.81 | 8.95 | 9.08 | 9.21 | 9.35 | 9.48 | 9.62 | 9.75 | 9.88 | 10.02 | 10.15 | 10.29 | 10.42 | 10.56
15.31 | 5.19 | 5.33 | 5.46 | 5.59 | 5.73 | 5.86 | 6 | 6.13 | 6.26 | 6.4 | 6.53 | 6.67 | 6.8 | 6.93 | 7.07 | 7.2 | 7.34 | 7.47 | 7.6 | 7.74 | 7.87 | 8.01 | 8.14 | 8.27 | 8.41 | 8.54 | 8.68 | 8.81 | 8.94 | 9.08 | 9.21 | 9.35 | 9.48 | 9.61 | 9.75 | 9.88 | 10.02 | 10.15 | 10.28 | 10.42 | 10.55
15.32 | 5.19 | 5.32 | 5.46 | 5.59 | 5.73 | 5.86 | 5.99 | 6.13 | 6.26 | 6.4 | 6.53 | 6.66 | 6.8 | 6.93 | 7.07 | 7.2 | 7.33 | 7.47 | 7.6 | 7.74 | 7.87 | 8 | 8.14 | 8.27 | 8.41 | 8.54 | 8.67 | 8.81 | 8.94 | 9.08 | 9.21 | 9.34 | 9.48 | 9.61 | 9.75 | 9.88 | 10.01 | 10.15 | 10.28 | 10.42 | 10.55
15.33 | 5.19 | 5.32 | 5.46 | 5.59 | 5.72 | 5.86 | 5.99 | 6.13 | 6.26 | 6.39 | 6.53 | 6.66 | 6.8 | 6.93 | 7.07 | 7.2 | 7.33 | 7.47 | 7.6 | 7.74 | 7.87 | 8 | 8.14 | 8.27 | 8.4 | 8.54 | 8.67 | 8.81 | 8.94 | 9.07 | 9.21 | 9.34 | 9.48 | 9.61 | 9.74 | 9.88 | 10.01 | 10.15 | 10.28 | 10.41 | 10.55
15.34 | 5.19 | 5.32 | 5.46 | 5.59 | 5.72 | 5.86 | 5.99 | 6.13 | 6.26 | 6.39 | 6.53 | 6.66 | 6.8 | 6.93 | 7.06 | 7.2 | 7.33 | 7.47 | 7.6 | 7.73 | 7.87 | 8 | 8.14 | 8.27 | 8.4 | 8.54 | 8.67 | 8.8 | 8.94 | 9.07 | 9.21 | 9.34 | 9.47 | 9.61 | 9.74 | 9.88 | 10.01 | 10.14 | 10.28 | 10.41 | 10.55
15.35 | 5.19 | 5.32 | 5.45 | 5.59 | 5.72 | 5.86 | 5.99 | 6.12 | 6.26 | 6.39 | 6.53 | 6.66 | 6.79 | 6.93 | 7.06 | 7.2 | 7.33 | 7.46 | 7.6 | 7.73 | 7.87 | 8 | 8.13 | 8.27 | 8.4 | 8.54 | 8.67 | 8.8 | 8.94 | 9.07 | 9.2 | 9.34 | 9.47 | 9.61 | 9.74 | 9.87 | 10.01 | 10.14 | 10.28 | 10.41 | 10.54
15.36 | 5.19 | 5.32 | 5.45 | 5.59 | 5.72 | 5.85 | 5.99 | 6.12 | 6.26 | 6.39 | 6.52 | 6.66 | 6.79 | 6.93 | 7.06 | 7.19 | 7.33 | 7.46 | 7.6 | 7.73 | 7.86 | 8 | 8.13 | 8.27 | 8.4 | 8.53 | 8.67 | 8.8 | 8.94 | 9.07 | 9.2 | 9.34 | 9.47 | 9.6 | 9.74 | 9.87 | 10.01 | 10.14 | 10.27 | 10.41 | 10.54
15.37 | 5.18 | 5.32 | 5.45 | 5.59 | 5.72 | 5.85 | 5.99 | 6.12 | 6.26 | 6.39 | 6.52 | 6.66 | 6.79 | 6.92 | 7.06 | 7.19 | 7.33 | 7.46 | 7.59 | 7.73 | 7.86 | 8 | 8.13 | 8.26 | 8.4 | 8.53 | 8.67 | 8.8 | 8.93 | 9.07 | 9.2 | 9.33 | 9.47 | 9.6 | 9.74 | 9.87 | 10 | 10.14 | 10.27 | 10.41 | 10.54
15.38 | 5.18 | 5.32 | 5.45 | 5.58 | 5.72 | 5.85 | 5.99 | 6.12 | 6.25 | 6.39 | 6.52 | 6.66 | 6.79 | 6.92 | 7.06 | 7.19 | 7.33 | 7.46 | 7.59 | 7.73 | 7.86 | 7.99 | 8.13 | 8.26 | 8.4 | 8.53 | 8.66 | 8.8 | 8.93 | 9.07 | 9.2 | 9.33 | 9.47 | 9.6 | 9.73 | 9.87 | 10 | 10.14 | 10.27 | 10.4 | 10.54
15.39 | 5.18 | 5.32 | 5.45 | 5.58 | 5.72 | 5.85 | 5.98 | 6.12 | 6.25 | 6.39 | 6.52 | 6.65 | 6.79 | 6.92 | 7.06 | 7.19 | 7.32 | 7.46 | 7.59 | 7.72 | 7.86 | 7.99 | 8.13 | 8.26 | 8.39 | 8.53 | 8.66 | 8.8 | 8.93 | 9.06 | 9.2 | 9.33 | 9.46 | 9.6 | 9.73 | 9.87 | 10 | 10.13 | 10.27 | 10.4 | 10.53
15.4 | 5.18 | 5.31 | 5.45 | 5.58 | 5.72 | 5.85 | 5.98 | 6.12 | 6.25 | 6.39 | 6.52 | 6.65 | 6.79 | 6.92 | 7.05 | 7.19 | 7.32 | 7.46 | 7.59 | 7.72 | 7.86 | 7.99 | 8.12 | 8.26 | 8.39 | 8.53 | 8.66 | 8.79 | 8.93 | 9.06 | 9.19 | 9.33 | 9.46 | 9.6 | 9.73 | 9.86 | 10 | 10.13 | 10.26 | 10.4 | 10.53
15.41 | 5.18 | 5.31 | 5.45 | 5.58 | 5.71 | 5.85 | 5.98 | 6.12 | 6.25 | 6.38 | 6.52 | 6.65 | 6.79 | 6.92 | 7.05 | 7.19 | 7.32 | 7.45 | 7.59 | 7.72 | 7.86 | 7.99 | 8.12 | 8.26 | 8.39 | 8.52 | 8.66 | 8.79 | 8.93 | 9.06 | 9.19 | 9.33 | 9.46 | 9.59 | 9.73 | 9.86 | 10 | 10.13 | 10.26 | 10.4 | 10.53
15.42 | 5.18 | 5.31 | 5.45 | 5.58 | 5.71 | 5.85 | 5.98 | 6.11 | 6.25 | 6.38 | 6.52 | 6.65 | 6.78 | 6.92 | 7.05 | 7.18 | 7.32 | 7.45 | 7.59 | 7.72 | 7.85 | 7.99 | 8.12 | 8.25 | 8.39 | 8.52 | 8.66 | 8.79 | 8.92 | 9.06 | 9.19 | 9.32 | 9.46 | 9.59 | 9.73 | 9.86 | 9.99 | 10.13 | 10.26 | 10.39 | 10.53
15.43 | 5.18 | 5.31 | 5.44 | 5.58 | 5.71 | 5.85 | 5.98 | 6.11 | 6.25 | 6.38 | 6.51 | 6.65 | 6.78 | 6.92 | 7.05 | 7.18 | 7.32 | 7.45 | 7.58 | 7.72 | 7.85 | 7.99 | 8.12 | 8.25 | 8.39 | 8.52 | 8.65 | 8.79 | 8.92 | 9.06 | 9.19 | 9.32 | 9.46 | 9.59 | 9.72 | 9.86 | 9.99 | 10.12 | 10.26 | 10.39 | 10.53
15.44 | 5.18 | 5.31 | 5.44 | 5.58 | 5.71 | 5.84 | 5.98 | 6.11 | 6.25 | 6.38 | 6.51 | 6.65 | 6.78 | 6.91 | 7.05 | 7.18 | 7.32 | 7.45 | 7.58 | 7.72 | 7.85 | 7.98 | 8.12 | 8.25 | 8.38 | 8.52 | 8.65 | 8.79 | 8.92 | 9.05 | 9.19 | 9.32 | 9.45 | 9.59 | 9.72 | 9.86 | 9.99 | 10.12 | 10.26 | 10.39 | 10.52
15.45 | 5.17 | 5.31 | 5.44 | 5.58 | 5.71 | 5.84 | 5.98 | 6.11 | 6.24 | 6.38 | 6.51 | 6.65 | 6.78 | 6.91 | 7.05 | 7.18 | 7.31 | 7.45 | 7.58 | 7.71 | 7.85 | 7.98 | 8.12 | 8.25 | 8.38 | 8.52 | 8.65 | 8.78 | 8.92 | 9.05 | 9.19 | 9.32 | 9.45 | 9.59 | 9.72 | 9.85 | 9.99 | 10.12 | 10.25 | 10.39 | 10.52
15.46 | 5.17 | 5.31 | 5.44 | 5.57 | 5.71 | 5.84 | 5.98 | 6.11 | 6.24 | 6.38 | 6.51 | 6.64 | 6.78 | 6.91 | 7.04 | 7.18 | 7.31 | 7.45 | 7.58 | 7.71 | 7.85 | 7.98 | 8.11 | 8.25 | 8.38 | 8.51 | 8.65 | 8.78 | 8.92 | 9.05 | 9.18 | 9.32 | 9.45 | 9.58 | 9.72 | 9.85 | 9.98 | 10.12 | 10.25 | 10.39 | 10.52
15.47 | 5.17 | 5.31 | 5.44 | 5.57 | 5.71 | 5.84 | 5.97 | 6.11 | 6.24 | 6.38 | 6.51 | 6.64 | 6.78 | 6.91 | 7.04 | 7.18 | 7.31 | 7.44 | 7.58 | 7.71 | 7.85 | 7.98 | 8.11 | 8.25 | 8.38 | 8.51 | 8.65 | 8.78 | 8.91 | 9.05 | 9.18 | 9.31 | 9.45 | 9.58 | 9.72 | 9.85 | 9.98 | 10.12 | 10.25 | 10.38 | 10.52
15.48 | 5.17 | 5.3 | 5.44 | 5.57 | 5.71 | 5.84 | 5.97 | 6.11 | 6.24 | 6.37 | 6.51 | 6.64 | 6.77 | 6.91 | 7.04 | 7.18 | 7.31 | 7.44 | 7.58 | 7.71 | 7.84 | 7.98 | 8.11 | 8.24 | 8.38 | 8.51 | 8.64 | 8.78 | 8.91 | 9.05 | 9.18 | 9.31 | 9.45 | 9.58 | 9.71 | 9.85 | 9.98 | 10.11 | 10.25 | 10.38 | 10.51
15.49 | 5.17 | 5.3 | 5.44 | 5.57 | 5.7 | 5.84 | 5.97 | 6.11 | 6.24 | 6.37 | 6.51 | 6.64 | 6.77 | 6.91 | 7.04 | 7.17 | 7.31 | 7.44 | 7.57 | 7.71 | 7.84 | 7.98 | 8.11 | 8.24 | 8.38 | 8.51 | 8.64 | 8.78 | 8.91 | 9.04 | 9.18 | 9.31 | 9.44 | 9.58 | 9.71 | 9.84 | 9.98 | 10.11 | 10.25 | 10.38 | 10.51
15.5 | 5.17 | 5.3 | 5.44 | 5.57 | 5.7 | 5.84 | 5.97 | 6.1 | 6.24 | 6.37 | 6.5 | 6.64 | 6.77 | 6.91 | 7.04 | 7.17 | 7.31 | 7.44 | 7.57 | 7.71 | 7.84 | 7.97 | 8.11 | 8.24 | 8.37 | 8.51 | 8.64 | 8.77 | 8.91 | 9.04 | 9.18 | 9.31 | 9.44 | 9.58 | 9.71 | 9.84 | 9.98 | 10.11 | 10.24 | 10.38 | 10.51
15.51 | 5.17 | 5.3 | 5.43 | 5.57 | 5.7 | 5.84 | 5.97 | 6.1 | 6.24 | 6.37 | 6.5 | 6.64 | 6.77 | 6.9 | 7.04 | 7.17 | 7.3 | 7.44 | 7.57 | 7.7 | 7.84 | 7.97 | 8.11 | 8.24 | 8.37 | 8.51 | 8.64 | 8.77 | 8.91 | 9.04 | 9.17 | 9.31 | 9.44 | 9.57 | 9.71 | 9.84 | 9.97 | 10.11 | 10.24 | 10.37 | 10.51
15.52 | 5.17 | 5.3 | 5.43 | 5.57 | 5.7 | 5.83 | 5.97 | 6.1 | 6.23 | 6.37 | 6.5 | 6.64 | 6.77 | 6.9 | 7.04 | 7.17 | 7.3 | 7.44 | 7.57 | 7.7 | 7.84 | 7.97 | 8.1 | 8.24 | 8.37 | 8.5 | 8.64 | 8.77 | 8.9 | 9.04 | 9.17 | 9.3 | 9.44 | 9.57 | 9.7 | 9.84 | 9.97 | 10.11 | 10.24 | 10.37 | 10.51
15.53 | 5.17 | 5.3 | 5.43 | 5.57 | 5.7 | 5.83 | 5.97 | 6.1 | 6.23 | 6.37 | 6.5 | 6.63 | 6.77 | 6.9 | 7.03 | 7.17 | 7.3 | 7.43 | 7.57 | 7.7 | 7.83 | 7.97 | 8.1 | 8.24 | 8.37 | 8.5 | 8.64 | 8.77 | 8.9 | 9.04 | 9.17 | 9.3 | 9.44 | 9.57 | 9.7 | 9.84 | 9.97 | 10.1 | 10.24 | 10.37 | 10.5
15.54 | 5.16 | 5.3 | 5.43 | 5.56 | 5.7 | 5.83 | 5.97 | 6.1 | 6.23 | 6.37 | 6.5 | 6.63 | 6.77 | 6.9 | 7.03 | 7.17 | 7.3 | 7.43 | 7.57 | 7.7 | 7.83 | 7.97 | 8.1 | 8.23 | 8.37 | 8.5 | 8.63 | 8.77 | 8.9 | 9.03 | 9.17 | 9.3 | 9.43 | 9.57 | 9.7 | 9.83 | 9.97 | 10.1 | 10.23 | 10.37 | 10.5
15.55 | 5.16 | 5.3 | 5.43 | 5.56 | 5.7 | 5.83 | 5.96 | 6.1 | 6.23 | 6.36 | 6.5 | 6.63 | 6.76 | 6.9 | 7.03 | 7.16 | 7.3 | 7.43 | 7.56 | 7.7 | 7.83 | 7.96 | 8.1 | 8.23 | 8.37 | 8.5 | 8.63 | 8.77 | 8.9 | 9.03 | 9.17 | 9.3 | 9.43 | 9.57 | 9.7 | 9.83 | 9.97 | 10.1 | 10.23 | 10.37 | 10.5
15.56 | 5.16 | 5.3 | 5.43 | 5.56 | 5.7 | 5.83 | 5.96 | 6.1 | 6.23 | 6.36 | 6.5 | 6.63 | 6.76 | 6.9 | 7.03 | 7.16 | 7.3 | 7.43 | 7.56 | 7.7 | 7.83 | 7.96 | 8.1 | 8.23 | 8.36 | 8.5 | 8.63 | 8.76 | 8.9 | 9.03 | 9.16 | 9.3 | 9.43 | 9.56 | 9.7 | 9.83 | 9.96 | 10.1 | 10.23 | 10.36 | 10.5
15.57 | 5.16 | 5.29 | 5.43 | 5.56 | 5.69 | 5.83 | 5.96 | 6.09 | 6.23 | 6.36 | 6.49 | 6.63 | 6.76 | 6.89 | 7.03 | 7.16 | 7.29 | 7.43 | 7.56 | 7.69 | 7.83 | 7.96 | 8.09 | 8.23 | 8.36 | 8.49 | 8.63 | 8.76 | 8.89 | 9.03 | 9.16 | 9.29 | 9.43 | 9.56 | 9.69 | 9.83 | 9.96 | 10.09 | 10.23 | 10.36 | 10.49
15.58 | 5.16 | 5.29 | 5.43 | 5.56 | 5.69 | 5.83 | 5.96 | 6.09 | 6.23 | 6.36 | 6.49 | 6.63 | 6.76 | 6.89 | 7.03 | 7.16 | 7.29 | 7.43 | 7.56 | 7.69 | 7.83 | 7.96 | 8.09 | 8.23 | 8.36 | 8.49 | 8.63 | 8.76 | 8.89 | 9.03 | 9.16 | 9.29 | 9.43 | 9.56 | 9.69 | 9.83 | 9.96 | 10.09 | 10.23 | 10.36 | 10.49
15.59 | 5.16 | 5.29 | 5.43 | 5.56 | 5.69 | 5.83 | 5.96 | 6.09 | 6.23 | 6.36 | 6.49 | 6.63 | 6.76 | 6.89 | 7.02 | 7.16 | 7.29 | 7.42 | 7.56 | 7.69 | 7.82 | 7.96 | 8.09 | 8.22 | 8.36 | 8.49 | 8.62 | 8.76 | 8.89 | 9.02 | 9.16 | 9.29 | 9.42 | 9.56 | 9.69 | 9.82 | 9.96 | 10.09 | 10.22 | 10.36 | 10.49
15.6 | 5.16 | 5.29 | 5.42 | 5.56 | 5.69 | 5.82 | 5.96 | 6.09 | 6.22 | 6.36 | 6.49 | 6.62 | 6.76 | 6.89 | 7.02 | 7.16 | 7.29 | 7.42 | 7.56 | 7.69 | 7.82 | 7.96 | 8.09 | 8.22 | 8.36 | 8.49 | 8.62 | 8.76 | 8.89 | 9.02 | 9.16 | 9.29 | 9.42 | 9.56 | 9.69 | 9.82 | 9.95 | 10.09 | 10.22 | 10.35 | 10.49
15.61 | 5.16 | 5.29 | 5.42 | 5.56 | 5.69 | 5.82 | 5.96 | 6.09 | 6.22 | 6.36 | 6.49 | 6.62 | 6.76 | 6.89 | 7.02 | 7.16 | 7.29 | 7.42 | 7.55 | 7.69 | 7.82 | 7.95 | 8.09 | 8.22 | 8.35 | 8.49 | 8.62 | 8.75 | 8.89 | 9.02 | 9.15 | 9.29 | 9.42 | 9.55 | 9.69 | 9.82 | 9.95 | 10.09 | 10.22 | 10.35 | 10.49
15.62 | 5.15 | 5.29 | 5.42 | 5.55 | 5.69 | 5.82 | 5.95 | 6.09 | 6.22 | 6.35 | 6.49 | 6.62 | 6.75 | 6.89 | 7.02 | 7.15 | 7.29 | 7.42 | 7.55 | 7.69 | 7.82 | 7.95 | 8.09 | 8.22 | 8.35 | 8.49 | 8.62 | 8.75 | 8.89 | 9.02 | 9.15 | 9.28 | 9.42 | 9.55 | 9.68 | 9.82 | 9.95 | 10.08 | 10.22 | 10.35 | 10.48
15.63 | 5.15 | 5.29 | 5.42 | 5.55 | 5.69 | 5.82 | 5.95 | 6.09 | 6.22 | 6.35 | 6.49 | 6.62 | 6.75 | 6.89 | 7.02 | 7.15 | 7.29 | 7.42 | 7.55 | 7.68 | 7.82 | 7.95 | 8.08 | 8.22 | 8.35 | 8.48 | 8.62 | 8.75 | 8.88 | 9.02 | 9.15 | 9.28 | 9.42 | 9.55 | 9.68 | 9.82 | 9.95 | 10.08 | 10.21 | 10.35 | 10.48
15.64 | 5.15 | 5.29 | 5.42 | 5.55 | 5.69 | 5.82 | 5.95 | 6.09 | 6.22 | 6.35 | 6.48 | 6.62 | 6.75 | 6.88 | 7.02 | 7.15 | 7.28 | 7.42 | 7.55 | 7.68 | 7.82 | 7.95 | 8.08 | 8.22 | 8.35 | 8.48 | 8.62 | 8.75 | 8.88 | 9.01 | 9.15 | 9.28 | 9.41 | 9.55 | 9.68 | 9.81 | 9.95 | 10.08 | 10.21 | 10.35 | 10.48
15.65 | 5.15 | 5.28 | 5.42 | 5.55 | 5.68 | 5.82 | 5.95 | 6.08 | 6.22 | 6.35 | 6.48 | 6.62 | 6.75 | 6.88 | 7.02 | 7.15 | 7.28 | 7.42 | 7.55 | 7.68 | 7.81 | 7.95 | 8.08 | 8.21 | 8.35 | 8.48 | 8.61 | 8.75 | 8.88 | 9.01 | 9.15 | 9.28 | 9.41 | 9.54 | 9.68 | 9.81 | 9.94 | 10.08 | 10.21 | 10.34 | 10.48
15.66 | 5.15 | 5.28 | 5.42 | 5.55 | 5.68 | 5.82 | 5.95 | 6.08 | 6.22 | 6.35 | 6.48 | 6.61 | 6.75 | 6.88 | 7.01 | 7.15 | 7.28 | 7.41 | 7.55 | 7.68 | 7.81 | 7.95 | 8.08 | 8.21 | 8.35 | 8.48 | 8.61 | 8.74 | 8.88 | 9.01 | 9.14 | 9.28 | 9.41 | 9.54 | 9.68 | 9.81 | 9.94 | 10.08 | 10.21 | 10.34 | 10.47
15.67 | 5.15 | 5.28 | 5.42 | 5.55 | 5.68 | 5.81 | 5.95 | 6.08 | 6.21 | 6.35 | 6.48 | 6.61 | 6.75 | 6.88 | 7.01 | 7.15 | 7.28 | 7.41 | 7.54 | 7.68 | 7.81 | 7.94 | 8.08 | 8.21 | 8.34 | 8.48 | 8.61 | 8.74 | 8.88 | 9.01 | 9.14 | 9.27 | 9.41 | 9.54 | 9.67 | 9.81 | 9.94 | 10.07 | 10.21 | 10.34 | 10.47
15.68 | 5.15 | 5.28 | 5.41 | 5.55 | 5.68 | 5.81 | 5.95 | 6.08 | 6.21 | 6.35 | 6.48 | 6.61 | 6.75 | 6.88 | 7.01 | 7.14 | 7.28 | 7.41 | 7.54 | 7.68 | 7.81 | 7.94 | 8.08 | 8.21 | 8.34 | 8.47 | 8.61 | 8.74 | 8.87 | 9.01 | 9.14 | 9.27 | 9.41 | 9.54 | 9.67 | 9.8 | 9.94 | 10.07 | 10.2 | 10.34 | 10.47
15.69 | 5.15 | 5.28 | 5.41 | 5.55 | 5.68 | 5.81 | 5.95 | 6.08 | 6.21 | 6.34 | 6.48 | 6.61 | 6.74 | 6.88 | 7.01 | 7.14 | 7.28 | 7.41 | 7.54 | 7.67 | 7.81 | 7.94 | 8.07 | 8.21 | 8.34 | 8.47 | 8.61 | 8.74 | 8.87 | 9 | 9.14 | 9.27 | 9.4 | 9.54 | 9.67 | 9.8 | 9.94 | 10.07 | 10.2 | 10.33 | 10.47
15.7 | 5.15 | 5.28 | 5.41 | 5.54 | 5.68 | 5.81 | 5.94 | 6.08 | 6.21 | 6.34 | 6.48 | 6.61 | 6.74 | 6.88 | 7.01 | 7.14 | 7.27 | 7.41 | 7.54 | 7.67 | 7.81 | 7.94 | 8.07 | 8.21 | 8.34 | 8.47 | 8.6 | 8.74 | 8.87 | 9 | 9.14 | 9.27 | 9.4 | 9.53 | 9.67 | 9.8 | 9.93 | 10.07 | 10.2 | 10.33 | 10.47
15.71 | 5.14 | 5.28 | 5.41 | 5.54 | 5.68 | 5.81 | 5.94 | 6.08 | 6.21 | 6.34 | 6.47 | 6.61 | 6.74 | 6.87 | 7.01 | 7.14 | 7.27 | 7.41 | 7.54 | 7.67 | 7.8 | 7.94 | 8.07 | 8.2 | 8.34 | 8.47 | 8.6 | 8.74 | 8.87 | 9 | 9.13 | 9.27 | 9.4 | 9.53 | 9.67 | 9.8 | 9.93 | 10.06 | 10.2 | 10.33 | 10.46
15.72 | 5.14 | 5.28 | 5.41 | 5.54 | 5.68 | 5.81 | 5.94 | 6.07 | 6.21 | 6.34 | 6.47 | 6.61 | 6.74 | 6.87 | 7.01 | 7.14 | 7.27 | 7.4 | 7.54 | 7.67 | 7.8 | 7.94 | 8.07 | 8.2 | 8.33 | 8.47 | 8.6 | 8.73 | 8.87 | 9 | 9.13 | 9.26 | 9.4 | 9.53 | 9.66 | 9.8 | 9.93 | 10.06 | 10.2 | 10.33 | 10.46
15.73 | 5.14 | 5.28 | 5.41 | 5.54 | 5.67 | 5.81 | 5.94 | 6.07 | 6.21 | 6.34 | 6.47 | 6.6 | 6.74 | 6.87 | 7 | 7.14 | 7.27 | 7.4 | 7.54 | 7.67 | 7.8 | 7.93 | 8.07 | 8.2 | 8.33 | 8.47 | 8.6 | 8.73 | 8.86 | 9 | 9.13 | 9.26 | 9.4 | 9.53 | 9.66 | 9.79 | 9.93 | 10.06 | 10.19 | 10.33 | 10.46
15.74 | 5.14 | 5.27 | 5.41 | 5.54 | 5.67 | 5.81 | 5.94 | 6.07 | 6.2 | 6.34 | 6.47 | 6.6 | 6.74 | 6.87 | 7 | 7.13 | 7.27 | 7.4 | 7.53 | 7.67 | 7.8 | 7.93 | 8.07 | 8.2 | 8.33 | 8.46 | 8.6 | 8.73 | 8.86 | 9 | 9.13 | 9.26 | 9.39 | 9.53 | 9.66 | 9.79 | 9.93 | 10.06 | 10.19 | 10.32 | 10.46
15.75 | 5.14 | 5.27 | 5.41 | 5.54 | 5.67 | 5.8 | 5.94 | 6.07 | 6.2 | 6.34 | 6.47 | 6.6 | 6.73 | 6.87 | 7 | 7.13 | 7.27 | 7.4 | 7.53 | 7.66 | 7.8 | 7.93 | 8.06 | 8.2 | 8.33 | 8.46 | 8.59 | 8.73 | 8.86 | 8.99 | 9.13 | 9.26 | 9.39 | 9.52 | 9.66 | 9.79 | 9.92 | 10.06 | 10.19 | 10.32 | 10.45
15.76 | 5.14 | 5.27 | 5.4 | 5.54 | 5.67 | 5.8 | 5.94 | 6.07 | 6.2 | 6.33 | 6.47 | 6.6 | 6.73 | 6.87 | 7 | 7.13 | 7.26 | 7.4 | 7.53 | 7.66 | 7.8 | 7.93 | 8.06 | 8.19 | 8.33 | 8.46 | 8.59 | 8.73 | 8.86 | 8.99 | 9.12 | 9.26 | 9.39 | 9.52 | 9.66 | 9.79 | 9.92 | 10.05 | 10.19 | 10.32 | 10.45
15.77 | 5.14 | 5.27 | 5.4 | 5.54 | 5.67 | 5.8 | 5.93 | 6.07 | 6.2 | 6.33 | 6.47 | 6.6 | 6.73 | 6.86 | 7 | 7.13 | 7.26 | 7.4 | 7.53 | 7.66 | 7.79 | 7.93 | 8.06 | 8.19 | 8.33 | 8.46 | 8.59 | 8.72 | 8.86 | 8.99 | 9.12 | 9.25 | 9.39 | 9.52 | 9.65 | 9.79 | 9.92 | 10.05 | 10.18 | 10.32 | 10.45
15.78 | 5.14 | 5.27 | 5.4 | 5.54 | 5.67 | 5.8 | 5.93 | 6.07 | 6.2 | 6.33 | 6.46 | 6.6 | 6.73 | 6.86 | 7 | 7.13 | 7.26 | 7.39 | 7.53 | 7.66 | 7.79 | 7.93 | 8.06 | 8.19 | 8.32 | 8.46 | 8.59 | 8.72 | 8.85 | 8.99 | 9.12 | 9.25 | 9.39 | 9.52 | 9.65 | 9.78 | 9.92 | 10.05 | 10.18 | 10.31 | 10.45
15.79 | 5.14 | 5.27 | 5.4 | 5.53 | 5.67 | 5.8 | 5.93 | 6.06 | 6.2 | 6.33 | 6.46 | 6.6 | 6.73 | 6.86 | 6.99 | 7.13 | 7.26 | 7.39 | 7.53 | 7.66 | 7.79 | 7.92 | 8.06 | 8.19 | 8.32 | 8.45 | 8.59 | 8.72 | 8.85 | 8.99 | 9.12 | 9.25 | 9.38 | 9.52 | 9.65 | 9.78 | 9.91 | 10.05 | 10.18 | 10.31 | 10.45
15.8 | 5.13 | 5.27 | 5.4 | 5.53 | 5.67 | 5.8 | 5.93 | 6.06 | 6.2 | 6.33 | 6.46 | 6.59 | 6.73 | 6.86 | 6.99 | 7.13 | 7.26 | 7.39 | 7.52 | 7.66 | 7.79 | 7.92 | 8.05 | 8.19 | 8.32 | 8.45 | 8.59 | 8.72 | 8.85 | 8.98 | 9.12 | 9.25 | 9.38 | 9.51 | 9.65 | 9.78 | 9.91 | 10.05 | 10.18 | 10.31 | 10.44
15.81 | 5.13 | 5.27 | 5.4 | 5.53 | 5.66 | 5.8 | 5.93 | 6.06 | 6.2 | 6.33 | 6.46 | 6.59 | 6.73 | 6.86 | 6.99 | 7.12 | 7.26 | 7.39 | 7.52 | 7.65 | 7.79 | 7.92 | 8.05 | 8.19 | 8.32 | 8.45 | 8.58 | 8.72 | 8.85 | 8.98 | 9.11 | 9.25 | 9.38 | 9.51 | 9.64 | 9.78 | 9.91 | 10.04 | 10.18 | 10.31 | 10.44
15.82 | 5.13 | 5.26 | 5.4 | 5.53 | 5.66 | 5.8 | 5.93 | 6.06 | 6.19 | 6.33 | 6.46 | 6.59 | 6.72 | 6.86 | 6.99 | 7.12 | 7.26 | 7.39 | 7.52 | 7.65 | 7.79 | 7.92 | 8.05 | 8.18 | 8.32 | 8.45 | 8.58 | 8.71 | 8.85 | 8.98 | 9.11 | 9.25 | 9.38 | 9.51 | 9.64 | 9.78 | 9.91 | 10.04 | 10.17 | 10.31 | 10.44
15.83 | 5.13 | 5.26 | 5.4 | 5.53 | 5.66 | 5.79 | 5.93 | 6.06 | 6.19 | 6.32 | 6.46 | 6.59 | 6.72 | 6.86 | 6.99 | 7.12 | 7.25 | 7.39 | 7.52 | 7.65 | 7.78 | 7.92 | 8.05 | 8.18 | 8.31 | 8.45 | 8.58 | 8.71 | 8.85 | 8.98 | 9.11 | 9.24 | 9.38 | 9.51 | 9.64 | 9.77 | 9.91 | 10.04 | 10.17 | 10.3 | 10.44
15.84 | 5.13 | 5.26 | 5.39 | 5.53 | 5.66 | 5.79 | 5.93 | 6.06 | 6.19 | 6.32 | 6.46 | 6.59 | 6.72 | 6.85 | 6.99 | 7.12 | 7.25 | 7.38 | 7.52 | 7.65 | 7.78 | 7.92 | 8.05 | 8.18 | 8.31 | 8.45 | 8.58 | 8.71 | 8.84 | 8.98 | 9.11 | 9.24 | 9.37 | 9.51 | 9.64 | 9.77 | 9.9 | 10.04 | 10.17 | 10.3 | 10.43
15.85 | 5.13 | 5.26 | 5.39 | 5.53 | 5.66 | 5.79 | 5.92 | 6.06 | 6.19 | 6.32 | 6.45 | 6.59 | 6.72 | 6.85 | 6.99 | 7.12 | 7.25 | 7.38 | 7.52 | 7.65 | 7.78 | 7.91 | 8.05 | 8.18 | 8.31 | 8.44 | 8.58 | 8.71 | 8.84 | 8.97 | 9.11 | 9.24 | 9.37 | 9.5 | 9.64 | 9.77 | 9.9 | 10.03 | 10.17 | 10.3 | 10.43
15.86 | 5.13 | 5.26 | 5.39 | 5.53 | 5.66 | 5.79 | 5.92 | 6.06 | 6.19 | 6.32 | 6.45 | 6.59 | 6.72 | 6.85 | 6.98 | 7.12 | 7.25 | 7.38 | 7.51 | 7.65 | 7.78 | 7.91 | 8.04 | 8.18 | 8.31 | 8.44 | 8.57 | 8.71 | 8.84 | 8.97 | 9.1 | 9.24 | 9.37 | 9.5 | 9.63 | 9.77 | 9.9 | 10.03 | 10.16 | 10.3 | 10.43
15.87 | 5.13 | 5.26 | 5.39 | 5.52 | 5.66 | 5.79 | 5.92 | 6.05 | 6.19 | 6.32 | 6.45 | 6.58 | 6.72 | 6.85 | 6.98 | 7.11 | 7.25 | 7.38 | 7.51 | 7.64 | 7.78 | 7.91 | 8.04 | 8.18 | 8.31 | 8.44 | 8.57 | 8.71 | 8.84 | 8.97 | 9.1 | 9.24 | 9.37 | 9.5 | 9.63 | 9.77 | 9.9 | 10.03 | 10.16 | 10.3 | 10.43
15.88 | 5.13 | 5.26 | 5.39 | 5.52 | 5.66 | 5.79 | 5.92 | 6.05 | 6.19 | 6.32 | 6.45 | 6.58 | 6.72 | 6.85 | 6.98 | 7.11 | 7.25 | 7.38 | 7.51 | 7.64 | 7.78 | 7.91 | 8.04 | 8.17 | 8.31 | 8.44 | 8.57 | 8.7 | 8.84 | 8.97 | 9.1 | 9.23 | 9.37 | 9.5 | 9.63 | 9.76 | 9.9 | 10.03 | 10.16 | 10.29 | 10.43
15.89 | 5.12 | 5.26 | 5.39 | 5.52 | 5.65 | 5.79 | 5.92 | 6.05 | 6.18 | 6.32 | 6.45 | 6.58 | 6.71 | 6.85 | 6.98 | 7.11 | 7.24 | 7.38 | 7.51 | 7.64 | 7.77 | 7.91 | 8.04 | 8.17 | 8.3 | 8.44 | 8.57 | 8.7 | 8.83 | 8.97 | 9.1 | 9.23 | 9.36 | 9.5 | 9.63 | 9.76 | 9.89 | 10.03 | 10.16 | 10.29 | 10.42
15.9 | 5.12 | 5.26 | 5.39 | 5.52 | 5.65 | 5.79 | 5.92 | 6.05 | 6.18 | 6.32 | 6.45 | 6.58 | 6.71 | 6.85 | 6.98 | 7.11 | 7.24 | 7.38 | 7.51 | 7.64 | 7.77 | 7.9 | 8.04 | 8.17 | 8.3 | 8.43 | 8.57 | 8.7 | 8.83 | 8.96 | 9.1 | 9.23 | 9.36 | 9.49 | 9.63 | 9.76 | 9.89 | 10.02 | 10.16 | 10.29 | 10.42
15.91 | 5.12 | 5.25 | 5.39 | 5.52 | 5.65 | 5.78 | 5.92 | 6.05 | 6.18 | 6.31 | 6.45 | 6.58 | 6.71 | 6.84 | 6.98 | 7.11 | 7.24 | 7.37 | 7.51 | 7.64 | 7.77 | 7.9 | 8.04 | 8.17 | 8.3 | 8.43 | 8.57 | 8.7 | 8.83 | 8.96 | 9.09 | 9.23 | 9.36 | 9.49 | 9.62 | 9.76 | 9.89 | 10.02 | 10.15 | 10.29 | 10.42
15.92 | 5.12 | 5.25 | 5.39 | 5.52 | 5.65 | 5.78 | 5.92 | 6.05 | 6.18 | 6.31 | 6.44 | 6.58 | 6.71 | 6.84 | 6.97 | 7.11 | 7.24 | 7.37 | 7.5 | 7.64 | 7.77 | 7.9 | 8.03 | 8.17 | 8.3 | 8.43 | 8.56 | 8.7 | 8.83 | 8.96 | 9.09 | 9.23 | 9.36 | 9.49 | 9.62 | 9.75 | 9.89 | 10.02 | 10.15 | 10.28 | 10.42
15.93 | 5.12 | 5.25 | 5.38 | 5.52 | 5.65 | 5.78 | 5.91 | 6.05 | 6.18 | 6.31 | 6.44 | 6.58 | 6.71 | 6.84 | 6.97 | 7.11 | 7.24 | 7.37 | 7.5 | 7.64 | 7.77 | 7.9 | 8.03 | 8.16 | 8.3 | 8.43 | 8.56 | 8.69 | 8.83 | 8.96 | 9.09 | 9.22 | 9.36 | 9.49 | 9.62 | 9.75 | 9.89 | 10.02 | 10.15 | 10.28 | 10.41
15.94 | 5.12 | 5.25 | 5.38 | 5.52 | 5.65 | 5.78 | 5.91 | 6.04 | 6.18 | 6.31 | 6.44 | 6.57 | 6.71 | 6.84 | 6.97 | 7.1 | 7.24 | 7.37 | 7.5 | 7.63 | 7.77 | 7.9 | 8.03 | 8.16 | 8.3 | 8.43 | 8.56 | 8.69 | 8.82 | 8.96 | 9.09 | 9.22 | 9.35 | 9.49 | 9.62 | 9.75 | 9.88 | 10.02 | 10.15 | 10.28 | 10.41
15.95 | 5.12 | 5.25 | 5.38 | 5.51 | 5.65 | 5.78 | 5.91 | 6.04 | 6.18 | 6.31 | 6.44 | 6.57 | 6.71 | 6.84 | 6.97 | 7.1 | 7.23 | 7.37 | 7.5 | 7.63 | 7.76 | 7.9 | 8.03 | 8.16 | 8.29 | 8.43 | 8.56 | 8.69 | 8.82 | 8.95 | 9.09 | 9.22 | 9.35 | 9.48 | 9.62 | 9.75 | 9.88 | 10.01 | 10.15 | 10.28 | 10.41
15.96 | 5.12 | 5.25 | 5.38 | 5.51 | 5.65 | 5.78 | 5.91 | 6.04 | 6.17 | 6.31 | 6.44 | 6.57 | 6.7 | 6.84 | 6.97 | 7.1 | 7.23 | 7.37 | 7.5 | 7.63 | 7.76 | 7.89 | 8.03 | 8.16 | 8.29 | 8.42 | 8.56 | 8.69 | 8.82 | 8.95 | 9.09 | 9.22 | 9.35 | 9.48 | 9.61 | 9.75 | 9.88 | 10.01 | 10.14 | 10.28 | 10.41
15.97 | 5.11 | 5.25 | 5.38 | 5.51 | 5.64 | 5.78 | 5.91 | 6.04 | 6.17 | 6.31 | 6.44 | 6.57 | 6.7 | 6.83 | 6.97 | 7.1 | 7.23 | 7.36 | 7.5 | 7.63 | 7.76 | 7.89 | 8.03 | 8.16 | 8.29 | 8.42 | 8.55 | 8.69 | 8.82 | 8.95 | 9.08 | 9.22 | 9.35 | 9.48 | 9.61 | 9.74 | 9.88 | 10.01 | 10.14 | 10.27 | 10.41
15.98 | 5.11 | 5.25 | 5.38 | 5.51 | 5.64 | 5.78 | 5.91 | 6.04 | 6.17 | 6.3 | 6.44 | 6.57 | 6.7 | 6.83 | 6.97 | 7.1 | 7.23 | 7.36 | 7.49 | 7.63 | 7.76 | 7.89 | 8.02 | 8.16 | 8.29 | 8.42 | 8.55 | 8.68 | 8.82 | 8.95 | 9.08 | 9.21 | 9.35 | 9.48 | 9.61 | 9.74 | 9.87 | 10.01 | 10.14 | 10.27 | 10.4
15.99 | 5.11 | 5.24 | 5.38 | 5.51 | 5.64 | 5.77 | 5.91 | 6.04 | 6.17 | 6.3 | 6.44 | 6.57 | 6.7 | 6.83 | 6.96 | 7.1 | 7.23 | 7.36 | 7.49 | 7.63 | 7.76 | 7.89 | 8.02 | 8.15 | 8.29 | 8.42 | 8.55 | 8.68 | 8.81 | 8.95 | 9.08 | 9.21 | 9.34 | 9.48 | 9.61 | 9.74 | 9.87 | 10 | 10.14 | 10.27 | 10.4
16 | 5.11 | 5.24 | 5.38 | 5.51 | 5.64 | 5.77 | 5.9 | 6.04 | 6.17 | 6.3 | 6.43 | 6.57 | 6.7 | 6.83 | 6.96 | 7.09 | 7.23 | 7.36 | 7.49 | 7.62 | 7.76 | 7.89 | 8.02 | 8.15 | 8.28 | 8.42 | 8.55 | 8.68 | 8.81 | 8.95 | 9.08 | 9.21 | 9.34 | 9.47 | 9.61 | 9.74 | 9.87 | 10 | 10.13 | 10.27 | 10.4
16.01 | 5.11 | 5.24 | 5.37 | 5.51 | 5.64 | 5.77 | 5.9 | 6.04 | 6.17 | 6.3 | 6.43 | 6.56 | 6.7 | 6.83 | 6.96 | 7.09 | 7.23 | 7.36 | 7.49 | 7.62 | 7.75 | 7.89 | 8.02 | 8.15 | 8.28 | 8.41 | 8.55 | 8.68 | 8.81 | 8.94 | 9.08 | 9.21 | 9.34 | 9.47 | 9.6 | 9.74 | 9.87 | 10 | 10.13 | 10.26 | 10.4
16.02 | 5.11 | 5.24 | 5.37 | 5.51 | 5.64 | 5.77 | 5.9 | 6.03 | 6.17 | 6.3 | 6.43 | 6.56 | 6.7 | 6.83 | 6.96 | 7.09 | 7.22 | 7.36 | 7.49 | 7.62 | 7.75 | 7.88 | 8.02 | 8.15 | 8.28 | 8.41 | 8.55 | 8.68 | 8.81 | 8.94 | 9.07 | 9.21 | 9.34 | 9.47 | 9.6 | 9.73 | 9.87 | 10 | 10.13 | 10.26 | 10.39
16.03 | 5.11 | 5.24 | 5.37 | 5.5 | 5.64 | 5.77 | 5.9 | 6.03 | 6.17 | 6.3 | 6.43 | 6.56 | 6.69 | 6.83 | 6.96 | 7.09 | 7.22 | 7.35 | 7.49 | 7.62 | 7.75 | 7.88 | 8.01 | 8.15 | 8.28 | 8.41 | 8.54 | 8.68 | 8.81 | 8.94 | 9.07 | 9.2 | 9.34 | 9.47 | 9.6 | 9.73 | 9.86 | 10 | 10.13 | 10.26 | 10.39
16.04 | 5.11 | 5.24 | 5.37 | 5.5 | 5.64 | 5.77 | 5.9 | 6.03 | 6.16 | 6.3 | 6.43 | 6.56 | 6.69 | 6.82 | 6.96 | 7.09 | 7.22 | 7.35 | 7.48 | 7.62 | 7.75 | 7.88 | 8.01 | 8.15 | 8.28 | 8.41 | 8.54 | 8.67 | 8.81 | 8.94 | 9.07 | 9.2 | 9.33 | 9.47 | 9.6 | 9.73 | 9.86 | 9.99 | 10.13 | 10.26 | 10.39
16.05 | 5.11 | 5.24 | 5.37 | 5.5 | 5.63 | 5.77 | 5.9 | 6.03 | 6.16 | 6.29 | 6.43 | 6.56 | 6.69 | 6.82 | 6.95 | 7.09 | 7.22 | 7.35 | 7.48 | 7.62 | 7.75 | 7.88 | 8.01 | 8.14 | 8.28 | 8.41 | 8.54 | 8.67 | 8.8 | 8.94 | 9.07 | 9.2 | 9.33 | 9.46 | 9.6 | 9.73 | 9.86 | 9.99 | 10.12 | 10.26 | 10.39
16.06 | 5.1 | 5.24 | 5.37 | 5.5 | 5.63 | 5.76 | 5.9 | 6.03 | 6.16 | 6.29 | 6.43 | 6.56 | 6.69 | 6.82 | 6.95 | 7.09 | 7.22 | 7.35 | 7.48 | 7.61 | 7.75 | 7.88 | 8.01 | 8.14 | 8.27 | 8.41 | 8.54 | 8.67 | 8.8 | 8.93 | 9.07 | 9.2 | 9.33 | 9.46 | 9.59 | 9.73 | 9.86 | 9.99 | 10.12 | 10.25 | 10.39
16.07 | 5.1 | 5.24 | 5.37 | 5.5 | 5.63 | 5.76 | 5.9 | 6.03 | 6.16 | 6.29 | 6.42 | 6.56 | 6.69 | 6.82 | 6.95 | 7.08 | 7.22 | 7.35 | 7.48 | 7.61 | 7.74 | 7.88 | 8.01 | 8.14 | 8.27 | 8.4 | 8.54 | 8.67 | 8.8 | 8.93 | 9.06 | 9.2 | 9.33 | 9.46 | 9.59 | 9.72 | 9.86 | 9.99 | 10.12 | 10.25 | 10.38
16.08 | 5.1 | 5.23 | 5.37 | 5.5 | 5.63 | 5.76 | 5.89 | 6.03 | 6.16 | 6.29 | 6.42 | 6.55 | 6.69 | 6.82 | 6.95 | 7.08 | 7.21 | 7.35 | 7.48 | 7.61 | 7.74 | 7.87 | 8.01 | 8.14 | 8.27 | 8.4 | 8.53 | 8.67 | 8.8 | 8.93 | 9.06 | 9.19 | 9.33 | 9.46 | 9.59 | 9.72 | 9.85 | 9.99 | 10.12 | 10.25 | 10.38
16.09 | 5.1 | 5.23 | 5.37 | 5.5 | 5.63 | 5.76 | 5.89 | 6.03 | 6.16 | 6.29 | 6.42 | 6.55 | 6.68 | 6.82 | 6.95 | 7.08 | 7.21 | 7.34 | 7.48 | 7.61 | 7.74 | 7.87 | 8 | 8.14 | 8.27 | 8.4 | 8.53 | 8.66 | 8.8 | 8.93 | 9.06 | 9.19 | 9.32 | 9.46 | 9.59 | 9.72 | 9.85 | 9.98 | 10.12 | 10.25 | 10.38
16.1 | 5.1 | 5.23 | 5.36 | 5.5 | 5.63 | 5.76 | 5.89 | 6.02 | 6.16 | 6.29 | 6.42 | 6.55 | 6.68 | 6.82 | 6.95 | 7.08 | 7.21 | 7.34 | 7.48 | 7.61 | 7.74 | 7.87 | 8 | 8.13 | 8.27 | 8.4 | 8.53 | 8.66 | 8.79 | 8.93 | 9.06 | 9.19 | 9.32 | 9.45 | 9.59 | 9.72 | 9.85 | 9.98 | 10.11 | 10.25 | 10.38
16.11 | 5.1 | 5.23 | 5.36 | 5.49 | 5.63 | 5.76 | 5.89 | 6.02 | 6.15 | 6.29 | 6.42 | 6.55 | 6.68 | 6.81 | 6.95 | 7.08 | 7.21 | 7.34 | 7.47 | 7.61 | 7.74 | 7.87 | 8 | 8.13 | 8.26 | 8.4 | 8.53 | 8.66 | 8.79 | 8.92 | 9.06 | 9.19 | 9.32 | 9.45 | 9.58 | 9.72 | 9.85 | 9.98 | 10.11 | 10.24 | 10.37
16.12 | 5.1 | 5.23 | 5.36 | 5.49 | 5.63 | 5.76 | 5.89 | 6.02 | 6.15 | 6.28 | 6.42 | 6.55 | 6.68 | 6.81 | 6.94 | 7.08 | 7.21 | 7.34 | 7.47 | 7.6 | 7.74 | 7.87 | 8 | 8.13 | 8.26 | 8.4 | 8.53 | 8.66 | 8.79 | 8.92 | 9.05 | 9.19 | 9.32 | 9.45 | 9.58 | 9.71 | 9.85 | 9.98 | 10.11 | 10.24 | 10.37
16.13 | 5.1 | 5.23 | 5.36 | 5.49 | 5.62 | 5.76 | 5.89 | 6.02 | 6.15 | 6.28 | 6.42 | 6.55 | 6.68 | 6.81 | 6.94 | 7.07 | 7.21 | 7.34 | 7.47 | 7.6 | 7.73 | 7.87 | 8 | 8.13 | 8.26 | 8.39 | 8.53 | 8.66 | 8.79 | 8.92 | 9.05 | 9.18 | 9.32 | 9.45 | 9.58 | 9.71 | 9.84 | 9.98 | 10.11 | 10.24 | 10.37
16.14 | 5.1 | 5.23 | 5.36 | 5.49 | 5.62 | 5.75 | 5.89 | 6.02 | 6.15 | 6.28 | 6.41 | 6.55 | 6.68 | 6.81 | 6.94 | 7.07 | 7.21 | 7.34 | 7.47 | 7.6 | 7.73 | 7.86 | 8 | 8.13 | 8.26 | 8.39 | 8.52 | 8.66 | 8.79 | 8.92 | 9.05 | 9.18 | 9.31 | 9.45 | 9.58 | 9.71 | 9.84 | 9.97 | 10.1 | 10.24 | 10.37
16.15 | 5.09 | 5.23 | 5.36 | 5.49 | 5.62 | 5.75 | 5.89 | 6.02 | 6.15 | 6.28 | 6.41 | 6.54 | 6.68 | 6.81 | 6.94 | 7.07 | 7.2 | 7.34 | 7.47 | 7.6 | 7.73 | 7.86 | 7.99 | 8.13 | 8.26 | 8.39 | 8.52 | 8.65 | 8.78 | 8.92 | 9.05 | 9.18 | 9.31 | 9.44 | 9.58 | 9.71 | 9.84 | 9.97 | 10.1 | 10.23 | 10.37
16.16 | 5.09 | 5.22 | 5.36 | 5.49 | 5.62 | 5.75 | 5.88 | 6.02 | 6.15 | 6.28 | 6.41 | 6.54 | 6.67 | 6.81 | 6.94 | 7.07 | 7.2 | 7.33 | 7.47 | 7.6 | 7.73 | 7.86 | 7.99 | 8.12 | 8.26 | 8.39 | 8.52 | 8.65 | 8.78 | 8.91 | 9.05 | 9.18 | 9.31 | 9.44 | 9.57 | 9.71 | 9.84 | 9.97 | 10.1 | 10.23 | 10.36
16.17 | 5.09 | 5.22 | 5.36 | 5.49 | 5.62 | 5.75 | 5.88 | 6.01 | 6.15 | 6.28 | 6.41 | 6.54 | 6.67 | 6.81 | 6.94 | 7.07 | 7.2 | 7.33 | 7.46 | 7.6 | 7.73 | 7.86 | 7.99 | 8.12 | 8.25 | 8.39 | 8.52 | 8.65 | 8.78 | 8.91 | 9.04 | 9.18 | 9.31 | 9.44 | 9.57 | 9.7 | 9.83 | 9.97 | 10.1 | 10.23 | 10.36
16.18 | 5.09 | 5.22 | 5.35 | 5.49 | 5.62 | 5.75 | 5.88 | 6.01 | 6.14 | 6.28 | 6.41 | 6.54 | 6.67 | 6.8 | 6.94 | 7.07 | 7.2 | 7.33 | 7.46 | 7.59 | 7.73 | 7.86 | 7.99 | 8.12 | 8.25 | 8.38 | 8.52 | 8.65 | 8.78 | 8.91 | 9.04 | 9.17 | 9.31 | 9.44 | 9.57 | 9.7 | 9.83 | 9.96 | 10.1 | 10.23 | 10.36
16.19 | 5.09 | 5.22 | 5.35 | 5.48 | 5.62 | 5.75 | 5.88 | 6.01 | 6.14 | 6.28 | 6.41 | 6.54 | 6.67 | 6.8 | 6.93 | 7.07 | 7.2 | 7.33 | 7.46 | 7.59 | 7.72 | 7.86 | 7.99 | 8.12 | 8.25 | 8.38 | 8.51 | 8.65 | 8.78 | 8.91 | 9.04 | 9.17 | 9.3 | 9.44 | 9.57 | 9.7 | 9.83 | 9.96 | 10.09 | 10.23 | 10.36
16.2 | 5.09 | 5.22 | 5.35 | 5.48 | 5.62 | 5.75 | 5.88 | 6.01 | 6.14 | 6.27 | 6.41 | 6.54 | 6.67 | 6.8 | 6.93 | 7.06 | 7.2 | 7.33 | 7.46 | 7.59 | 7.72 | 7.85 | 7.99 | 8.12 | 8.25 | 8.38 | 8.51 | 8.64 | 8.78 | 8.91 | 9.04 | 9.17 | 9.3 | 9.43 | 9.57 | 9.7 | 9.83 | 9.96 | 10.09 | 10.22 | 10.36
16.21 | 5.09 | 5.22 | 5.35 | 5.48 | 5.61 | 5.75 | 5.88 | 6.01 | 6.14 | 6.27 | 6.4 | 6.54 | 6.67 | 6.8 | 6.93 | 7.06 | 7.19 | 7.33 | 7.46 | 7.59 | 7.72 | 7.85 | 7.98 | 8.12 | 8.25 | 8.38 | 8.51 | 8.64 | 8.77 | 8.91 | 9.04 | 9.17 | 9.3 | 9.43 | 9.56 | 9.69 | 9.83 | 9.96 | 10.09 | 10.22 | 10.35
16.22 | 5.09 | 5.22 | 5.35 | 5.48 | 5.61 | 5.74 | 5.88 | 6.01 | 6.14 | 6.27 | 6.4 | 6.53 | 6.67 | 6.8 | 6.93 | 7.06 | 7.19 | 7.32 | 7.46 | 7.59 | 7.72 | 7.85 | 7.98 | 8.11 | 8.25 | 8.38 | 8.51 | 8.64 | 8.77 | 8.9 | 9.04 | 9.17 | 9.3 | 9.43 | 9.56 | 9.69 | 9.82 | 9.96 | 10.09 | 10.22 | 10.35
16.23 | 5.09 | 5.22 | 5.35 | 5.48 | 5.61 | 5.74 | 5.88 | 6.01 | 6.14 | 6.27 | 6.4 | 6.53 | 6.66 | 6.8 | 6.93 | 7.06 | 7.19 | 7.32 | 7.45 | 7.59 | 7.72 | 7.85 | 7.98 | 8.11 | 8.24 | 8.38 | 8.51 | 8.64 | 8.77 | 8.9 | 9.03 | 9.16 | 9.3 | 9.43 | 9.56 | 9.69 | 9.82 | 9.95 | 10.09 | 10.22 | 10.35
16.24 | 5.08 | 5.22 | 5.35 | 5.48 | 5.61 | 5.74 | 5.87 | 6.01 | 6.14 | 6.27 | 6.4 | 6.53 | 6.66 | 6.79 | 6.93 | 7.06 | 7.19 | 7.32 | 7.45 | 7.58 | 7.72 | 7.85 | 7.98 | 8.11 | 8.24 | 8.37 | 8.51 | 8.64 | 8.77 | 8.9 | 9.03 | 9.16 | 9.29 | 9.43 | 9.56 | 9.69 | 9.82 | 9.95 | 10.08 | 10.21 | 10.35
16.25 | 5.08 | 5.21 | 5.35 | 5.48 | 5.61 | 5.74 | 5.87 | 6 | 6.14 | 6.27 | 6.4 | 6.53 | 6.66 | 6.79 | 6.92 | 7.06 | 7.19 | 7.32 | 7.45 | 7.58 | 7.71 | 7.85 | 7.98 | 8.11 | 8.24 | 8.37 | 8.5 | 8.63 | 8.77 | 8.9 | 9.03 | 9.16 | 9.29 | 9.42 | 9.56 | 9.69 | 9.82 | 9.95 | 10.08 | 10.21 | 10.34
16.26 | 5.08 | 5.21 | 5.34 | 5.48 | 5.61 | 5.74 | 5.87 | 6 | 6.13 | 6.27 | 6.4 | 6.53 | 6.66 | 6.79 | 6.92 | 7.05 | 7.19 | 7.32 | 7.45 | 7.58 | 7.71 | 7.84 | 7.98 | 8.11 | 8.24 | 8.37 | 8.5 | 8.63 | 8.76 | 8.9 | 9.03 | 9.16 | 9.29 | 9.42 | 9.55 | 9.68 | 9.82 | 9.95 | 10.08 | 10.21 | 10.34
16.27 | 5.08 | 5.21 | 5.34 | 5.48 | 5.61 | 5.74 | 5.87 | 6 | 6.13 | 6.26 | 6.4 | 6.53 | 6.66 | 6.79 | 6.92 | 7.05 | 7.18 | 7.32 | 7.45 | 7.58 | 7.71 | 7.84 | 7.97 | 8.11 | 8.24 | 8.37 | 8.5 | 8.63 | 8.76 | 8.89 | 9.03 | 9.16 | 9.29 | 9.42 | 9.55 | 9.68 | 9.81 | 9.95 | 10.08 | 10.21 | 10.34
16.28 | 5.08 | 5.21 | 5.34 | 5.47 | 5.61 | 5.74 | 5.87 | 6 | 6.13 | 6.26 | 6.39 | 6.53 | 6.66 | 6.79 | 6.92 | 7.05 | 7.18 | 7.31 | 7.45 | 7.58 | 7.71 | 7.84 | 7.97 | 8.1 | 8.23 | 8.37 | 8.5 | 8.63 | 8.76 | 8.89 | 9.02 | 9.15 | 9.29 | 9.42 | 9.55 | 9.68 | 9.81 | 9.94 | 10.07 | 10.21 | 10.34
16.29 | 5.08 | 5.21 | 5.34 | 5.47 | 5.6 | 5.74 | 5.87 | 6 | 6.13 | 6.26 | 6.39 | 6.52 | 6.66 | 6.79 | 6.92 | 7.05 | 7.18 | 7.31 | 7.44 | 7.58 | 7.71 | 7.84 | 7.97 | 8.1 | 8.23 | 8.36 | 8.5 | 8.63 | 8.76 | 8.89 | 9.02 | 9.15 | 9.28 | 9.42 | 9.55 | 9.68 | 9.81 | 9.94 | 10.07 | 10.2 | 10.34
16.3 | 5.08 | 5.21 | 5.34 | 5.47 | 5.6 | 5.73 | 5.87 | 6 | 6.13 | 6.26 | 6.39 | 6.52 | 6.65 | 6.79 | 6.92 | 7.05 | 7.18 | 7.31 | 7.44 | 7.57 | 7.71 | 7.84 | 7.97 | 8.1 | 8.23 | 8.36 | 8.49 | 8.63 | 8.76 | 8.89 | 9.02 | 9.15 | 9.28 | 9.41 | 9.55 | 9.68 | 9.81 | 9.94 | 10.07 | 10.2 | 10.33
16.31 | 5.08 | 5.21 | 5.34 | 5.47 | 5.6 | 5.73 | 5.86 | 6 | 6.13 | 6.26 | 6.39 | 6.52 | 6.65 | 6.78 | 6.92 | 7.05 | 7.18 | 7.31 | 7.44 | 7.57 | 7.7 | 7.84 | 7.97 | 8.1 | 8.23 | 8.36 | 8.49 | 8.62 | 8.76 | 8.89 | 9.02 | 9.15 | 9.28 | 9.41 | 9.54 | 9.67 | 9.81 | 9.94 | 10.07 | 10.2 | 10.33
16.32 | 5.08 | 5.21 | 5.34 | 5.47 | 5.6 | 5.73 | 5.86 | 5.99 | 6.13 | 6.26 | 6.39 | 6.52 | 6.65 | 6.78 | 6.91 | 7.05 | 7.18 | 7.31 | 7.44 | 7.57 | 7.7 | 7.83 | 7.97 | 8.1 | 8.23 | 8.36 | 8.49 | 8.62 | 8.75 | 8.88 | 9.02 | 9.15 | 9.28 | 9.41 | 9.54 | 9.67 | 9.8 | 9.94 | 10.07 | 10.2 | 10.33
16.33 | 5.07 | 5.21 | 5.34 | 5.47 | 5.6 | 5.73 | 5.86 | 5.99 | 6.12 | 6.26 | 6.39 | 6.52 | 6.65 | 6.78 | 6.91 | 7.04 | 7.18 | 7.31 | 7.44 | 7.57 | 7.7 | 7.83 | 7.96 | 8.09 | 8.23 | 8.36 | 8.49 | 8.62 | 8.75 | 8.88 | 9.01 | 9.15 | 9.28 | 9.41 | 9.54 | 9.67 | 9.8 | 9.93 | 10.06 | 10.2 | 10.33
16.34 | 5.07 | 5.2 | 5.34 | 5.47 | 5.6 | 5.73 | 5.86 | 5.99 | 6.12 | 6.25 | 6.39 | 6.52 | 6.65 | 6.78 | 6.91 | 7.04 | 7.17 | 7.31 | 7.44 | 7.57 | 7.7 | 7.83 | 7.96 | 8.09 | 8.22 | 8.36 | 8.49 | 8.62 | 8.75 | 8.88 | 9.01 | 9.14 | 9.27 | 9.41 | 9.54 | 9.67 | 9.8 | 9.93 | 10.06 | 10.19 | 10.32
16.35 | 5.07 | 5.2 | 5.33 | 5.47 | 5.6 | 5.73 | 5.86 | 5.99 | 6.12 | 6.25 | 6.38 | 6.52 | 6.65 | 6.78 | 6.91 | 7.04 | 7.17 | 7.3 | 7.44 | 7.57 | 7.7 | 7.83 | 7.96 | 8.09 | 8.22 | 8.35 | 8.49 | 8.62 | 8.75 | 8.88 | 9.01 | 9.14 | 9.27 | 9.4 | 9.54 | 9.67 | 9.8 | 9.93 | 10.06 | 10.19 | 10.32
16.36 | 5.07 | 5.2 | 5.33 | 5.46 | 5.6 | 5.73 | 5.86 | 5.99 | 6.12 | 6.25 | 6.38 | 6.51 | 6.65 | 6.78 | 6.91 | 7.04 | 7.17 | 7.3 | 7.43 | 7.56 | 7.7 | 7.83 | 7.96 | 8.09 | 8.22 | 8.35 | 8.48 | 8.61 | 8.75 | 8.88 | 9.01 | 9.14 | 9.27 | 9.4 | 9.53 | 9.66 | 9.8 | 9.93 | 10.06 | 10.19 | 10.32
16.37 | 5.07 | 5.2 | 5.33 | 5.46 | 5.59 | 5.73 | 5.86 | 5.99 | 6.12 | 6.25 | 6.38 | 6.51 | 6.64 | 6.78 | 6.91 | 7.04 | 7.17 | 7.3 | 7.43 | 7.56 | 7.69 | 7.83 | 7.96 | 8.09 | 8.22 | 8.35 | 8.48 | 8.61 | 8.74 | 8.88 | 9.01 | 9.14 | 9.27 | 9.4 | 9.53 | 9.66 | 9.79 | 9.92 | 10.06 | 10.19 | 10.32
16.38 | 5.07 | 5.2 | 5.33 | 5.46 | 5.59 | 5.72 | 5.86 | 5.99 | 6.12 | 6.25 | 6.38 | 6.51 | 6.64 | 6.77 | 6.91 | 7.04 | 7.17 | 7.3 | 7.43 | 7.56 | 7.69 | 7.82 | 7.96 | 8.09 | 8.22 | 8.35 | 8.48 | 8.61 | 8.74 | 8.87 | 9 | 9.14 | 9.27 | 9.4 | 9.53 | 9.66 | 9.79 | 9.92 | 10.05 | 10.18 | 10.32
16.39 | 5.07 | 5.2 | 5.33 | 5.46 | 5.59 | 5.72 | 5.85 | 5.99 | 6.12 | 6.25 | 6.38 | 6.51 | 6.64 | 6.77 | 6.9 | 7.04 | 7.17 | 7.3 | 7.43 | 7.56 | 7.69 | 7.82 | 7.95 | 8.08 | 8.22 | 8.35 | 8.48 | 8.61 | 8.74 | 8.87 | 9 | 9.13 | 9.26 | 9.4 | 9.53 | 9.66 | 9.79 | 9.92 | 10.05 | 10.18 | 10.31
16.4 | 5.07 | 5.2 | 5.33 | 5.46 | 5.59 | 5.72 | 5.85 | 5.98 | 6.12 | 6.25 | 6.38 | 6.51 | 6.64 | 6.77 | 6.9 | 7.03 | 7.16 | 7.3 | 7.43 | 7.56 | 7.69 | 7.82 | 7.95 | 8.08 | 8.21 | 8.35 | 8.48 | 8.61 | 8.74 | 8.87 | 9 | 9.13 | 9.26 | 9.39 | 9.53 | 9.66 | 9.79 | 9.92 | 10.05 | 10.18 | 10.31
16.41 | 5.06 | 5.2 | 5.33 | 5.46 | 5.59 | 5.72 | 5.85 | 5.98 | 6.11 | 6.25 | 6.38 | 6.51 | 6.64 | 6.77 | 6.9 | 7.03 | 7.16 | 7.29 | 7.43 | 7.56 | 7.69 | 7.82 | 7.95 | 8.08 | 8.21 | 8.34 | 8.47 | 8.61 | 8.74 | 8.87 | 9 | 9.13 | 9.26 | 9.39 | 9.52 | 9.65 | 9.79 | 9.92 | 10.05 | 10.18 | 10.31
16.42 | 5.06 | 5.19 | 5.33 | 5.46 | 5.59 | 5.72 | 5.85 | 5.98 | 6.11 | 6.24 | 6.38 | 6.51 | 6.64 | 6.77 | 6.9 | 7.03 | 7.16 | 7.29 | 7.42 | 7.55 | 7.69 | 7.82 | 7.95 | 8.08 | 8.21 | 8.34 | 8.47 | 8.6 | 8.73 | 8.87 | 9 | 9.13 | 9.26 | 9.39 | 9.52 | 9.65 | 9.78 | 9.91 | 10.05 | 10.18 | 10.31
16.43 | 5.06 | 5.19 | 5.32 | 5.46 | 5.59 | 5.72 | 5.85 | 5.98 | 6.11 | 6.24 | 6.37 | 6.5 | 6.64 | 6.77 | 6.9 | 7.03 | 7.16 | 7.29 | 7.42 | 7.55 | 7.68 | 7.82 | 7.95 | 8.08 | 8.21 | 8.34 | 8.47 | 8.6 | 8.73 | 8.86 | 8.99 | 9.13 | 9.26 | 9.39 | 9.52 | 9.65 | 9.78 | 9.91 | 10.04 | 10.17 | 10.31
16.44 | 5.06 | 5.19 | 5.32 | 5.45 | 5.59 | 5.72 | 5.85 | 5.98 | 6.11 | 6.24 | 6.37 | 6.5 | 6.63 | 6.77 | 6.9 | 7.03 | 7.16 | 7.29 | 7.42 | 7.55 | 7.68 | 7.81 | 7.94 | 8.08 | 8.21 | 8.34 | 8.47 | 8.6 | 8.73 | 8.86 | 8.99 | 9.12 | 9.26 | 9.39 | 9.52 | 9.65 | 9.78 | 9.91 | 10.04 | 10.17 | 10.3
16.45 | 5.06 | 5.19 | 5.32 | 5.45 | 5.58 | 5.72 | 5.85 | 5.98 | 6.11 | 6.24 | 6.37 | 6.5 | 6.63 | 6.76 | 6.9 | 7.03 | 7.16 | 7.29 | 7.42 | 7.55 | 7.68 | 7.81 | 7.94 | 8.07 | 8.21 | 8.34 | 8.47 | 8.6 | 8.73 | 8.86 | 8.99 | 9.12 | 9.25 | 9.38 | 9.51 | 9.65 | 9.78 | 9.91 | 10.04 | 10.17 | 10.3
16.46 | 5.06 | 5.19 | 5.32 | 5.45 | 5.58 | 5.71 | 5.85 | 5.98 | 6.11 | 6.24 | 6.37 | 6.5 | 6.63 | 6.76 | 6.89 | 7.02 | 7.16 | 7.29 | 7.42 | 7.55 | 7.68 | 7.81 | 7.94 | 8.07 | 8.2 | 8.33 | 8.47 | 8.6 | 8.73 | 8.86 | 8.99 | 9.12 | 9.25 | 9.38 | 9.51 | 9.64 | 9.77 | 9.91 | 10.04 | 10.17 | 10.3
16.47 | 5.06 | 5.19 | 5.32 | 5.45 | 5.58 | 5.71 | 5.84 | 5.98 | 6.11 | 6.24 | 6.37 | 6.5 | 6.63 | 6.76 | 6.89 | 7.02 | 7.15 | 7.28 | 7.42 | 7.55 | 7.68 | 7.81 | 7.94 | 8.07 | 8.2 | 8.33 | 8.46 | 8.59 | 8.73 | 8.86 | 8.99 | 9.12 | 9.25 | 9.38 | 9.51 | 9.64 | 9.77 | 9.9 | 10.03 | 10.17 | 10.3
16.48 | 5.06 | 5.19 | 5.32 | 5.45 | 5.58 | 5.71 | 5.84 | 5.97 | 6.1 | 6.24 | 6.37 | 6.5 | 6.63 | 6.76 | 6.89 | 7.02 | 7.15 | 7.28 | 7.41 | 7.55 | 7.68 | 7.81 | 7.94 | 8.07 | 8.2 | 8.33 | 8.46 | 8.59 | 8.72 | 8.85 | 8.99 | 9.12 | 9.25 | 9.38 | 9.51 | 9.64 | 9.77 | 9.9 | 10.03 | 10.16 | 10.29
16.49 | 5.06 | 5.19 | 5.32 | 5.45 | 5.58 | 5.71 | 5.84 | 5.97 | 6.1 | 6.23 | 6.37 | 6.5 | 6.63 | 6.76 | 6.89 | 7.02 | 7.15 | 7.28 | 7.41 | 7.54 | 7.67 | 7.81 | 7.94 | 8.07 | 8.2 | 8.33 | 8.46 | 8.59 | 8.72 | 8.85 | 8.98 | 9.11 | 9.25 | 9.38 | 9.51 | 9.64 | 9.77 | 9.9 | 10.03 | 10.16 | 10.29
16.5 | 5.05 | 5.19 | 5.32 | 5.45 | 5.58 | 5.71 | 5.84 | 5.97 | 6.1 | 6.23 | 6.36 | 6.49 | 6.63 | 6.76 | 6.89 | 7.02 | 7.15 | 7.28 | 7.41 | 7.54 | 7.67 | 7.8 | 7.93 | 8.07 | 8.2 | 8.33 | 8.46 | 8.59 | 8.72 | 8.85 | 8.98 | 9.11 | 9.24 | 9.37 | 9.5 | 9.64 | 9.77 | 9.9 | 10.03 | 10.16 | 10.29
16.51 | 5.05 | 5.18 | 5.32 | 5.45 | 5.58 | 5.71 | 5.84 | 5.97 | 6.1 | 6.23 | 6.36 | 6.49 | 6.62 | 6.76 | 6.89 | 7.02 | 7.15 | 7.28 | 7.41 | 7.54 | 7.67 | 7.8 | 7.93 | 8.06 | 8.19 | 8.33 | 8.46 | 8.59 | 8.72 | 8.85 | 8.98 | 9.11 | 9.24 | 9.37 | 9.5 | 9.63 | 9.76 | 9.9 | 10.03 | 10.16 | 10.29
16.52 | 5.05 | 5.18 | 5.31 | 5.45 | 5.58 | 5.71 | 5.84 | 5.97 | 6.1 | 6.23 | 6.36 | 6.49 | 6.62 | 6.75 | 6.88 | 7.02 | 7.15 | 7.28 | 7.41 | 7.54 | 7.67 | 7.8 | 7.93 | 8.06 | 8.19 | 8.32 | 8.45 | 8.59 | 8.72 | 8.85 | 8.98 | 9.11 | 9.24 | 9.37 | 9.5 | 9.63 | 9.76 | 9.89 | 10.02 | 10.15 | 10.29
16.53 | 5.05 | 5.18 | 5.31 | 5.44 | 5.57 | 5.71 | 5.84 | 5.97 | 6.1 | 6.23 | 6.36 | 6.49 | 6.62 | 6.75 | 6.88 | 7.01 | 7.14 | 7.28 | 7.41 | 7.54 | 7.67 | 7.8 | 7.93 | 8.06 | 8.19 | 8.32 | 8.45 | 8.58 | 8.71 | 8.85 | 8.98 | 9.11 | 9.24 | 9.37 | 9.5 | 9.63 | 9.76 | 9.89 | 10.02 | 10.15 | 10.28
16.54 | 5.05 | 5.18 | 5.31 | 5.44 | 5.57 | 5.7 | 5.84 | 5.97 | 6.1 | 6.23 | 6.36 | 6.49 | 6.62 | 6.75 | 6.88 | 7.01 | 7.14 | 7.27 | 7.4 | 7.54 | 7.67 | 7.8 | 7.93 | 8.06 | 8.19 | 8.32 | 8.45 | 8.58 | 8.71 | 8.84 | 8.97 | 9.1 | 9.24 | 9.37 | 9.5 | 9.63 | 9.76 | 9.89 | 10.02 | 10.15 | 10.28
16.55 | 5.05 | 5.18 | 5.31 | 5.44 | 5.57 | 5.7 | 5.83 | 5.96 | 6.1 | 6.23 | 6.36 | 6.49 | 6.62 | 6.75 | 6.88 | 7.01 | 7.14 | 7.27 | 7.4 | 7.53 | 7.66 | 7.8 | 7.93 | 8.06 | 8.19 | 8.32 | 8.45 | 8.58 | 8.71 | 8.84 | 8.97 | 9.1 | 9.23 | 9.36 | 9.49 | 9.63 | 9.76 | 9.89 | 10.02 | 10.15 | 10.28
16.56 | 5.05 | 5.18 | 5.31 | 5.44 | 5.57 | 5.7 | 5.83 | 5.96 | 6.09 | 6.23 | 6.36 | 6.49 | 6.62 | 6.75 | 6.88 | 7.01 | 7.14 | 7.27 | 7.4 | 7.53 | 7.66 | 7.79 | 7.92 | 8.06 | 8.19 | 8.32 | 8.45 | 8.58 | 8.71 | 8.84 | 8.97 | 9.1 | 9.23 | 9.36 | 9.49 | 9.62 | 9.75 | 9.89 | 10.02 | 10.15 | 10.28
16.57 | 5.05 | 5.18 | 5.31 | 5.44 | 5.57 | 5.7 | 5.83 | 5.96 | 6.09 | 6.22 | 6.35 | 6.49 | 6.62 | 6.75 | 6.88 | 7.01 | 7.14 | 7.27 | 7.4 | 7.53 | 7.66 | 7.79 | 7.92 | 8.05 | 8.18 | 8.31 | 8.45 | 8.58 | 8.71 | 8.84 | 8.97 | 9.1 | 9.23 | 9.36 | 9.49 | 9.62 | 9.75 | 9.88 | 10.01 | 10.14 | 10.27
16.58 | 5.05 | 5.18 | 5.31 | 5.44 | 5.57 | 5.7 | 5.83 | 5.96 | 6.09 | 6.22 | 6.35 | 6.48 | 6.61 | 6.75 | 6.88 | 7.01 | 7.14 | 7.27 | 7.4 | 7.53 | 7.66 | 7.79 | 7.92 | 8.05 | 8.18 | 8.31 | 8.44 | 8.57 | 8.71 | 8.84 | 8.97 | 9.1 | 9.23 | 9.36 | 9.49 | 9.62 | 9.75 | 9.88 | 10.01 | 10.14 | 10.27
16.59 | 5.04 | 5.18 | 5.31 | 5.44 | 5.57 | 5.7 | 5.83 | 5.96 | 6.09 | 6.22 | 6.35 | 6.48 | 6.61 | 6.74 | 6.87 | 7 | 7.14 | 7.27 | 7.4 | 7.53 | 7.66 | 7.79 | 7.92 | 8.05 | 8.18 | 8.31 | 8.44 | 8.57 | 8.7 | 8.83 | 8.96 | 9.1 | 9.23 | 9.36 | 9.49 | 9.62 | 9.75 | 9.88 | 10.01 | 10.14 | 10.27
16.6 | 5.04 | 5.17 | 5.31 | 5.44 | 5.57 | 5.7 | 5.83 | 5.96 | 6.09 | 6.22 | 6.35 | 6.48 | 6.61 | 6.74 | 6.87 | 7 | 7.13 | 7.26 | 7.4 | 7.53 | 7.66 | 7.79 | 7.92 | 8.05 | 8.18 | 8.31 | 8.44 | 8.57 | 8.7 | 8.83 | 8.96 | 9.09 | 9.22 | 9.35 | 9.48 | 9.62 | 9.75 | 9.88 | 10.01 | 10.14 | 10.27
16.61 | 5.04 | 5.17 | 5.3 | 5.43 | 5.57 | 5.7 | 5.83 | 5.96 | 6.09 | 6.22 | 6.35 | 6.48 | 6.61 | 6.74 | 6.87 | 7 | 7.13 | 7.26 | 7.39 | 7.52 | 7.65 | 7.79 | 7.92 | 8.05 | 8.18 | 8.31 | 8.44 | 8.57 | 8.7 | 8.83 | 8.96 | 9.09 | 9.22 | 9.35 | 9.48 | 9.61 | 9.74 | 9.87 | 10.01 | 10.14 | 10.27
16.62 | 5.04 | 5.17 | 5.3 | 5.43 | 5.56 | 5.69 | 5.83 | 5.96 | 6.09 | 6.22 | 6.35 | 6.48 | 6.61 | 6.74 | 6.87 | 7 | 7.13 | 7.26 | 7.39 | 7.52 | 7.65 | 7.78 | 7.91 | 8.04 | 8.18 | 8.31 | 8.44 | 8.57 | 8.7 | 8.83 | 8.96 | 9.09 | 9.22 | 9.35 | 9.48 | 9.61 | 9.74 | 9.87 | 10 | 10.13 | 10.26
16.63 | 5.04 | 5.17 | 5.3 | 5.43 | 5.56 | 5.69 | 5.82 | 5.95 | 6.08 | 6.22 | 6.35 | 6.48 | 6.61 | 6.74 | 6.87 | 7 | 7.13 | 7.26 | 7.39 | 7.52 | 7.65 | 7.78 | 7.91 | 8.04 | 8.17 | 8.3 | 8.43 | 8.57 | 8.7 | 8.83 | 8.96 | 9.09 | 9.22 | 9.35 | 9.48 | 9.61 | 9.74 | 9.87 | 10 | 10.13 | 10.26
16.64 | 5.04 | 5.17 | 5.3 | 5.43 | 5.56 | 5.69 | 5.82 | 5.95 | 6.08 | 6.21 | 6.34 | 6.48 | 6.61 | 6.74 | 6.87 | 7 | 7.13 | 7.26 | 7.39 | 7.52 | 7.65 | 7.78 | 7.91 | 8.04 | 8.17 | 8.3 | 8.43 | 8.56 | 8.69 | 8.82 | 8.96 | 9.09 | 9.22 | 9.35 | 9.48 | 9.61 | 9.74 | 9.87 | 10 | 10.13 | 10.26
16.65 | 5.04 | 5.17 | 5.3 | 5.43 | 5.56 | 5.69 | 5.82 | 5.95 | 6.08 | 6.21 | 6.34 | 6.47 | 6.6 | 6.73 | 6.87 | 7 | 7.13 | 7.26 | 7.39 | 7.52 | 7.65 | 7.78 | 7.91 | 8.04 | 8.17 | 8.3 | 8.43 | 8.56 | 8.69 | 8.82 | 8.95 | 9.08 | 9.21 | 9.34 | 9.48 | 9.61 | 9.74 | 9.87 | 10 | 10.13 | 10.26
16.66 | 5.04 | 5.17 | 5.3 | 5.43 | 5.56 | 5.69 | 5.82 | 5.95 | 6.08 | 6.21 | 6.34 | 6.47 | 6.6 | 6.73 | 6.86 | 6.99 | 7.12 | 7.26 | 7.39 | 7.52 | 7.65 | 7.78 | 7.91 | 8.04 | 8.17 | 8.3 | 8.43 | 8.56 | 8.69 | 8.82 | 8.95 | 9.08 | 9.21 | 9.34 | 9.47 | 9.6 | 9.73 | 9.86 | 9.99 | 10.13 | 10.26
16.67 | 5.04 | 5.17 | 5.3 | 5.43 | 5.56 | 5.69 | 5.82 | 5.95 | 6.08 | 6.21 | 6.34 | 6.47 | 6.6 | 6.73 | 6.86 | 6.99 | 7.12 | 7.25 | 7.38 | 7.51 | 7.65 | 7.78 | 7.91 | 8.04 | 8.17 | 8.3 | 8.43 | 8.56 | 8.69 | 8.82 | 8.95 | 9.08 | 9.21 | 9.34 | 9.47 | 9.6 | 9.73 | 9.86 | 9.99 | 10.12 | 10.25
16.68 | 5.03 | 5.17 | 5.3 | 5.43 | 5.56 | 5.69 | 5.82 | 5.95 | 6.08 | 6.21 | 6.34 | 6.47 | 6.6 | 6.73 | 6.86 | 6.99 | 7.12 | 7.25 | 7.38 | 7.51 | 7.64 | 7.77 | 7.9 | 8.03 | 8.17 | 8.3 | 8.43 | 8.56 | 8.69 | 8.82 | 8.95 | 9.08 | 9.21 | 9.34 | 9.47 | 9.6 | 9.73 | 9.86 | 9.99 | 10.12 | 10.25
16.69 | 5.03 | 5.16 | 5.29 | 5.42 | 5.56 | 5.69 | 5.82 | 5.95 | 6.08 | 6.21 | 6.34 | 6.47 | 6.6 | 6.73 | 6.86 | 6.99 | 7.12 | 7.25 | 7.38 | 7.51 | 7.64 | 7.77 | 7.9 | 8.03 | 8.16 | 8.29 | 8.42 | 8.55 | 8.68 | 8.82 | 8.95 | 9.08 | 9.21 | 9.34 | 9.47 | 9.6 | 9.73 | 9.86 | 9.99 | 10.12 | 10.25
16.7 | 5.03 | 5.16 | 5.29 | 5.42 | 5.55 | 5.68 | 5.81 | 5.95 | 6.08 | 6.21 | 6.34 | 6.47 | 6.6 | 6.73 | 6.86 | 6.99 | 7.12 | 7.25 | 7.38 | 7.51 | 7.64 | 7.77 | 7.9 | 8.03 | 8.16 | 8.29 | 8.42 | 8.55 | 8.68 | 8.81 | 8.94 | 9.07 | 9.2 | 9.33 | 9.47 | 9.6 | 9.73 | 9.86 | 9.99 | 10.12 | 10.25
16.71 | 5.03 | 5.16 | 5.29 | 5.42 | 5.55 | 5.68 | 5.81 | 5.94 | 6.07 | 6.2 | 6.34 | 6.47 | 6.6 | 6.73 | 6.86 | 6.99 | 7.12 | 7.25 | 7.38 | 7.51 | 7.64 | 7.77 | 7.9 | 8.03 | 8.16 | 8.29 | 8.42 | 8.55 | 8.68 | 8.81 | 8.94 | 9.07 | 9.2 | 9.33 | 9.46 | 9.59 | 9.72 | 9.85 | 9.98 | 10.11 | 10.24
16.72 | 5.03 | 5.16 | 5.29 | 5.42 | 5.55 | 5.68 | 5.81 | 5.94 | 6.07 | 6.2 | 6.33 | 6.46 | 6.59 | 6.72 | 6.86 | 6.99 | 7.12 | 7.25 | 7.38 | 7.51 | 7.64 | 7.77 | 7.9 | 8.03 | 8.16 | 8.29 | 8.42 | 8.55 | 8.68 | 8.81 | 8.94 | 9.07 | 9.2 | 9.33 | 9.46 | 9.59 | 9.72 | 9.85 | 9.98 | 10.11 | 10.24
16.73 | 5.03 | 5.16 | 5.29 | 5.42 | 5.55 | 5.68 | 5.81 | 5.94 | 6.07 | 6.2 | 6.33 | 6.46 | 6.59 | 6.72 | 6.85 | 6.98 | 7.11 | 7.24 | 7.37 | 7.51 | 7.64 | 7.77 | 7.9 | 8.03 | 8.16 | 8.29 | 8.42 | 8.55 | 8.68 | 8.81 | 8.94 | 9.07 | 9.2 | 9.33 | 9.46 | 9.59 | 9.72 | 9.85 | 9.98 | 10.11 | 10.24
16.74 | 5.03 | 5.16 | 5.29 | 5.42 | 5.55 | 5.68 | 5.81 | 5.94 | 6.07 | 6.2 | 6.33 | 6.46 | 6.59 | 6.72 | 6.85 | 6.98 | 7.11 | 7.24 | 7.37 | 7.5 | 7.63 | 7.76 | 7.89 | 8.02 | 8.15 | 8.29 | 8.42 | 8.55 | 8.68 | 8.81 | 8.94 | 9.07 | 9.2 | 9.33 | 9.46 | 9.59 | 9.72 | 9.85 | 9.98 | 10.11 | 10.24
16.75 | 5.03 | 5.16 | 5.29 | 5.42 | 5.55 | 5.68 | 5.81 | 5.94 | 6.07 | 6.2 | 6.33 | 6.46 | 6.59 | 6.72 | 6.85 | 6.98 | 7.11 | 7.24 | 7.37 | 7.5 | 7.63 | 7.76 | 7.89 | 8.02 | 8.15 | 8.28 | 8.41 | 8.54 | 8.67 | 8.8 | 8.93 | 9.06 | 9.19 | 9.32 | 9.46 | 9.59 | 9.72 | 9.85 | 9.98 | 10.11 | 10.24
16.76 | 5.03 | 5.16 | 5.29 | 5.42 | 5.55 | 5.68 | 5.81 | 5.94 | 6.07 | 6.2 | 6.33 | 6.46 | 6.59 | 6.72 | 6.85 | 6.98 | 7.11 | 7.24 | 7.37 | 7.5 | 7.63 | 7.76 | 7.89 | 8.02 | 8.15 | 8.28 | 8.41 | 8.54 | 8.67 | 8.8 | 8.93 | 9.06 | 9.19 | 9.32 | 9.45 | 9.58 | 9.71 | 9.84 | 9.97 | 10.1 | 10.23
16.77 | 5.02 | 5.15 | 5.29 | 5.42 | 5.55 | 5.68 | 5.81 | 5.94 | 6.07 | 6.2 | 6.33 | 6.46 | 6.59 | 6.72 | 6.85 | 6.98 | 7.11 | 7.24 | 7.37 | 7.5 | 7.63 | 7.76 | 7.89 | 8.02 | 8.15 | 8.28 | 8.41 | 8.54 | 8.67 | 8.8 | 8.93 | 9.06 | 9.19 | 9.32 | 9.45 | 9.58 | 9.71 | 9.84 | 9.97 | 10.1 | 10.23
16.78 | 5.02 | 5.15 | 5.28 | 5.41 | 5.54 | 5.67 | 5.8 | 5.94 | 6.07 | 6.2 | 6.33 | 6.46 | 6.59 | 6.72 | 6.85 | 6.98 | 7.11 | 7.24 | 7.37 | 7.5 | 7.63 | 7.76 | 7.89 | 8.02 | 8.15 | 8.28 | 8.41 | 8.54 | 8.67 | 8.8 | 8.93 | 9.06 | 9.19 | 9.32 | 9.45 | 9.58 | 9.71 | 9.84 | 9.97 | 10.1 | 10.23
16.79 | 5.02 | 5.15 | 5.28 | 5.41 | 5.54 | 5.67 | 5.8 | 5.93 | 6.06 | 6.19 | 6.32 | 6.45 | 6.58 | 6.71 | 6.84 | 6.97 | 7.11 | 7.24 | 7.37 | 7.5 | 7.63 | 7.76 | 7.89 | 8.02 | 8.15 | 8.28 | 8.41 | 8.54 | 8.67 | 8.8 | 8.93 | 9.06 | 9.19 | 9.32 | 9.45 | 9.58 | 9.71 | 9.84 | 9.97 | 10.1 | 10.23
16.8 | 5.02 | 5.15 | 5.28 | 5.41 | 5.54 | 5.67 | 5.8 | 5.93 | 6.06 | 6.19 | 6.32 | 6.45 | 6.58 | 6.71 | 6.84 | 6.97 | 7.1 | 7.23 | 7.36 | 7.49 | 7.62 | 7.75 | 7.88 | 8.01 | 8.14 | 8.27 | 8.4 | 8.53 | 8.66 | 8.79 | 8.92 | 9.06 | 9.19 | 9.32 | 9.45 | 9.58 | 9.71 | 9.84 | 9.97 | 10.1 | 10.23
16.81 | 5.02 | 5.15 | 5.28 | 5.41 | 5.54 | 5.67 | 5.8 | 5.93 | 6.06 | 6.19 | 6.32 | 6.45 | 6.58 | 6.71 | 6.84 | 6.97 | 7.1 | 7.23 | 7.36 | 7.49 | 7.62 | 7.75 | 7.88 | 8.01 | 8.14 | 8.27 | 8.4 | 8.53 | 8.66 | 8.79 | 8.92 | 9.05 | 9.18 | 9.31 | 9.44 | 9.57 | 9.7 | 9.83 | 9.96 | 10.09 | 10.22
16.82 | 5.02 | 5.15 | 5.28 | 5.41 | 5.54 | 5.67 | 5.8 | 5.93 | 6.06 | 6.19 | 6.32 | 6.45 | 6.58 | 6.71 | 6.84 | 6.97 | 7.1 | 7.23 | 7.36 | 7.49 | 7.62 | 7.75 | 7.88 | 8.01 | 8.14 | 8.27 | 8.4 | 8.53 | 8.66 | 8.79 | 8.92 | 9.05 | 9.18 | 9.31 | 9.44 | 9.57 | 9.7 | 9.83 | 9.96 | 10.09 | 10.22
16.83 | 5.02 | 5.15 | 5.28 | 5.41 | 5.54 | 5.67 | 5.8 | 5.93 | 6.06 | 6.19 | 6.32 | 6.45 | 6.58 | 6.71 | 6.84 | 6.97 | 7.1 | 7.23 | 7.36 | 7.49 | 7.62 | 7.75 | 7.88 | 8.01 | 8.14 | 8.27 | 8.4 | 8.53 | 8.66 | 8.79 | 8.92 | 9.05 | 9.18 | 9.31 | 9.44 | 9.57 | 9.7 | 9.83 | 9.96 | 10.09 | 10.22
16.84 | 5.02 | 5.15 | 5.28 | 5.41 | 5.54 | 5.67 | 5.8 | 5.93 | 6.06 | 6.19 | 6.32 | 6.45 | 6.58 | 6.71 | 6.84 | 6.97 | 7.1 | 7.23 | 7.36 | 7.49 | 7.62 | 7.75 | 7.88 | 8.01 | 8.14 | 8.27 | 8.4 | 8.53 | 8.66 | 8.79 | 8.92 | 9.05 | 9.18 | 9.31 | 9.44 | 9.57 | 9.7 | 9.83 | 9.96 | 10.09 | 10.22
16.85 | 5.02 | 5.15 | 5.28 | 5.41 | 5.54 | 5.67 | 5.8 | 5.93 | 6.06 | 6.19 | 6.32 | 6.45 | 6.58 | 6.71 | 6.84 | 6.97 | 7.1 | 7.23 | 7.36 | 7.49 | 7.62 | 7.75 | 7.88 | 8.01 | 8.14 | 8.27 | 8.4 | 8.53 | 8.66 | 8.79 | 8.92 | 9.05 | 9.18 | 9.31 | 9.44 | 9.57 | 9.7 | 9.83 | 9.96 | 10.09 | 10.21
16.86 | 5.01 | 5.14 | 5.27 | 5.4 | 5.53 | 5.66 | 5.79 | 5.92 | 6.05 | 6.18 | 6.31 | 6.44 | 6.57 | 6.7 | 6.83 | 6.96 | 7.09 | 7.22 | 7.35 | 7.48 | 7.61 | 7.74 | 7.87 | 8 | 8.13 | 8.26 | 8.39 | 8.52 | 8.65 | 8.78 | 8.91 | 9.04 | 9.17 | 9.3 | 9.43 | 9.56 | 9.69 | 9.82 | 9.95 | 10.08 | 10.21
16.87 | 5.01 | 5.14 | 5.27 | 5.4 | 5.53 | 5.66 | 5.79 | 5.92 | 6.05 | 6.18 | 6.31 | 6.44 | 6.57 | 6.7 | 6.83 | 6.96 | 7.09 | 7.22 | 7.35 | 7.48 | 7.61 | 7.74 | 7.87 | 8 | 8.13 | 8.26 | 8.39 | 8.52 | 8.65 | 8.78 | 8.91 | 9.04 | 9.17 | 9.3 | 9.43 | 9.56 | 9.69 | 9.82 | 9.95 | 10.08 | 10.21
16.88 | 5.01 | 5.14 | 5.27 | 5.4 | 5.53 | 5.66 | 5.79 | 5.92 | 6.05 | 6.18 | 6.31 | 6.44 | 6.57 | 6.7 | 6.83 | 6.96 | 7.09 | 7.22 | 7.35 | 7.48 | 7.61 | 7.74 | 7.87 | 8 | 8.13 | 8.26 | 8.39 | 8.52 | 8.65 | 8.78 | 8.91 | 9.04 | 9.17 | 9.3 | 9.43 | 9.56 | 9.69 | 9.82 | 9.95 | 10.08 | 10.21
16.89 | 5.01 | 5.14 | 5.27 | 5.4 | 5.53 | 5.66 | 5.79 | 5.92 | 6.05 | 6.18 | 6.31 | 6.44 | 6.57 | 6.7 | 6.83 | 6.96 | 7.09 | 7.22 | 7.35 | 7.48 | 7.61 | 7.74 | 7.87 | 8 | 8.13 | 8.26 | 8.39 | 8.52 | 8.65 | 8.78 | 8.91 | 9.04 | 9.17 | 9.3 | 9.43 | 9.56 | 9.69 | 9.82 | 9.95 | 10.08 | 10.21
16.9 | 5.01 | 5.14 | 5.27 | 5.4 | 5.53 | 5.66 | 5.79 | 5.92 | 6.05 | 6.18 | 6.31 | 6.44 | 6.57 | 6.7 | 6.83 | 6.96 | 7.09 | 7.22 | 7.35 | 7.48 | 7.61 | 7.74 | 7.87 | 8 | 8.13 | 8.26 | 8.39 | 8.52 | 8.65 | 8.78 | 8.91 | 9.04 | 9.17 | 9.3 | 9.43 | 9.56 | 9.69 | 9.81 | 9.94 | 10.07 | 10.2
16.91 | 5.01 | 5.14 | 5.27 | 5.4 | 5.53 | 5.66 | 5.79 | 5.92 | 6.05 | 6.18 | 6.31 | 6.44 | 6.57 | 6.7 | 6.83 | 6.96 | 7.09 | 7.22 | 7.35 | 7.48 | 7.61 | 7.74 | 7.87 | 8 | 8.13 | 8.26 | 8.39 | 8.51 | 8.64 | 8.77 | 8.9 | 9.03 | 9.16 | 9.29 | 9.42 | 9.55 | 9.68 | 9.81 | 9.94 | 10.07 | 10.2
16.92 | 5.01 | 5.14 | 5.27 | 5.4 | 5.53 | 5.66 | 5.79 | 5.92 | 6.05 | 6.18 | 6.31 | 6.44 | 6.57 | 6.7 | 6.83 | 6.96 | 7.09 | 7.22 | 7.35 | 7.47 | 7.6 | 7.73 | 7.86 | 7.99 | 8.12 | 8.25 | 8.38 | 8.51 | 8.64 | 8.77 | 8.9 | 9.03 | 9.16 | 9.29 | 9.42 | 9.55 | 9.68 | 9.81 | 9.94 | 10.07 | 10.2
16.93 | 5.01 | 5.14 | 5.27 | 5.4 | 5.53 | 5.66 | 5.79 | 5.92 | 6.05 | 6.18 | 6.31 | 6.43 | 6.56 | 6.69 | 6.82 | 6.95 | 7.08 | 7.21 | 7.34 | 7.47 | 7.6 | 7.73 | 7.86 | 7.99 | 8.12 | 8.25 | 8.38 | 8.51 | 8.64 | 8.77 | 8.9 | 9.03 | 9.16 | 9.29 | 9.42 | 9.55 | 9.68 | 9.81 | 9.94 | 10.07 | 10.2
16.94 | 5.01 | 5.14 | 5.27 | 5.4 | 5.53 | 5.65 | 5.78 | 5.91 | 6.04 | 6.17 | 6.3 | 6.43 | 6.56 | 6.69 | 6.82 | 6.95 | 7.08 | 7.21 | 7.34 | 7.47 | 7.6 | 7.73 | 7.86 | 7.99 | 8.12 | 8.25 | 8.38 | 8.51 | 8.64 | 8.77 | 8.9 | 9.03 | 9.16 | 9.29 | 9.42 | 9.55 | 9.68 | 9.81 | 9.94 | 10.07 | 10.2
16.95 | 5 | 5.13 | 5.26 | 5.39 | 5.52 | 5.65 | 5.78 | 5.91 | 6.04 | 6.17 | 6.3 | 6.43 | 6.56 | 6.69 | 6.82 | 6.95 | 7.08 | 7.21 | 7.34 | 7.47 | 7.6 | 7.73 | 7.86 | 7.99 | 8.12 | 8.25 | 8.38 | 8.51 | 8.64 | 8.77 | 8.9 | 9.03 | 9.16 | 9.29 | 9.42 | 9.55 | 9.67 | 9.8 | 9.93 | 10.06 | 10.19
16.96 | 5 | 5.13 | 5.26 | 5.39 | 5.52 | 5.65 | 5.78 | 5.91 | 6.04 | 6.17 | 6.3 | 6.43 | 6.56 | 6.69 | 6.82 | 6.95 | 7.08 | 7.21 | 7.34 | 7.47 | 7.6 | 7.73 | 7.86 | 7.99 | 8.12 | 8.25 | 8.38 | 8.51 | 8.64 | 8.77 | 8.89 | 9.02 | 9.15 | 9.28 | 9.41 | 9.54 | 9.67 | 9.8 | 9.93 | 10.06 | 10.19
16.97 | 5 | 5.13 | 5.26 | 5.39 | 5.52 | 5.65 | 5.78 | 5.91 | 6.04 | 6.17 | 6.3 | 6.43 | 6.56 | 6.69 | 6.82 | 6.95 | 7.08 | 7.21 | 7.34 | 7.47 | 7.6 | 7.73 | 7.86 | 7.99 | 8.12 | 8.24 | 8.37 | 8.5 | 8.63 | 8.76 | 8.89 | 9.02 | 9.15 | 9.28 | 9.41 | 9.54 | 9.67 | 9.8 | 9.93 | 10.06 | 10.19
16.98 | 5 | 5.13 | 5.26 | 5.39 | 5.52 | 5.65 | 5.78 | 5.91 | 6.04 | 6.17 | 6.3 | 6.43 | 6.56 | 6.69 | 6.82 | 6.95 | 7.08 | 7.21 | 7.34 | 7.47 | 7.59 | 7.72 | 7.85 | 7.98 | 8.11 | 8.24 | 8.37 | 8.5 | 8.63 | 8.76 | 8.89 | 9.02 | 9.15 | 9.28 | 9.41 | 9.54 | 9.67 | 9.8 | 9.93 | 10.06 | 10.19
16.99 | 5 | 5.13 | 5.26 | 5.39 | 5.52 | 5.65 | 5.78 | 5.91 | 6.04 | 6.17 | 6.3 | 6.43 | 6.56 | 6.69 | 6.82 | 6.95 | 7.07 | 7.2 | 7.33 | 7.46 | 7.59 | 7.72 | 7.85 | 7.98 | 8.11 | 8.24 | 8.37 | 8.5 | 8.63 | 8.76 | 8.89 | 9.02 | 9.15 | 9.28 | 9.41 | 9.54 | 9.67 | 9.8 | 9.93 | 10.06 | 10.19
17 | 5 | 5.13 | 5.26 | 5.39 | 5.52 | 5.65 | 5.78 | 5.91 | 6.04 | 6.17 | 6.3 | 6.43 | 6.55 | 6.68 | 6.81 | 6.94 | 7.07 | 7.2 | 7.33 | 7.46 | 7.59 | 7.72 | 7.85 | 7.98 | 8.11 | 8.24 | 8.37 | 8.5 | 8.63 | 8.76 | 8.89 | 9.02 | 9.15 | 9.28 | 9.41 | 9.54 | 9.66 | 9.79 | 9.92 | 10.05 | 10.18
17.01 | 5 | 5.13 | 5.26 | 5.39 | 5.52 | 5.65 | 5.78 | 5.91 | 6.04 | 6.16 | 6.29 | 6.42 | 6.55 | 6.68 | 6.81 | 6.94 | 7.07 | 7.2 | 7.33 | 7.46 | 7.59 | 7.72 | 7.85 | 7.98 | 8.11 | 8.24 | 8.37 | 8.5 | 8.63 | 8.76 | 8.89 | 9.02 | 9.14 | 9.27 | 9.4 | 9.53 | 9.66 | 9.79 | 9.92 | 10.05 | 10.18
17.02 | 5 | 5.13 | 5.26 | 5.39 | 5.52 | 5.65 | 5.77 | 5.9 | 6.03 | 6.16 | 6.29 | 6.42 | 6.55 | 6.68 | 6.81 | 6.94 | 7.07 | 7.2 | 7.33 | 7.46 | 7.59 | 7.72 | 7.85 | 7.98 | 8.11 | 8.24 | 8.37 | 8.5 | 8.62 | 8.75 | 8.88 | 9.01 | 9.14 | 9.27 | 9.4 | 9.53 | 9.66 | 9.79 | 9.92 | 10.05 | 10.18
17.03 | 5 | 5.13 | 5.26 | 5.38 | 5.51 | 5.64 | 5.77 | 5.9 | 6.03 | 6.16 | 6.29 | 6.42 | 6.55 | 6.68 | 6.81 | 6.94 | 7.07 | 7.2 | 7.33 | 7.46 | 7.59 | 7.72 | 7.85 | 7.98 | 8.1 | 8.23 | 8.36 | 8.49 | 8.62 | 8.75 | 8.88 | 9.01 | 9.14 | 9.27 | 9.4 | 9.53 | 9.66 | 9.79 | 9.92 | 10.05 | 10.18
17.04 | 4.99 | 5.12 | 5.25 | 5.38 | 5.51 | 5.64 | 5.77 | 5.9 | 6.03 | 6.16 | 6.29 | 6.42 | 6.55 | 6.68 | 6.81 | 6.94 | 7.07 | 7.2 | 7.33 | 7.46 | 7.59 | 7.71 | 7.84 | 7.97 | 8.1 | 8.23 | 8.36 | 8.49 | 8.62 | 8.75 | 8.88 | 9.01 | 9.14 | 9.27 | 9.4 | 9.53 | 9.66 | 9.79 | 9.92 | 10.05 | 10.17
17.05 | 4.99 | 5.12 | 5.25 | 5.38 | 5.51 | 5.64 | 5.77 | 5.9 | 6.03 | 6.16 | 6.29 | 6.42 | 6.55 | 6.68 | 6.81 | 6.94 | 7.07 | 7.2 | 7.32 | 7.45 | 7.58 | 7.71 | 7.84 | 7.97 | 8.1 | 8.23 | 8.36 | 8.49 | 8.62 | 8.75 | 8.88 | 9.01 | 9.14 | 9.27 | 9.4 | 9.53 | 9.65 | 9.78 | 9.91 | 10.04 | 10.17
17.06 | 4.99 | 5.12 | 5.25 | 5.38 | 5.51 | 5.64 | 5.77 | 5.9 | 6.03 | 6.16 | 6.29 | 6.42 | 6.55 | 6.68 | 6.81 | 6.93 | 7.06 | 7.19 | 7.32 | 7.45 | 7.58 | 7.71 | 7.84 | 7.97 | 8.1 | 8.23 | 8.36 | 8.49 | 8.62 | 8.75 | 8.88 | 9.01 | 9.14 | 9.26 | 9.39 | 9.52 | 9.65 | 9.78 | 9.91 | 10.04 | 10.17
17.07 | 4.99 | 5.12 | 5.25 | 5.38 | 5.51 | 5.64 | 5.77 | 5.9 | 6.03 | 6.16 | 6.29 | 6.42 | 6.55 | 6.67 | 6.8 | 6.93 | 7.06 | 7.19 | 7.32 | 7.45 | 7.58 | 7.71 | 7.84 | 7.97 | 8.1 | 8.23 | 8.36 | 8.49 | 8.62 | 8.75 | 8.87 | 9 | 9.13 | 9.26 | 9.39 | 9.52 | 9.65 | 9.78 | 9.91 | 10.04 | 10.17
17.08 | 4.99 | 5.12 | 5.25 | 5.38 | 5.51 | 5.64 | 5.77 | 5.9 | 6.03 | 6.16 | 6.28 | 6.41 | 6.54 | 6.67 | 6.8 | 6.93 | 7.06 | 7.19 | 7.32 | 7.45 | 7.58 | 7.71 | 7.84 | 7.97 | 8.1 | 8.23 | 8.36 | 8.48 | 8.61 | 8.74 | 8.87 | 9 | 9.13 | 9.26 | 9.39 | 9.52 | 9.65 | 9.78 | 9.91 | 10.04 | 10.17
17.09 | 4.99 | 5.12 | 5.25 | 5.38 | 5.51 | 5.64 | 5.77 | 5.9 | 6.02 | 6.15 | 6.28 | 6.41 | 6.54 | 6.67 | 6.8 | 6.93 | 7.06 | 7.19 | 7.32 | 7.45 | 7.58 | 7.71 | 7.84 | 7.97 | 8.09 | 8.22 | 8.35 | 8.48 | 8.61 | 8.74 | 8.87 | 9 | 9.13 | 9.26 | 9.39 | 9.52 | 9.65 | 9.78 | 9.91 | 10.03 | 10.16
17.1 | 4.99 | 5.12 | 5.25 | 5.38 | 5.51 | 5.64 | 5.76 | 5.89 | 6.02 | 6.15 | 6.28 | 6.41 | 6.54 | 6.67 | 6.8 | 6.93 | 7.06 | 7.19 | 7.32 | 7.45 | 7.58 | 7.7 | 7.83 | 7.96 | 8.09 | 8.22 | 8.35 | 8.48 | 8.61 | 8.74 | 8.87 | 9 | 9.13 | 9.26 | 9.39 | 9.52 | 9.64 | 9.77 | 9.9 | 10.03 | 10.16
17.11 | 4.99 | 5.12 | 5.25 | 5.38 | 5.5 | 5.63 | 5.76 | 5.89 | 6.02 | 6.15 | 6.28 | 6.41 | 6.54 | 6.67 | 6.8 | 6.93 | 7.06 | 7.19 | 7.32 | 7.44 | 7.57 | 7.7 | 7.83 | 7.96 | 8.09 | 8.22 | 8.35 | 8.48 | 8.61 | 8.74 | 8.87 | 9 | 9.13 | 9.25 | 9.38 | 9.51 | 9.64 | 9.77 | 9.9 | 10.03 | 10.16
17.12 | 4.99 | 5.12 | 5.24 | 5.37 | 5.5 | 5.63 | 5.76 | 5.89 | 6.02 | 6.15 | 6.28 | 6.41 | 6.54 | 6.67 | 6.8 | 6.93 | 7.06 | 7.18 | 7.31 | 7.44 | 7.57 | 7.7 | 7.83 | 7.96 | 8.09 | 8.22 | 8.35 | 8.48 | 8.61 | 8.74 | 8.87 | 8.99 | 9.12 | 9.25 | 9.38 | 9.51 | 9.64 | 9.77 | 9.9 | 10.03 | 10.16
17.13 | 4.99 | 5.11 | 5.24 | 5.37 | 5.5 | 5.63 | 5.76 | 5.89 | 6.02 | 6.15 | 6.28 | 6.41 | 6.54 | 6.67 | 6.8 | 6.92 | 7.05 | 7.18 | 7.31 | 7.44 | 7.57 | 7.7 | 7.83 | 7.96 | 8.09 | 8.22 | 8.35 | 8.48 | 8.6 | 8.73 | 8.86 | 8.99 | 9.12 | 9.25 | 9.38 | 9.51 | 9.64 | 9.77 | 9.9 | 10.03 | 10.16
17.14 | 4.98 | 5.11 | 5.24 | 5.37 | 5.5 | 5.63 | 5.76 | 5.89 | 6.02 | 6.15 | 6.28 | 6.41 | 6.54 | 6.66 | 6.79 | 6.92 | 7.05 | 7.18 | 7.31 | 7.44 | 7.57 | 7.7 | 7.83 | 7.96 | 8.09 | 8.22 | 8.34 | 8.47 | 8.6 | 8.73 | 8.86 | 8.99 | 9.12 | 9.25 | 9.38 | 9.51 | 9.64 | 9.77 | 9.9 | 10.02 | 10.15
17.15 | 4.98 | 5.11 | 5.24 | 5.37 | 5.5 | 5.63 | 5.76 | 5.89 | 6.02 | 6.15 | 6.28 | 6.4 | 6.53 | 6.66 | 6.79 | 6.92 | 7.05 | 7.18 | 7.31 | 7.44 | 7.57 | 7.7 | 7.83 | 7.96 | 8.08 | 8.21 | 8.34 | 8.47 | 8.6 | 8.73 | 8.86 | 8.99 | 9.12 | 9.25 | 9.38 | 9.51 | 9.63 | 9.76 | 9.89 | 10.02 | 10.15
17.16 | 4.98 | 5.11 | 5.24 | 5.37 | 5.5 | 5.63 | 5.76 | 5.89 | 6.02 | 6.14 | 6.27 | 6.4 | 6.53 | 6.66 | 6.79 | 6.92 | 7.05 | 7.18 | 7.31 | 7.44 | 7.57 | 7.7 | 7.82 | 7.95 | 8.08 | 8.21 | 8.34 | 8.47 | 8.6 | 8.73 | 8.86 | 8.99 | 9.12 | 9.25 | 9.37 | 9.5 | 9.63 | 9.76 | 9.89 | 10.02 | 10.15
17.17 | 4.98 | 5.11 | 5.24 | 5.37 | 5.5 | 5.63 | 5.76 | 5.89 | 6.01 | 6.14 | 6.27 | 6.4 | 6.53 | 6.66 | 6.79 | 6.92 | 7.05 | 7.18 | 7.31 | 7.44 | 7.56 | 7.69 | 7.82 | 7.95 | 8.08 | 8.21 | 8.34 | 8.47 | 8.6 | 8.73 | 8.86 | 8.99 | 9.11 | 9.24 | 9.37 | 9.5 | 9.63 | 9.76 | 9.89 | 10.02 | 10.15
17.18 | 4.98 | 5.11 | 5.24 | 5.37 | 5.5 | 5.63 | 5.75 | 5.88 | 6.01 | 6.14 | 6.27 | 6.4 | 6.53 | 6.66 | 6.79 | 6.92 | 7.05 | 7.18 | 7.3 | 7.43 | 7.56 | 7.69 | 7.82 | 7.95 | 8.08 | 8.21 | 8.34 | 8.47 | 8.6 | 8.72 | 8.85 | 8.98 | 9.11 | 9.24 | 9.37 | 9.5 | 9.63 | 9.76 | 9.89 | 10.02 | 10.15
17.19 | 4.98 | 5.11 | 5.24 | 5.37 | 5.5 | 5.62 | 5.75 | 5.88 | 6.01 | 6.14 | 6.27 | 6.4 | 6.53 | 6.66 | 6.79 | 6.92 | 7.04 | 7.17 | 7.3 | 7.43 | 7.56 | 7.69 | 7.82 | 7.95 | 8.08 | 8.21 | 8.34 | 8.46 | 8.59 | 8.72 | 8.85 | 8.98 | 9.11 | 9.24 | 9.37 | 9.5 | 9.63 | 9.76 | 9.88 | 10.01 | 10.14
17.2 | 4.98 | 5.11 | 5.24 | 5.36 | 5.49 | 5.62 | 5.75 | 5.88 | 6.01 | 6.14 | 6.27 | 6.4 | 6.53 | 6.66 | 6.79 | 6.91 | 7.04 | 7.17 | 7.3 | 7.43 | 7.56 | 7.69 | 7.82 | 7.95 | 8.08 | 8.21 | 8.33 | 8.46 | 8.59 | 8.72 | 8.85 | 8.98 | 9.11 | 9.24 | 9.37 | 9.5 | 9.62 | 9.75 | 9.88 | 10.01 | 10.14
17.21 | 4.98 | 5.11 | 5.23 | 5.36 | 5.49 | 5.62 | 5.75 | 5.88 | 6.01 | 6.14 | 6.27 | 6.4 | 6.53 | 6.65 | 6.78 | 6.91 | 7.04 | 7.17 | 7.3 | 7.43 | 7.56 | 7.69 | 7.82 | 7.95 | 8.07 | 8.2 | 8.33 | 8.46 | 8.59 | 8.72 | 8.85 | 8.98 | 9.11 | 9.24 | 9.36 | 9.49 | 9.62 | 9.75 | 9.88 | 10.01 | 10.14
17.22 | 4.98 | 5.1 | 5.23 | 5.36 | 5.49 | 5.62 | 5.75 | 5.88 | 6.01 | 6.14 | 6.27 | 6.4 | 6.52 | 6.65 | 6.78 | 6.91 | 7.04 | 7.17 | 7.3 | 7.43 | 7.56 | 7.69 | 7.81 | 7.94 | 8.07 | 8.2 | 8.33 | 8.46 | 8.59 | 8.72 | 8.85 | 8.98 | 9.1 | 9.23 | 9.36 | 9.49 | 9.62 | 9.75 | 9.88 | 10.01 | 10.14
17.23 | 4.97 | 5.1 | 5.23 | 5.36 | 5.49 | 5.62 | 5.75 | 5.88 | 6.01 | 6.14 | 6.26 | 6.39 | 6.52 | 6.65 | 6.78 | 6.91 | 7.04 | 7.17 | 7.3 | 7.43 | 7.55 | 7.68 | 7.81 | 7.94 | 8.07 | 8.2 | 8.33 | 8.46 | 8.59 | 8.72 | 8.84 | 8.97 | 9.1 | 9.23 | 9.36 | 9.49 | 9.62 | 9.75 | 9.88 | 10.01 | 10.13
17.24 | 4.97 | 5.1 | 5.23 | 5.36 | 5.49 | 5.62 | 5.75 | 5.88 | 6.01 | 6.13 | 6.26 | 6.39 | 6.52 | 6.65 | 6.78 | 6.91 | 7.04 | 7.17 | 7.3 | 7.42 | 7.55 | 7.68 | 7.81 | 7.94 | 8.07 | 8.2 | 8.33 | 8.46 | 8.59 | 8.71 | 8.84 | 8.97 | 9.1 | 9.23 | 9.36 | 9.49 | 9.62 | 9.75 | 9.87 | 10 | 10.13
17.25 | 4.97 | 5.1 | 5.23 | 5.36 | 5.49 | 5.62 | 5.75 | 5.87 | 6 | 6.13 | 6.26 | 6.39 | 6.52 | 6.65 | 6.78 | 6.91 | 7.04 | 7.16 | 7.29 | 7.42 | 7.55 | 7.68 | 7.81 | 7.94 | 8.07 | 8.2 | 8.33 | 8.45 | 8.58 | 8.71 | 8.84 | 8.97 | 9.1 | 9.23 | 9.36 | 9.49 | 9.61 | 9.74 | 9.87 | 10 | 10.13
17.26 | 4.97 | 5.1 | 5.23 | 5.36 | 5.49 | 5.62 | 5.74 | 5.87 | 6 | 6.13 | 6.26 | 6.39 | 6.52 | 6.65 | 6.78 | 6.91 | 7.03 | 7.16 | 7.29 | 7.42 | 7.55 | 7.68 | 7.81 | 7.94 | 8.07 | 8.19 | 8.32 | 8.45 | 8.58 | 8.71 | 8.84 | 8.97 | 9.1 | 9.23 | 9.35 | 9.48 | 9.61 | 9.74 | 9.87 | 10 | 10.13
17.27 | 4.97 | 5.1 | 5.23 | 5.36 | 5.49 | 5.61 | 5.74 | 5.87 | 6 | 6.13 | 6.26 | 6.39 | 6.52 | 6.65 | 6.77 | 6.9 | 7.03 | 7.16 | 7.29 | 7.42 | 7.55 | 7.68 | 7.81 | 7.94 | 8.06 | 8.19 | 8.32 | 8.45 | 8.58 | 8.71 | 8.84 | 8.97 | 9.1 | 9.22 | 9.35 | 9.48 | 9.61 | 9.74 | 9.87 | 10 | 10.13
17.28 | 4.97 | 5.1 | 5.23 | 5.36 | 5.48 | 5.61 | 5.74 | 5.87 | 6 | 6.13 | 6.26 | 6.39 | 6.52 | 6.64 | 6.77 | 6.9 | 7.03 | 7.16 | 7.29 | 7.42 | 7.55 | 7.68 | 7.8 | 7.93 | 8.06 | 8.19 | 8.32 | 8.45 | 8.58 | 8.71 | 8.84 | 8.96 | 9.09 | 9.22 | 9.35 | 9.48 | 9.61 | 9.74 | 9.87 | 10 | 10.12
17.29 | 4.97 | 5.1 | 5.23 | 5.35 | 5.48 | 5.61 | 5.74 | 5.87 | 6 | 6.13 | 6.26 | 6.39 | 6.51 | 6.64 | 6.77 | 6.9 | 7.03 | 7.16 | 7.29 | 7.42 | 7.55 | 7.67 | 7.8 | 7.93 | 8.06 | 8.19 | 8.32 | 8.45 | 8.58 | 8.7 | 8.83 | 8.96 | 9.09 | 9.22 | 9.35 | 9.48 | 9.61 | 9.74 | 9.86 | 9.99 | 10.12
17.3 | 4.97 | 5.1 | 5.22 | 5.35 | 5.48 | 5.61 | 5.74 | 5.87 | 6 | 6.13 | 6.26 | 6.38 | 6.51 | 6.64 | 6.77 | 6.9 | 7.03 | 7.16 | 7.29 | 7.41 | 7.54 | 7.67 | 7.8 | 7.93 | 8.06 | 8.19 | 8.32 | 8.45 | 8.57 | 8.7 | 8.83 | 8.96 | 9.09 | 9.22 | 9.35 | 9.48 | 9.6 | 9.73 | 9.86 | 9.99 | 10.12
17.31 | 4.97 | 5.09 | 5.22 | 5.35 | 5.48 | 5.61 | 5.74 | 5.87 | 6 | 6.13 | 6.25 | 6.38 | 6.51 | 6.64 | 6.77 | 6.9 | 7.03 | 7.16 | 7.28 | 7.41 | 7.54 | 7.67 | 7.8 | 7.93 | 8.06 | 8.19 | 8.31 | 8.44 | 8.57 | 8.7 | 8.83 | 8.96 | 9.09 | 9.22 | 9.35 | 9.47 | 9.6 | 9.73 | 9.86 | 9.99 | 10.12
17.32 | 4.96 | 5.09 | 5.22 | 5.35 | 5.48 | 5.61 | 5.74 | 5.87 | 5.99 | 6.12 | 6.25 | 6.38 | 6.51 | 6.64 | 6.77 | 6.9 | 7.03 | 7.15 | 7.28 | 7.41 | 7.54 | 7.67 | 7.8 | 7.93 | 8.06 | 8.18 | 8.31 | 8.44 | 8.57 | 8.7 | 8.83 | 8.96 | 9.09 | 9.21 | 9.34 | 9.47 | 9.6 | 9.73 | 9.86 | 9.99 | 10.12
17.33 | 4.96 | 5.09 | 5.22 | 5.35 | 5.48 | 5.61 | 5.74 | 5.86 | 5.99 | 6.12 | 6.25 | 6.38 | 6.51 | 6.64 | 6.77 | 6.9 | 7.02 | 7.15 | 7.28 | 7.41 | 7.54 | 7.67 | 7.8 | 7.93 | 8.05 | 8.18 | 8.31 | 8.44 | 8.57 | 8.7 | 8.83 | 8.96 | 9.08 | 9.21 | 9.34 | 9.47 | 9.6 | 9.73 | 9.86 | 9.98 | 10.11
17.34 | 4.96 | 5.09 | 5.22 | 5.35 | 5.48 | 5.61 | 5.73 | 5.86 | 5.99 | 6.12 | 6.25 | 6.38 | 6.51 | 6.64 | 6.76 | 6.89 | 7.02 | 7.15 | 7.28 | 7.41 | 7.54 | 7.67 | 7.79 | 7.92 | 8.05 | 8.18 | 8.31 | 8.44 | 8.57 | 8.7 | 8.82 | 8.95 | 9.08 | 9.21 | 9.34 | 9.47 | 9.6 | 9.73 | 9.85 | 9.98 | 10.11
17.35 | 4.96 | 5.09 | 5.22 | 5.35 | 5.48 | 5.6 | 5.73 | 5.86 | 5.99 | 6.12 | 6.25 | 6.38 | 6.51 | 6.63 | 6.76 | 6.89 | 7.02 | 7.15 | 7.28 | 7.41 | 7.54 | 7.66 | 7.79 | 7.92 | 8.05 | 8.18 | 8.31 | 8.44 | 8.57 | 8.69 | 8.82 | 8.95 | 9.08 | 9.21 | 9.34 | 9.47 | 9.59 | 9.72 | 9.85 | 9.98 | 10.11
17.36 | 4.96 | 5.09 | 5.22 | 5.35 | 5.47 | 5.6 | 5.73 | 5.86 | 5.99 | 6.12 | 6.25 | 6.38 | 6.5 | 6.63 | 6.76 | 6.89 | 7.02 | 7.15 | 7.28 | 7.41 | 7.53 | 7.66 | 7.79 | 7.92 | 8.05 | 8.18 | 8.31 | 8.43 | 8.56 | 8.69 | 8.82 | 8.95 | 9.08 | 9.21 | 9.34 | 9.46 | 9.59 | 9.72 | 9.85 | 9.98 | 10.11
17.37 | 4.96 | 5.09 | 5.22 | 5.34 | 5.47 | 5.6 | 5.73 | 5.86 | 5.99 | 6.12 | 6.25 | 6.37 | 6.5 | 6.63 | 6.76 | 6.89 | 7.02 | 7.15 | 7.28 | 7.4 | 7.53 | 7.66 | 7.79 | 7.92 | 8.05 | 8.18 | 8.3 | 8.43 | 8.56 | 8.69 | 8.82 | 8.95 | 9.08 | 9.2 | 9.33 | 9.46 | 9.59 | 9.72 | 9.85 | 9.98 | 10.11
17.38 | 4.96 | 5.09 | 5.22 | 5.34 | 5.47 | 5.6 | 5.73 | 5.86 | 5.99 | 6.12 | 6.24 | 6.37 | 6.5 | 6.63 | 6.76 | 6.89 | 7.02 | 7.15 | 7.27 | 7.4 | 7.53 | 7.66 | 7.79 | 7.92 | 8.05 | 8.17 | 8.3 | 8.43 | 8.56 | 8.69 | 8.82 | 8.95 | 9.07 | 9.2 | 9.33 | 9.46 | 9.59 | 9.72 | 9.85 | 9.97 | 10.1
17.39 | 4.96 | 5.09 | 5.21 | 5.34 | 5.47 | 5.6 | 5.73 | 5.86 | 5.99 | 6.11 | 6.24 | 6.37 | 6.5 | 6.63 | 6.76 | 6.89 | 7.01 | 7.14 | 7.27 | 7.4 | 7.53 | 7.66 | 7.79 | 7.92 | 8.04 | 8.17 | 8.3 | 8.43 | 8.56 | 8.69 | 8.82 | 8.94 | 9.07 | 9.2 | 9.33 | 9.46 | 9.59 | 9.72 | 9.84 | 9.97 | 10.1
17.4 | 4.96 | 5.08 | 5.21 | 5.34 | 5.47 | 5.6 | 5.73 | 5.86 | 5.98 | 6.11 | 6.24 | 6.37 | 6.5 | 6.63 | 6.76 | 6.88 | 7.01 | 7.14 | 7.27 | 7.4 | 7.53 | 7.66 | 7.78 | 7.91 | 8.04 | 8.17 | 8.3 | 8.43 | 8.56 | 8.68 | 8.81 | 8.94 | 9.07 | 9.2 | 9.33 | 9.46 | 9.58 | 9.71 | 9.84 | 9.97 | 10.1
17.41 | 4.95 | 5.08 | 5.21 | 5.34 | 5.47 | 5.6 | 5.73 | 5.85 | 5.98 | 6.11 | 6.24 | 6.37 | 6.5 | 6.63 | 6.75 | 6.88 | 7.01 | 7.14 | 7.27 | 7.4 | 7.53 | 7.65 | 7.78 | 7.91 | 8.04 | 8.17 | 8.3 | 8.43 | 8.55 | 8.68 | 8.81 | 8.94 | 9.07 | 9.2 | 9.33 | 9.45 | 9.58 | 9.71 | 9.84 | 9.97 | 10.1
17.42 | 4.95 | 5.08 | 5.21 | 5.34 | 5.47 | 5.6 | 5.72 | 5.85 | 5.98 | 6.11 | 6.24 | 6.37 | 6.5 | 6.62 | 6.75 | 6.88 | 7.01 | 7.14 | 7.27 | 7.4 | 7.52 | 7.65 | 7.78 | 7.91 | 8.04 | 8.17 | 8.3 | 8.42 | 8.55 | 8.68 | 8.81 | 8.94 | 9.07 | 9.2 | 9.32 | 9.45 | 9.58 | 9.71 | 9.84 | 9.97 | 10.09
17.43 | 4.95 | 5.08 | 5.21 | 5.34 | 5.47 | 5.6 | 5.72 | 5.85 | 5.98 | 6.11 | 6.24 | 6.37 | 6.49 | 6.62 | 6.75 | 6.88 | 7.01 | 7.14 | 7.27 | 7.39 | 7.52 | 7.65 | 7.78 | 7.91 | 8.04 | 8.17 | 8.29 | 8.42 | 8.55 | 8.68 | 8.81 | 8.94 | 9.06 | 9.19 | 9.32 | 9.45 | 9.58 | 9.71 | 9.84 | 9.96 | 10.09
17.44 | 4.95 | 5.08 | 5.21 | 5.34 | 5.47 | 5.59 | 5.72 | 5.85 | 5.98 | 6.11 | 6.24 | 6.36 | 6.49 | 6.62 | 6.75 | 6.88 | 7.01 | 7.14 | 7.26 | 7.39 | 7.52 | 7.65 | 7.78 | 7.91 | 8.04 | 8.16 | 8.29 | 8.42 | 8.55 | 8.68 | 8.81 | 8.93 | 9.06 | 9.19 | 9.32 | 9.45 | 9.58 | 9.71 | 9.83 | 9.96 | 10.09
17.45 | 4.95 | 5.08 | 5.21 | 5.34 | 5.46 | 5.59 | 5.72 | 5.85 | 5.98 | 6.11 | 6.24 | 6.36 | 6.49 | 6.62 | 6.75 | 6.88 | 7.01 | 7.13 | 7.26 | 7.39 | 7.52 | 7.65 | 7.78 | 7.91 | 8.03 | 8.16 | 8.29 | 8.42 | 8.55 | 8.68 | 8.8 | 8.93 | 9.06 | 9.19 | 9.32 | 9.45 | 9.57 | 9.7 | 9.83 | 9.96 | 10.09
17.46 | 4.95 | 5.08 | 5.21 | 5.33 | 5.46 | 5.59 | 5.72 | 5.85 | 5.98 | 6.11 | 6.23 | 6.36 | 6.49 | 6.62 | 6.75 | 6.88 | 7 | 7.13 | 7.26 | 7.39 | 7.52 | 7.65 | 7.78 | 7.9 | 8.03 | 8.16 | 8.29 | 8.42 | 8.55 | 8.67 | 8.8 | 8.93 | 9.06 | 9.19 | 9.32 | 9.44 | 9.57 | 9.7 | 9.83 | 9.96 | 10.09
17.47 | 4.95 | 5.08 | 5.2 | 5.33 | 5.46 | 5.59 | 5.72 | 5.85 | 5.98 | 6.1 | 6.23 | 6.36 | 6.49 | 6.62 | 6.75 | 6.87 | 7 | 7.13 | 7.26 | 7.39 | 7.52 | 7.65 | 7.77 | 7.9 | 8.03 | 8.16 | 8.29 | 8.42 | 8.54 | 8.67 | 8.8 | 8.93 | 9.06 | 9.19 | 9.31 | 9.44 | 9.57 | 9.7 | 9.83 | 9.96 | 10.08
17.48 | 4.95 | 5.08 | 5.2 | 5.33 | 5.46 | 5.59 | 5.72 | 5.85 | 5.97 | 6.1 | 6.23 | 6.36 | 6.49 | 6.62 | 6.74 | 6.87 | 7 | 7.13 | 7.26 | 7.39 | 7.52 | 7.64 | 7.77 | 7.9 | 8.03 | 8.16 | 8.29 | 8.41 | 8.54 | 8.67 | 8.8 | 8.93 | 9.06 | 9.18 | 9.31 | 9.44 | 9.57 | 9.7 | 9.83 | 9.95 | 10.08
17.49 | 4.95 | 5.07 | 5.2 | 5.33 | 5.46 | 5.59 | 5.72 | 5.84 | 5.97 | 6.1 | 6.23 | 6.36 | 6.49 | 6.61 | 6.74 | 6.87 | 7 | 7.13 | 7.26 | 7.39 | 7.51 | 7.64 | 7.77 | 7.9 | 8.03 | 8.16 | 8.28 | 8.41 | 8.54 | 8.67 | 8.8 | 8.93 | 9.05 | 9.18 | 9.31 | 9.44 | 9.57 | 9.7 | 9.82 | 9.95 | 10.08
17.5 | 4.94 | 5.07 | 5.2 | 5.33 | 5.46 | 5.59 | 5.72 | 5.84 | 5.97 | 6.1 | 6.23 | 6.36 | 6.49 | 6.61 | 6.74 | 6.87 | 7 | 7.13 | 7.26 | 7.38 | 7.51 | 7.64 | 7.77 | 7.9 | 8.03 | 8.15 | 8.28 | 8.41 | 8.54 | 8.67 | 8.8 | 8.92 | 9.05 | 9.18 | 9.31 | 9.44 | 9.57 | 9.69 | 9.82 | 9.95 | 10.08
17.51 | 4.94 | 5.07 | 5.2 | 5.33 | 5.46 | 5.59 | 5.71 | 5.84 | 5.97 | 6.1 | 6.23 | 6.36 | 6.48 | 6.61 | 6.74 | 6.87 | 7 | 7.13 | 7.25 | 7.38 | 7.51 | 7.64 | 7.77 | 7.9 | 8.02 | 8.15 | 8.28 | 8.41 | 8.54 | 8.67 | 8.79 | 8.92 | 9.05 | 9.18 | 9.31 | 9.43 | 9.56 | 9.69 | 9.82 | 9.95 | 10.08
17.52 | 4.94 | 5.07 | 5.2 | 5.33 | 5.46 | 5.58 | 5.71 | 5.84 | 5.97 | 6.1 | 6.23 | 6.35 | 6.48 | 6.61 | 6.74 | 6.87 | 7 | 7.12 | 7.25 | 7.38 | 7.51 | 7.64 | 7.77 | 7.89 | 8.02 | 8.15 | 8.28 | 8.41 | 8.54 | 8.66 | 8.79 | 8.92 | 9.05 | 9.18 | 9.3 | 9.43 | 9.56 | 9.69 | 9.82 | 9.95 | 10.07
17.53 | 4.94 | 5.07 | 5.2 | 5.33 | 5.45 | 5.58 | 5.71 | 5.84 | 5.97 | 6.1 | 6.22 | 6.35 | 6.48 | 6.61 | 6.74 | 6.87 | 6.99 | 7.12 | 7.25 | 7.38 | 7.51 | 7.64 | 7.76 | 7.89 | 8.02 | 8.15 | 8.28 | 8.4 | 8.53 | 8.66 | 8.79 | 8.92 | 9.05 | 9.17 | 9.3 | 9.43 | 9.56 | 9.69 | 9.82 | 9.94 | 10.07
17.54 | 4.94 | 5.07 | 5.2 | 5.33 | 5.45 | 5.58 | 5.71 | 5.84 | 5.97 | 6.09 | 6.22 | 6.35 | 6.48 | 6.61 | 6.74 | 6.86 | 6.99 | 7.12 | 7.25 | 7.38 | 7.51 | 7.63 | 7.76 | 7.89 | 8.02 | 8.15 | 8.28 | 8.4 | 8.53 | 8.66 | 8.79 | 8.92 | 9.04 | 9.17 | 9.3 | 9.43 | 9.56 | 9.69 | 9.81 | 9.94 | 10.07
17.55 | 4.94 | 5.07 | 5.2 | 5.32 | 5.45 | 5.58 | 5.71 | 5.84 | 5.97 | 6.09 | 6.22 | 6.35 | 6.48 | 6.61 | 6.73 | 6.86 | 6.99 | 7.12 | 7.25 | 7.38 | 7.5 | 7.63 | 7.76 | 7.89 | 8.02 | 8.15 | 8.27 | 8.4 | 8.53 | 8.66 | 8.79 | 8.91 | 9.04 | 9.17 | 9.3 | 9.43 | 9.56 | 9.68 | 9.81 | 9.94 | 10.07
17.56 | 4.94 | 5.07 | 5.19 | 5.32 | 5.45 | 5.58 | 5.71 | 5.84 | 5.96 | 6.09 | 6.22 | 6.35 | 6.48 | 6.61 | 6.73 | 6.86 | 6.99 | 7.12 | 7.25 | 7.37 | 7.5 | 7.63 | 7.76 | 7.89 | 8.02 | 8.14 | 8.27 | 8.4 | 8.53 | 8.66 | 8.78 | 8.91 | 9.04 | 9.17 | 9.3 | 9.43 | 9.55 | 9.68 | 9.81 | 9.94 | 10.07
17.57 | 4.94 | 5.07 | 5.19 | 5.32 | 5.45 | 5.58 | 5.71 | 5.83 | 5.96 | 6.09 | 6.22 | 6.35 | 6.48 | 6.6 | 6.73 | 6.86 | 6.99 | 7.12 | 7.24 | 7.37 | 7.5 | 7.63 | 7.76 | 7.89 | 8.01 | 8.14 | 8.27 | 8.4 | 8.53 | 8.65 | 8.78 | 8.91 | 9.04 | 9.17 | 9.29 | 9.42 | 9.55 | 9.68 | 9.81 | 9.94 | 10.06
17.58 | 4.94 | 5.06 | 5.19 | 5.32 | 5.45 | 5.58 | 5.71 | 5.83 | 5.96 | 6.09 | 6.22 | 6.35 | 6.47 | 6.6 | 6.73 | 6.86 | 6.99 | 7.11 | 7.24 | 7.37 | 7.5 | 7.63 | 7.76 | 7.88 | 8.01 | 8.14 | 8.27 | 8.4 | 8.52 | 8.65 | 8.78 | 8.91 | 9.04 | 9.16 | 9.29 | 9.42 | 9.55 | 9.68 | 9.81 | 9.93 | 10.06
17.59 | 4.94 | 5.06 | 5.19 | 5.32 | 5.45 | 5.58 | 5.7 | 5.83 | 5.96 | 6.09 | 6.22 | 6.34 | 6.47 | 6.6 | 6.73 | 6.86 | 6.99 | 7.11 | 7.24 | 7.37 | 7.5 | 7.63 | 7.75 | 7.88 | 8.01 | 8.14 | 8.27 | 8.39 | 8.52 | 8.65 | 8.78 | 8.91 | 9.03 | 9.16 | 9.29 | 9.42 | 9.55 | 9.68 | 9.8 | 9.93 | 10.06
17.6 | 4.93 | 5.06 | 5.19 | 5.32 | 5.45 | 5.57 | 5.7 | 5.83 | 5.96 | 6.09 | 6.22 | 6.34 | 6.47 | 6.6 | 6.73 | 6.86 | 6.98 | 7.11 | 7.24 | 7.37 | 7.5 | 7.62 | 7.75 | 7.88 | 8.01 | 8.14 | 8.26 | 8.39 | 8.52 | 8.65 | 8.78 | 8.91 | 9.03 | 9.16 | 9.29 | 9.42 | 9.55 | 9.67 | 9.8 | 9.93 | 10.06
17.61 | 4.93 | 5.06 | 5.19 | 5.32 | 5.45 | 5.57 | 5.7 | 5.83 | 5.96 | 6.09 | 6.21 | 6.34 | 6.47 | 6.6 | 6.73 | 6.85 | 6.98 | 7.11 | 7.24 | 7.37 | 7.49 | 7.62 | 7.75 | 7.88 | 8.01 | 8.13 | 8.26 | 8.39 | 8.52 | 8.65 | 8.78 | 8.9 | 9.03 | 9.16 | 9.29 | 9.42 | 9.54 | 9.67 | 9.8 | 9.93 | 10.06
17.62 | 4.93 | 5.06 | 5.19 | 5.32 | 5.44 | 5.57 | 5.7 | 5.83 | 5.96 | 6.08 | 6.21 | 6.34 | 6.47 | 6.6 | 6.72 | 6.85 | 6.98 | 7.11 | 7.24 | 7.37 | 7.49 | 7.62 | 7.75 | 7.88 | 8.01 | 8.13 | 8.26 | 8.39 | 8.52 | 8.65 | 8.77 | 8.9 | 9.03 | 9.16 | 9.29 | 9.41 | 9.54 | 9.67 | 9.8 | 9.93 | 10.05
17.63 | 4.93 | 5.06 | 5.19 | 5.31 | 5.44 | 5.57 | 5.7 | 5.83 | 5.96 | 6.08 | 6.21 | 6.34 | 6.47 | 6.6 | 6.72 | 6.85 | 6.98 | 7.11 | 7.24 | 7.36 | 7.49 | 7.62 | 7.75 | 7.88 | 8 | 8.13 | 8.26 | 8.39 | 8.52 | 8.64 | 8.77 | 8.9 | 9.03 | 9.16 | 9.28 | 9.41 | 9.54 | 9.67 | 9.8 | 9.92 | 10.05
17.64 | 4.93 | 5.06 | 5.19 | 5.31 | 5.44 | 5.57 | 5.7 | 5.83 | 5.95 | 6.08 | 6.21 | 6.34 | 6.47 | 6.59 | 6.72 | 6.85 | 6.98 | 7.11 | 7.23 | 7.36 | 7.49 | 7.62 | 7.75 | 7.87 | 8 | 8.13 | 8.26 | 8.39 | 8.51 | 8.64 | 8.77 | 8.9 | 9.03 | 9.15 | 9.28 | 9.41 | 9.54 | 9.67 | 9.79 | 9.92 | 10.05
17.65 | 4.93 | 5.06 | 5.18 | 5.31 | 5.44 | 5.57 | 5.7 | 5.82 | 5.95 | 6.08 | 6.21 | 6.34 | 6.46 | 6.59 | 6.72 | 6.85 | 6.98 | 7.1 | 7.23 | 7.36 | 7.49 | 7.62 | 7.74 | 7.87 | 8 | 8.13 | 8.26 | 8.38 | 8.51 | 8.64 | 8.77 | 8.9 | 9.02 | 9.15 | 9.28 | 9.41 | 9.54 | 9.66 | 9.79 | 9.92 | 10.05
17.66 | 4.93 | 5.06 | 5.18 | 5.31 | 5.44 | 5.57 | 5.7 | 5.82 | 5.95 | 6.08 | 6.21 | 6.34 | 6.46 | 6.59 | 6.72 | 6.85 | 6.97 | 7.1 | 7.23 | 7.36 | 7.49 | 7.61 | 7.74 | 7.87 | 8 | 8.13 | 8.25 | 8.38 | 8.51 | 8.64 | 8.77 | 8.89 | 9.02 | 9.15 | 9.28 | 9.41 | 9.53 | 9.66 | 9.79 | 9.92 | 10.05
17.67 | 4.93 | 5.05 | 5.18 | 5.31 | 5.44 | 5.57 | 5.69 | 5.82 | 5.95 | 6.08 | 6.21 | 6.33 | 6.46 | 6.59 | 6.72 | 6.85 | 6.97 | 7.1 | 7.23 | 7.36 | 7.49 | 7.61 | 7.74 | 7.87 | 8 | 8.12 | 8.25 | 8.38 | 8.51 | 8.64 | 8.76 | 8.89 | 9.02 | 9.15 | 9.28 | 9.4 | 9.53 | 9.66 | 9.79 | 9.92 | 10.04
17.68 | 4.93 | 5.05 | 5.18 | 5.31 | 5.44 | 5.57 | 5.69 | 5.82 | 5.95 | 6.08 | 6.2 | 6.33 | 6.46 | 6.59 | 6.72 | 6.84 | 6.97 | 7.1 | 7.23 | 7.36 | 7.48 | 7.61 | 7.74 | 7.87 | 8 | 8.12 | 8.25 | 8.38 | 8.51 | 8.63 | 8.76 | 8.89 | 9.02 | 9.15 | 9.27 | 9.4 | 9.53 | 9.66 | 9.79 | 9.91 | 10.04
17.69 | 4.92 | 5.05 | 5.18 | 5.31 | 5.44 | 5.56 | 5.69 | 5.82 | 5.95 | 6.08 | 6.2 | 6.33 | 6.46 | 6.59 | 6.71 | 6.84 | 6.97 | 7.1 | 7.23 | 7.35 | 7.48 | 7.61 | 7.74 | 7.87 | 7.99 | 8.12 | 8.25 | 8.38 | 8.5 | 8.63 | 8.76 | 8.89 | 9.02 | 9.14 | 9.27 | 9.4 | 9.53 | 9.66 | 9.78 | 9.91 | 10.04
17.7 | 4.92 | 5.05 | 5.18 | 5.31 | 5.43 | 5.56 | 5.69 | 5.82 | 5.95 | 6.07 | 6.2 | 6.33 | 6.46 | 6.59 | 6.71 | 6.84 | 6.97 | 7.1 | 7.22 | 7.35 | 7.48 | 7.61 | 7.74 | 7.86 | 7.99 | 8.12 | 8.25 | 8.38 | 8.5 | 8.63 | 8.76 | 8.89 | 9.01 | 9.14 | 9.27 | 9.4 | 9.53 | 9.65 | 9.78 | 9.91 | 10.04
17.71 | 4.92 | 5.05 | 5.18 | 5.31 | 5.43 | 5.56 | 5.69 | 5.82 | 5.94 | 6.07 | 6.2 | 6.33 | 6.46 | 6.58 | 6.71 | 6.84 | 6.97 | 7.1 | 7.22 | 7.35 | 7.48 | 7.61 | 7.73 | 7.86 | 7.99 | 8.12 | 8.25 | 8.37 | 8.5 | 8.63 | 8.76 | 8.88 | 9.01 | 9.14 | 9.27 | 9.4 | 9.52 | 9.65 | 9.78 | 9.91 | 10.03
17.72 | 4.92 | 5.05 | 5.18 | 5.3 | 5.43 | 5.56 | 5.69 | 5.82 | 5.94 | 6.07 | 6.2 | 6.33 | 6.45 | 6.58 | 6.71 | 6.84 | 6.97 | 7.09 | 7.22 | 7.35 | 7.48 | 7.61 | 7.73 | 7.86 | 7.99 | 8.12 | 8.24 | 8.37 | 8.5 | 8.63 | 8.76 | 8.88 | 9.01 | 9.14 | 9.27 | 9.39 | 9.52 | 9.65 | 9.78 | 9.9 | 10.03
17.73 | 4.92 | 5.05 | 5.18 | 5.3 | 5.43 | 5.56 | 5.69 | 5.81 | 5.94 | 6.07 | 6.2 | 6.33 | 6.45 | 6.58 | 6.71 | 6.84 | 6.96 | 7.09 | 7.22 | 7.35 | 7.48 | 7.6 | 7.73 | 7.86 | 7.99 | 8.11 | 8.24 | 8.37 | 8.5 | 8.63 | 8.75 | 8.88 | 9.01 | 9.14 | 9.26 | 9.39 | 9.52 | 9.65 | 9.78 | 9.9 | 10.03
17.74 | 4.92 | 5.05 | 5.17 | 5.3 | 5.43 | 5.56 | 5.69 | 5.81 | 5.94 | 6.07 | 6.2 | 6.32 | 6.45 | 6.58 | 6.71 | 6.84 | 6.96 | 7.09 | 7.22 | 7.35 | 7.47 | 7.6 | 7.73 | 7.86 | 7.99 | 8.11 | 8.24 | 8.37 | 8.5 | 8.62 | 8.75 | 8.88 | 9.01 | 9.13 | 9.26 | 9.39 | 9.52 | 9.65 | 9.77 | 9.9 | 10.03
17.75 | 4.92 | 5.05 | 5.17 | 5.3 | 5.43 | 5.56 | 5.68 | 5.81 | 5.94 | 6.07 | 6.2 | 6.32 | 6.45 | 6.58 | 6.71 | 6.83 | 6.96 | 7.09 | 7.22 | 7.34 | 7.47 | 7.6 | 7.73 | 7.86 | 7.98 | 8.11 | 8.24 | 8.37 | 8.49 | 8.62 | 8.75 | 8.88 | 9.01 | 9.13 | 9.26 | 9.39 | 9.52 | 9.64 | 9.77 | 9.9 | 10.03
17.76 | 4.92 | 5.04 | 5.17 | 5.3 | 5.43 | 5.56 | 5.68 | 5.81 | 5.94 | 6.07 | 6.19 | 6.32 | 6.45 | 6.58 | 6.7 | 6.83 | 6.96 | 7.09 | 7.22 | 7.34 | 7.47 | 7.6 | 7.73 | 7.85 | 7.98 | 8.11 | 8.24 | 8.36 | 8.49 | 8.62 | 8.75 | 8.88 | 9 | 9.13 | 9.26 | 9.39 | 9.51 | 9.64 | 9.77 | 9.9 | 10.02
17.77 | 4.92 | 5.04 | 5.17 | 5.3 | 5.43 | 5.55 | 5.68 | 5.81 | 5.94 | 6.07 | 6.19 | 6.32 | 6.45 | 6.58 | 6.7 | 6.83 | 6.96 | 7.09 | 7.21 | 7.34 | 7.47 | 7.6 | 7.72 | 7.85 | 7.98 | 8.11 | 8.24 | 8.36 | 8.49 | 8.62 | 8.75 | 8.87 | 9 | 9.13 | 9.26 | 9.38 | 9.51 | 9.64 | 9.77 | 9.89 | 10.02
17.78 | 4.91 | 5.04 | 5.17 | 5.3 | 5.43 | 5.55 | 5.68 | 5.81 | 5.94 | 6.06 | 6.19 | 6.32 | 6.45 | 6.57 | 6.7 | 6.83 | 6.96 | 7.09 | 7.21 | 7.34 | 7.47 | 7.6 | 7.72 | 7.85 | 7.98 | 8.11 | 8.23 | 8.36 | 8.49 | 8.62 | 8.74 | 8.87 | 9 | 9.13 | 9.25 | 9.38 | 9.51 | 9.64 | 9.77 | 9.89 | 10.02
17.79 | 4.91 | 5.04 | 5.17 | 5.3 | 5.42 | 5.55 | 5.68 | 5.81 | 5.93 | 6.06 | 6.19 | 6.32 | 6.45 | 6.57 | 6.7 | 6.83 | 6.96 | 7.08 | 7.21 | 7.34 | 7.47 | 7.59 | 7.72 | 7.85 | 7.98 | 8.1 | 8.23 | 8.36 | 8.49 | 8.61 | 8.74 | 8.87 | 9 | 9.13 | 9.25 | 9.38 | 9.51 | 9.64 | 9.76 | 9.89 | 10.02
17.8 | 4.91 | 5.04 | 5.17 | 5.3 | 5.42 | 5.55 | 5.68 | 5.81 | 5.93 | 6.06 | 6.19 | 6.32 | 6.44 | 6.57 | 6.7 | 6.83 | 6.95 | 7.08 | 7.21 | 7.34 | 7.46 | 7.59 | 7.72 | 7.85 | 7.98 | 8.1 | 8.23 | 8.36 | 8.49 | 8.61 | 8.74 | 8.87 | 9 | 9.12 | 9.25 | 9.38 | 9.51 | 9.63 | 9.76 | 9.89 | 10.02
17.81 | 4.91 | 5.04 | 5.17 | 5.29 | 5.42 | 5.55 | 5.68 | 5.8 | 5.93 | 6.06 | 6.19 | 6.32 | 6.44 | 6.57 | 6.7 | 6.83 | 6.95 | 7.08 | 7.21 | 7.34 | 7.46 | 7.59 | 7.72 | 7.85 | 7.97 | 8.1 | 8.23 | 8.36 | 8.48 | 8.61 | 8.74 | 8.87 | 8.99 | 9.12 | 9.25 | 9.38 | 9.5 | 9.63 | 9.76 | 9.89 | 10.01
17.82 | 4.91 | 5.04 | 5.17 | 5.29 | 5.42 | 5.55 | 5.68 | 5.8 | 5.93 | 6.06 | 6.19 | 6.31 | 6.44 | 6.57 | 6.7 | 6.82 | 6.95 | 7.08 | 7.21 | 7.33 | 7.46 | 7.59 | 7.72 | 7.84 | 7.97 | 8.1 | 8.23 | 8.35 | 8.48 | 8.61 | 8.74 | 8.86 | 8.99 | 9.12 | 9.25 | 9.37 | 9.5 | 9.63 | 9.76 | 9.88 | 10.01
17.83 | 4.91 | 5.04 | 5.16 | 5.29 | 5.42 | 5.55 | 5.67 | 5.8 | 5.93 | 6.06 | 6.18 | 6.31 | 6.44 | 6.57 | 6.69 | 6.82 | 6.95 | 7.08 | 7.21 | 7.33 | 7.46 | 7.59 | 7.72 | 7.84 | 7.97 | 8.1 | 8.23 | 8.35 | 8.48 | 8.61 | 8.74 | 8.86 | 8.99 | 9.12 | 9.25 | 9.37 | 9.5 | 9.63 | 9.76 | 9.88 | 10.01
17.84 | 4.91 | 5.04 | 5.16 | 5.29 | 5.42 | 5.55 | 5.67 | 5.8 | 5.93 | 6.06 | 6.18 | 6.31 | 6.44 | 6.57 | 6.69 | 6.82 | 6.95 | 7.08 | 7.2 | 7.33 | 7.46 | 7.59 | 7.71 | 7.84 | 7.97 | 8.1 | 8.22 | 8.35 | 8.48 | 8.61 | 8.73 | 8.86 | 8.99 | 9.12 | 9.24 | 9.37 | 9.5 | 9.63 | 9.75 | 9.88 | 10.01
17.85 | 4.91 | 5.03 | 5.16 | 5.29 | 5.42 | 5.54 | 5.67 | 5.8 | 5.93 | 6.05 | 6.18 | 6.31 | 6.44 | 6.56 | 6.69 | 6.82 | 6.95 | 7.07 | 7.2 | 7.33 | 7.46 | 7.58 | 7.71 | 7.84 | 7.97 | 8.09 | 8.22 | 8.35 | 8.48 | 8.6 | 8.73 | 8.86 | 8.99 | 9.11 | 9.24 | 9.37 | 9.5 | 9.62 | 9.75 | 9.88 | 10.01
17.86 | 4.91 | 5.03 | 5.16 | 5.29 | 5.42 | 5.54 | 5.67 | 5.8 | 5.93 | 6.05 | 6.18 | 6.31 | 6.44 | 6.56 | 6.69 | 6.82 | 6.95 | 7.07 | 7.2 | 7.33 | 7.46 | 7.58 | 7.71 | 7.84 | 7.97 | 8.09 | 8.22 | 8.35 | 8.47 | 8.6 | 8.73 | 8.86 | 8.98 | 9.11 | 9.24 | 9.37 | 9.49 | 9.62 | 9.75 | 9.88 | 10
17.87 | 4.9 | 5.03 | 5.16 | 5.29 | 5.41 | 5.54 | 5.67 | 5.8 | 5.92 | 6.05 | 6.18 | 6.31 | 6.43 | 6.56 | 6.69 | 6.82 | 6.94 | 7.07 | 7.2 | 7.33 | 7.45 | 7.58 | 7.71 | 7.84 | 7.96 | 8.09 | 8.22 | 8.35 | 8.47 | 8.6 | 8.73 | 8.86 | 8.98 | 9.11 | 9.24 | 9.36 | 9.49 | 9.62 | 9.75 | 9.87 | 10
17.88 | 4.9 | 5.03 | 5.16 | 5.29 | 5.41 | 5.54 | 5.67 | 5.8 | 5.92 | 6.05 | 6.18 | 6.31 | 6.43 | 6.56 | 6.69 | 6.82 | 6.94 | 7.07 | 7.2 | 7.32 | 7.45 | 7.58 | 7.71 | 7.83 | 7.96 | 8.09 | 8.22 | 8.34 | 8.47 | 8.6 | 8.73 | 8.85 | 8.98 | 9.11 | 9.24 | 9.36 | 9.49 | 9.62 | 9.75 | 9.87 | 10
17.89 | 4.9 | 5.03 | 5.16 | 5.29 | 5.41 | 5.54 | 5.67 | 5.79 | 5.92 | 6.05 | 6.18 | 6.3 | 6.43 | 6.56 | 6.69 | 6.81 | 6.94 | 7.07 | 7.2 | 7.32 | 7.45 | 7.58 | 7.71 | 7.83 | 7.96 | 8.09 | 8.21 | 8.34 | 8.47 | 8.6 | 8.72 | 8.85 | 8.98 | 9.11 | 9.23 | 9.36 | 9.49 | 9.62 | 9.74 | 9.87 | 10
17.9 | 4.9 | 5.03 | 5.16 | 5.28 | 5.41 | 5.54 | 5.67 | 5.79 | 5.92 | 6.05 | 6.18 | 6.3 | 6.43 | 6.56 | 6.69 | 6.81 | 6.94 | 7.07 | 7.19 | 7.32 | 7.45 | 7.58 | 7.7 | 7.83 | 7.96 | 8.09 | 8.21 | 8.34 | 8.47 | 8.6 | 8.72 | 8.85 | 8.98 | 9.1 | 9.23 | 9.36 | 9.49 | 9.61 | 9.74 | 9.87 | 10
17.91 | 4.9 | 5.03 | 5.16 | 5.28 | 5.41 | 5.54 | 5.66 | 5.79 | 5.92 | 6.05 | 6.17 | 6.3 | 6.43 | 6.56 | 6.68 | 6.81 | 6.94 | 7.07 | 7.19 | 7.32 | 7.45 | 7.57 | 7.7 | 7.83 | 7.96 | 8.08 | 8.21 | 8.34 | 8.47 | 8.59 | 8.72 | 8.85 | 8.98 | 9.1 | 9.23 | 9.36 | 9.48 | 9.61 | 9.74 | 9.87 | 9.99
17.92 | 4.9 | 5.03 | 5.15 | 5.28 | 5.41 | 5.54 | 5.66 | 5.79 | 5.92 | 6.05 | 6.17 | 6.3 | 6.43 | 6.55 | 6.68 | 6.81 | 6.94 | 7.06 | 7.19 | 7.32 | 7.45 | 7.57 | 7.7 | 7.83 | 7.96 | 8.08 | 8.21 | 8.34 | 8.46 | 8.59 | 8.72 | 8.85 | 8.97 | 9.1 | 9.23 | 9.36 | 9.48 | 9.61 | 9.74 | 9.86 | 9.99
17.93 | 4.9 | 5.03 | 5.15 | 5.28 | 5.41 | 5.54 | 5.66 | 5.79 | 5.92 | 6.04 | 6.17 | 6.3 | 6.43 | 6.55 | 6.68 | 6.81 | 6.94 | 7.06 | 7.19 | 7.32 | 7.44 | 7.57 | 7.7 | 7.83 | 7.95 | 8.08 | 8.21 | 8.34 | 8.46 | 8.59 | 8.72 | 8.84 | 8.97 | 9.1 | 9.23 | 9.35 | 9.48 | 9.61 | 9.74 | 9.86 | 9.99
17.94 | 4.9 | 5.02 | 5.15 | 5.28 | 5.41 | 5.53 | 5.66 | 5.79 | 5.92 | 6.04 | 6.17 | 6.3 | 6.42 | 6.55 | 6.68 | 6.81 | 6.93 | 7.06 | 7.19 | 7.32 | 7.44 | 7.57 | 7.7 | 7.82 | 7.95 | 8.08 | 8.21 | 8.33 | 8.46 | 8.59 | 8.72 | 8.84 | 8.97 | 9.1 | 9.22 | 9.35 | 9.48 | 9.61 | 9.73 | 9.86 | 9.99
17.95 | 4.9 | 5.02 | 5.15 | 5.28 | 5.41 | 5.53 | 5.66 | 5.79 | 5.91 | 6.04 | 6.17 | 6.3 | 6.42 | 6.55 | 6.68 | 6.81 | 6.93 | 7.06 | 7.19 | 7.31 | 7.44 | 7.57 | 7.7 | 7.82 | 7.95 | 8.08 | 8.2 | 8.33 | 8.46 | 8.59 | 8.71 | 8.84 | 8.97 | 9.1 | 9.22 | 9.35 | 9.48 | 9.6 | 9.73 | 9.86 | 9.99
17.96 | 4.9 | 5.02 | 5.15 | 5.28 | 5.4 | 5.53 | 5.66 | 5.79 | 5.91 | 6.04 | 6.17 | 6.29 | 6.42 | 6.55 | 6.68 | 6.8 | 6.93 | 7.06 | 7.19 | 7.31 | 7.44 | 7.57 | 7.69 | 7.82 | 7.95 | 8.08 | 8.2 | 8.33 | 8.46 | 8.58 | 8.71 | 8.84 | 8.97 | 9.09 | 9.22 | 9.35 | 9.47 | 9.6 | 9.73 | 9.86 | 9.98
17.97 | 4.89 | 5.02 | 5.15 | 5.28 | 5.4 | 5.53 | 5.66 | 5.78 | 5.91 | 6.04 | 6.17 | 6.29 | 6.42 | 6.55 | 6.68 | 6.8 | 6.93 | 7.06 | 7.18 | 7.31 | 7.44 | 7.57 | 7.69 | 7.82 | 7.95 | 8.07 | 8.2 | 8.33 | 8.46 | 8.58 | 8.71 | 8.84 | 8.96 | 9.09 | 9.22 | 9.35 | 9.47 | 9.6 | 9.73 | 9.85 | 9.98
17.98 | 4.89 | 5.02 | 5.15 | 5.27 | 5.4 | 5.53 | 5.66 | 5.78 | 5.91 | 6.04 | 6.17 | 6.29 | 6.42 | 6.55 | 6.67 | 6.8 | 6.93 | 7.06 | 7.18 | 7.31 | 7.44 | 7.56 | 7.69 | 7.82 | 7.95 | 8.07 | 8.2 | 8.33 | 8.45 | 8.58 | 8.71 | 8.84 | 8.96 | 9.09 | 9.22 | 9.34 | 9.47 | 9.6 | 9.73 | 9.85 | 9.98
17.99 | 4.89 | 5.02 | 5.15 | 5.27 | 5.4 | 5.53 | 5.66 | 5.78 | 5.91 | 6.04 | 6.16 | 6.29 | 6.42 | 6.55 | 6.67 | 6.8 | 6.93 | 7.05 | 7.18 | 7.31 | 7.44 | 7.56 | 7.69 | 7.82 | 7.94 | 8.07 | 8.2 | 8.33 | 8.45 | 8.58 | 8.71 | 8.83 | 8.96 | 9.09 | 9.21 | 9.34 | 9.47 | 9.6 | 9.72 | 9.85 | 9.98
18 | 4.89 | 5.02 | 5.15 | 5.27 | 5.4 | 5.53 | 5.65 | 5.78 | 5.91 | 6.04 | 6.16 | 6.29 | 6.42 | 6.54 | 6.67 | 6.8 | 6.93 | 7.05 | 7.18 | 7.31 | 7.43 | 7.56 | 7.69 | 7.81 | 7.94 | 8.07 | 8.2 | 8.32 | 8.45 | 8.58 | 8.7 | 8.83 | 8.96 | 9.09 | 9.21 | 9.34 | 9.47 | 9.59 | 9.72 | 9.85 | 9.98
18.01 | 4.89 | 5.02 | 5.14 | 5.27 | 5.4 | 5.53 | 5.65 | 5.78 | 5.91 | 6.03 | 6.16 | 6.29 | 6.42 | 6.54 | 6.67 | 6.8 | 6.92 | 7.05 | 7.18 | 7.31 | 7.43 | 7.56 | 7.69 | 7.81 | 7.94 | 8.07 | 8.19 | 8.32 | 8.45 | 8.58 | 8.7 | 8.83 | 8.96 | 9.08 | 9.21 | 9.34 | 9.47 | 9.59 | 9.72 | 9.85 | 9.97
18.02 | 4.89 | 5.02 | 5.14 | 5.27 | 5.4 | 5.52 | 5.65 | 5.78 | 5.91 | 6.03 | 6.16 | 6.29 | 6.41 | 6.54 | 6.67 | 6.8 | 6.92 | 7.05 | 7.18 | 7.3 | 7.43 | 7.56 | 7.68 | 7.81 | 7.94 | 8.07 | 8.19 | 8.32 | 8.45 | 8.57 | 8.7 | 8.83 | 8.96 | 9.08 | 9.21 | 9.34 | 9.46 | 9.59 | 9.72 | 9.84 | 9.97
18.03 | 4.89 | 5.02 | 5.14 | 5.27 | 5.4 | 5.52 | 5.65 | 5.78 | 5.9 | 6.03 | 6.16 | 6.29 | 6.41 | 6.54 | 6.67 | 6.79 | 6.92 | 7.05 | 7.17 | 7.3 | 7.43 | 7.56 | 7.68 | 7.81 | 7.94 | 8.06 | 8.19 | 8.32 | 8.45 | 8.57 | 8.7 | 8.83 | 8.95 | 9.08 | 9.21 | 9.33 | 9.46 | 9.59 | 9.72 | 9.84 | 9.97
18.04 | 4.89 | 5.01 | 5.14 | 5.27 | 5.4 | 5.52 | 5.65 | 5.78 | 5.9 | 6.03 | 6.16 | 6.28 | 6.41 | 6.54 | 6.67 | 6.79 | 6.92 | 7.05 | 7.17 | 7.3 | 7.43 | 7.55 | 7.68 | 7.81 | 7.94 | 8.06 | 8.19 | 8.32 | 8.44 | 8.57 | 8.7 | 8.82 | 8.95 | 9.08 | 9.21 | 9.33 | 9.46 | 9.59 | 9.71 | 9.84 | 9.97
18.05 | 4.89 | 5.01 | 5.14 | 5.27 | 5.39 | 5.52 | 5.65 | 5.77 | 5.9 | 6.03 | 6.16 | 6.28 | 6.41 | 6.54 | 6.66 | 6.79 | 6.92 | 7.04 | 7.17 | 7.3 | 7.43 | 7.55 | 7.68 | 7.81 | 7.93 | 8.06 | 8.19 | 8.31 | 8.44 | 8.57 | 8.7 | 8.82 | 8.95 | 9.08 | 9.2 | 9.33 | 9.46 | 9.58 | 9.71 | 9.84 | 9.97
18.06 | 4.88 | 5.01 | 5.14 | 5.27 | 5.39 | 5.52 | 5.65 | 5.77 | 5.9 | 6.03 | 6.15 | 6.28 | 6.41 | 6.54 | 6.66 | 6.79 | 6.92 | 7.04 | 7.17 | 7.3 | 7.42 | 7.55 | 7.68 | 7.81 | 7.93 | 8.06 | 8.19 | 8.31 | 8.44 | 8.57 | 8.69 | 8.82 | 8.95 | 9.07 | 9.2 | 9.33 | 9.46 | 9.58 | 9.71 | 9.84 | 9.96
18.07 | 4.88 | 5.01 | 5.14 | 5.26 | 5.39 | 5.52 | 5.65 | 5.77 | 5.9 | 6.03 | 6.15 | 6.28 | 6.41 | 6.53 | 6.66 | 6.79 | 6.92 | 7.04 | 7.17 | 7.3 | 7.42 | 7.55 | 7.68 | 7.8 | 7.93 | 8.06 | 8.18 | 8.31 | 8.44 | 8.57 | 8.69 | 8.82 | 8.95 | 9.07 | 9.2 | 9.33 | 9.45 | 9.58 | 9.71 | 9.83 | 9.96
18.08 | 4.88 | 5.01 | 5.14 | 5.26 | 5.39 | 5.52 | 5.64 | 5.77 | 5.9 | 6.03 | 6.15 | 6.28 | 6.41 | 6.53 | 6.66 | 6.79 | 6.91 | 7.04 | 7.17 | 7.29 | 7.42 | 7.55 | 7.68 | 7.8 | 7.93 | 8.06 | 8.18 | 8.31 | 8.44 | 8.56 | 8.69 | 8.82 | 8.94 | 9.07 | 9.2 | 9.32 | 9.45 | 9.58 | 9.71 | 9.83 | 9.96
18.09 | 4.88 | 5.01 | 5.14 | 5.26 | 5.39 | 5.52 | 5.64 | 5.77 | 5.9 | 6.02 | 6.15 | 6.28 | 6.4 | 6.53 | 6.66 | 6.79 | 6.91 | 7.04 | 7.17 | 7.29 | 7.42 | 7.55 | 7.67 | 7.8 | 7.93 | 8.05 | 8.18 | 8.31 | 8.43 | 8.56 | 8.69 | 8.82 | 8.94 | 9.07 | 9.2 | 9.32 | 9.45 | 9.58 | 9.7 | 9.83 | 9.96
18.1 | 4.88 | 5.01 | 5.13 | 5.26 | 5.39 | 5.52 | 5.64 | 5.77 | 5.9 | 6.02 | 6.15 | 6.28 | 6.4 | 6.53 | 6.66 | 6.78 | 6.91 | 7.04 | 7.16 | 7.29 | 7.42 | 7.55 | 7.67 | 7.8 | 7.93 | 8.05 | 8.18 | 8.31 | 8.43 | 8.56 | 8.69 | 8.81 | 8.94 | 9.07 | 9.19 | 9.32 | 9.45 | 9.57 | 9.7 | 9.83 | 9.95
18.11 | 4.88 | 5.01 | 5.13 | 5.26 | 5.39 | 5.51 | 5.64 | 5.77 | 5.89 | 6.02 | 6.15 | 6.28 | 6.4 | 6.53 | 6.66 | 6.78 | 6.91 | 7.04 | 7.16 | 7.29 | 7.42 | 7.54 | 7.67 | 7.8 | 7.92 | 8.05 | 8.18 | 8.3 | 8.43 | 8.56 | 8.68 | 8.81 | 8.94 | 9.07 | 9.19 | 9.32 | 9.45 | 9.57 | 9.7 | 9.83 | 9.95
18.12 | 4.88 | 5.01 | 5.13 | 5.26 | 5.39 | 5.51 | 5.64 | 5.77 | 5.89 | 6.02 | 6.15 | 6.27 | 6.4 | 6.53 | 6.65 | 6.78 | 6.91 | 7.03 | 7.16 | 7.29 | 7.42 | 7.54 | 7.67 | 7.8 | 7.92 | 8.05 | 8.18 | 8.3 | 8.43 | 8.56 | 8.68 | 8.81 | 8.94 | 9.06 | 9.19 | 9.32 | 9.44 | 9.57 | 9.7 | 9.82 | 9.95
18.13 | 4.88 | 5 | 5.13 | 5.26 | 5.38 | 5.51 | 5.64 | 5.77 | 5.89 | 6.02 | 6.15 | 6.27 | 6.4 | 6.53 | 6.65 | 6.78 | 6.91 | 7.03 | 7.16 | 7.29 | 7.41 | 7.54 | 7.67 | 7.79 | 7.92 | 8.05 | 8.17 | 8.3 | 8.43 | 8.55 | 8.68 | 8.81 | 8.93 | 9.06 | 9.19 | 9.32 | 9.44 | 9.57 | 9.7 | 9.82 | 9.95
18.14 | 4.88 | 5 | 5.13 | 5.26 | 5.38 | 5.51 | 5.64 | 5.76 | 5.89 | 6.02 | 6.14 | 6.27 | 6.4 | 6.52 | 6.65 | 6.78 | 6.9 | 7.03 | 7.16 | 7.29 | 7.41 | 7.54 | 7.67 | 7.79 | 7.92 | 8.05 | 8.17 | 8.3 | 8.43 | 8.55 | 8.68 | 8.81 | 8.93 | 9.06 | 9.19 | 9.31 | 9.44 | 9.57 | 9.69 | 9.82 | 9.95
18.15 | 4.88 | 5 | 5.13 | 5.26 | 5.38 | 5.51 | 5.64 | 5.76 | 5.89 | 6.02 | 6.14 | 6.27 | 6.4 | 6.52 | 6.65 | 6.78 | 6.9 | 7.03 | 7.16 | 7.28 | 7.41 | 7.54 | 7.66 | 7.79 | 7.92 | 8.04 | 8.17 | 8.3 | 8.42 | 8.55 | 8.68 | 8.8 | 8.93 | 9.06 | 9.18 | 9.31 | 9.44 | 9.56 | 9.69 | 9.82 | 9.94
18.16 | 4.87 | 5 | 5.13 | 5.25 | 5.38 | 5.51 | 5.63 | 5.76 | 5.89 | 6.01 | 6.14 | 6.27 | 6.4 | 6.52 | 6.65 | 6.78 | 6.9 | 7.03 | 7.16 | 7.28 | 7.41 | 7.54 | 7.66 | 7.79 | 7.92 | 8.04 | 8.17 | 8.3 | 8.42 | 8.55 | 8.68 | 8.8 | 8.93 | 9.06 | 9.18 | 9.31 | 9.44 | 9.56 | 9.69 | 9.82 | 9.94
18.17 | 4.87 | 5 | 5.13 | 5.25 | 5.38 | 5.51 | 5.63 | 5.76 | 5.89 | 6.01 | 6.14 | 6.27 | 6.39 | 6.52 | 6.65 | 6.77 | 6.9 | 7.03 | 7.15 | 7.28 | 7.41 | 7.53 | 7.66 | 7.79 | 7.91 | 8.04 | 8.17 | 8.29 | 8.42 | 8.55 | 8.67 | 8.8 | 8.93 | 9.05 | 9.18 | 9.31 | 9.43 | 9.56 | 9.69 | 9.81 | 9.94
18.18 | 4.87 | 5 | 5.13 | 5.25 | 5.38 | 5.51 | 5.63 | 5.76 | 5.89 | 6.01 | 6.14 | 6.27 | 6.39 | 6.52 | 6.65 | 6.77 | 6.9 | 7.03 | 7.15 | 7.28 | 7.41 | 7.53 | 7.66 | 7.79 | 7.91 | 8.04 | 8.17 | 8.29 | 8.42 | 8.55 | 8.67 | 8.8 | 8.93 | 9.05 | 9.18 | 9.31 | 9.43 | 9.56 | 9.69 | 9.81 | 9.94
18.19 | 4.87 | 5 | 5.12 | 5.25 | 5.38 | 5.5 | 5.63 | 5.76 | 5.88 | 6.01 | 6.14 | 6.26 | 6.39 | 6.52 | 6.64 | 6.77 | 6.9 | 7.02 | 7.15 | 7.28 | 7.4 | 7.53 | 7.66 | 7.78 | 7.91 | 8.04 | 8.16 | 8.29 | 8.42 | 8.54 | 8.67 | 8.8 | 8.92 | 9.05 | 9.18 | 9.3 | 9.43 | 9.56 | 9.68 | 9.81 | 9.94
18.2 | 4.87 | 5 | 5.12 | 5.25 | 5.38 | 5.5 | 5.63 | 5.76 | 5.88 | 6.01 | 6.14 | 6.26 | 6.39 | 6.52 | 6.64 | 6.77 | 6.9 | 7.02 | 7.15 | 7.28 | 7.4 | 7.53 | 7.66 | 7.78 | 7.91 | 8.04 | 8.16 | 8.29 | 8.42 | 8.54 | 8.67 | 8.8 | 8.92 | 9.05 | 9.18 | 9.3 | 9.43 | 9.55 | 9.68 | 9.81 | 9.93
18.21 | 4.87 | 5 | 5.12 | 5.25 | 5.38 | 5.5 | 5.63 | 5.76 | 5.88 | 6.01 | 6.14 | 6.26 | 6.39 | 6.51 | 6.64 | 6.77 | 6.89 | 7.02 | 7.15 | 7.27 | 7.4 | 7.53 | 7.65 | 7.78 | 7.91 | 8.03 | 8.16 | 8.29 | 8.41 | 8.54 | 8.67 | 8.79 | 8.92 | 9.05 | 9.17 | 9.3 | 9.43 | 9.55 | 9.68 | 9.81 | 9.93
18.22 | 4.87 | 4.99 | 5.12 | 5.25 | 5.37 | 5.5 | 5.63 | 5.75 | 5.88 | 6.01 | 6.13 | 6.26 | 6.39 | 6.51 | 6.64 | 6.77 | 6.89 | 7.02 | 7.15 | 7.27 | 7.4 | 7.53 | 7.65 | 7.78 | 7.91 | 8.03 | 8.16 | 8.29 | 8.41 | 8.54 | 8.67 | 8.79 | 8.92 | 9.04 | 9.17 | 9.3 | 9.42 | 9.55 | 9.68 | 9.8 | 9.93
18.23 | 4.87 | 4.99 | 5.12 | 5.25 | 5.37 | 5.5 | 5.63 | 5.75 | 5.88 | 6.01 | 6.13 | 6.26 | 6.39 | 6.51 | 6.64 | 6.77 | 6.89 | 7.02 | 7.15 | 7.27 | 7.4 | 7.52 | 7.65 | 7.78 | 7.9 | 8.03 | 8.16 | 8.28 | 8.41 | 8.54 | 8.66 | 8.79 | 8.92 | 9.04 | 9.17 | 9.3 | 9.42 | 9.55 | 9.68 | 9.8 | 9.93
18.24 | 4.87 | 4.99 | 5.12 | 5.25 | 5.37 | 5.5 | 5.63 | 5.75 | 5.88 | 6 | 6.13 | 6.26 | 6.38 | 6.51 | 6.64 | 6.76 | 6.89 | 7.02 | 7.14 | 7.27 | 7.4 | 7.52 | 7.65 | 7.78 | 7.9 | 8.03 | 8.16 | 8.28 | 8.41 | 8.54 | 8.66 | 8.79 | 8.91 | 9.04 | 9.17 | 9.29 | 9.42 | 9.55 | 9.67 | 9.8 | 9.93
18.25 | 4.86 | 4.99 | 5.12 | 5.24 | 5.37 | 5.5 | 5.62 | 5.75 | 5.88 | 6 | 6.13 | 6.26 | 6.38 | 6.51 | 6.64 | 6.76 | 6.89 | 7.02 | 7.14 | 7.27 | 7.4 | 7.52 | 7.65 | 7.77 | 7.9 | 8.03 | 8.15 | 8.28 | 8.41 | 8.53 | 8.66 | 8.79 | 8.91 | 9.04 | 9.17 | 9.29 | 9.42 | 9.55 | 9.67 | 9.8 | 9.92
18.26 | 4.86 | 4.99 | 5.12 | 5.24 | 5.37 | 5.5 | 5.62 | 5.75 | 5.88 | 6 | 6.13 | 6.26 | 6.38 | 6.51 | 6.63 | 6.76 | 6.89 | 7.01 | 7.14 | 7.27 | 7.39 | 7.52 | 7.65 | 7.77 | 7.9 | 8.03 | 8.15 | 8.28 | 8.41 | 8.53 | 8.66 | 8.78 | 8.91 | 9.04 | 9.16 | 9.29 | 9.42 | 9.54 | 9.67 | 9.8 | 9.92
18.27 | 4.86 | 4.99 | 5.12 | 5.24 | 5.37 | 5.49 | 5.62 | 5.75 | 5.87 | 6 | 6.13 | 6.25 | 6.38 | 6.51 | 6.63 | 6.76 | 6.89 | 7.01 | 7.14 | 7.27 | 7.39 | 7.52 | 7.64 | 7.77 | 7.9 | 8.02 | 8.15 | 8.28 | 8.4 | 8.53 | 8.66 | 8.78 | 8.91 | 9.04 | 9.16 | 9.29 | 9.41 | 9.54 | 9.67 | 9.79 | 9.92
18.28 | 4.86 | 4.99 | 5.11 | 5.24 | 5.37 | 5.49 | 5.62 | 5.75 | 5.87 | 6 | 6.13 | 6.25 | 6.38 | 6.51 | 6.63 | 6.76 | 6.88 | 7.01 | 7.14 | 7.26 | 7.39 | 7.52 | 7.64 | 7.77 | 7.9 | 8.02 | 8.15 | 8.28 | 8.4 | 8.53 | 8.65 | 8.78 | 8.91 | 9.03 | 9.16 | 9.29 | 9.41 | 9.54 | 9.67 | 9.79 | 9.92
18.29 | 4.86 | 4.99 | 5.11 | 5.24 | 5.37 | 5.49 | 5.62 | 5.75 | 5.87 | 6 | 6.12 | 6.25 | 6.38 | 6.5 | 6.63 | 6.76 | 6.88 | 7.01 | 7.14 | 7.26 | 7.39 | 7.52 | 7.64 | 7.77 | 7.89 | 8.02 | 8.15 | 8.27 | 8.4 | 8.53 | 8.65 | 8.78 | 8.91 | 9.03 | 9.16 | 9.28 | 9.41 | 9.54 | 9.66 | 9.79 | 9.92
18.3 | 4.86 | 4.99 | 5.11 | 5.24 | 5.37 | 5.49 | 5.62 | 5.74 | 5.87 | 6 | 6.12 | 6.25 | 6.38 | 6.5 | 6.63 | 6.76 | 6.88 | 7.01 | 7.13 | 7.26 | 7.39 | 7.51 | 7.64 | 7.77 | 7.89 | 8.02 | 8.15 | 8.27 | 8.4 | 8.52 | 8.65 | 8.78 | 8.9 | 9.03 | 9.16 | 9.28 | 9.41 | 9.54 | 9.66 | 9.79 | 9.91
18.31 | 4.86 | 4.98 | 5.11 | 5.24 | 5.36 | 5.49 | 5.62 | 5.74 | 5.87 | 6 | 6.12 | 6.25 | 6.37 | 6.5 | 6.63 | 6.75 | 6.88 | 7.01 | 7.13 | 7.26 | 7.39 | 7.51 | 7.64 | 7.76 | 7.89 | 8.02 | 8.14 | 8.27 | 8.4 | 8.52 | 8.65 | 8.78 | 8.9 | 9.03 | 9.15 | 9.28 | 9.41 | 9.53 | 9.66 | 9.79 | 9.91
18.32 | 4.86 | 4.98 | 5.11 | 5.24 | 5.36 | 5.49 | 5.62 | 5.74 | 5.87 | 5.99 | 6.12 | 6.25 | 6.37 | 6.5 | 6.63 | 6.75 | 6.88 | 7.01 | 7.13 | 7.26 | 7.38 | 7.51 | 7.64 | 7.76 | 7.89 | 8.02 | 8.14 | 8.27 | 8.39 | 8.52 | 8.65 | 8.77 | 8.9 | 9.03 | 9.15 | 9.28 | 9.41 | 9.53 | 9.66 | 9.78 | 9.91
18.33 | 4.86 | 4.98 | 5.11 | 5.24 | 5.36 | 5.49 | 5.61 | 5.74 | 5.87 | 5.99 | 6.12 | 6.25 | 6.37 | 6.5 | 6.62 | 6.75 | 6.88 | 7 | 7.13 | 7.26 | 7.38 | 7.51 | 7.64 | 7.76 | 7.89 | 8.01 | 8.14 | 8.27 | 8.39 | 8.52 | 8.65 | 8.77 | 8.9 | 9.02 | 9.15 | 9.28 | 9.4 | 9.53 | 9.66 | 9.78 | 9.91
18.34 | 4.86 | 4.98 | 5.11 | 5.23 | 5.36 | 5.49 | 5.61 | 5.74 | 5.87 | 5.99 | 6.12 | 6.24 | 6.37 | 6.5 | 6.62 | 6.75 | 6.88 | 7 | 7.13 | 7.25 | 7.38 | 7.51 | 7.63 | 7.76 | 7.89 | 8.01 | 8.14 | 8.27 | 8.39 | 8.52 | 8.64 | 8.77 | 8.9 | 9.02 | 9.15 | 9.28 | 9.4 | 9.53 | 9.65 | 9.78 | 9.91
18.35 | 4.85 | 4.98 | 5.11 | 5.23 | 5.36 | 5.49 | 5.61 | 5.74 | 5.86 | 5.99 | 6.12 | 6.24 | 6.37 | 6.5 | 6.62 | 6.75 | 6.87 | 7 | 7.13 | 7.25 | 7.38 | 7.51 | 7.63 | 7.76 | 7.88 | 8.01 | 8.14 | 8.26 | 8.39 | 8.52 | 8.64 | 8.77 | 8.89 | 9.02 | 9.15 | 9.27 | 9.4 | 9.53 | 9.65 | 9.78 | 9.9
18.36 | 4.85 | 4.98 | 5.11 | 5.23 | 5.36 | 5.48 | 5.61 | 5.74 | 5.86 | 5.99 | 6.12 | 6.24 | 6.37 | 6.49 | 6.62 | 6.75 | 6.87 | 7 | 7.13 | 7.25 | 7.38 | 7.5 | 7.63 | 7.76 | 7.88 | 8.01 | 8.14 | 8.26 | 8.39 | 8.51 | 8.64 | 8.77 | 8.89 | 9.02 | 9.15 | 9.27 | 9.4 | 9.52 | 9.65 | 9.78 | 9.9
18.37 | 4.85 | 4.98 | 5.1 | 5.23 | 5.36 | 5.48 | 5.61 | 5.74 | 5.86 | 5.99 | 6.11 | 6.24 | 6.37 | 6.49 | 6.62 | 6.75 | 6.87 | 7 | 7.12 | 7.25 | 7.38 | 7.5 | 7.63 | 7.76 | 7.88 | 8.01 | 8.13 | 8.26 | 8.39 | 8.51 | 8.64 | 8.76 | 8.89 | 9.02 | 9.14 | 9.27 | 9.4 | 9.52 | 9.65 | 9.77 | 9.9
18.38 | 4.85 | 4.98 | 5.1 | 5.23 | 5.36 | 5.48 | 5.61 | 5.73 | 5.86 | 5.99 | 6.11 | 6.24 | 6.37 | 6.49 | 6.62 | 6.74 | 6.87 | 7 | 7.12 | 7.25 | 7.38 | 7.5 | 7.63 | 7.75 | 7.88 | 8.01 | 8.13 | 8.26 | 8.38 | 8.51 | 8.64 | 8.76 | 8.89 | 9.02 | 9.14 | 9.27 | 9.39 | 9.52 | 9.65 | 9.77 | 9.9
18.39 | 4.85 | 4.98 | 5.1 | 5.23 | 5.35 | 5.48 | 5.61 | 5.73 | 5.86 | 5.99 | 6.11 | 6.24 | 6.36 | 6.49 | 6.62 | 6.74 | 6.87 | 7 | 7.12 | 7.25 | 7.37 | 7.5 | 7.63 | 7.75 | 7.88 | 8 | 8.13 | 8.26 | 8.38 | 8.51 | 8.64 | 8.76 | 8.89 | 9.01 | 9.14 | 9.27 | 9.39 | 9.52 | 9.64 | 9.77 | 9.9
18.4 | 4.85 | 4.98 | 5.1 | 5.23 | 5.35 | 5.48 | 5.61 | 5.73 | 5.86 | 5.98 | 6.11 | 6.24 | 6.36 | 6.49 | 6.62 | 6.74 | 6.87 | 6.99 | 7.12 | 7.25 | 7.37 | 7.5 | 7.62 | 7.75 | 7.88 | 8 | 8.13 | 8.25 | 8.38 | 8.51 | 8.63 | 8.76 | 8.89 | 9.01 | 9.14 | 9.26 | 9.39 | 9.52 | 9.64 | 9.77 | 9.89
18.41 | 4.85 | 4.97 | 5.1 | 5.23 | 5.35 | 5.48 | 5.6 | 5.73 | 5.86 | 5.98 | 6.11 | 6.24 | 6.36 | 6.49 | 6.61 | 6.74 | 6.87 | 6.99 | 7.12 | 7.24 | 7.37 | 7.5 | 7.62 | 7.75 | 7.87 | 8 | 8.13 | 8.25 | 8.38 | 8.51 | 8.63 | 8.76 | 8.88 | 9.01 | 9.14 | 9.26 | 9.39 | 9.51 | 9.64 | 9.77 | 9.89
18.42 | 4.85 | 4.97 | 5.1 | 5.23 | 5.35 | 5.48 | 5.6 | 5.73 | 5.86 | 5.98 | 6.11 | 6.23 | 6.36 | 6.49 | 6.61 | 6.74 | 6.86 | 6.99 | 7.12 | 7.24 | 7.37 | 7.5 | 7.62 | 7.75 | 7.87 | 8 | 8.13 | 8.25 | 8.38 | 8.5 | 8.63 | 8.76 | 8.88 | 9.01 | 9.13 | 9.26 | 9.39 | 9.51 | 9.64 | 9.76 | 9.89
18.43 | 4.85 | 4.97 | 5.1 | 5.22 | 5.35 | 5.48 | 5.6 | 5.73 | 5.85 | 5.98 | 6.11 | 6.23 | 6.36 | 6.48 | 6.61 | 6.74 | 6.86 | 6.99 | 7.12 | 7.24 | 7.37 | 7.49 | 7.62 | 7.75 | 7.87 | 8 | 8.12 | 8.25 | 8.38 | 8.5 | 8.63 | 8.75 | 8.88 | 9.01 | 9.13 | 9.26 | 9.38 | 9.51 | 9.64 | 9.76 | 9.89
18.44 | 4.84 | 4.97 | 5.1 | 5.22 | 5.35 | 5.48 | 5.6 | 5.73 | 5.85 | 5.98 | 6.11 | 6.23 | 6.36 | 6.48 | 6.61 | 6.74 | 6.86 | 6.99 | 7.11 | 7.24 | 7.37 | 7.49 | 7.62 | 7.74 | 7.87 | 8 | 8.12 | 8.25 | 8.37 | 8.5 | 8.63 | 8.75 | 8.88 | 9 | 9.13 | 9.26 | 9.38 | 9.51 | 9.63 | 9.76 | 9.89
18.45 | 4.84 | 4.97 | 5.1 | 5.22 | 5.35 | 5.47 | 5.6 | 5.73 | 5.85 | 5.98 | 6.1 | 6.23 | 6.36 | 6.48 | 6.61 | 6.73 | 6.86 | 6.99 | 7.11 | 7.24 | 7.36 | 7.49 | 7.62 | 7.74 | 7.87 | 7.99 | 8.12 | 8.25 | 8.37 | 8.5 | 8.62 | 8.75 | 8.88 | 9 | 9.13 | 9.25 | 9.38 | 9.51 | 9.63 | 9.76 | 9.88
18.46 | 4.84 | 4.97 | 5.09 | 5.22 | 5.35 | 5.47 | 5.6 | 5.72 | 5.85 | 5.98 | 6.1 | 6.23 | 6.35 | 6.48 | 6.61 | 6.73 | 6.86 | 6.98 | 7.11 | 7.24 | 7.36 | 7.49 | 7.61 | 7.74 | 7.87 | 7.99 | 8.12 | 8.24 | 8.37 | 8.5 | 8.62 | 8.75 | 8.87 | 9 | 9.13 | 9.25 | 9.38 | 9.5 | 9.63 | 9.76 | 9.88
18.47 | 4.84 | 4.97 | 5.09 | 5.22 | 5.35 | 5.47 | 5.6 | 5.72 | 5.85 | 5.98 | 6.1 | 6.23 | 6.35 | 6.48 | 6.61 | 6.73 | 6.86 | 6.98 | 7.11 | 7.24 | 7.36 | 7.49 | 7.61 | 7.74 | 7.87 | 7.99 | 8.12 | 8.24 | 8.37 | 8.49 | 8.62 | 8.75 | 8.87 | 9 | 9.12 | 9.25 | 9.38 | 9.5 | 9.63 | 9.75 | 9.88
18.48 | 4.84 | 4.97 | 5.09 | 5.22 | 5.34 | 5.47 | 5.6 | 5.72 | 5.85 | 5.97 | 6.1 | 6.23 | 6.35 | 6.48 | 6.6 | 6.73 | 6.86 | 6.98 | 7.11 | 7.23 | 7.36 | 7.49 | 7.61 | 7.74 | 7.86 | 7.99 | 8.12 | 8.24 | 8.37 | 8.49 | 8.62 | 8.75 | 8.87 | 9 | 9.12 | 9.25 | 9.37 | 9.5 | 9.63 | 9.75 | 9.88
18.49 | 4.84 | 4.97 | 5.09 | 5.22 | 5.34 | 5.47 | 5.6 | 5.72 | 5.85 | 5.97 | 6.1 | 6.22 | 6.35 | 6.48 | 6.6 | 6.73 | 6.85 | 6.98 | 7.11 | 7.23 | 7.36 | 7.48 | 7.61 | 7.74 | 7.86 | 7.99 | 8.11 | 8.24 | 8.37 | 8.49 | 8.62 | 8.74 | 8.87 | 9 | 9.12 | 9.25 | 9.37 | 9.5 | 9.62 | 9.75 | 9.88
18.5 | 4.84 | 4.96 | 5.09 | 5.22 | 5.34 | 5.47 | 5.59 | 5.72 | 5.85 | 5.97 | 6.1 | 6.22 | 6.35 | 6.48 | 6.6 | 6.73 | 6.85 | 6.98 | 7.1 | 7.23 | 7.36 | 7.48 | 7.61 | 7.73 | 7.86 | 7.99 | 8.11 | 8.24 | 8.36 | 8.49 | 8.62 | 8.74 | 8.87 | 8.99 | 9.12 | 9.24 | 9.37 | 9.5 | 9.62 | 9.75 | 9.87
18.51 | 4.84 | 4.96 | 5.09 | 5.22 | 5.34 | 5.47 | 5.59 | 5.72 | 5.84 | 5.97 | 6.1 | 6.22 | 6.35 | 6.47 | 6.6 | 6.73 | 6.85 | 6.98 | 7.1 | 7.23 | 7.36 | 7.48 | 7.61 | 7.73 | 7.86 | 7.98 | 8.11 | 8.24 | 8.36 | 8.49 | 8.61 | 8.74 | 8.87 | 8.99 | 9.12 | 9.24 | 9.37 | 9.49 | 9.62 | 9.75 | 9.87
18.52 | 4.84 | 4.96 | 5.09 | 5.21 | 5.34 | 5.47 | 5.59 | 5.72 | 5.84 | 5.97 | 6.1 | 6.22 | 6.35 | 6.47 | 6.6 | 6.72 | 6.85 | 6.98 | 7.1 | 7.23 | 7.35 | 7.48 | 7.61 | 7.73 | 7.86 | 7.98 | 8.11 | 8.23 | 8.36 | 8.49 | 8.61 | 8.74 | 8.86 | 8.99 | 9.12 | 9.24 | 9.37 | 9.49 | 9.62 | 9.74 | 9.87
18.53 | 4.84 | 4.96 | 5.09 | 5.21 | 5.34 | 5.46 | 5.59 | 5.72 | 5.84 | 5.97 | 6.09 | 6.22 | 6.35 | 6.47 | 6.6 | 6.72 | 6.85 | 6.97 | 7.1 | 7.23 | 7.35 | 7.48 | 7.6 | 7.73 | 7.86 | 7.98 | 8.11 | 8.23 | 8.36 | 8.48 | 8.61 | 8.74 | 8.86 | 8.99 | 9.11 | 9.24 | 9.37 | 9.49 | 9.62 | 9.74 | 9.87
18.54 | 4.83 | 4.96 | 5.09 | 5.21 | 5.34 | 5.46 | 5.59 | 5.72 | 5.84 | 5.97 | 6.09 | 6.22 | 6.34 | 6.47 | 6.6 | 6.72 | 6.85 | 6.97 | 7.1 | 7.22 | 7.35 | 7.48 | 7.6 | 7.73 | 7.85 | 7.98 | 8.11 | 8.23 | 8.36 | 8.48 | 8.61 | 8.73 | 8.86 | 8.99 | 9.11 | 9.24 | 9.36 | 9.49 | 9.61 | 9.74 | 9.87
18.55 | 4.83 | 4.96 | 5.08 | 5.21 | 5.34 | 5.46 | 5.59 | 5.71 | 5.84 | 5.97 | 6.09 | 6.22 | 6.34 | 6.47 | 6.59 | 6.72 | 6.85 | 6.97 | 7.1 | 7.22 | 7.35 | 7.47 | 7.6 | 7.73 | 7.85 | 7.98 | 8.1 | 8.23 | 8.36 | 8.48 | 8.61 | 8.73 | 8.86 | 8.98 | 9.11 | 9.24 | 9.36 | 9.49 | 9.61 | 9.74 | 9.86
18.56 | 4.83 | 4.96 | 5.08 | 5.21 | 5.34 | 5.46 | 5.59 | 5.71 | 5.84 | 5.96 | 6.09 | 6.22 | 6.34 | 6.47 | 6.59 | 6.72 | 6.84 | 6.97 | 7.1 | 7.22 | 7.35 | 7.47 | 7.6 | 7.72 | 7.85 | 7.98 | 8.1 | 8.23 | 8.35 | 8.48 | 8.61 | 8.73 | 8.86 | 8.98 | 9.11 | 9.23 | 9.36 | 9.49 | 9.61 | 9.74 | 9.86
18.57 | 4.83 | 4.96 | 5.08 | 5.21 | 5.33 | 5.46 | 5.59 | 5.71 | 5.84 | 5.96 | 6.09 | 6.21 | 6.34 | 6.47 | 6.59 | 6.72 | 6.84 | 6.97 | 7.09 | 7.22 | 7.35 | 7.47 | 7.6 | 7.72 | 7.85 | 7.97 | 8.1 | 8.23 | 8.35 | 8.48 | 8.6 | 8.73 | 8.85 | 8.98 | 9.11 | 9.23 | 9.36 | 9.48 | 9.61 | 9.73 | 9.86
18.58 | 4.83 | 4.96 | 5.08 | 5.21 | 5.33 | 5.46 | 5.58 | 5.71 | 5.84 | 5.96 | 6.09 | 6.21 | 6.34 | 6.46 | 6.59 | 6.72 | 6.84 | 6.97 | 7.09 | 7.22 | 7.34 | 7.47 | 7.6 | 7.72 | 7.85 | 7.97 | 8.1 | 8.22 | 8.35 | 8.48 | 8.6 | 8.73 | 8.85 | 8.98 | 9.1 | 9.23 | 9.36 | 9.48 | 9.61 | 9.73 | 9.86
18.59 | 4.83 | 4.95 | 5.08 | 5.21 | 5.33 | 5.46 | 5.58 | 5.71 | 5.83 | 5.96 | 6.09 | 6.21 | 6.34 | 6.46 | 6.59 | 6.71 | 6.84 | 6.97 | 7.09 | 7.22 | 7.34 | 7.47 | 7.59 | 7.72 | 7.85 | 7.97 | 8.1 | 8.22 | 8.35 | 8.47 | 8.6 | 8.73 | 8.85 | 8.98 | 9.1 | 9.23 | 9.35 | 9.48 | 9.6 | 9.73 | 9.86
18.6 | 4.83 | 4.95 | 5.08 | 5.21 | 5.33 | 5.46 | 5.58 | 5.71 | 5.83 | 5.96 | 6.08 | 6.21 | 6.34 | 6.46 | 6.59 | 6.71 | 6.84 | 6.96 | 7.09 | 7.22 | 7.34 | 7.47 | 7.59 | 7.72 | 7.84 | 7.97 | 8.1 | 8.22 | 8.35 | 8.47 | 8.6 | 8.72 | 8.85 | 8.97 | 9.1 | 9.23 | 9.35 | 9.48 | 9.6 | 9.73 | 9.85
18.61 | 4.83 | 4.95 | 5.08 | 5.2 | 5.33 | 5.46 | 5.58 | 5.71 | 5.83 | 5.96 | 6.08 | 6.21 | 6.33 | 6.46 | 6.59 | 6.71 | 6.84 | 6.96 | 7.09 | 7.21 | 7.34 | 7.47 | 7.59 | 7.72 | 7.84 | 7.97 | 8.09 | 8.22 | 8.35 | 8.47 | 8.6 | 8.72 | 8.85 | 8.97 | 9.1 | 9.22 | 9.35 | 9.48 | 9.6 | 9.73 | 9.85
18.62 | 4.83 | 4.95 | 5.08 | 5.2 | 5.33 | 5.45 | 5.58 | 5.71 | 5.83 | 5.96 | 6.08 | 6.21 | 6.33 | 6.46 | 6.58 | 6.71 | 6.84 | 6.96 | 7.09 | 7.21 | 7.34 | 7.46 | 7.59 | 7.72 | 7.84 | 7.97 | 8.09 | 8.22 | 8.34 | 8.47 | 8.59 | 8.72 | 8.85 | 8.97 | 9.1 | 9.22 | 9.35 | 9.47 | 9.6 | 9.72 | 9.85
18.63 | 4.82 | 4.95 | 5.08 | 5.2 | 5.33 | 5.45 | 5.58 | 5.7 | 5.83 | 5.96 | 6.08 | 6.21 | 6.33 | 6.46 | 6.58 | 6.71 | 6.83 | 6.96 | 7.09 | 7.21 | 7.34 | 7.46 | 7.59 | 7.71 | 7.84 | 7.96 | 8.09 | 8.22 | 8.34 | 8.47 | 8.59 | 8.72 | 8.84 | 8.97 | 9.09 | 9.22 | 9.35 | 9.47 | 9.6 | 9.72 | 9.85
18.64 | 4.82 | 4.95 | 5.07 | 5.2 | 5.33 | 5.45 | 5.58 | 5.7 | 5.83 | 5.95 | 6.08 | 6.21 | 6.33 | 6.46 | 6.58 | 6.71 | 6.83 | 6.96 | 7.08 | 7.21 | 7.34 | 7.46 | 7.59 | 7.71 | 7.84 | 7.96 | 8.09 | 8.21 | 8.34 | 8.47 | 8.59 | 8.72 | 8.84 | 8.97 | 9.09 | 9.22 | 9.34 | 9.47 | 9.6 | 9.72 | 9.85
18.65 | 4.82 | 4.95 | 5.07 | 5.2 | 5.33 | 5.45 | 5.58 | 5.7 | 5.83 | 5.95 | 6.08 | 6.2 | 6.33 | 6.46 | 6.58 | 6.71 | 6.83 | 6.96 | 7.08 | 7.21 | 7.33 | 7.46 | 7.59 | 7.71 | 7.84 | 7.96 | 8.09 | 8.21 | 8.34 | 8.46 | 8.59 | 8.71 | 8.84 | 8.97 | 9.09 | 9.22 | 9.34 | 9.47 | 9.59 | 9.72 | 9.84
18.66 | 4.82 | 4.95 | 5.07 | 5.2 | 5.32 | 5.45 | 5.58 | 5.7 | 5.83 | 5.95 | 6.08 | 6.2 | 6.33 | 6.45 | 6.58 | 6.7 | 6.83 | 6.96 | 7.08 | 7.21 | 7.33 | 7.46 | 7.58 | 7.71 | 7.83 | 7.96 | 8.09 | 8.21 | 8.34 | 8.46 | 8.59 | 8.71 | 8.84 | 8.96 | 9.09 | 9.21 | 9.34 | 9.47 | 9.59 | 9.72 | 9.84
18.67 | 4.82 | 4.95 | 5.07 | 5.2 | 5.32 | 5.45 | 5.57 | 5.7 | 5.82 | 5.95 | 6.08 | 6.2 | 6.33 | 6.45 | 6.58 | 6.7 | 6.83 | 6.95 | 7.08 | 7.21 | 7.33 | 7.46 | 7.58 | 7.71 | 7.83 | 7.96 | 8.08 | 8.21 | 8.33 | 8.46 | 8.59 | 8.71 | 8.84 | 8.96 | 9.09 | 9.21 | 9.34 | 9.46 | 9.59 | 9.71 | 9.84
18.68 | 4.82 | 4.95 | 5.07 | 5.2 | 5.32 | 5.45 | 5.57 | 5.7 | 5.82 | 5.95 | 6.07 | 6.2 | 6.33 | 6.45 | 6.58 | 6.7 | 6.83 | 6.95 | 7.08 | 7.2 | 7.33 | 7.45 | 7.58 | 7.71 | 7.83 | 7.96 | 8.08 | 8.21 | 8.33 | 8.46 | 8.58 | 8.71 | 8.83 | 8.96 | 9.09 | 9.21 | 9.34 | 9.46 | 9.59 | 9.71 | 9.84
18.69 | 4.82 | 4.94 | 5.07 | 5.2 | 5.32 | 5.45 | 5.57 | 5.7 | 5.82 | 5.95 | 6.07 | 6.2 | 6.32 | 6.45 | 6.58 | 6.7 | 6.83 | 6.95 | 7.08 | 7.2 | 7.33 | 7.45 | 7.58 | 7.7 | 7.83 | 7.96 | 8.08 | 8.21 | 8.33 | 8.46 | 8.58 | 8.71 | 8.83 | 8.96 | 9.08 | 9.21 | 9.33 | 9.46 | 9.59 | 9.71 | 9.84
18.7 | 4.82 | 4.94 | 5.07 | 5.19 | 5.32 | 5.44 | 5.57 | 5.7 | 5.82 | 5.95 | 6.07 | 6.2 | 6.32 | 6.45 | 6.57 | 6.7 | 6.82 | 6.95 | 7.08 | 7.2 | 7.33 | 7.45 | 7.58 | 7.7 | 7.83 | 7.95 | 8.08 | 8.2 | 8.33 | 8.46 | 8.58 | 8.71 | 8.83 | 8.96 | 9.08 | 9.21 | 9.33 | 9.46 | 9.58 | 9.71 | 9.83
18.71 | 4.82 | 4.94 | 5.07 | 5.19 | 5.32 | 5.44 | 5.57 | 5.69 | 5.82 | 5.95 | 6.07 | 6.2 | 6.32 | 6.45 | 6.57 | 6.7 | 6.82 | 6.95 | 7.07 | 7.2 | 7.32 | 7.45 | 7.58 | 7.7 | 7.83 | 7.95 | 8.08 | 8.2 | 8.33 | 8.45 | 8.58 | 8.7 | 8.83 | 8.95 | 9.08 | 9.21 | 9.33 | 9.46 | 9.58 | 9.71 | 9.83
18.72 | 4.82 | 4.94 | 5.07 | 5.19 | 5.32 | 5.44 | 5.57 | 5.69 | 5.82 | 5.94 | 6.07 | 6.19 | 6.32 | 6.45 | 6.57 | 6.7 | 6.82 | 6.95 | 7.07 | 7.2 | 7.32 | 7.45 | 7.57 | 7.7 | 7.82 | 7.95 | 8.08 | 8.2 | 8.33 | 8.45 | 8.58 | 8.7 | 8.83 | 8.95 | 9.08 | 9.2 | 9.33 | 9.45 | 9.58 | 9.7 | 9.83
18.73 | 4.81 | 4.94 | 5.07 | 5.19 | 5.32 | 5.44 | 5.57 | 5.69 | 5.82 | 5.94 | 6.07 | 6.19 | 6.32 | 6.44 | 6.57 | 6.7 | 6.82 | 6.95 | 7.07 | 7.2 | 7.32 | 7.45 | 7.57 | 7.7 | 7.82 | 7.95 | 8.07 | 8.2 | 8.32 | 8.45 | 8.58 | 8.7 | 8.83 | 8.95 | 9.08 | 9.2 | 9.33 | 9.45 | 9.58 | 9.7 | 9.83
18.74 | 4.81 | 4.94 | 5.06 | 5.19 | 5.31 | 5.44 | 5.57 | 5.69 | 5.82 | 5.94 | 6.07 | 6.19 | 6.32 | 6.44 | 6.57 | 6.69 | 6.82 | 6.94 | 7.07 | 7.2 | 7.32 | 7.45 | 7.57 | 7.7 | 7.82 | 7.95 | 8.07 | 8.2 | 8.32 | 8.45 | 8.57 | 8.7 | 8.82 | 8.95 | 9.07 | 9.2 | 9.33 | 9.45 | 9.58 | 9.7 | 9.83
18.75 | 4.81 | 4.94 | 5.06 | 5.19 | 5.31 | 5.44 | 5.56 | 5.69 | 5.82 | 5.94 | 6.07 | 6.19 | 6.32 | 6.44 | 6.57 | 6.69 | 6.82 | 6.94 | 7.07 | 7.19 | 7.32 | 7.44 | 7.57 | 7.69 | 7.82 | 7.95 | 8.07 | 8.2 | 8.32 | 8.45 | 8.57 | 8.7 | 8.82 | 8.95 | 9.07 | 9.2 | 9.32 | 9.45 | 9.57 | 9.7 | 9.82
18.76 | 4.81 | 4.94 | 5.06 | 5.19 | 5.31 | 5.44 | 5.56 | 5.69 | 5.81 | 5.94 | 6.06 | 6.19 | 6.32 | 6.44 | 6.57 | 6.69 | 6.82 | 6.94 | 7.07 | 7.19 | 7.32 | 7.44 | 7.57 | 7.69 | 7.82 | 7.94 | 8.07 | 8.19 | 8.32 | 8.44 | 8.57 | 8.7 | 8.82 | 8.95 | 9.07 | 9.2 | 9.32 | 9.45 | 9.57 | 9.7 | 9.82
18.77 | 4.81 | 4.94 | 5.06 | 5.19 | 5.31 | 5.44 | 5.56 | 5.69 | 5.81 | 5.94 | 6.06 | 6.19 | 6.31 | 6.44 | 6.56 | 6.69 | 6.81 | 6.94 | 7.07 | 7.19 | 7.32 | 7.44 | 7.57 | 7.69 | 7.82 | 7.94 | 8.07 | 8.19 | 8.32 | 8.44 | 8.57 | 8.69 | 8.82 | 8.94 | 9.07 | 9.19 | 9.32 | 9.44 | 9.57 | 9.7 | 9.82
18.78 | 4.81 | 4.93 | 5.06 | 5.19 | 5.31 | 5.44 | 5.56 | 5.69 | 5.81 | 5.94 | 6.06 | 6.19 | 6.31 | 6.44 | 6.56 | 6.69 | 6.81 | 6.94 | 7.06 | 7.19 | 7.31 | 7.44 | 7.56 | 7.69 | 7.82 | 7.94 | 8.07 | 8.19 | 8.32 | 8.44 | 8.57 | 8.69 | 8.82 | 8.94 | 9.07 | 9.19 | 9.32 | 9.44 | 9.57 | 9.69 | 9.82
18.79 | 4.81 | 4.93 | 5.06 | 5.18 | 5.31 | 5.43 | 5.56 | 5.68 | 5.81 | 5.94 | 6.06 | 6.19 | 6.31 | 6.44 | 6.56 | 6.69 | 6.81 | 6.94 | 7.06 | 7.19 | 7.31 | 7.44 | 7.56 | 7.69 | 7.81 | 7.94 | 8.06 | 8.19 | 8.31 | 8.44 | 8.56 | 8.69 | 8.82 | 8.94 | 9.07 | 9.19 | 9.32 | 9.44 | 9.57 | 9.69 | 9.82
18.8 | 4.81 | 4.93 | 5.06 | 5.18 | 5.31 | 5.43 | 5.56 | 5.68 | 5.81 | 5.93 | 6.06 | 6.18 | 6.31 | 6.43 | 6.56 | 6.69 | 6.81 | 6.94 | 7.06 | 7.19 | 7.31 | 7.44 | 7.56 | 7.69 | 7.81 | 7.94 | 8.06 | 8.19 | 8.31 | 8.44 | 8.56 | 8.69 | 8.81 | 8.94 | 9.06 | 9.19 | 9.31 | 9.44 | 9.56 | 9.69 | 9.81
18.81 | 4.81 | 4.93 | 5.06 | 5.18 | 5.31 | 5.43 | 5.56 | 5.68 | 5.81 | 5.93 | 6.06 | 6.18 | 6.31 | 6.43 | 6.56 | 6.68 | 6.81 | 6.93 | 7.06 | 7.18 | 7.31 | 7.43 | 7.56 | 7.69 | 7.81 | 7.94 | 8.06 | 8.19 | 8.31 | 8.44 | 8.56 | 8.69 | 8.81 | 8.94 | 9.06 | 9.19 | 9.31 | 9.44 | 9.56 | 9.69 | 9.81
18.82 | 4.81 | 4.93 | 5.06 | 5.18 | 5.31 | 5.43 | 5.56 | 5.68 | 5.81 | 5.93 | 6.06 | 6.18 | 6.31 | 6.43 | 6.56 | 6.68 | 6.81 | 6.93 | 7.06 | 7.18 | 7.31 | 7.43 | 7.56 | 7.68 | 7.81 | 7.93 | 8.06 | 8.18 | 8.31 | 8.43 | 8.56 | 8.68 | 8.81 | 8.93 | 9.06 | 9.19 | 9.31 | 9.44 | 9.56 | 9.69 | 9.81
18.83 | 4.8 | 4.93 | 5.05 | 5.18 | 5.3 | 5.43 | 5.55 | 5.68 | 5.81 | 5.93 | 6.06 | 6.18 | 6.31 | 6.43 | 6.56 | 6.68 | 6.81 | 6.93 | 7.06 | 7.18 | 7.31 | 7.43 | 7.56 | 7.68 | 7.81 | 7.93 | 8.06 | 8.18 | 8.31 | 8.43 | 8.56 | 8.68 | 8.81 | 8.93 | 9.06 | 9.18 | 9.31 | 9.43 | 9.56 | 9.68 | 9.81
18.84 | 4.8 | 4.93 | 5.05 | 5.18 | 5.3 | 5.43 | 5.55 | 5.68 | 5.8 | 5.93 | 6.05 | 6.18 | 6.3 | 6.43 | 6.55 | 6.68 | 6.8 | 6.93 | 7.06 | 7.18 | 7.31 | 7.43 | 7.56 | 7.68 | 7.81 | 7.93 | 8.06 | 8.18 | 8.31 | 8.43 | 8.56 | 8.68 | 8.81 | 8.93 | 9.06 | 9.18 | 9.31 | 9.43 | 9.56 | 9.68 | 9.81
18.85 | 4.8 | 4.93 | 5.05 | 5.18 | 5.3 | 5.43 | 5.55 | 5.68 | 5.8 | 5.93 | 6.05 | 6.18 | 6.3 | 6.43 | 6.55 | 6.68 | 6.8 | 6.93 | 7.05 | 7.18 | 7.3 | 7.43 | 7.55 | 7.68 | 7.8 | 7.93 | 8.05 | 8.18 | 8.3 | 8.43 | 8.55 | 8.68 | 8.8 | 8.93 | 9.05 | 9.18 | 9.3 | 9.43 | 9.55 | 9.68 | 9.8
18.86 | 4.8 | 4.93 | 5.05 | 5.18 | 5.3 | 5.43 | 5.55 | 5.68 | 5.8 | 5.93 | 6.05 | 6.18 | 6.3 | 6.43 | 6.55 | 6.68 | 6.8 | 6.93 | 7.05 | 7.18 | 7.3 | 7.43 | 7.55 | 7.68 | 7.8 | 7.93 | 8.05 | 8.18 | 8.3 | 8.43 | 8.55 | 8.68 | 8.8 | 8.93 | 9.05 | 9.18 | 9.3 | 9.43 | 9.55 | 9.68 | 9.8
18.87 | 4.8 | 4.92 | 5.05 | 5.18 | 5.3 | 5.43 | 5.55 | 5.68 | 5.8 | 5.93 | 6.05 | 6.18 | 6.3 | 6.43 | 6.55 | 6.68 | 6.8 | 6.93 | 7.05 | 7.18 | 7.3 | 7.43 | 7.55 | 7.68 | 7.8 | 7.93 | 8.05 | 8.18 | 8.3 | 8.43 | 8.55 | 8.68 | 8.8 | 8.93 | 9.05 | 9.18 | 9.3 | 9.43 | 9.55 | 9.68 | 9.8
18.88 | 4.8 | 4.92 | 5.05 | 5.17 | 5.3 | 5.42 | 5.55 | 5.67 | 5.8 | 5.92 | 6.05 | 6.17 | 6.3 | 6.42 | 6.55 | 6.67 | 6.8 | 6.92 | 7.05 | 7.17 | 7.3 | 7.42 | 7.55 | 7.67 | 7.8 | 7.92 | 8.05 | 8.17 | 8.3 | 8.42 | 8.55 | 8.67 | 8.8 | 8.92 | 9.05 | 9.17 | 9.3 | 9.42 | 9.55 | 9.67 | 9.8
18.89 | 4.8 | 4.92 | 5.05 | 5.17 | 5.3 | 5.42 | 5.55 | 5.67 | 5.8 | 5.92 | 6.05 | 6.17 | 6.3 | 6.42 | 6.55 | 6.67 | 6.8 | 6.92 | 7.05 | 7.17 | 7.3 | 7.42 | 7.55 | 7.67 | 7.8 | 7.92 | 8.05 | 8.17 | 8.3 | 8.42 | 8.55 | 8.67 | 8.8 | 8.92 | 9.05 | 9.17 | 9.3 | 9.42 | 9.55 | 9.67 | 9.8
18.9 | 4.8 | 4.92 | 5.05 | 5.17 | 5.3 | 5.42 | 5.55 | 5.67 | 5.8 | 5.92 | 6.05 | 6.17 | 6.3 | 6.42 | 6.55 | 6.67 | 6.8 | 6.92 | 7.05 | 7.17 | 7.3 | 7.42 | 7.55 | 7.67 | 7.8 | 7.92 | 8.05 | 8.17 | 8.3 | 8.42 | 8.55 | 8.67 | 8.8 | 8.92 | 9.05 | 9.17 | 9.3 | 9.42 | 9.54 | 9.67 | 9.79
18.91 | 4.8 | 4.92 | 5.05 | 5.17 | 5.3 | 5.42 | 5.55 | 5.67 | 5.8 | 5.92 | 6.05 | 6.17 | 6.3 | 6.42 | 6.55 | 6.67 | 6.8 | 6.92 | 7.04 | 7.17 | 7.29 | 7.42 | 7.54 | 7.67 | 7.79 | 7.92 | 8.04 | 8.17 | 8.29 | 8.42 | 8.54 | 8.67 | 8.79 | 8.92 | 9.04 | 9.17 | 9.29 | 9.42 | 9.54 | 9.67 | 9.79
18.92 | 4.79 | 4.92 | 5.04 | 5.17 | 5.29 | 5.42 | 5.54 | 5.67 | 5.79 | 5.92 | 6.04 | 6.17 | 6.29 | 6.42 | 6.54 | 6.67 | 6.79 | 6.92 | 7.04 | 7.17 | 7.29 | 7.42 | 7.54 | 7.67 | 7.79 | 7.92 | 8.04 | 8.17 | 8.29 | 8.42 | 8.54 | 8.67 | 8.79 | 8.92 | 9.04 | 9.17 | 9.29 | 9.42 | 9.54 | 9.67 | 9.79
18.93 | 4.79 | 4.92 | 5.04 | 5.17 | 5.29 | 5.42 | 5.54 | 5.67 | 5.79 | 5.92 | 6.04 | 6.17 | 6.29 | 6.42 | 6.54 | 6.67 | 6.79 | 6.92 | 7.04 | 7.17 | 7.29 | 7.42 | 7.54 | 7.67 | 7.79 | 7.92 | 8.04 | 8.17 | 8.29 | 8.42 | 8.54 | 8.67 | 8.79 | 8.91 | 9.04 | 9.16 | 9.29 | 9.41 | 9.54 | 9.66 | 9.79
18.94 | 4.79 | 4.92 | 5.04 | 5.17 | 5.29 | 5.42 | 5.54 | 5.67 | 5.79 | 5.92 | 6.04 | 6.17 | 6.29 | 6.42 | 6.54 | 6.67 | 6.79 | 6.92 | 7.04 | 7.17 | 7.29 | 7.42 | 7.54 | 7.66 | 7.79 | 7.91 | 8.04 | 8.16 | 8.29 | 8.41 | 8.54 | 8.66 | 8.79 | 8.91 | 9.04 | 9.16 | 9.29 | 9.41 | 9.54 | 9.66 | 9.79
18.95 | 4.79 | 4.92 | 5.04 | 5.17 | 5.29 | 5.42 | 5.54 | 5.67 | 5.79 | 5.92 | 6.04 | 6.17 | 6.29 | 6.41 | 6.54 | 6.66 | 6.79 | 6.91 | 7.04 | 7.16 | 7.29 | 7.41 | 7.54 | 7.66 | 7.79 | 7.91 | 8.04 | 8.16 | 8.29 | 8.41 | 8.54 | 8.66 | 8.79 | 8.91 | 9.04 | 9.16 | 9.29 | 9.41 | 9.54 | 9.66 | 9.78
18.96 | 4.79 | 4.92 | 5.04 | 5.17 | 5.29 | 5.41 | 5.54 | 5.66 | 5.79 | 5.91 | 6.04 | 6.16 | 6.29 | 6.41 | 6.54 | 6.66 | 6.79 | 6.91 | 7.04 | 7.16 | 7.29 | 7.41 | 7.54 | 7.66 | 7.79 | 7.91 | 8.04 | 8.16 | 8.29 | 8.41 | 8.54 | 8.66 | 8.78 | 8.91 | 9.03 | 9.16 | 9.28 | 9.41 | 9.53 | 9.66 | 9.78
18.97 | 4.79 | 4.91 | 5.04 | 5.16 | 5.29 | 5.41 | 5.54 | 5.66 | 5.79 | 5.91 | 6.04 | 6.16 | 6.29 | 6.41 | 6.54 | 6.66 | 6.79 | 6.91 | 7.04 | 7.16 | 7.29 | 7.41 | 7.54 | 7.66 | 7.78 | 7.91 | 8.03 | 8.16 | 8.28 | 8.41 | 8.53 | 8.66 | 8.78 | 8.91 | 9.03 | 9.16 | 9.28 | 9.41 | 9.53 | 9.66 | 9.78
18.98 | 4.79 | 4.91 | 5.04 | 5.16 | 5.29 | 5.41 | 5.54 | 5.66 | 5.79 | 5.91 | 6.04 | 6.16 | 6.29 | 6.41 | 6.54 | 6.66 | 6.79 | 6.91 | 7.03 | 7.16 | 7.28 | 7.41 | 7.53 | 7.66 | 7.78 | 7.91 | 8.03 | 8.16 | 8.28 | 8.41 | 8.53 | 8.66 | 8.78 | 8.91 | 9.03 | 9.16 | 9.28 | 9.4 | 9.53 | 9.65 | 9.78
18.99 | 4.79 | 4.91 | 5.04 | 5.16 | 5.29 | 5.41 | 5.54 | 5.66 | 5.79 | 5.91 | 6.04 | 6.16 | 6.28 | 6.41 | 6.53 | 6.66 | 6.78 | 6.91 | 7.03 | 7.16 | 7.28 | 7.41 | 7.53 | 7.66 | 7.78 | 7.91 | 8.03 | 8.16 | 8.28 | 8.41 | 8.53 | 8.65 | 8.78 | 8.9 | 9.03 | 9.15 | 9.28 | 9.4 | 9.53 | 9.65 | 9.78
19 | 4.79 | 4.91 | 5.04 | 5.16 | 5.29 | 5.41 | 5.54 | 5.66 | 5.78 | 5.91 | 6.03 | 6.16 | 6.28 | 6.41 | 6.53 | 6.66 | 6.78 | 6.91 | 7.03 | 7.16 | 7.28 | 7.41 | 7.53 | 7.66 | 7.78 | 7.9 | 8.03 | 8.15 | 8.28 | 8.4 | 8.53 | 8.65 | 8.78 | 8.9 | 9.03 | 9.15 | 9.28 | 9.4 | 9.53 | 9.65 | 9.77
19.01 | 4.79 | 4.91 | 5.03 | 5.16 | 5.28 | 5.41 | 5.53 | 5.66 | 5.78 | 5.91 | 6.03 | 6.16 | 6.28 | 6.41 | 6.53 | 6.66 | 6.78 | 6.91 | 7.03 | 7.16 | 7.28 | 7.4 | 7.53 | 7.65 | 7.78 | 7.9 | 8.03 | 8.15 | 8.28 | 8.4 | 8.53 | 8.65 | 8.78 | 8.9 | 9.03 | 9.15 | 9.27 | 9.4 | 9.52 | 9.65 | 9.77
19.02 | 4.78 | 4.91 | 5.03 | 5.16 | 5.28 | 5.41 | 5.53 | 5.66 | 5.78 | 5.91 | 6.03 | 6.16 | 6.28 | 6.41 | 6.53 | 6.65 | 6.78 | 6.9 | 7.03 | 7.15 | 7.28 | 7.4 | 7.53 | 7.65 | 7.78 | 7.9 | 8.03 | 8.15 | 8.28 | 8.4 | 8.52 | 8.65 | 8.77 | 8.9 | 9.02 | 9.15 | 9.27 | 9.4 | 9.52 | 9.65 | 9.77
19.03 | 4.78 | 4.91 | 5.03 | 5.16 | 5.28 | 5.41 | 5.53 | 5.66 | 5.78 | 5.91 | 6.03 | 6.15 | 6.28 | 6.4 | 6.53 | 6.65 | 6.78 | 6.9 | 7.03 | 7.15 | 7.28 | 7.4 | 7.53 | 7.65 | 7.78 | 7.9 | 8.02 | 8.15 | 8.27 | 8.4 | 8.52 | 8.65 | 8.77 | 8.9 | 9.02 | 9.15 | 9.27 | 9.4 | 9.52 | 9.64 | 9.77
19.04 | 4.78 | 4.91 | 5.03 | 5.16 | 5.28 | 5.41 | 5.53 | 5.66 | 5.78 | 5.9 | 6.03 | 6.15 | 6.28 | 6.4 | 6.53 | 6.65 | 6.78 | 6.9 | 7.03 | 7.15 | 7.28 | 7.4 | 7.52 | 7.65 | 7.77 | 7.9 | 8.02 | 8.15 | 8.27 | 8.4 | 8.52 | 8.65 | 8.77 | 8.9 | 9.02 | 9.14 | 9.27 | 9.39 | 9.52 | 9.64 | 9.77
19.05 | 4.78 | 4.91 | 5.03 | 5.16 | 5.28 | 5.4 | 5.53 | 5.65 | 5.78 | 5.9 | 6.03 | 6.15 | 6.28 | 6.4 | 6.53 | 6.65 | 6.78 | 6.9 | 7.02 | 7.15 | 7.27 | 7.4 | 7.52 | 7.65 | 7.77 | 7.9 | 8.02 | 8.15 | 8.27 | 8.4 | 8.52 | 8.64 | 8.77 | 8.89 | 9.02 | 9.14 | 9.27 | 9.39 | 9.52 | 9.64 | 9.77
19.06 | 4.78 | 4.9 | 5.03 | 5.15 | 5.28 | 5.4 | 5.53 | 5.65 | 5.78 | 5.9 | 6.03 | 6.15 | 6.28 | 6.4 | 6.52 | 6.65 | 6.77 | 6.9 | 7.02 | 7.15 | 7.27 | 7.4 | 7.52 | 7.65 | 7.77 | 7.9 | 8.02 | 8.14 | 8.27 | 8.39 | 8.52 | 8.64 | 8.77 | 8.89 | 9.02 | 9.14 | 9.27 | 9.39 | 9.51 | 9.64 | 9.76
19.07 | 4.78 | 4.9 | 5.03 | 5.15 | 5.28 | 5.4 | 5.53 | 5.65 | 5.78 | 5.9 | 6.03 | 6.15 | 6.27 | 6.4 | 6.52 | 6.65 | 6.77 | 6.9 | 7.02 | 7.15 | 7.27 | 7.4 | 7.52 | 7.64 | 7.77 | 7.89 | 8.02 | 8.14 | 8.27 | 8.39 | 8.52 | 8.64 | 8.77 | 8.89 | 9.01 | 9.14 | 9.26 | 9.39 | 9.51 | 9.64 | 9.76
19.08 | 4.78 | 4.9 | 5.03 | 5.15 | 5.28 | 5.4 | 5.53 | 5.65 | 5.77 | 5.9 | 6.02 | 6.15 | 6.27 | 6.4 | 6.52 | 6.65 | 6.77 | 6.9 | 7.02 | 7.14 | 7.27 | 7.39 | 7.52 | 7.64 | 7.77 | 7.89 | 8.02 | 8.14 | 8.27 | 8.39 | 8.51 | 8.64 | 8.76 | 8.89 | 9.01 | 9.14 | 9.26 | 9.39 | 9.51 | 9.63 | 9.76
19.09 | 4.78 | 4.9 | 5.03 | 5.15 | 5.28 | 5.4 | 5.52 | 5.65 | 5.77 | 5.9 | 6.02 | 6.15 | 6.27 | 6.4 | 6.52 | 6.65 | 6.77 | 6.89 | 7.02 | 7.14 | 7.27 | 7.39 | 7.52 | 7.64 | 7.77 | 7.89 | 8.01 | 8.14 | 8.26 | 8.39 | 8.51 | 8.64 | 8.76 | 8.89 | 9.01 | 9.14 | 9.26 | 9.38 | 9.51 | 9.63 | 9.76
19.1 | 4.78 | 4.9 | 5.03 | 5.15 | 5.27 | 5.4 | 5.52 | 5.65 | 5.77 | 5.9 | 6.02 | 6.15 | 6.27 | 6.39 | 6.52 | 6.64 | 6.77 | 6.89 | 7.02 | 7.14 | 7.27 | 7.39 | 7.52 | 7.64 | 7.76 | 7.89 | 8.01 | 8.14 | 8.26 | 8.39 | 8.51 | 8.64 | 8.76 | 8.88 | 9.01 | 9.13 | 9.26 | 9.38 | 9.51 | 9.63 | 9.76
19.11 | 4.78 | 4.9 | 5.02 | 5.15 | 5.27 | 5.4 | 5.52 | 5.65 | 5.77 | 5.9 | 6.02 | 6.14 | 6.27 | 6.39 | 6.52 | 6.64 | 6.77 | 6.89 | 7.02 | 7.14 | 7.26 | 7.39 | 7.51 | 7.64 | 7.76 | 7.89 | 8.01 | 8.14 | 8.26 | 8.38 | 8.51 | 8.63 | 8.76 | 8.88 | 9.01 | 9.13 | 9.26 | 9.38 | 9.5 | 9.63 | 9.75
19.12 | 4.77 | 4.9 | 5.02 | 5.15 | 5.27 | 5.4 | 5.52 | 5.65 | 5.77 | 5.89 | 6.02 | 6.14 | 6.27 | 6.39 | 6.52 | 6.64 | 6.77 | 6.89 | 7.01 | 7.14 | 7.26 | 7.39 | 7.51 | 7.64 | 7.76 | 7.89 | 8.01 | 8.13 | 8.26 | 8.38 | 8.51 | 8.63 | 8.76 | 8.88 | 9.01 | 9.13 | 9.25 | 9.38 | 9.5 | 9.63 | 9.75
19.13 | 4.77 | 4.9 | 5.02 | 5.15 | 5.27 | 5.4 | 5.52 | 5.64 | 5.77 | 5.89 | 6.02 | 6.14 | 6.27 | 6.39 | 6.52 | 6.64 | 6.76 | 6.89 | 7.01 | 7.14 | 7.26 | 7.39 | 7.51 | 7.64 | 7.76 | 7.88 | 8.01 | 8.13 | 8.26 | 8.38 | 8.51 | 8.63 | 8.75 | 8.88 | 9 | 9.13 | 9.25 | 9.38 | 9.5 | 9.63 | 9.75
19.14 | 4.77 | 4.9 | 5.02 | 5.15 | 5.27 | 5.39 | 5.52 | 5.64 | 5.77 | 5.89 | 6.02 | 6.14 | 6.27 | 6.39 | 6.51 | 6.64 | 6.76 | 6.89 | 7.01 | 7.14 | 7.26 | 7.38 | 7.51 | 7.63 | 7.76 | 7.88 | 8.01 | 8.13 | 8.26 | 8.38 | 8.5 | 8.63 | 8.75 | 8.88 | 9 | 9.13 | 9.25 | 9.37 | 9.5 | 9.62 | 9.75
19.15 | 4.77 | 4.9 | 5.02 | 5.14 | 5.27 | 5.39 | 5.52 | 5.64 | 5.77 | 5.89 | 6.02 | 6.14 | 6.26 | 6.39 | 6.51 | 6.64 | 6.76 | 6.89 | 7.01 | 7.13 | 7.26 | 7.38 | 7.51 | 7.63 | 7.76 | 7.88 | 8 | 8.13 | 8.25 | 8.38 | 8.5 | 8.63 | 8.75 | 8.88 | 9 | 9.12 | 9.25 | 9.37 | 9.5 | 9.62 | 9.75
19.16 | 4.77 | 4.89 | 5.02 | 5.14 | 5.27 | 5.39 | 5.52 | 5.64 | 5.77 | 5.89 | 6.01 | 6.14 | 6.26 | 6.39 | 6.51 | 6.64 | 6.76 | 6.88 | 7.01 | 7.13 | 7.26 | 7.38 | 7.51 | 7.63 | 7.75 | 7.88 | 8 | 8.13 | 8.25 | 8.38 | 8.5 | 8.62 | 8.75 | 8.87 | 9 | 9.12 | 9.25 | 9.37 | 9.5 | 9.62 | 9.74
19.17 | 4.77 | 4.89 | 5.02 | 5.14 | 5.27 | 5.39 | 5.52 | 5.64 | 5.76 | 5.89 | 6.01 | 6.14 | 6.26 | 6.39 | 6.51 | 6.63 | 6.76 | 6.88 | 7.01 | 7.13 | 7.26 | 7.38 | 7.5 | 7.63 | 7.75 | 7.88 | 8 | 8.13 | 8.25 | 8.37 | 8.5 | 8.62 | 8.75 | 8.87 | 9 | 9.12 | 9.24 | 9.37 | 9.49 | 9.62 | 9.74
19.18 | 4.77 | 4.89 | 5.02 | 5.14 | 5.27 | 5.39 | 5.51 | 5.64 | 5.76 | 5.89 | 6.01 | 6.14 | 6.26 | 6.38 | 6.51 | 6.63 | 6.76 | 6.88 | 7.01 | 7.13 | 7.25 | 7.38 | 7.5 | 7.63 | 7.75 | 7.88 | 8 | 8.12 | 8.25 | 8.37 | 8.5 | 8.62 | 8.75 | 8.87 | 8.99 | 9.12 | 9.24 | 9.37 | 9.49 | 9.62 | 9.74
19.19 | 4.77 | 4.89 | 5.02 | 5.14 | 5.26 | 5.39 | 5.51 | 5.64 | 5.76 | 5.89 | 6.01 | 6.13 | 6.26 | 6.38 | 6.51 | 6.63 | 6.76 | 6.88 | 7 | 7.13 | 7.25 | 7.38 | 7.5 | 7.63 | 7.75 | 7.87 | 8 | 8.12 | 8.25 | 8.37 | 8.5 | 8.62 | 8.74 | 8.87 | 8.99 | 9.12 | 9.24 | 9.37 | 9.49 | 9.61 | 9.74
19.2 | 4.77 | 4.89 | 5.01 | 5.14 | 5.26 | 5.39 | 5.51 | 5.64 | 5.76 | 5.88 | 6.01 | 6.13 | 6.26 | 6.38 | 6.51 | 6.63 | 6.75 | 6.88 | 7 | 7.13 | 7.25 | 7.38 | 7.5 | 7.62 | 7.75 | 7.87 | 8 | 8.12 | 8.25 | 8.37 | 8.49 | 8.62 | 8.74 | 8.87 | 8.99 | 9.11 | 9.24 | 9.36 | 9.49 | 9.61 | 9.74
19.21 | 4.77 | 4.89 | 5.01 | 5.14 | 5.26 | 5.39 | 5.51 | 5.63 | 5.76 | 5.88 | 6.01 | 6.13 | 6.26 | 6.38 | 6.5 | 6.63 | 6.75 | 6.88 | 7 | 7.13 | 7.25 | 7.37 | 7.5 | 7.62 | 7.75 | 7.87 | 8 | 8.12 | 8.24 | 8.37 | 8.49 | 8.62 | 8.74 | 8.86 | 8.99 | 9.11 | 9.24 | 9.36 | 9.49 | 9.61 | 9.73
19.22 | 4.76 | 4.89 | 5.01 | 5.14 | 5.26 | 5.39 | 5.51 | 5.63 | 5.76 | 5.88 | 6.01 | 6.13 | 6.25 | 6.38 | 6.5 | 6.63 | 6.75 | 6.88 | 7 | 7.12 | 7.25 | 7.37 | 7.5 | 7.62 | 7.75 | 7.87 | 7.99 | 8.12 | 8.24 | 8.37 | 8.49 | 8.61 | 8.74 | 8.86 | 8.99 | 9.11 | 9.24 | 9.36 | 9.48 | 9.61 | 9.73
19.23 | 4.76 | 4.89 | 5.01 | 5.14 | 5.26 | 5.38 | 5.51 | 5.63 | 5.76 | 5.88 | 6.01 | 6.13 | 6.25 | 6.38 | 6.5 | 6.63 | 6.75 | 6.87 | 7 | 7.12 | 7.25 | 7.37 | 7.5 | 7.62 | 7.74 | 7.87 | 7.99 | 8.12 | 8.24 | 8.36 | 8.49 | 8.61 | 8.74 | 8.86 | 8.99 | 9.11 | 9.23 | 9.36 | 9.48 | 9.61 | 9.73
19.24 | 4.76 | 4.89 | 5.01 | 5.13 | 5.26 | 5.38 | 5.51 | 5.63 | 5.76 | 5.88 | 6 | 6.13 | 6.25 | 6.38 | 6.5 | 6.62 | 6.75 | 6.87 | 7 | 7.12 | 7.25 | 7.37 | 7.49 | 7.62 | 7.74 | 7.87 | 7.99 | 8.11 | 8.24 | 8.36 | 8.49 | 8.61 | 8.74 | 8.86 | 8.98 | 9.11 | 9.23 | 9.36 | 9.48 | 9.6 | 9.73
19.25 | 4.76 | 4.89 | 5.01 | 5.13 | 5.26 | 5.38 | 5.51 | 5.63 | 5.75 | 5.88 | 6 | 6.13 | 6.25 | 6.38 | 6.5 | 6.62 | 6.75 | 6.87 | 7 | 7.12 | 7.24 | 7.37 | 7.49 | 7.62 | 7.74 | 7.86 | 7.99 | 8.11 | 8.24 | 8.36 | 8.49 | 8.61 | 8.73 | 8.86 | 8.98 | 9.11 | 9.23 | 9.35 | 9.48 | 9.6 | 9.73
19.26 | 4.76 | 4.88 | 5.01 | 5.13 | 5.26 | 5.38 | 5.5 | 5.63 | 5.75 | 5.88 | 6 | 6.13 | 6.25 | 6.37 | 6.5 | 6.62 | 6.75 | 6.87 | 6.99 | 7.12 | 7.24 | 7.37 | 7.49 | 7.61 | 7.74 | 7.86 | 7.99 | 8.11 | 8.24 | 8.36 | 8.48 | 8.61 | 8.73 | 8.86 | 8.98 | 9.1 | 9.23 | 9.35 | 9.48 | 9.6 | 9.72
19.27 | 4.76 | 4.88 | 5.01 | 5.13 | 5.26 | 5.38 | 5.5 | 5.63 | 5.75 | 5.88 | 6 | 6.12 | 6.25 | 6.37 | 6.5 | 6.62 | 6.74 | 6.87 | 6.99 | 7.12 | 7.24 | 7.37 | 7.49 | 7.61 | 7.74 | 7.86 | 7.99 | 8.11 | 8.23 | 8.36 | 8.48 | 8.61 | 8.73 | 8.85 | 8.98 | 9.1 | 9.23 | 9.35 | 9.47 | 9.6 | 9.72
19.28 | 4.76 | 4.88 | 5.01 | 5.13 | 5.25 | 5.38 | 5.5 | 5.63 | 5.75 | 5.87 | 6 | 6.12 | 6.25 | 6.37 | 6.5 | 6.62 | 6.74 | 6.87 | 6.99 | 7.12 | 7.24 | 7.36 | 7.49 | 7.61 | 7.74 | 7.86 | 7.98 | 8.11 | 8.23 | 8.36 | 8.48 | 8.6 | 8.73 | 8.85 | 8.98 | 9.1 | 9.22 | 9.35 | 9.47 | 9.6 | 9.72
19.29 | 4.76 | 4.88 | 5.01 | 5.13 | 5.25 | 5.38 | 5.5 | 5.63 | 5.75 | 5.87 | 6 | 6.12 | 6.25 | 6.37 | 6.49 | 6.62 | 6.74 | 6.87 | 6.99 | 7.11 | 7.24 | 7.36 | 7.49 | 7.61 | 7.73 | 7.86 | 7.98 | 8.11 | 8.23 | 8.35 | 8.48 | 8.6 | 8.73 | 8.85 | 8.97 | 9.1 | 9.22 | 9.35 | 9.47 | 9.59 | 9.72
19.3 | 4.76 | 4.88 | 5 | 5.13 | 5.25 | 5.38 | 5.5 | 5.62 | 5.75 | 5.87 | 6 | 6.12 | 6.24 | 6.37 | 6.49 | 6.62 | 6.74 | 6.86 | 6.99 | 7.11 | 7.24 | 7.36 | 7.48 | 7.61 | 7.73 | 7.86 | 7.98 | 8.1 | 8.23 | 8.35 | 8.48 | 8.6 | 8.72 | 8.85 | 8.97 | 9.1 | 9.22 | 9.34 | 9.47 | 9.59 | 9.72
19.31 | 4.75 | 4.88 | 5 | 5.13 | 5.25 | 5.37 | 5.5 | 5.62 | 5.75 | 5.87 | 6 | 6.12 | 6.24 | 6.37 | 6.49 | 6.62 | 6.74 | 6.86 | 6.99 | 7.11 | 7.24 | 7.36 | 7.48 | 7.61 | 7.73 | 7.85 | 7.98 | 8.1 | 8.23 | 8.35 | 8.47 | 8.6 | 8.72 | 8.85 | 8.97 | 9.09 | 9.22 | 9.34 | 9.47 | 9.59 | 9.71
19.32 | 4.75 | 4.88 | 5 | 5.13 | 5.25 | 5.37 | 5.5 | 5.62 | 5.75 | 5.87 | 5.99 | 6.12 | 6.24 | 6.37 | 6.49 | 6.61 | 6.74 | 6.86 | 6.99 | 7.11 | 7.23 | 7.36 | 7.48 | 7.61 | 7.73 | 7.85 | 7.98 | 8.1 | 8.23 | 8.35 | 8.47 | 8.6 | 8.72 | 8.84 | 8.97 | 9.09 | 9.22 | 9.34 | 9.46 | 9.59 | 9.71
19.33 | 4.75 | 4.88 | 5 | 5.12 | 5.25 | 5.37 | 5.5 | 5.62 | 5.74 | 5.87 | 5.99 | 6.12 | 6.24 | 6.36 | 6.49 | 6.61 | 6.74 | 6.86 | 6.98 | 7.11 | 7.23 | 7.36 | 7.48 | 7.6 | 7.73 | 7.85 | 7.98 | 8.1 | 8.22 | 8.35 | 8.47 | 8.6 | 8.72 | 8.84 | 8.97 | 9.09 | 9.21 | 9.34 | 9.46 | 9.59 | 9.71
19.34 | 4.75 | 4.88 | 5 | 5.12 | 5.25 | 5.37 | 5.5 | 5.62 | 5.74 | 5.87 | 5.99 | 6.12 | 6.24 | 6.36 | 6.49 | 6.61 | 6.73 | 6.86 | 6.98 | 7.11 | 7.23 | 7.35 | 7.48 | 7.6 | 7.73 | 7.85 | 7.97 | 8.1 | 8.22 | 8.35 | 8.47 | 8.59 | 8.72 | 8.84 | 8.97 | 9.09 | 9.21 | 9.34 | 9.46 | 9.58 | 9.71
19.35 | 4.75 | 4.87 | 5 | 5.12 | 5.25 | 5.37 | 5.49 | 5.62 | 5.74 | 5.87 | 5.99 | 6.11 | 6.24 | 6.36 | 6.49 | 6.61 | 6.73 | 6.86 | 6.98 | 7.11 | 7.23 | 7.35 | 7.48 | 7.6 | 7.72 | 7.85 | 7.97 | 8.1 | 8.22 | 8.34 | 8.47 | 8.59 | 8.72 | 8.84 | 8.96 | 9.09 | 9.21 | 9.34 | 9.46 | 9.58 | 9.71
19.36 | 4.75 | 4.87 | 5 | 5.12 | 5.25 | 5.37 | 5.49 | 5.62 | 5.74 | 5.86 | 5.99 | 6.11 | 6.24 | 6.36 | 6.48 | 6.61 | 6.73 | 6.86 | 6.98 | 7.1 | 7.23 | 7.35 | 7.48 | 7.6 | 7.72 | 7.85 | 7.97 | 8.09 | 8.22 | 8.34 | 8.47 | 8.59 | 8.71 | 8.84 | 8.96 | 9.09 | 9.21 | 9.33 | 9.46 | 9.58 | 9.7
19.37 | 4.75 | 4.87 | 5 | 5.12 | 5.24 | 5.37 | 5.49 | 5.62 | 5.74 | 5.86 | 5.99 | 6.11 | 6.24 | 6.36 | 6.48 | 6.61 | 6.73 | 6.85 | 6.98 | 7.1 | 7.23 | 7.35 | 7.47 | 7.6 | 7.72 | 7.85 | 7.97 | 8.09 | 8.22 | 8.34 | 8.46 | 8.59 | 8.71 | 8.84 | 8.96 | 9.08 | 9.21 | 9.33 | 9.46 | 9.58 | 9.7
19.38 | 4.75 | 4.87 | 5 | 5.12 | 5.24 | 5.37 | 5.49 | 5.61 | 5.74 | 5.86 | 5.99 | 6.11 | 6.23 | 6.36 | 6.48 | 6.61 | 6.73 | 6.85 | 6.98 | 7.1 | 7.22 | 7.35 | 7.47 | 7.6 | 7.72 | 7.84 | 7.97 | 8.09 | 8.22 | 8.34 | 8.46 | 8.59 | 8.71 | 8.83 | 8.96 | 9.08 | 9.21 | 9.33 | 9.45 | 9.58 | 9.7
19.39 | 4.75 | 4.87 | 4.99 | 5.12 | 5.24 | 5.37 | 5.49 | 5.61 | 5.74 | 5.86 | 5.99 | 6.11 | 6.23 | 6.36 | 6.48 | 6.6 | 6.73 | 6.85 | 6.98 | 7.1 | 7.22 | 7.35 | 7.47 | 7.59 | 7.72 | 7.84 | 7.97 | 8.09 | 8.21 | 8.34 | 8.46 | 8.58 | 8.71 | 8.83 | 8.96 | 9.08 | 9.2 | 9.33 | 9.45 | 9.58 | 9.7
19.4 | 4.75 | 4.87 | 4.99 | 5.12 | 5.24 | 5.36 | 5.49 | 5.61 | 5.74 | 5.86 | 5.98 | 6.11 | 6.23 | 6.36 | 6.48 | 6.6 | 6.73 | 6.85 | 6.97 | 7.1 | 7.22 | 7.35 | 7.47 | 7.59 | 7.72 | 7.84 | 7.96 | 8.09 | 8.21 | 8.34 | 8.46 | 8.58 | 8.71 | 8.83 | 8.95 | 9.08 | 9.2 | 9.33 | 9.45 | 9.57 | 9.7
19.41 | 4.74 | 4.87 | 4.99 | 5.12 | 5.24 | 5.36 | 5.49 | 5.61 | 5.74 | 5.86 | 5.98 | 6.11 | 6.23 | 6.35 | 6.48 | 6.6 | 6.73 | 6.85 | 6.97 | 7.1 | 7.22 | 7.34 | 7.47 | 7.59 | 7.72 | 7.84 | 7.96 | 8.09 | 8.21 | 8.33 | 8.46 | 8.58 | 8.71 | 8.83 | 8.95 | 9.08 | 9.2 | 9.32 | 9.45 | 9.57 | 9.69
19.42 | 4.74 | 4.87 | 4.99 | 5.11 | 5.24 | 5.36 | 5.49 | 5.61 | 5.73 | 5.86 | 5.98 | 6.11 | 6.23 | 6.35 | 6.48 | 6.6 | 6.72 | 6.85 | 6.97 | 7.1 | 7.22 | 7.34 | 7.47 | 7.59 | 7.71 | 7.84 | 7.96 | 8.08 | 8.21 | 8.33 | 8.46 | 8.58 | 8.7 | 8.83 | 8.95 | 9.07 | 9.2 | 9.32 | 9.45 | 9.57 | 9.69
19.43 | 4.74 | 4.87 | 4.99 | 5.11 | 5.24 | 5.36 | 5.49 | 5.61 | 5.73 | 5.86 | 5.98 | 6.1 | 6.23 | 6.35 | 6.48 | 6.6 | 6.72 | 6.85 | 6.97 | 7.09 | 7.22 | 7.34 | 7.46 | 7.59 | 7.71 | 7.84 | 7.96 | 8.08 | 8.21 | 8.33 | 8.45 | 8.58 | 8.7 | 8.83 | 8.95 | 9.07 | 9.2 | 9.32 | 9.44 | 9.57 | 9.69
19.44 | 4.74 | 4.87 | 4.99 | 5.11 | 5.24 | 5.36 | 5.48 | 5.61 | 5.73 | 5.86 | 5.98 | 6.1 | 6.23 | 6.35 | 6.47 | 6.6 | 6.72 | 6.84 | 6.97 | 7.09 | 7.22 | 7.34 | 7.46 | 7.59 | 7.71 | 7.83 | 7.96 | 8.08 | 8.21 | 8.33 | 8.45 | 8.58 | 8.7 | 8.82 | 8.95 | 9.07 | 9.19 | 9.32 | 9.44 | 9.57 | 9.69
19.45 | 4.74 | 4.86 | 4.99 | 5.11 | 5.24 | 5.36 | 5.48 | 5.61 | 5.73 | 5.85 | 5.98 | 6.1 | 6.22 | 6.35 | 6.47 | 6.6 | 6.72 | 6.84 | 6.97 | 7.09 | 7.21 | 7.34 | 7.46 | 7.59 | 7.71 | 7.83 | 7.96 | 8.08 | 8.2 | 8.33 | 8.45 | 8.57 | 8.7 | 8.82 | 8.95 | 9.07 | 9.19 | 9.32 | 9.44 | 9.56 | 9.69
19.46 | 4.74 | 4.86 | 4.99 | 5.11 | 5.23 | 5.36 | 5.48 | 5.61 | 5.73 | 5.85 | 5.98 | 6.1 | 6.22 | 6.35 | 6.47 | 6.59 | 6.72 | 6.84 | 6.97 | 7.09 | 7.21 | 7.34 | 7.46 | 7.58 | 7.71 | 7.83 | 7.95 | 8.08 | 8.2 | 8.33 | 8.45 | 8.57 | 8.7 | 8.82 | 8.94 | 9.07 | 9.19 | 9.31 | 9.44 | 9.56 | 9.69
19.47 | 4.74 | 4.86 | 4.99 | 5.11 | 5.23 | 5.36 | 5.48 | 5.6 | 5.73 | 5.85 | 5.98 | 6.1 | 6.22 | 6.35 | 6.47 | 6.59 | 6.72 | 6.84 | 6.96 | 7.09 | 7.21 | 7.34 | 7.46 | 7.58 | 7.71 | 7.83 | 7.95 | 8.08 | 8.2 | 8.32 | 8.45 | 8.57 | 8.69 | 8.82 | 8.94 | 9.07 | 9.19 | 9.31 | 9.44 | 9.56 | 9.68
19.48 | 4.74 | 4.86 | 4.98 | 5.11 | 5.23 | 5.36 | 5.48 | 5.6 | 5.73 | 5.85 | 5.97 | 6.1 | 6.22 | 6.34 | 6.47 | 6.59 | 6.72 | 6.84 | 6.96 | 7.09 | 7.21 | 7.33 | 7.46 | 7.58 | 7.7 | 7.83 | 7.95 | 8.08 | 8.2 | 8.32 | 8.45 | 8.57 | 8.69 | 8.82 | 8.94 | 9.06 | 9.19 | 9.31 | 9.43 | 9.56 | 9.68
19.49 | 4.74 | 4.86 | 4.98 | 5.11 | 5.23 | 5.35 | 5.48 | 5.6 | 5.73 | 5.85 | 5.97 | 6.1 | 6.22 | 6.34 | 6.47 | 6.59 | 6.71 | 6.84 | 6.96 | 7.08 | 7.21 | 7.33 | 7.46 | 7.58 | 7.7 | 7.83 | 7.95 | 8.07 | 8.2 | 8.32 | 8.44 | 8.57 | 8.69 | 8.81 | 8.94 | 9.06 | 9.19 | 9.31 | 9.43 | 9.56 | 9.68
19.5 | 4.74 | 4.86 | 4.98 | 5.11 | 5.23 | 5.35 | 5.48 | 5.6 | 5.72 | 5.85 | 5.97 | 6.09 | 6.22 | 6.34 | 6.47 | 6.59 | 6.71 | 6.84 | 6.96 | 7.08 | 7.21 | 7.33 | 7.45 | 7.58 | 7.7 | 7.82 | 7.95 | 8.07 | 8.2 | 8.32 | 8.44 | 8.57 | 8.69 | 8.81 | 8.94 | 9.06 | 9.18 | 9.31 | 9.43 | 9.55 | 9.68
19.51 | 4.73 | 4.86 | 4.98 | 5.11 | 5.23 | 5.35 | 5.48 | 5.6 | 5.72 | 5.85 | 5.97 | 6.09 | 6.22 | 6.34 | 6.46 | 6.59 | 6.71 | 6.83 | 6.96 | 7.08 | 7.21 | 7.33 | 7.45 | 7.58 | 7.7 | 7.82 | 7.95 | 8.07 | 8.19 | 8.32 | 8.44 | 8.56 | 8.69 | 8.81 | 8.93 | 9.06 | 9.18 | 9.31 | 9.43 | 9.55 | 9.68
19.52 | 4.73 | 4.86 | 4.98 | 5.1 | 5.23 | 5.35 | 5.47 | 5.6 | 5.72 | 5.85 | 5.97 | 6.09 | 6.22 | 6.34 | 6.46 | 6.59 | 6.71 | 6.83 | 6.96 | 7.08 | 7.2 | 7.33 | 7.45 | 7.57 | 7.7 | 7.82 | 7.95 | 8.07 | 8.19 | 8.32 | 8.44 | 8.56 | 8.69 | 8.81 | 8.93 | 9.06 | 9.18 | 9.3 | 9.43 | 9.55 | 9.67
19.53 | 4.73 | 4.86 | 4.98 | 5.1 | 5.23 | 5.35 | 5.47 | 5.6 | 5.72 | 5.84 | 5.97 | 6.09 | 6.21 | 6.34 | 6.46 | 6.59 | 6.71 | 6.83 | 6.96 | 7.08 | 7.2 | 7.33 | 7.45 | 7.57 | 7.7 | 7.82 | 7.94 | 8.07 | 8.19 | 8.31 | 8.44 | 8.56 | 8.68 | 8.81 | 8.93 | 9.05 | 9.18 | 9.3 | 9.42 | 9.55 | 9.67
19.54 | 4.73 | 4.85 | 4.98 | 5.1 | 5.23 | 5.35 | 5.47 | 5.6 | 5.72 | 5.84 | 5.97 | 6.09 | 6.21 | 6.34 | 6.46 | 6.58 | 6.71 | 6.83 | 6.95 | 7.08 | 7.2 | 7.32 | 7.45 | 7.57 | 7.69 | 7.82 | 7.94 | 8.07 | 8.19 | 8.31 | 8.44 | 8.56 | 8.68 | 8.81 | 8.93 | 9.05 | 9.18 | 9.3 | 9.42 | 9.55 | 9.67
19.55 | 4.73 | 4.85 | 4.98 | 5.1 | 5.22 | 5.35 | 5.47 | 5.59 | 5.72 | 5.84 | 5.97 | 6.09 | 6.21 | 6.34 | 6.46 | 6.58 | 6.71 | 6.83 | 6.95 | 7.08 | 7.2 | 7.32 | 7.45 | 7.57 | 7.69 | 7.82 | 7.94 | 8.06 | 8.19 | 8.31 | 8.43 | 8.56 | 8.68 | 8.8 | 8.93 | 9.05 | 9.17 | 9.3 | 9.42 | 9.54 | 9.67
19.56 | 4.73 | 4.85 | 4.98 | 5.1 | 5.22 | 5.35 | 5.47 | 5.59 | 5.72 | 5.84 | 5.96 | 6.09 | 6.21 | 6.33 | 6.46 | 6.58 | 6.7 | 6.83 | 6.95 | 7.07 | 7.2 | 7.32 | 7.45 | 7.57 | 7.69 | 7.82 | 7.94 | 8.06 | 8.19 | 8.31 | 8.43 | 8.56 | 8.68 | 8.8 | 8.93 | 9.05 | 9.17 | 9.3 | 9.42 | 9.54 | 9.67
19.57 | 4.73 | 4.85 | 4.98 | 5.1 | 5.22 | 5.35 | 5.47 | 5.59 | 5.72 | 5.84 | 5.96 | 6.09 | 6.21 | 6.33 | 6.46 | 6.58 | 6.7 | 6.83 | 6.95 | 7.07 | 7.2 | 7.32 | 7.44 | 7.57 | 7.69 | 7.81 | 7.94 | 8.06 | 8.18 | 8.31 | 8.43 | 8.55 | 8.68 | 8.8 | 8.92 | 9.05 | 9.17 | 9.29 | 9.42 | 9.54 | 9.66
19.58 | 4.73 | 4.85 | 4.97 | 5.1 | 5.22 | 5.34 | 5.47 | 5.59 | 5.71 | 5.84 | 5.96 | 6.08 | 6.21 | 6.33 | 6.46 | 6.58 | 6.7 | 6.83 | 6.95 | 7.07 | 7.2 | 7.32 | 7.44 | 7.57 | 7.69 | 7.81 | 7.94 | 8.06 | 8.18 | 8.31 | 8.43 | 8.55 | 8.68 | 8.8 | 8.92 | 9.05 | 9.17 | 9.29 | 9.42 | 9.54 | 9.66
19.59 | 4.73 | 4.85 | 4.97 | 5.1 | 5.22 | 5.34 | 5.47 | 5.59 | 5.71 | 5.84 | 5.96 | 6.08 | 6.21 | 6.33 | 6.45 | 6.58 | 6.7 | 6.82 | 6.95 | 7.07 | 7.19 | 7.32 | 7.44 | 7.56 | 7.69 | 7.81 | 7.93 | 8.06 | 8.18 | 8.3 | 8.43 | 8.55 | 8.67 | 8.8 | 8.92 | 9.04 | 9.17 | 9.29 | 9.41 | 9.54 | 9.66
19.6 | 4.73 | 4.85 | 4.97 | 5.1 | 5.22 | 5.34 | 5.47 | 5.59 | 5.71 | 5.84 | 5.96 | 6.08 | 6.21 | 6.33 | 6.45 | 6.58 | 6.7 | 6.82 | 6.95 | 7.07 | 7.19 | 7.32 | 7.44 | 7.56 | 7.69 | 7.81 | 7.93 | 8.06 | 8.18 | 8.3 | 8.43 | 8.55 | 8.67 | 8.8 | 8.92 | 9.04 | 9.17 | 9.29 | 9.41 | 9.54 | 9.66
19.61 | 4.72 | 4.85 | 4.97 | 5.09 | 5.22 | 5.34 | 5.46 | 5.59 | 5.71 | 5.83 | 5.96 | 6.08 | 6.2 | 6.33 | 6.45 | 6.57 | 6.7 | 6.82 | 6.94 | 7.07 | 7.19 | 7.31 | 7.44 | 7.56 | 7.68 | 7.81 | 7.93 | 8.05 | 8.18 | 8.3 | 8.42 | 8.55 | 8.67 | 8.79 | 8.92 | 9.04 | 9.16 | 9.29 | 9.41 | 9.53 | 9.66
19.62 | 4.72 | 4.85 | 4.97 | 5.09 | 5.22 | 5.34 | 5.46 | 5.59 | 5.71 | 5.83 | 5.96 | 6.08 | 6.2 | 6.33 | 6.45 | 6.57 | 6.7 | 6.82 | 6.94 | 7.07 | 7.19 | 7.31 | 7.44 | 7.56 | 7.68 | 7.81 | 7.93 | 8.05 | 8.18 | 8.3 | 8.42 | 8.55 | 8.67 | 8.79 | 8.92 | 9.04 | 9.16 | 9.28 | 9.41 | 9.53 | 9.65
19.63 | 4.72 | 4.85 | 4.97 | 5.09 | 5.22 | 5.34 | 5.46 | 5.59 | 5.71 | 5.83 | 5.96 | 6.08 | 6.2 | 6.33 | 6.45 | 6.57 | 6.69 | 6.82 | 6.94 | 7.06 | 7.19 | 7.31 | 7.43 | 7.56 | 7.68 | 7.8 | 7.93 | 8.05 | 8.17 | 8.3 | 8.42 | 8.54 | 8.67 | 8.79 | 8.91 | 9.04 | 9.16 | 9.28 | 9.41 | 9.53 | 9.65
19.64 | 4.72 | 4.84 | 4.97 | 5.09 | 5.21 | 5.34 | 5.46 | 5.58 | 5.71 | 5.83 | 5.95 | 6.08 | 6.2 | 6.32 | 6.45 | 6.57 | 6.69 | 6.82 | 6.94 | 7.06 | 7.19 | 7.31 | 7.43 | 7.56 | 7.68 | 7.8 | 7.93 | 8.05 | 8.17 | 8.3 | 8.42 | 8.54 | 8.67 | 8.79 | 8.91 | 9.03 | 9.16 | 9.28 | 9.4 | 9.53 | 9.65
19.65 | 4.72 | 4.84 | 4.97 | 5.09 | 5.21 | 5.34 | 5.46 | 5.58 | 5.71 | 5.83 | 5.95 | 6.08 | 6.2 | 6.32 | 6.45 | 6.57 | 6.69 | 6.82 | 6.94 | 7.06 | 7.19 | 7.31 | 7.43 | 7.55 | 7.68 | 7.8 | 7.92 | 8.05 | 8.17 | 8.29 | 8.42 | 8.54 | 8.66 | 8.79 | 8.91 | 9.03 | 9.16 | 9.28 | 9.4 | 9.53 | 9.65
19.66 | 4.72 | 4.84 | 4.97 | 5.09 | 5.21 | 5.34 | 5.46 | 5.58 | 5.71 | 5.83 | 5.95 | 6.07 | 6.2 | 6.32 | 6.44 | 6.57 | 6.69 | 6.81 | 6.94 | 7.06 | 7.18 | 7.31 | 7.43 | 7.55 | 7.68 | 7.8 | 7.92 | 8.05 | 8.17 | 8.29 | 8.42 | 8.54 | 8.66 | 8.78 | 8.91 | 9.03 | 9.15 | 9.28 | 9.4 | 9.52 | 9.65
19.67 | 4.72 | 4.84 | 4.96 | 5.09 | 5.21 | 5.33 | 5.46 | 5.58 | 5.7 | 5.83 | 5.95 | 6.07 | 6.2 | 6.32 | 6.44 | 6.57 | 6.69 | 6.81 | 6.94 | 7.06 | 7.18 | 7.31 | 7.43 | 7.55 | 7.67 | 7.8 | 7.92 | 8.04 | 8.17 | 8.29 | 8.41 | 8.54 | 8.66 | 8.78 | 8.91 | 9.03 | 9.15 | 9.28 | 9.4 | 9.52 | 9.64
19.68 | 4.72 | 4.84 | 4.96 | 5.09 | 5.21 | 5.33 | 5.46 | 5.58 | 5.7 | 5.83 | 5.95 | 6.07 | 6.2 | 6.32 | 6.44 | 6.56 | 6.69 | 6.81 | 6.93 | 7.06 | 7.18 | 7.3 | 7.43 | 7.55 | 7.67 | 7.8 | 7.92 | 8.04 | 8.17 | 8.29 | 8.41 | 8.54 | 8.66 | 8.78 | 8.9 | 9.03 | 9.15 | 9.27 | 9.4 | 9.52 | 9.64
19.69 | 4.72 | 4.84 | 4.96 | 5.09 | 5.21 | 5.33 | 5.46 | 5.58 | 5.7 | 5.82 | 5.95 | 6.07 | 6.19 | 6.32 | 6.44 | 6.56 | 6.69 | 6.81 | 6.93 | 7.06 | 7.18 | 7.3 | 7.43 | 7.55 | 7.67 | 7.79 | 7.92 | 8.04 | 8.16 | 8.29 | 8.41 | 8.53 | 8.66 | 8.78 | 8.9 | 9.03 | 9.15 | 9.27 | 9.39 | 9.52 | 9.64
19.7 | 4.72 | 4.84 | 4.96 | 5.08 | 5.21 | 5.33 | 5.45 | 5.58 | 5.7 | 5.82 | 5.95 | 6.07 | 6.19 | 6.32 | 6.44 | 6.56 | 6.69 | 6.81 | 6.93 | 7.05 | 7.18 | 7.3 | 7.42 | 7.55 | 7.67 | 7.79 | 7.92 | 8.04 | 8.16 | 8.29 | 8.41 | 8.53 | 8.65 | 8.78 | 8.9 | 9.02 | 9.15 | 9.27 | 9.39 | 9.52 | 9.64
19.71 | 4.71 | 4.84 | 4.96 | 5.08 | 5.21 | 5.33 | 5.45 | 5.58 | 5.7 | 5.82 | 5.95 | 6.07 | 6.19 | 6.31 | 6.44 | 6.56 | 6.68 | 6.81 | 6.93 | 7.05 | 7.18 | 7.3 | 7.42 | 7.55 | 7.67 | 7.79 | 7.91 | 8.04 | 8.16 | 8.28 | 8.41 | 8.53 | 8.65 | 8.78 | 8.9 | 9.02 | 9.15 | 9.27 | 9.39 | 9.51 | 9.64
19.72 | 4.71 | 4.84 | 4.96 | 5.08 | 5.21 | 5.33 | 5.45 | 5.57 | 5.7 | 5.82 | 5.94 | 6.07 | 6.19 | 6.31 | 6.44 | 6.56 | 6.68 | 6.81 | 6.93 | 7.05 | 7.17 | 7.3 | 7.42 | 7.54 | 7.67 | 7.79 | 7.91 | 8.04 | 8.16 | 8.28 | 8.41 | 8.53 | 8.65 | 8.77 | 8.9 | 9.02 | 9.14 | 9.27 | 9.39 | 9.51 | 9.64
19.73 | 4.71 | 4.84 | 4.96 | 5.08 | 5.2 | 5.33 | 5.45 | 5.57 | 5.7 | 5.82 | 5.94 | 6.07 | 6.19 | 6.31 | 6.44 | 6.56 | 6.68 | 6.8 | 6.93 | 7.05 | 7.17 | 7.3 | 7.42 | 7.54 | 7.67 | 7.79 | 7.91 | 8.03 | 8.16 | 8.28 | 8.4 | 8.53 | 8.65 | 8.77 | 8.9 | 9.02 | 9.14 | 9.26 | 9.39 | 9.51 | 9.63
19.74 | 4.71 | 4.83 | 4.96 | 5.08 | 5.2 | 5.33 | 5.45 | 5.57 | 5.7 | 5.82 | 5.94 | 6.06 | 6.19 | 6.31 | 6.43 | 6.56 | 6.68 | 6.8 | 6.93 | 7.05 | 7.17 | 7.29 | 7.42 | 7.54 | 7.66 | 7.79 | 7.91 | 8.03 | 8.16 | 8.28 | 8.4 | 8.52 | 8.65 | 8.77 | 8.89 | 9.02 | 9.14 | 9.26 | 9.39 | 9.51 | 9.63
19.75 | 4.71 | 4.83 | 4.96 | 5.08 | 5.2 | 5.33 | 5.45 | 5.57 | 5.69 | 5.82 | 5.94 | 6.06 | 6.19 | 6.31 | 6.43 | 6.56 | 6.68 | 6.8 | 6.92 | 7.05 | 7.17 | 7.29 | 7.42 | 7.54 | 7.66 | 7.79 | 7.91 | 8.03 | 8.15 | 8.28 | 8.4 | 8.52 | 8.65 | 8.77 | 8.89 | 9.01 | 9.14 | 9.26 | 9.38 | 9.51 | 9.63
19.76 | 4.71 | 4.83 | 4.96 | 5.08 | 5.2 | 5.32 | 5.45 | 5.57 | 5.69 | 5.82 | 5.94 | 6.06 | 6.19 | 6.31 | 6.43 | 6.55 | 6.68 | 6.8 | 6.92 | 7.05 | 7.17 | 7.29 | 7.41 | 7.54 | 7.66 | 7.78 | 7.91 | 8.03 | 8.15 | 8.28 | 8.4 | 8.52 | 8.64 | 8.77 | 8.89 | 9.01 | 9.14 | 9.26 | 9.38 | 9.5 | 9.63
19.77 | 4.71 | 4.83 | 4.95 | 5.08 | 5.2 | 5.32 | 5.45 | 5.57 | 5.69 | 5.82 | 5.94 | 6.06 | 6.18 | 6.31 | 6.43 | 6.55 | 6.68 | 6.8 | 6.92 | 7.04 | 7.17 | 7.29 | 7.41 | 7.54 | 7.66 | 7.78 | 7.91 | 8.03 | 8.15 | 8.27 | 8.4 | 8.52 | 8.64 | 8.77 | 8.89 | 9.01 | 9.13 | 9.26 | 9.38 | 9.5 | 9.63
19.78 | 4.71 | 4.83 | 4.95 | 5.08 | 5.2 | 5.32 | 5.45 | 5.57 | 5.69 | 5.81 | 5.94 | 6.06 | 6.18 | 6.31 | 6.43 | 6.55 | 6.67 | 6.8 | 6.92 | 7.04 | 7.17 | 7.29 | 7.41 | 7.53 | 7.66 | 7.78 | 7.9 | 8.03 | 8.15 | 8.27 | 8.4 | 8.52 | 8.64 | 8.76 | 8.89 | 9.01 | 9.13 | 9.26 | 9.38 | 9.5 | 9.62
19.79 | 4.71 | 4.83 | 4.95 | 5.08 | 5.2 | 5.32 | 5.44 | 5.57 | 5.69 | 5.81 | 5.94 | 6.06 | 6.18 | 6.3 | 6.43 | 6.55 | 6.67 | 6.8 | 6.92 | 7.04 | 7.16 | 7.29 | 7.41 | 7.53 | 7.66 | 7.78 | 7.9 | 8.02 | 8.15 | 8.27 | 8.39 | 8.52 | 8.64 | 8.76 | 8.88 | 9.01 | 9.13 | 9.25 | 9.38 | 9.5 | 9.62
19.8 | 4.71 | 4.83 | 4.95 | 5.07 | 5.2 | 5.32 | 5.44 | 5.57 | 5.69 | 5.81 | 5.93 | 6.06 | 6.18 | 6.3 | 6.43 | 6.55 | 6.67 | 6.79 | 6.92 | 7.04 | 7.16 | 7.29 | 7.41 | 7.53 | 7.65 | 7.78 | 7.9 | 8.02 | 8.15 | 8.27 | 8.39 | 8.51 | 8.64 | 8.76 | 8.88 | 9.01 | 9.13 | 9.25 | 9.37 | 9.5 | 9.62
19.81 | 4.7 | 4.83 | 4.95 | 5.07 | 5.2 | 5.32 | 5.44 | 5.56 | 5.69 | 5.81 | 5.93 | 6.06 | 6.18 | 6.3 | 6.42 | 6.55 | 6.67 | 6.79 | 6.92 | 7.04 | 7.16 | 7.28 | 7.41 | 7.53 | 7.65 | 7.78 | 7.9 | 8.02 | 8.14 | 8.27 | 8.39 | 8.51 | 8.64 | 8.76 | 8.88 | 9 | 9.13 | 9.25 | 9.37 | 9.5 | 9.62
19.82 | 4.7 | 4.83 | 4.95 | 5.07 | 5.19 | 5.32 | 5.44 | 5.56 | 5.69 | 5.81 | 5.93 | 6.05 | 6.18 | 6.3 | 6.42 | 6.55 | 6.67 | 6.79 | 6.91 | 7.04 | 7.16 | 7.28 | 7.41 | 7.53 | 7.65 | 7.77 | 7.9 | 8.02 | 8.14 | 8.27 | 8.39 | 8.51 | 8.63 | 8.76 | 8.88 | 9 | 9.13 | 9.25 | 9.37 | 9.49 | 9.62
19.83 | 4.7 | 4.83 | 4.95 | 5.07 | 5.19 | 5.32 | 5.44 | 5.56 | 5.68 | 5.81 | 5.93 | 6.05 | 6.18 | 6.3 | 6.42 | 6.54 | 6.67 | 6.79 | 6.91 | 7.04 | 7.16 | 7.28 | 7.4 | 7.53 | 7.65 | 7.77 | 7.9 | 8.02 | 8.14 | 8.26 | 8.39 | 8.51 | 8.63 | 8.75 | 8.88 | 9 | 9.12 | 9.25 | 9.37 | 9.49 | 9.61
19.84 | 4.7 | 4.82 | 4.95 | 5.07 | 5.19 | 5.32 | 5.44 | 5.56 | 5.68 | 5.81 | 5.93 | 6.05 | 6.18 | 6.3 | 6.42 | 6.54 | 6.67 | 6.79 | 6.91 | 7.03 | 7.16 | 7.28 | 7.4 | 7.53 | 7.65 | 7.77 | 7.89 | 8.02 | 8.14 | 8.26 | 8.38 | 8.51 | 8.63 | 8.75 | 8.88 | 9 | 9.12 | 9.24 | 9.37 | 9.49 | 9.61
19.85 | 4.7 | 4.82 | 4.95 | 5.07 | 5.19 | 5.31 | 5.44 | 5.56 | 5.68 | 5.81 | 5.93 | 6.05 | 6.17 | 6.3 | 6.42 | 6.54 | 6.66 | 6.79 | 6.91 | 7.03 | 7.16 | 7.28 | 7.4 | 7.52 | 7.65 | 7.77 | 7.89 | 8.02 | 8.14 | 8.26 | 8.38 | 8.51 | 8.63 | 8.75 | 8.87 | 9 | 9.12 | 9.24 | 9.37 | 9.49 | 9.61
19.86 | 4.7 | 4.82 | 4.94 | 5.07 | 5.19 | 5.31 | 5.44 | 5.56 | 5.68 | 5.8 | 5.93 | 6.05 | 6.17 | 6.3 | 6.42 | 6.54 | 6.66 | 6.79 | 6.91 | 7.03 | 7.15 | 7.28 | 7.4 | 7.52 | 7.65 | 7.77 | 7.89 | 8.01 | 8.14 | 8.26 | 8.38 | 8.5 | 8.63 | 8.75 | 8.87 | 9 | 9.12 | 9.24 | 9.36 | 9.49 | 9.61
19.87 | 4.7 | 4.82 | 4.94 | 5.07 | 5.19 | 5.31 | 5.43 | 5.56 | 5.68 | 5.8 | 5.93 | 6.05 | 6.17 | 6.29 | 6.42 | 6.54 | 6.66 | 6.78 | 6.91 | 7.03 | 7.15 | 7.28 | 7.4 | 7.52 | 7.64 | 7.77 | 7.89 | 8.01 | 8.13 | 8.26 | 8.38 | 8.5 | 8.63 | 8.75 | 8.87 | 8.99 | 9.12 | 9.24 | 9.36 | 9.48 | 9.61
19.88 | 4.7 | 4.82 | 4.94 | 5.07 | 5.19 | 5.31 | 5.43 | 5.56 | 5.68 | 5.8 | 5.92 | 6.05 | 6.17 | 6.29 | 6.42 | 6.54 | 6.66 | 6.78 | 6.91 | 7.03 | 7.15 | 7.27 | 7.4 | 7.52 | 7.64 | 7.76 | 7.89 | 8.01 | 8.13 | 8.26 | 8.38 | 8.5 | 8.62 | 8.75 | 8.87 | 8.99 | 9.11 | 9.24 | 9.36 | 9.48 | 9.6
19.89 | 4.7 | 4.82 | 4.94 | 5.06 | 5.19 | 5.31 | 5.43 | 5.56 | 5.68 | 5.8 | 5.92 | 6.05 | 6.17 | 6.29 | 6.41 | 6.54 | 6.66 | 6.78 | 6.9 | 7.03 | 7.15 | 7.27 | 7.4 | 7.52 | 7.64 | 7.76 | 7.89 | 8.01 | 8.13 | 8.25 | 8.38 | 8.5 | 8.62 | 8.74 | 8.87 | 8.99 | 9.11 | 9.23 | 9.36 | 9.48 | 9.6
19.9 | 4.7 | 4.82 | 4.94 | 5.06 | 5.19 | 5.31 | 5.43 | 5.55 | 5.68 | 5.8 | 5.92 | 6.04 | 6.17 | 6.29 | 6.41 | 6.54 | 6.66 | 6.78 | 6.9 | 7.03 | 7.15 | 7.27 | 7.39 | 7.52 | 7.64 | 7.76 | 7.88 | 8.01 | 8.13 | 8.25 | 8.37 | 8.5 | 8.62 | 8.74 | 8.87 | 8.99 | 9.11 | 9.23 | 9.36 | 9.48 | 9.6
19.91 | 4.69 | 4.82 | 4.94 | 5.06 | 5.18 | 5.31 | 5.43 | 5.55 | 5.68 | 5.8 | 5.92 | 6.04 | 6.17 | 6.29 | 6.41 | 6.53 | 6.66 | 6.78 | 6.9 | 7.02 | 7.15 | 7.27 | 7.39 | 7.51 | 7.64 | 7.76 | 7.88 | 8.01 | 8.13 | 8.25 | 8.37 | 8.5 | 8.62 | 8.74 | 8.86 | 8.99 | 9.11 | 9.23 | 9.35 | 9.48 | 9.6
19.92 | 4.69 | 4.82 | 4.94 | 5.06 | 5.18 | 5.31 | 5.43 | 5.55 | 5.67 | 5.8 | 5.92 | 6.04 | 6.16 | 6.29 | 6.41 | 6.53 | 6.66 | 6.78 | 6.9 | 7.02 | 7.15 | 7.27 | 7.39 | 7.51 | 7.64 | 7.76 | 7.88 | 8 | 8.13 | 8.25 | 8.37 | 8.49 | 8.62 | 8.74 | 8.86 | 8.98 | 9.11 | 9.23 | 9.35 | 9.47 | 9.6
19.93 | 4.69 | 4.81 | 4.94 | 5.06 | 5.18 | 5.31 | 5.43 | 5.55 | 5.67 | 5.8 | 5.92 | 6.04 | 6.16 | 6.29 | 6.41 | 6.53 | 6.65 | 6.78 | 6.9 | 7.02 | 7.14 | 7.27 | 7.39 | 7.51 | 7.63 | 7.76 | 7.88 | 8 | 8.12 | 8.25 | 8.37 | 8.49 | 8.61 | 8.74 | 8.86 | 8.98 | 9.11 | 9.23 | 9.35 | 9.47 | 9.6
19.94 | 4.69 | 4.81 | 4.94 | 5.06 | 5.18 | 5.3 | 5.43 | 5.55 | 5.67 | 5.79 | 5.92 | 6.04 | 6.16 | 6.28 | 6.41 | 6.53 | 6.65 | 6.78 | 6.9 | 7.02 | 7.14 | 7.27 | 7.39 | 7.51 | 7.63 | 7.76 | 7.88 | 8 | 8.12 | 8.25 | 8.37 | 8.49 | 8.61 | 8.74 | 8.86 | 8.98 | 9.1 | 9.23 | 9.35 | 9.47 | 9.59
19.95 | 4.69 | 4.81 | 4.94 | 5.06 | 5.18 | 5.3 | 5.43 | 5.55 | 5.67 | 5.79 | 5.92 | 6.04 | 6.16 | 6.28 | 6.41 | 6.53 | 6.65 | 6.77 | 6.9 | 7.02 | 7.14 | 7.26 | 7.39 | 7.51 | 7.63 | 7.75 | 7.88 | 8 | 8.12 | 8.24 | 8.37 | 8.49 | 8.61 | 8.73 | 8.86 | 8.98 | 9.1 | 9.22 | 9.35 | 9.47 | 9.59
19.96 | 4.69 | 4.81 | 4.93 | 5.06 | 5.18 | 5.3 | 5.42 | 5.55 | 5.67 | 5.79 | 5.91 | 6.04 | 6.16 | 6.28 | 6.4 | 6.53 | 6.65 | 6.77 | 6.89 | 7.02 | 7.14 | 7.26 | 7.38 | 7.51 | 7.63 | 7.75 | 7.87 | 8 | 8.12 | 8.24 | 8.36 | 8.49 | 8.61 | 8.73 | 8.85 | 8.98 | 9.1 | 9.22 | 9.34 | 9.47 | 9.59
19.97 | 4.69 | 4.81 | 4.93 | 5.06 | 5.18 | 5.3 | 5.42 | 5.55 | 5.67 | 5.79 | 5.91 | 6.04 | 6.16 | 6.28 | 6.4 | 6.53 | 6.65 | 6.77 | 6.89 | 7.02 | 7.14 | 7.26 | 7.38 | 7.51 | 7.63 | 7.75 | 7.87 | 8 | 8.12 | 8.24 | 8.36 | 8.49 | 8.61 | 8.73 | 8.85 | 8.98 | 9.1 | 9.22 | 9.34 | 9.47 | 9.59
19.98 | 4.69 | 4.81 | 4.93 | 5.05 | 5.18 | 5.3 | 5.42 | 5.54 | 5.67 | 5.79 | 5.91 | 6.03 | 6.16 | 6.28 | 6.4 | 6.52 | 6.65 | 6.77 | 6.89 | 7.01 | 7.14 | 7.26 | 7.38 | 7.5 | 7.63 | 7.75 | 7.87 | 7.99 | 8.12 | 8.24 | 8.36 | 8.48 | 8.61 | 8.73 | 8.85 | 8.97 | 9.1 | 9.22 | 9.34 | 9.46 | 9.59
19.99 | 4.69 | 4.81 | 4.93 | 5.05 | 5.18 | 5.3 | 5.42 | 5.54 | 5.67 | 5.79 | 5.91 | 6.03 | 6.16 | 6.28 | 6.4 | 6.52 | 6.65 | 6.77 | 6.89 | 7.01 | 7.14 | 7.26 | 7.38 | 7.5 | 7.63 | 7.75 | 7.87 | 7.99 | 8.11 | 8.24 | 8.36 | 8.48 | 8.6 | 8.73 | 8.85 | 8.97 | 9.09 | 9.22 | 9.34 | 9.46 | 9.58
20 | 4.69 | 4.81 | 4.93 | 5.05 | 5.18 | 5.3 | 5.42 | 5.54 | 5.66 | 5.79 | 5.91 | 6.03 | 6.15 | 6.28 | 6.4 | 6.52 | 6.64 | 6.77 | 6.89 | 7.01 | 7.13 | 7.26 | 7.38 | 7.5 | 7.62 | 7.75 | 7.87 | 7.99 | 8.11 | 8.24 | 8.36 | 8.48 | 8.6 | 8.73 | 8.85 | 8.97 | 9.09 | 9.21 | 9.34 | 9.46 | 9.58
20.01 | 4.68 | 4.81 | 4.93 | 5.05 | 5.17 | 5.3 | 5.42 | 5.54 | 5.66 | 5.79 | 5.91 | 6.03 | 6.15 | 6.28 | 6.4 | 6.52 | 6.64 | 6.77 | 6.89 | 7.01 | 7.13 | 7.26 | 7.38 | 7.5 | 7.62 | 7.74 | 7.87 | 7.99 | 8.11 | 8.23 | 8.36 | 8.48 | 8.6 | 8.72 | 8.85 | 8.97 | 9.09 | 9.21 | 9.34 | 9.46 | 9.58
20.02 | 4.68 | 4.81 | 4.93 | 5.05 | 5.17 | 5.3 | 5.42 | 5.54 | 5.66 | 5.79 | 5.91 | 6.03 | 6.15 | 6.27 | 6.4 | 6.52 | 6.64 | 6.76 | 6.89 | 7.01 | 7.13 | 7.25 | 7.38 | 7.5 | 7.62 | 7.74 | 7.87 | 7.99 | 8.11 | 8.23 | 8.35 | 8.48 | 8.6 | 8.72 | 8.84 | 8.97 | 9.09 | 9.21 | 9.33 | 9.46 | 9.58
20.03 | 4.68 | 4.8 | 4.93 | 5.05 | 5.17 | 5.29 | 5.42 | 5.54 | 5.66 | 5.78 | 5.91 | 6.03 | 6.15 | 6.27 | 6.4 | 6.52 | 6.64 | 6.76 | 6.89 | 7.01 | 7.13 | 7.25 | 7.37 | 7.5 | 7.62 | 7.74 | 7.86 | 7.99 | 8.11 | 8.23 | 8.35 | 8.48 | 8.6 | 8.72 | 8.84 | 8.96 | 9.09 | 9.21 | 9.33 | 9.45 | 9.58
20.04 | 4.68 | 4.8 | 4.93 | 5.05 | 5.17 | 5.29 | 5.42 | 5.54 | 5.66 | 5.78 | 5.9 | 6.03 | 6.15 | 6.27 | 6.39 | 6.52 | 6.64 | 6.76 | 6.88 | 7.01 | 7.13 | 7.25 | 7.37 | 7.5 | 7.62 | 7.74 | 7.86 | 7.98 | 8.11 | 8.23 | 8.35 | 8.47 | 8.6 | 8.72 | 8.84 | 8.96 | 9.09 | 9.21 | 9.33 | 9.45 | 9.57
20.05 | 4.68 | 4.8 | 4.93 | 5.05 | 5.17 | 5.29 | 5.41 | 5.54 | 5.66 | 5.78 | 5.9 | 6.03 | 6.15 | 6.27 | 6.39 | 6.52 | 6.64 | 6.76 | 6.88 | 7 | 7.13 | 7.25 | 7.37 | 7.49 | 7.62 | 7.74 | 7.86 | 7.98 | 8.11 | 8.23 | 8.35 | 8.47 | 8.59 | 8.72 | 8.84 | 8.96 | 9.08 | 9.21 | 9.33 | 9.45 | 9.57
20.06 | 4.68 | 4.8 | 4.92 | 5.05 | 5.17 | 5.29 | 5.41 | 5.54 | 5.66 | 5.78 | 5.9 | 6.02 | 6.15 | 6.27 | 6.39 | 6.51 | 6.64 | 6.76 | 6.88 | 7 | 7.13 | 7.25 | 7.37 | 7.49 | 7.61 | 7.74 | 7.86 | 7.98 | 8.1 | 8.23 | 8.35 | 8.47 | 8.59 | 8.71 | 8.84 | 8.96 | 9.08 | 9.2 | 9.33 | 9.45 | 9.57
20.07 | 4.68 | 4.8 | 4.92 | 5.05 | 5.17 | 5.29 | 5.41 | 5.53 | 5.66 | 5.78 | 5.9 | 6.02 | 6.15 | 6.27 | 6.39 | 6.51 | 6.63 | 6.76 | 6.88 | 7 | 7.12 | 7.25 | 7.37 | 7.49 | 7.61 | 7.74 | 7.86 | 7.98 | 8.1 | 8.22 | 8.35 | 8.47 | 8.59 | 8.71 | 8.84 | 8.96 | 9.08 | 9.2 | 9.32 | 9.45 | 9.57
20.08 | 4.68 | 4.8 | 4.92 | 5.04 | 5.17 | 5.29 | 5.41 | 5.53 | 5.66 | 5.78 | 5.9 | 6.02 | 6.14 | 6.27 | 6.39 | 6.51 | 6.63 | 6.76 | 6.88 | 7 | 7.12 | 7.24 | 7.37 | 7.49 | 7.61 | 7.73 | 7.86 | 7.98 | 8.1 | 8.22 | 8.34 | 8.47 | 8.59 | 8.71 | 8.83 | 8.96 | 9.08 | 9.2 | 9.32 | 9.44 | 9.57
20.09 | 4.68 | 4.8 | 4.92 | 5.04 | 5.17 | 5.29 | 5.41 | 5.53 | 5.65 | 5.78 | 5.9 | 6.02 | 6.14 | 6.27 | 6.39 | 6.51 | 6.63 | 6.75 | 6.88 | 7 | 7.12 | 7.24 | 7.37 | 7.49 | 7.61 | 7.73 | 7.85 | 7.98 | 8.1 | 8.22 | 8.34 | 8.47 | 8.59 | 8.71 | 8.83 | 8.95 | 9.08 | 9.2 | 9.32 | 9.44 | 9.56
20.1 | 4.68 | 4.8 | 4.92 | 5.04 | 5.16 | 5.29 | 5.41 | 5.53 | 5.65 | 5.78 | 5.9 | 6.02 | 6.14 | 6.26 | 6.39 | 6.51 | 6.63 | 6.75 | 6.88 | 7 | 7.12 | 7.24 | 7.36 | 7.49 | 7.61 | 7.73 | 7.85 | 7.97 | 8.1 | 8.22 | 8.34 | 8.46 | 8.59 | 8.71 | 8.83 | 8.95 | 9.07 | 9.2 | 9.32 | 9.44 | 9.56
20.11 | 4.67 | 4.8 | 4.92 | 5.04 | 5.16 | 5.29 | 5.41 | 5.53 | 5.65 | 5.77 | 5.9 | 6.02 | 6.14 | 6.26 | 6.39 | 6.51 | 6.63 | 6.75 | 6.87 | 7 | 7.12 | 7.24 | 7.36 | 7.48 | 7.61 | 7.73 | 7.85 | 7.97 | 8.1 | 8.22 | 8.34 | 8.46 | 8.58 | 8.71 | 8.83 | 8.95 | 9.07 | 9.19 | 9.32 | 9.44 | 9.56
20.12 | 4.67 | 4.8 | 4.92 | 5.04 | 5.16 | 5.28 | 5.41 | 5.53 | 5.65 | 5.77 | 5.9 | 6.02 | 6.14 | 6.26 | 6.38 | 6.51 | 6.63 | 6.75 | 6.87 | 6.99 | 7.12 | 7.24 | 7.36 | 7.48 | 7.61 | 7.73 | 7.85 | 7.97 | 8.09 | 8.22 | 8.34 | 8.46 | 8.58 | 8.7 | 8.83 | 8.95 | 9.07 | 9.19 | 9.32 | 9.44 | 9.56
20.13 | 4.67 | 4.79 | 4.92 | 5.04 | 5.16 | 5.28 | 5.41 | 5.53 | 5.65 | 5.77 | 5.89 | 6.02 | 6.14 | 6.26 | 6.38 | 6.5 | 6.63 | 6.75 | 6.87 | 6.99 | 7.12 | 7.24 | 7.36 | 7.48 | 7.6 | 7.73 | 7.85 | 7.97 | 8.09 | 8.21 | 8.34 | 8.46 | 8.58 | 8.7 | 8.82 | 8.95 | 9.07 | 9.19 | 9.31 | 9.44 | 9.56
20.14 | 4.67 | 4.79 | 4.92 | 5.04 | 5.16 | 5.28 | 5.4 | 5.53 | 5.65 | 5.77 | 5.89 | 6.01 | 6.14 | 6.26 | 6.38 | 6.5 | 6.63 | 6.75 | 6.87 | 6.99 | 7.11 | 7.24 | 7.36 | 7.48 | 7.6 | 7.72 | 7.85 | 7.97 | 8.09 | 8.21 | 8.33 | 8.46 | 8.58 | 8.7 | 8.82 | 8.95 | 9.07 | 9.19 | 9.31 | 9.43 | 9.56
20.15 | 4.67 | 4.79 | 4.91 | 5.04 | 5.16 | 5.28 | 5.4 | 5.53 | 5.65 | 5.77 | 5.89 | 6.01 | 6.14 | 6.26 | 6.38 | 6.5 | 6.62 | 6.75 | 6.87 | 6.99 | 7.11 | 7.23 | 7.36 | 7.48 | 7.6 | 7.72 | 7.84 | 7.97 | 8.09 | 8.21 | 8.33 | 8.46 | 8.58 | 8.7 | 8.82 | 8.94 | 9.07 | 9.19 | 9.31 | 9.43 | 9.55
20.16 | 4.67 | 4.79 | 4.91 | 5.04 | 5.16 | 5.28 | 5.4 | 5.52 | 5.65 | 5.77 | 5.89 | 6.01 | 6.13 | 6.26 | 6.38 | 6.5 | 6.62 | 6.74 | 6.87 | 6.99 | 7.11 | 7.23 | 7.36 | 7.48 | 7.6 | 7.72 | 7.84 | 7.97 | 8.09 | 8.21 | 8.33 | 8.45 | 8.58 | 8.7 | 8.82 | 8.94 | 9.06 | 9.19 | 9.31 | 9.43 | 9.55
20.17 | 4.67 | 4.79 | 4.91 | 5.03 | 5.16 | 5.28 | 5.4 | 5.52 | 5.65 | 5.77 | 5.89 | 6.01 | 6.13 | 6.26 | 6.38 | 6.5 | 6.62 | 6.74 | 6.87 | 6.99 | 7.11 | 7.23 | 7.35 | 7.48 | 7.6 | 7.72 | 7.84 | 7.96 | 8.09 | 8.21 | 8.33 | 8.45 | 8.57 | 8.7 | 8.82 | 8.94 | 9.06 | 9.18 | 9.31 | 9.43 | 9.55
20.18 | 4.67 | 4.79 | 4.91 | 5.03 | 5.16 | 5.28 | 5.4 | 5.52 | 5.64 | 5.77 | 5.89 | 6.01 | 6.13 | 6.25 | 6.38 | 6.5 | 6.62 | 6.74 | 6.86 | 6.99 | 7.11 | 7.23 | 7.35 | 7.47 | 7.6 | 7.72 | 7.84 | 7.96 | 8.08 | 8.21 | 8.33 | 8.45 | 8.57 | 8.69 | 8.82 | 8.94 | 9.06 | 9.18 | 9.3 | 9.43 | 9.55
20.19 | 4.67 | 4.79 | 4.91 | 5.03 | 5.15 | 5.28 | 5.4 | 5.52 | 5.64 | 5.76 | 5.89 | 6.01 | 6.13 | 6.25 | 6.37 | 6.5 | 6.62 | 6.74 | 6.86 | 6.98 | 7.11 | 7.23 | 7.35 | 7.47 | 7.59 | 7.72 | 7.84 | 7.96 | 8.08 | 8.2 | 8.33 | 8.45 | 8.57 | 8.69 | 8.81 | 8.94 | 9.06 | 9.18 | 9.3 | 9.42 | 9.55
20.2 | 4.67 | 4.79 | 4.91 | 5.03 | 5.15 | 5.28 | 5.4 | 5.52 | 5.64 | 5.76 | 5.89 | 6.01 | 6.13 | 6.25 | 6.37 | 6.5 | 6.62 | 6.74 | 6.86 | 6.98 | 7.11 | 7.23 | 7.35 | 7.47 | 7.59 | 7.72 | 7.84 | 7.96 | 8.08 | 8.2 | 8.32 | 8.45 | 8.57 | 8.69 | 8.81 | 8.93 | 9.06 | 9.18 | 9.3 | 9.42 | 9.54
20.21 | 4.66 | 4.79 | 4.91 | 5.03 | 5.15 | 5.27 | 5.4 | 5.52 | 5.64 | 5.76 | 5.88 | 6.01 | 6.13 | 6.25 | 6.37 | 6.49 | 6.62 | 6.74 | 6.86 | 6.98 | 7.1 | 7.23 | 7.35 | 7.47 | 7.59 | 7.71 | 7.84 | 7.96 | 8.08 | 8.2 | 8.32 | 8.45 | 8.57 | 8.69 | 8.81 | 8.93 | 9.05 | 9.18 | 9.3 | 9.42 | 9.54
20.22 | 4.66 | 4.79 | 4.91 | 5.03 | 5.15 | 5.27 | 5.4 | 5.52 | 5.64 | 5.76 | 5.88 | 6.01 | 6.13 | 6.25 | 6.37 | 6.49 | 6.61 | 6.74 | 6.86 | 6.98 | 7.1 | 7.22 | 7.35 | 7.47 | 7.59 | 7.71 | 7.83 | 7.96 | 8.08 | 8.2 | 8.32 | 8.44 | 8.57 | 8.69 | 8.81 | 8.93 | 9.05 | 9.17 | 9.3 | 9.42 | 9.54
20.23 | 4.66 | 4.78 | 4.91 | 5.03 | 5.15 | 5.27 | 5.39 | 5.52 | 5.64 | 5.76 | 5.88 | 6 | 6.13 | 6.25 | 6.37 | 6.49 | 6.61 | 6.74 | 6.86 | 6.98 | 7.1 | 7.22 | 7.34 | 7.47 | 7.59 | 7.71 | 7.83 | 7.95 | 8.08 | 8.2 | 8.32 | 8.44 | 8.56 | 8.69 | 8.81 | 8.93 | 9.05 | 9.17 | 9.29 | 9.42 | 9.54
20.24 | 4.66 | 4.78 | 4.91 | 5.03 | 5.15 | 5.27 | 5.39 | 5.51 | 5.64 | 5.76 | 5.88 | 6 | 6.12 | 6.25 | 6.37 | 6.49 | 6.61 | 6.73 | 6.86 | 6.98 | 7.1 | 7.22 | 7.34 | 7.47 | 7.59 | 7.71 | 7.83 | 7.95 | 8.07 | 8.2 | 8.32 | 8.44 | 8.56 | 8.68 | 8.81 | 8.93 | 9.05 | 9.17 | 9.29 | 9.41 | 9.54
20.25 | 4.66 | 4.78 | 4.9 | 5.03 | 5.15 | 5.27 | 5.39 | 5.51 | 5.64 | 5.76 | 5.88 | 6 | 6.12 | 6.25 | 6.37 | 6.49 | 6.61 | 6.73 | 6.85 | 6.98 | 7.1 | 7.22 | 7.34 | 7.46 | 7.59 | 7.71 | 7.83 | 7.95 | 8.07 | 8.19 | 8.32 | 8.44 | 8.56 | 8.68 | 8.8 | 8.93 | 9.05 | 9.17 | 9.29 | 9.41 | 9.53
20.26 | 4.66 | 4.78 | 4.9 | 5.03 | 5.15 | 5.27 | 5.39 | 5.51 | 5.63 | 5.76 | 5.88 | 6 | 6.12 | 6.24 | 6.37 | 6.49 | 6.61 | 6.73 | 6.85 | 6.97 | 7.1 | 7.22 | 7.34 | 7.46 | 7.58 | 7.71 | 7.83 | 7.95 | 8.07 | 8.19 | 8.31 | 8.44 | 8.56 | 8.68 | 8.8 | 8.92 | 9.05 | 9.17 | 9.29 | 9.41 | 9.53
20.27 | 4.66 | 4.78 | 4.9 | 5.02 | 5.15 | 5.27 | 5.39 | 5.51 | 5.63 | 5.76 | 5.88 | 6 | 6.12 | 6.24 | 6.36 | 6.49 | 6.61 | 6.73 | 6.85 | 6.97 | 7.1 | 7.22 | 7.34 | 7.46 | 7.58 | 7.7 | 7.83 | 7.95 | 8.07 | 8.19 | 8.31 | 8.44 | 8.56 | 8.68 | 8.8 | 8.92 | 9.04 | 9.17 | 9.29 | 9.41 | 9.53
20.28 | 4.66 | 4.78 | 4.9 | 5.02 | 5.14 | 5.27 | 5.39 | 5.51 | 5.63 | 5.75 | 5.88 | 6 | 6.12 | 6.24 | 6.36 | 6.48 | 6.61 | 6.73 | 6.85 | 6.97 | 7.09 | 7.22 | 7.34 | 7.46 | 7.58 | 7.7 | 7.82 | 7.95 | 8.07 | 8.19 | 8.31 | 8.43 | 8.56 | 8.68 | 8.8 | 8.92 | 9.04 | 9.16 | 9.29 | 9.41 | 9.53
20.29 | 4.66 | 4.78 | 4.9 | 5.02 | 5.14 | 5.27 | 5.39 | 5.51 | 5.63 | 5.75 | 5.87 | 6 | 6.12 | 6.24 | 6.36 | 6.48 | 6.61 | 6.73 | 6.85 | 6.97 | 7.09 | 7.21 | 7.34 | 7.46 | 7.58 | 7.7 | 7.82 | 7.94 | 8.07 | 8.19 | 8.31 | 8.43 | 8.55 | 8.68 | 8.8 | 8.92 | 9.04 | 9.16 | 9.28 | 9.41 | 9.53
20.3 | 4.66 | 4.78 | 4.9 | 5.02 | 5.14 | 5.26 | 5.39 | 5.51 | 5.63 | 5.75 | 5.87 | 6 | 6.12 | 6.24 | 6.36 | 6.48 | 6.6 | 6.73 | 6.85 | 6.97 | 7.09 | 7.21 | 7.33 | 7.46 | 7.58 | 7.7 | 7.82 | 7.94 | 8.06 | 8.19 | 8.31 | 8.43 | 8.55 | 8.67 | 8.8 | 8.92 | 9.04 | 9.16 | 9.28 | 9.4 | 9.53
20.31 | 4.65 | 4.78 | 4.9 | 5.02 | 5.14 | 5.26 | 5.39 | 5.51 | 5.63 | 5.75 | 5.87 | 5.99 | 6.12 | 6.24 | 6.36 | 6.48 | 6.6 | 6.72 | 6.85 | 6.97 | 7.09 | 7.21 | 7.33 | 7.45 | 7.58 | 7.7 | 7.82 | 7.94 | 8.06 | 8.18 | 8.31 | 8.43 | 8.55 | 8.67 | 8.79 | 8.92 | 9.04 | 9.16 | 9.28 | 9.4 | 9.52
20.32 | 4.65 | 4.78 | 4.9 | 5.02 | 5.14 | 5.26 | 5.38 | 5.51 | 5.63 | 5.75 | 5.87 | 5.99 | 6.11 | 6.24 | 6.36 | 6.48 | 6.6 | 6.72 | 6.84 | 6.97 | 7.09 | 7.21 | 7.33 | 7.45 | 7.57 | 7.7 | 7.82 | 7.94 | 8.06 | 8.18 | 8.3 | 8.43 | 8.55 | 8.67 | 8.79 | 8.91 | 9.03 | 9.16 | 9.28 | 9.4 | 9.52
20.33 | 4.65 | 4.77 | 4.9 | 5.02 | 5.14 | 5.26 | 5.38 | 5.5 | 5.63 | 5.75 | 5.87 | 5.99 | 6.11 | 6.23 | 6.36 | 6.48 | 6.6 | 6.72 | 6.84 | 6.96 | 7.09 | 7.21 | 7.33 | 7.45 | 7.57 | 7.7 | 7.82 | 7.94 | 8.06 | 8.18 | 8.3 | 8.42 | 8.55 | 8.67 | 8.79 | 8.91 | 9.03 | 9.15 | 9.28 | 9.4 | 9.52
20.34 | 4.65 | 4.77 | 4.9 | 5.02 | 5.14 | 5.26 | 5.38 | 5.5 | 5.63 | 5.75 | 5.87 | 5.99 | 6.11 | 6.23 | 6.36 | 6.48 | 6.6 | 6.72 | 6.84 | 6.96 | 7.09 | 7.21 | 7.33 | 7.45 | 7.57 | 7.69 | 7.82 | 7.94 | 8.06 | 8.18 | 8.3 | 8.42 | 8.54 | 8.67 | 8.79 | 8.91 | 9.03 | 9.15 | 9.27 | 9.4 | 9.52
20.35 | 4.65 | 4.77 | 4.89 | 5.02 | 5.14 | 5.26 | 5.38 | 5.5 | 5.62 | 5.75 | 5.87 | 5.99 | 6.11 | 6.23 | 6.35 | 6.48 | 6.6 | 6.72 | 6.84 | 6.96 | 7.08 | 7.21 | 7.33 | 7.45 | 7.57 | 7.69 | 7.81 | 7.94 | 8.06 | 8.18 | 8.3 | 8.42 | 8.54 | 8.66 | 8.79 | 8.91 | 9.03 | 9.15 | 9.27 | 9.39 | 9.52
20.36 | 4.65 | 4.77 | 4.89 | 5.01 | 5.14 | 5.26 | 5.38 | 5.5 | 5.62 | 5.74 | 5.87 | 5.99 | 6.11 | 6.23 | 6.35 | 6.47 | 6.6 | 6.72 | 6.84 | 6.96 | 7.08 | 7.2 | 7.33 | 7.45 | 7.57 | 7.69 | 7.81 | 7.93 | 8.06 | 8.18 | 8.3 | 8.42 | 8.54 | 8.66 | 8.78 | 8.91 | 9.03 | 9.15 | 9.27 | 9.39 | 9.51
20.37 | 4.65 | 4.77 | 4.89 | 5.01 | 5.14 | 5.26 | 5.38 | 5.5 | 5.62 | 5.74 | 5.86 | 5.99 | 6.11 | 6.23 | 6.35 | 6.47 | 6.59 | 6.72 | 6.84 | 6.96 | 7.08 | 7.2 | 7.32 | 7.45 | 7.57 | 7.69 | 7.81 | 7.93 | 8.05 | 8.18 | 8.3 | 8.42 | 8.54 | 8.66 | 8.78 | 8.9 | 9.03 | 9.15 | 9.27 | 9.39 | 9.51
20.38 | 4.65 | 4.77 | 4.89 | 5.01 | 5.13 | 5.26 | 5.38 | 5.5 | 5.62 | 5.74 | 5.86 | 5.99 | 6.11 | 6.23 | 6.35 | 6.47 | 6.59 | 6.71 | 6.84 | 6.96 | 7.08 | 7.2 | 7.32 | 7.44 | 7.57 | 7.69 | 7.81 | 7.93 | 8.05 | 8.17 | 8.3 | 8.42 | 8.54 | 8.66 | 8.78 | 8.9 | 9.02 | 9.15 | 9.27 | 9.39 | 9.51
20.39 | 4.65 | 4.77 | 4.89 | 5.01 | 5.13 | 5.25 | 5.38 | 5.5 | 5.62 | 5.74 | 5.86 | 5.98 | 6.11 | 6.23 | 6.35 | 6.47 | 6.59 | 6.71 | 6.83 | 6.96 | 7.08 | 7.2 | 7.32 | 7.44 | 7.56 | 7.69 | 7.81 | 7.93 | 8.05 | 8.17 | 8.29 | 8.41 | 8.54 | 8.66 | 8.78 | 8.9 | 9.02 | 9.14 | 9.27 | 9.39 | 9.51
20.4 | 4.65 | 4.77 | 4.89 | 5.01 | 5.13 | 5.25 | 5.38 | 5.5 | 5.62 | 5.74 | 5.86 | 5.98 | 6.1 | 6.23 | 6.35 | 6.47 | 6.59 | 6.71 | 6.83 | 6.96 | 7.08 | 7.2 | 7.32 | 7.44 | 7.56 | 7.68 | 7.81 | 7.93 | 8.05 | 8.17 | 8.29 | 8.41 | 8.53 | 8.66 | 8.78 | 8.9 | 9.02 | 9.14 | 9.26 | 9.39 | 9.51
20.41 | 4.64 | 4.77 | 4.89 | 5.01 | 5.13 | 5.25 | 5.37 | 5.5 | 5.62 | 5.74 | 5.86 | 5.98 | 6.1 | 6.22 | 6.35 | 6.47 | 6.59 | 6.71 | 6.83 | 6.95 | 7.08 | 7.2 | 7.32 | 7.44 | 7.56 | 7.68 | 7.8 | 7.93 | 8.05 | 8.17 | 8.29 | 8.41 | 8.53 | 8.65 | 8.78 | 8.9 | 9.02 | 9.14 | 9.26 | 9.38 | 9.5
20.42 | 4.64 | 4.77 | 4.89 | 5.01 | 5.13 | 5.25 | 5.37 | 5.49 | 5.62 | 5.74 | 5.86 | 5.98 | 6.1 | 6.22 | 6.34 | 6.47 | 6.59 | 6.71 | 6.83 | 6.95 | 7.07 | 7.2 | 7.32 | 7.44 | 7.56 | 7.68 | 7.8 | 7.92 | 8.05 | 8.17 | 8.29 | 8.41 | 8.53 | 8.65 | 8.77 | 8.9 | 9.02 | 9.14 | 9.26 | 9.38 | 9.5
20.43 | 4.64 | 4.76 | 4.89 | 5.01 | 5.13 | 5.25 | 5.37 | 5.49 | 5.61 | 5.74 | 5.86 | 5.98 | 6.1 | 6.22 | 6.34 | 6.47 | 6.59 | 6.71 | 6.83 | 6.95 | 7.07 | 7.19 | 7.32 | 7.44 | 7.56 | 7.68 | 7.8 | 7.92 | 8.04 | 8.17 | 8.29 | 8.41 | 8.53 | 8.65 | 8.77 | 8.89 | 9.02 | 9.14 | 9.26 | 9.38 | 9.5
20.44 | 4.64 | 4.76 | 4.88 | 5.01 | 5.13 | 5.25 | 5.37 | 5.49 | 5.61 | 5.74 | 5.86 | 5.98 | 6.1 | 6.22 | 6.34 | 6.46 | 6.59 | 6.71 | 6.83 | 6.95 | 7.07 | 7.19 | 7.31 | 7.44 | 7.56 | 7.68 | 7.8 | 7.92 | 8.04 | 8.16 | 8.29 | 8.41 | 8.53 | 8.65 | 8.77 | 8.89 | 9.01 | 9.13 | 9.26 | 9.38 | 9.5
20.45 | 4.64 | 4.76 | 4.88 | 5.01 | 5.13 | 5.25 | 5.37 | 5.49 | 5.61 | 5.73 | 5.86 | 5.98 | 6.1 | 6.22 | 6.34 | 6.46 | 6.58 | 6.71 | 6.83 | 6.95 | 7.07 | 7.19 | 7.31 | 7.43 | 7.56 | 7.68 | 7.8 | 7.92 | 8.04 | 8.16 | 8.28 | 8.4 | 8.53 | 8.65 | 8.77 | 8.89 | 9.01 | 9.13 | 9.25 | 9.38 | 9.5
20.46 | 4.64 | 4.76 | 4.88 | 5 | 5.13 | 5.25 | 5.37 | 5.49 | 5.61 | 5.73 | 5.85 | 5.98 | 6.1 | 6.22 | 6.34 | 6.46 | 6.58 | 6.7 | 6.83 | 6.95 | 7.07 | 7.19 | 7.31 | 7.43 | 7.55 | 7.68 | 7.8 | 7.92 | 8.04 | 8.16 | 8.28 | 8.4 | 8.52 | 8.65 | 8.77 | 8.89 | 9.01 | 9.13 | 9.25 | 9.37 | 9.5
20.47 | 4.64 | 4.76 | 4.88 | 5 | 5.12 | 5.25 | 5.37 | 5.49 | 5.61 | 5.73 | 5.85 | 5.97 | 6.1 | 6.22 | 6.34 | 6.46 | 6.58 | 6.7 | 6.82 | 6.95 | 7.07 | 7.19 | 7.31 | 7.43 | 7.55 | 7.67 | 7.79 | 7.92 | 8.04 | 8.16 | 8.28 | 8.4 | 8.52 | 8.64 | 8.77 | 8.89 | 9.01 | 9.13 | 9.25 | 9.37 | 9.49
20.48 | 4.64 | 4.76 | 4.88 | 5 | 5.12 | 5.24 | 5.37 | 5.49 | 5.61 | 5.73 | 5.85 | 5.97 | 6.09 | 6.22 | 6.34 | 6.46 | 6.58 | 6.7 | 6.82 | 6.94 | 7.07 | 7.19 | 7.31 | 7.43 | 7.55 | 7.67 | 7.79 | 7.91 | 8.04 | 8.16 | 8.28 | 8.4 | 8.52 | 8.64 | 8.76 | 8.89 | 9.01 | 9.13 | 9.25 | 9.37 | 9.49
20.49 | 4.64 | 4.76 | 4.88 | 5 | 5.12 | 5.24 | 5.37 | 5.49 | 5.61 | 5.73 | 5.85 | 5.97 | 6.09 | 6.21 | 6.34 | 6.46 | 6.58 | 6.7 | 6.82 | 6.94 | 7.06 | 7.19 | 7.31 | 7.43 | 7.55 | 7.67 | 7.79 | 7.91 | 8.03 | 8.16 | 8.28 | 8.4 | 8.52 | 8.64 | 8.76 | 8.88 | 9 | 9.13 | 9.25 | 9.37 | 9.49
20.5 | 4.64 | 4.76 | 4.88 | 5 | 5.12 | 5.24 | 5.36 | 5.49 | 5.61 | 5.73 | 5.85 | 5.97 | 6.09 | 6.21 | 6.33 | 6.46 | 6.58 | 6.7 | 6.82 | 6.94 | 7.06 | 7.18 | 7.31 | 7.43 | 7.55 | 7.67 | 7.79 | 7.91 | 8.03 | 8.15 | 8.28 | 8.4 | 8.52 | 8.64 | 8.76 | 8.88 | 9 | 9.12 | 9.25 | 9.37 | 9.49
20.51 | 4.64 | 4.76 | 4.88 | 5 | 5.12 | 5.24 | 5.36 | 5.48 | 5.61 | 5.73 | 5.85 | 5.97 | 6.09 | 6.21 | 6.33 | 6.45 | 6.58 | 6.7 | 6.82 | 6.94 | 7.06 | 7.18 | 7.3 | 7.42 | 7.55 | 7.67 | 7.79 | 7.91 | 8.03 | 8.15 | 8.27 | 8.39 | 8.52 | 8.64 | 8.76 | 8.88 | 9 | 9.12 | 9.24 | 9.36 | 9.49
20.52 | 4.63 | 4.76 | 4.88 | 5 | 5.12 | 5.24 | 5.36 | 5.48 | 5.6 | 5.73 | 5.85 | 5.97 | 6.09 | 6.21 | 6.33 | 6.45 | 6.57 | 6.7 | 6.82 | 6.94 | 7.06 | 7.18 | 7.3 | 7.42 | 7.54 | 7.67 | 7.79 | 7.91 | 8.03 | 8.15 | 8.27 | 8.39 | 8.51 | 8.64 | 8.76 | 8.88 | 9 | 9.12 | 9.24 | 9.36 | 9.48
20.53 | 4.63 | 4.75 | 4.88 | 5 | 5.12 | 5.24 | 5.36 | 5.48 | 5.6 | 5.72 | 5.85 | 5.97 | 6.09 | 6.21 | 6.33 | 6.45 | 6.57 | 6.69 | 6.82 | 6.94 | 7.06 | 7.18 | 7.3 | 7.42 | 7.54 | 7.66 | 7.79 | 7.91 | 8.03 | 8.15 | 8.27 | 8.39 | 8.51 | 8.63 | 8.76 | 8.88 | 9 | 9.12 | 9.24 | 9.36 | 9.48
20.54 | 4.63 | 4.75 | 4.87 | 5 | 5.12 | 5.24 | 5.36 | 5.48 | 5.6 | 5.72 | 5.84 | 5.97 | 6.09 | 6.21 | 6.33 | 6.45 | 6.57 | 6.69 | 6.81 | 6.94 | 7.06 | 7.18 | 7.3 | 7.42 | 7.54 | 7.66 | 7.78 | 7.91 | 8.03 | 8.15 | 8.27 | 8.39 | 8.51 | 8.63 | 8.75 | 8.87 | 9 | 9.12 | 9.24 | 9.36 | 9.48
20.55 | 4.63 | 4.75 | 4.87 | 4.99 | 5.12 | 5.24 | 5.36 | 5.48 | 5.6 | 5.72 | 5.84 | 5.96 | 6.09 | 6.21 | 6.33 | 6.45 | 6.57 | 6.69 | 6.81 | 6.93 | 7.06 | 7.18 | 7.3 | 7.42 | 7.54 | 7.66 | 7.78 | 7.9 | 8.02 | 8.15 | 8.27 | 8.39 | 8.51 | 8.63 | 8.75 | 8.87 | 8.99 | 9.12 | 9.24 | 9.36 | 9.48
20.56 | 4.63 | 4.75 | 4.87 | 4.99 | 5.11 | 5.24 | 5.36 | 5.48 | 5.6 | 5.72 | 5.84 | 5.96 | 6.08 | 6.21 | 6.33 | 6.45 | 6.57 | 6.69 | 6.81 | 6.93 | 7.05 | 7.18 | 7.3 | 7.42 | 7.54 | 7.66 | 7.78 | 7.9 | 8.02 | 8.14 | 8.27 | 8.39 | 8.51 | 8.63 | 8.75 | 8.87 | 8.99 | 9.11 | 9.23 | 9.36 | 9.48
20.57 | 4.63 | 4.75 | 4.87 | 4.99 | 5.11 | 5.24 | 5.36 | 5.48 | 5.6 | 5.72 | 5.84 | 5.96 | 6.08 | 6.2 | 6.33 | 6.45 | 6.57 | 6.69 | 6.81 | 6.93 | 7.05 | 7.17 | 7.29 | 7.42 | 7.54 | 7.66 | 7.78 | 7.9 | 8.02 | 8.14 | 8.26 | 8.38 | 8.51 | 8.63 | 8.75 | 8.87 | 8.99 | 9.11 | 9.23 | 9.35 | 9.47
20.58 | 4.63 | 4.75 | 4.87 | 4.99 | 5.11 | 5.23 | 5.36 | 5.48 | 5.6 | 5.72 | 5.84 | 5.96 | 6.08 | 6.2 | 6.32 | 6.45 | 6.57 | 6.69 | 6.81 | 6.93 | 7.05 | 7.17 | 7.29 | 7.41 | 7.54 | 7.66 | 7.78 | 7.9 | 8.02 | 8.14 | 8.26 | 8.38 | 8.5 | 8.63 | 8.75 | 8.87 | 8.99 | 9.11 | 9.23 | 9.35 | 9.47
20.59 | 4.63 | 4.75 | 4.87 | 4.99 | 5.11 | 5.23 | 5.35 | 5.48 | 5.6 | 5.72 | 5.84 | 5.96 | 6.08 | 6.2 | 6.32 | 6.44 | 6.57 | 6.69 | 6.81 | 6.93 | 7.05 | 7.17 | 7.29 | 7.41 | 7.53 | 7.66 | 7.78 | 7.9 | 8.02 | 8.14 | 8.26 | 8.38 | 8.5 | 8.62 | 8.74 | 8.87 | 8.99 | 9.11 | 9.23 | 9.35 | 9.47
20.6 | 4.63 | 4.75 | 4.87 | 4.99 | 5.11 | 5.23 | 5.35 | 5.47 | 5.6 | 5.72 | 5.84 | 5.96 | 6.08 | 6.2 | 6.32 | 6.44 | 6.56 | 6.68 | 6.81 | 6.93 | 7.05 | 7.17 | 7.29 | 7.41 | 7.53 | 7.65 | 7.77 | 7.9 | 8.02 | 8.14 | 8.26 | 8.38 | 8.5 | 8.62 | 8.74 | 8.86 | 8.99 | 9.11 | 9.23 | 9.35 | 9.47
20.61 | 4.63 | 4.75 | 4.87 | 4.99 | 5.11 | 5.23 | 5.35 | 5.47 | 5.59 | 5.72 | 5.84 | 5.96 | 6.08 | 6.2 | 6.32 | 6.44 | 6.56 | 6.68 | 6.8 | 6.93 | 7.05 | 7.17 | 7.29 | 7.41 | 7.53 | 7.65 | 7.77 | 7.89 | 8.02 | 8.14 | 8.26 | 8.38 | 8.5 | 8.62 | 8.74 | 8.86 | 8.98 | 9.1 | 9.23 | 9.35 | 9.47
20.62 | 4.62 | 4.75 | 4.87 | 4.99 | 5.11 | 5.23 | 5.35 | 5.47 | 5.59 | 5.71 | 5.83 | 5.96 | 6.08 | 6.2 | 6.32 | 6.44 | 6.56 | 6.68 | 6.8 | 6.92 | 7.05 | 7.17 | 7.29 | 7.41 | 7.53 | 7.65 | 7.77 | 7.89 | 8.01 | 8.13 | 8.26 | 8.38 | 8.5 | 8.62 | 8.74 | 8.86 | 8.98 | 9.1 | 9.22 | 9.34 | 9.47
20.63 | 4.62 | 4.74 | 4.87 | 4.99 | 5.11 | 5.23 | 5.35 | 5.47 | 5.59 | 5.71 | 5.83 | 5.95 | 6.08 | 6.2 | 6.32 | 6.44 | 6.56 | 6.68 | 6.8 | 6.92 | 7.04 | 7.17 | 7.29 | 7.41 | 7.53 | 7.65 | 7.77 | 7.89 | 8.01 | 8.13 | 8.25 | 8.38 | 8.5 | 8.62 | 8.74 | 8.86 | 8.98 | 9.1 | 9.22 | 9.34 | 9.46
20.64 | 4.62 | 4.74 | 4.86 | 4.99 | 5.11 | 5.23 | 5.35 | 5.47 | 5.59 | 5.71 | 5.83 | 5.95 | 6.07 | 6.2 | 6.32 | 6.44 | 6.56 | 6.68 | 6.8 | 6.92 | 7.04 | 7.16 | 7.28 | 7.41 | 7.53 | 7.65 | 7.77 | 7.89 | 8.01 | 8.13 | 8.25 | 8.37 | 8.49 | 8.62 | 8.74 | 8.86 | 8.98 | 9.1 | 9.22 | 9.34 | 9.46
20.65 | 4.62 | 4.74 | 4.86 | 4.98 | 5.11 | 5.23 | 5.35 | 5.47 | 5.59 | 5.71 | 5.83 | 5.95 | 6.07 | 6.19 | 6.32 | 6.44 | 6.56 | 6.68 | 6.8 | 6.92 | 7.04 | 7.16 | 7.28 | 7.4 | 7.53 | 7.65 | 7.77 | 7.89 | 8.01 | 8.13 | 8.25 | 8.37 | 8.49 | 8.61 | 8.73 | 8.86 | 8.98 | 9.1 | 9.22 | 9.34 | 9.46
20.66 | 4.62 | 4.74 | 4.86 | 4.98 | 5.1 | 5.23 | 5.35 | 5.47 | 5.59 | 5.71 | 5.83 | 5.95 | 6.07 | 6.19 | 6.31 | 6.43 | 6.56 | 6.68 | 6.8 | 6.92 | 7.04 | 7.16 | 7.28 | 7.4 | 7.52 | 7.64 | 7.77 | 7.89 | 8.01 | 8.13 | 8.25 | 8.37 | 8.49 | 8.61 | 8.73 | 8.85 | 8.97 | 9.1 | 9.22 | 9.34 | 9.46
20.67 | 4.62 | 4.74 | 4.86 | 4.98 | 5.1 | 5.22 | 5.35 | 5.47 | 5.59 | 5.71 | 5.83 | 5.95 | 6.07 | 6.19 | 6.31 | 6.43 | 6.55 | 6.68 | 6.8 | 6.92 | 7.04 | 7.16 | 7.28 | 7.4 | 7.52 | 7.64 | 7.76 | 7.88 | 8.01 | 8.13 | 8.25 | 8.37 | 8.49 | 8.61 | 8.73 | 8.85 | 8.97 | 9.09 | 9.21 | 9.34 | 9.46
20.68 | 4.62 | 4.74 | 4.86 | 4.98 | 5.1 | 5.22 | 5.34 | 5.46 | 5.59 | 5.71 | 5.83 | 5.95 | 6.07 | 6.19 | 6.31 | 6.43 | 6.55 | 6.67 | 6.8 | 6.92 | 7.04 | 7.16 | 7.28 | 7.4 | 7.52 | 7.64 | 7.76 | 7.88 | 8 | 8.12 | 8.25 | 8.37 | 8.49 | 8.61 | 8.73 | 8.85 | 8.97 | 9.09 | 9.21 | 9.33 | 9.45
20.69 | 4.62 | 4.74 | 4.86 | 4.98 | 5.1 | 5.22 | 5.34 | 5.46 | 5.58 | 5.71 | 5.83 | 5.95 | 6.07 | 6.19 | 6.31 | 6.43 | 6.55 | 6.67 | 6.79 | 6.91 | 7.04 | 7.16 | 7.28 | 7.4 | 7.52 | 7.64 | 7.76 | 7.88 | 8 | 8.12 | 8.24 | 8.37 | 8.49 | 8.61 | 8.73 | 8.85 | 8.97 | 9.09 | 9.21 | 9.33 | 9.45
20.7 | 4.62 | 4.74 | 4.86 | 4.98 | 5.1 | 5.22 | 5.34 | 5.46 | 5.58 | 5.7 | 5.83 | 5.95 | 6.07 | 6.19 | 6.31 | 6.43 | 6.55 | 6.67 | 6.79 | 6.91 | 7.03 | 7.15 | 7.28 | 7.4 | 7.52 | 7.64 | 7.76 | 7.88 | 8 | 8.12 | 8.24 | 8.36 | 8.48 | 8.61 | 8.73 | 8.85 | 8.97 | 9.09 | 9.21 | 9.33 | 9.45
20.71 | 4.62 | 4.74 | 4.86 | 4.98 | 5.1 | 5.22 | 5.34 | 5.46 | 5.58 | 5.7 | 5.82 | 5.95 | 6.07 | 6.19 | 6.31 | 6.43 | 6.55 | 6.67 | 6.79 | 6.91 | 7.03 | 7.15 | 7.27 | 7.4 | 7.52 | 7.64 | 7.76 | 7.88 | 8 | 8.12 | 8.24 | 8.36 | 8.48 | 8.6 | 8.72 | 8.85 | 8.97 | 9.09 | 9.21 | 9.33 | 9.45
20.72 | 4.61 | 4.74 | 4.86 | 4.98 | 5.1 | 5.22 | 5.34 | 5.46 | 5.58 | 5.7 | 5.82 | 5.94 | 6.06 | 6.19 | 6.31 | 6.43 | 6.55 | 6.67 | 6.79 | 6.91 | 7.03 | 7.15 | 7.27 | 7.39 | 7.51 | 7.64 | 7.76 | 7.88 | 8 | 8.12 | 8.24 | 8.36 | 8.48 | 8.6 | 8.72 | 8.84 | 8.96 | 9.08 | 9.21 | 9.33 | 9.45
20.73 | 4.61 | 4.73 | 4.86 | 4.98 | 5.1 | 5.22 | 5.34 | 5.46 | 5.58 | 5.7 | 5.82 | 5.94 | 6.06 | 6.18 | 6.31 | 6.43 | 6.55 | 6.67 | 6.79 | 6.91 | 7.03 | 7.15 | 7.27 | 7.39 | 7.51 | 7.63 | 7.75 | 7.88 | 8 | 8.12 | 8.24 | 8.36 | 8.48 | 8.6 | 8.72 | 8.84 | 8.96 | 9.08 | 9.2 | 9.32 | 9.45
20.74 | 4.61 | 4.73 | 4.85 | 4.98 | 5.1 | 5.22 | 5.34 | 5.46 | 5.58 | 5.7 | 5.82 | 5.94 | 6.06 | 6.18 | 6.3 | 6.42 | 6.55 | 6.67 | 6.79 | 6.91 | 7.03 | 7.15 | 7.27 | 7.39 | 7.51 | 7.63 | 7.75 | 7.87 | 7.99 | 8.12 | 8.24 | 8.36 | 8.48 | 8.6 | 8.72 | 8.84 | 8.96 | 9.08 | 9.2 | 9.32 | 9.44
20.75 | 4.61 | 4.73 | 4.85 | 4.97 | 5.09 | 5.22 | 5.34 | 5.46 | 5.58 | 5.7 | 5.82 | 5.94 | 6.06 | 6.18 | 6.3 | 6.42 | 6.54 | 6.66 | 6.79 | 6.91 | 7.03 | 7.15 | 7.27 | 7.39 | 7.51 | 7.63 | 7.75 | 7.87 | 7.99 | 8.11 | 8.23 | 8.36 | 8.48 | 8.6 | 8.72 | 8.84 | 8.96 | 9.08 | 9.2 | 9.32 | 9.44
20.76 | 4.61 | 4.73 | 4.85 | 4.97 | 5.09 | 5.21 | 5.34 | 5.46 | 5.58 | 5.7 | 5.82 | 5.94 | 6.06 | 6.18 | 6.3 | 6.42 | 6.54 | 6.66 | 6.78 | 6.9 | 7.03 | 7.15 | 7.27 | 7.39 | 7.51 | 7.63 | 7.75 | 7.87 | 7.99 | 8.11 | 8.23 | 8.35 | 8.47 | 8.59 | 8.72 | 8.84 | 8.96 | 9.08 | 9.2 | 9.32 | 9.44
20.77 | 4.61 | 4.73 | 4.85 | 4.97 | 5.09 | 5.21 | 5.33 | 5.45 | 5.58 | 5.7 | 5.82 | 5.94 | 6.06 | 6.18 | 6.3 | 6.42 | 6.54 | 6.66 | 6.78 | 6.9 | 7.02 | 7.14 | 7.27 | 7.39 | 7.51 | 7.63 | 7.75 | 7.87 | 7.99 | 8.11 | 8.23 | 8.35 | 8.47 | 8.59 | 8.71 | 8.83 | 8.96 | 9.08 | 9.2 | 9.32 | 9.44
20.78 | 4.61 | 4.73 | 4.85 | 4.97 | 5.09 | 5.21 | 5.33 | 5.45 | 5.57 | 5.7 | 5.82 | 5.94 | 6.06 | 6.18 | 6.3 | 6.42 | 6.54 | 6.66 | 6.78 | 6.9 | 7.02 | 7.14 | 7.26 | 7.38 | 7.51 | 7.63 | 7.75 | 7.87 | 7.99 | 8.11 | 8.23 | 8.35 | 8.47 | 8.59 | 8.71 | 8.83 | 8.95 | 9.07 | 9.19 | 9.32 | 9.44
20.79 | 4.61 | 4.73 | 4.85 | 4.97 | 5.09 | 5.21 | 5.33 | 5.45 | 5.57 | 5.69 | 5.81 | 5.94 | 6.06 | 6.18 | 6.3 | 6.42 | 6.54 | 6.66 | 6.78 | 6.9 | 7.02 | 7.14 | 7.26 | 7.38 | 7.5 | 7.62 | 7.75 | 7.87 | 7.99 | 8.11 | 8.23 | 8.35 | 8.47 | 8.59 | 8.71 | 8.83 | 8.95 | 9.07 | 9.19 | 9.31 | 9.43
20.8 | 4.61 | 4.73 | 4.85 | 4.97 | 5.09 | 5.21 | 5.33 | 5.45 | 5.57 | 5.69 | 5.81 | 5.93 | 6.05 | 6.18 | 6.3 | 6.42 | 6.54 | 6.66 | 6.78 | 6.9 | 7.02 | 7.14 | 7.26 | 7.38 | 7.5 | 7.62 | 7.74 | 7.86 | 7.99 | 8.11 | 8.23 | 8.35 | 8.47 | 8.59 | 8.71 | 8.83 | 8.95 | 9.07 | 9.19 | 9.31 | 9.43
20.81 | 4.61 | 4.73 | 4.85 | 4.97 | 5.09 | 5.21 | 5.33 | 5.45 | 5.57 | 5.69 | 5.81 | 5.93 | 6.05 | 6.17 | 6.29 | 6.42 | 6.54 | 6.66 | 6.78 | 6.9 | 7.02 | 7.14 | 7.26 | 7.38 | 7.5 | 7.62 | 7.74 | 7.86 | 7.98 | 8.1 | 8.22 | 8.35 | 8.47 | 8.59 | 8.71 | 8.83 | 8.95 | 9.07 | 9.19 | 9.31 | 9.43
20.82 | 4.6 | 4.73 | 4.85 | 4.97 | 5.09 | 5.21 | 5.33 | 5.45 | 5.57 | 5.69 | 5.81 | 5.93 | 6.05 | 6.17 | 6.29 | 6.41 | 6.53 | 6.66 | 6.78 | 6.9 | 7.02 | 7.14 | 7.26 | 7.38 | 7.5 | 7.62 | 7.74 | 7.86 | 7.98 | 8.1 | 8.22 | 8.34 | 8.46 | 8.58 | 8.71 | 8.83 | 8.95 | 9.07 | 9.19 | 9.31 | 9.43
20.83 | 4.6 | 4.72 | 4.85 | 4.97 | 5.09 | 5.21 | 5.33 | 5.45 | 5.57 | 5.69 | 5.81 | 5.93 | 6.05 | 6.17 | 6.29 | 6.41 | 6.53 | 6.65 | 6.77 | 6.9 | 7.02 | 7.14 | 7.26 | 7.38 | 7.5 | 7.62 | 7.74 | 7.86 | 7.98 | 8.1 | 8.22 | 8.34 | 8.46 | 8.58 | 8.7 | 8.82 | 8.94 | 9.07 | 9.19 | 9.31 | 9.43
20.84 | 4.6 | 4.72 | 4.84 | 4.96 | 5.09 | 5.21 | 5.33 | 5.45 | 5.57 | 5.69 | 5.81 | 5.93 | 6.05 | 6.17 | 6.29 | 6.41 | 6.53 | 6.65 | 6.77 | 6.89 | 7.01 | 7.13 | 7.26 | 7.38 | 7.5 | 7.62 | 7.74 | 7.86 | 7.98 | 8.1 | 8.22 | 8.34 | 8.46 | 8.58 | 8.7 | 8.82 | 8.94 | 9.06 | 9.18 | 9.3 | 9.43
20.85 | 4.6 | 4.72 | 4.84 | 4.96 | 5.08 | 5.2 | 5.33 | 5.45 | 5.57 | 5.69 | 5.81 | 5.93 | 6.05 | 6.17 | 6.29 | 6.41 | 6.53 | 6.65 | 6.77 | 6.89 | 7.01 | 7.13 | 7.25 | 7.37 | 7.5 | 7.62 | 7.74 | 7.86 | 7.98 | 8.1 | 8.22 | 8.34 | 8.46 | 8.58 | 8.7 | 8.82 | 8.94 | 9.06 | 9.18 | 9.3 | 9.42
20.86 | 4.6 | 4.72 | 4.84 | 4.96 | 5.08 | 5.2 | 5.32 | 5.44 | 5.57 | 5.69 | 5.81 | 5.93 | 6.05 | 6.17 | 6.29 | 6.41 | 6.53 | 6.65 | 6.77 | 6.89 | 7.01 | 7.13 | 7.25 | 7.37 | 7.49 | 7.61 | 7.73 | 7.86 | 7.98 | 8.1 | 8.22 | 8.34 | 8.46 | 8.58 | 8.7 | 8.82 | 8.94 | 9.06 | 9.18 | 9.3 | 9.42
20.87 | 4.6 | 4.72 | 4.84 | 4.96 | 5.08 | 5.2 | 5.32 | 5.44 | 5.56 | 5.68 | 5.81 | 5.93 | 6.05 | 6.17 | 6.29 | 6.41 | 6.53 | 6.65 | 6.77 | 6.89 | 7.01 | 7.13 | 7.25 | 7.37 | 7.49 | 7.61 | 7.73 | 7.85 | 7.97 | 8.09 | 8.21 | 8.34 | 8.46 | 8.58 | 8.7 | 8.82 | 8.94 | 9.06 | 9.18 | 9.3 | 9.42
20.88 | 4.6 | 4.72 | 4.84 | 4.96 | 5.08 | 5.2 | 5.32 | 5.44 | 5.56 | 5.68 | 5.8 | 5.92 | 6.04 | 6.17 | 6.29 | 6.41 | 6.53 | 6.65 | 6.77 | 6.89 | 7.01 | 7.13 | 7.25 | 7.37 | 7.49 | 7.61 | 7.73 | 7.85 | 7.97 | 8.09 | 8.21 | 8.33 | 8.45 | 8.57 | 8.7 | 8.82 | 8.94 | 9.06 | 9.18 | 9.3 | 9.42
20.89 | 4.6 | 4.72 | 4.84 | 4.96 | 5.08 | 5.2 | 5.32 | 5.44 | 5.56 | 5.68 | 5.8 | 5.92 | 6.04 | 6.16 | 6.28 | 6.41 | 6.53 | 6.65 | 6.77 | 6.89 | 7.01 | 7.13 | 7.25 | 7.37 | 7.49 | 7.61 | 7.73 | 7.85 | 7.97 | 8.09 | 8.21 | 8.33 | 8.45 | 8.57 | 8.69 | 8.81 | 8.93 | 9.05 | 9.18 | 9.3 | 9.42
20.9 | 4.6 | 4.72 | 4.84 | 4.96 | 5.08 | 5.2 | 5.32 | 5.44 | 5.56 | 5.68 | 5.8 | 5.92 | 6.04 | 6.16 | 6.28 | 6.4 | 6.52 | 6.64 | 6.77 | 6.89 | 7.01 | 7.13 | 7.25 | 7.37 | 7.49 | 7.61 | 7.73 | 7.85 | 7.97 | 8.09 | 8.21 | 8.33 | 8.45 | 8.57 | 8.69 | 8.81 | 8.93 | 9.05 | 9.17 | 9.29 | 9.41
20.91 | 4.6 | 4.72 | 4.84 | 4.96 | 5.08 | 5.2 | 5.32 | 5.44 | 5.56 | 5.68 | 5.8 | 5.92 | 6.04 | 6.16 | 6.28 | 6.4 | 6.52 | 6.64 | 6.76 | 6.88 | 7 | 7.12 | 7.25 | 7.37 | 7.49 | 7.61 | 7.73 | 7.85 | 7.97 | 8.09 | 8.21 | 8.33 | 8.45 | 8.57 | 8.69 | 8.81 | 8.93 | 9.05 | 9.17 | 9.29 | 9.41
20.92 | 4.6 | 4.72 | 4.84 | 4.96 | 5.08 | 5.2 | 5.32 | 5.44 | 5.56 | 5.68 | 5.8 | 5.92 | 6.04 | 6.16 | 6.28 | 6.4 | 6.52 | 6.64 | 6.76 | 6.88 | 7 | 7.12 | 7.24 | 7.36 | 7.48 | 7.61 | 7.73 | 7.85 | 7.97 | 8.09 | 8.21 | 8.33 | 8.45 | 8.57 | 8.69 | 8.81 | 8.93 | 9.05 | 9.17 | 9.29 | 9.41
20.93 | 4.59 | 4.71 | 4.83 | 4.96 | 5.08 | 5.2 | 5.32 | 5.44 | 5.56 | 5.68 | 5.8 | 5.92 | 6.04 | 6.16 | 6.28 | 6.4 | 6.52 | 6.64 | 6.76 | 6.88 | 7 | 7.12 | 7.24 | 7.36 | 7.48 | 7.6 | 7.72 | 7.84 | 7.96 | 8.08 | 8.21 | 8.33 | 8.45 | 8.57 | 8.69 | 8.81 | 8.93 | 9.05 | 9.17 | 9.29 | 9.41
20.94 | 4.59 | 4.71 | 4.83 | 4.95 | 5.07 | 5.2 | 5.32 | 5.44 | 5.56 | 5.68 | 5.8 | 5.92 | 6.04 | 6.16 | 6.28 | 6.4 | 6.52 | 6.64 | 6.76 | 6.88 | 7 | 7.12 | 7.24 | 7.36 | 7.48 | 7.6 | 7.72 | 7.84 | 7.96 | 8.08 | 8.2 | 8.32 | 8.44 | 8.56 | 8.68 | 8.81 | 8.93 | 9.05 | 9.17 | 9.29 | 9.41
20.95 | 4.59 | 4.71 | 4.83 | 4.95 | 5.07 | 5.19 | 5.31 | 5.43 | 5.56 | 5.68 | 5.8 | 5.92 | 6.04 | 6.16 | 6.28 | 6.4 | 6.52 | 6.64 | 6.76 | 6.88 | 7 | 7.12 | 7.24 | 7.36 | 7.48 | 7.6 | 7.72 | 7.84 | 7.96 | 8.08 | 8.2 | 8.32 | 8.44 | 8.56 | 8.68 | 8.8 | 8.92 | 9.04 | 9.16 | 9.28 | 9.4
20.96 | 4.59 | 4.71 | 4.83 | 4.95 | 5.07 | 5.19 | 5.31 | 5.43 | 5.55 | 5.67 | 5.79 | 5.91 | 6.04 | 6.16 | 6.28 | 6.4 | 6.52 | 6.64 | 6.76 | 6.88 | 7 | 7.12 | 7.24 | 7.36 | 7.48 | 7.6 | 7.72 | 7.84 | 7.96 | 8.08 | 8.2 | 8.32 | 8.44 | 8.56 | 8.68 | 8.8 | 8.92 | 9.04 | 9.16 | 9.28 | 9.4
20.97 | 4.59 | 4.71 | 4.83 | 4.95 | 5.07 | 5.19 | 5.31 | 5.43 | 5.55 | 5.67 | 5.79 | 5.91 | 6.03 | 6.15 | 6.27 | 6.39 | 6.52 | 6.64 | 6.76 | 6.88 | 7 | 7.12 | 7.24 | 7.36 | 7.48 | 7.6 | 7.72 | 7.84 | 7.96 | 8.08 | 8.2 | 8.32 | 8.44 | 8.56 | 8.68 | 8.8 | 8.92 | 9.04 | 9.16 | 9.28 | 9.4
20.98 | 4.59 | 4.71 | 4.83 | 4.95 | 5.07 | 5.19 | 5.31 | 5.43 | 5.55 | 5.67 | 5.79 | 5.91 | 6.03 | 6.15 | 6.27 | 6.39 | 6.51 | 6.63 | 6.75 | 6.87 | 6.99 | 7.12 | 7.24 | 7.36 | 7.48 | 7.6 | 7.72 | 7.84 | 7.96 | 8.08 | 8.2 | 8.32 | 8.44 | 8.56 | 8.68 | 8.8 | 8.92 | 9.04 | 9.16 | 9.28 | 9.4
20.99 | 4.59 | 4.71 | 4.83 | 4.95 | 5.07 | 5.19 | 5.31 | 5.43 | 5.55 | 5.67 | 5.79 | 5.91 | 6.03 | 6.15 | 6.27 | 6.39 | 6.51 | 6.63 | 6.75 | 6.87 | 6.99 | 7.11 | 7.23 | 7.35 | 7.47 | 7.59 | 7.71 | 7.83 | 7.96 | 8.08 | 8.2 | 8.32 | 8.44 | 8.56 | 8.68 | 8.8 | 8.92 | 9.04 | 9.16 | 9.28 | 9.4
21 | 4.59 | 4.71 | 4.83 | 4.95 | 5.07 | 5.19 | 5.31 | 5.43 | 5.55 | 5.67 | 5.79 | 5.91 | 6.03 | 6.15 | 6.27 | 6.39 | 6.51 | 6.63 | 6.75 | 6.87 | 6.99 | 7.11 | 7.23 | 7.35 | 7.47 | 7.59 | 7.71 | 7.83 | 7.95 | 8.07 | 8.19 | 8.31 | 8.43 | 8.55 | 8.67 | 8.79 | 8.92 | 9.04 | 9.16 | 9.28 | 9.4
21.01 | 4.59 | 4.71 | 4.83 | 4.95 | 5.07 | 5.19 | 5.31 | 5.43 | 5.55 | 5.67 | 5.79 | 5.91 | 6.03 | 6.15 | 6.27 | 6.39 | 6.51 | 6.63 | 6.75 | 6.87 | 6.99 | 7.11 | 7.23 | 7.35 | 7.47 | 7.59 | 7.71 | 7.83 | 7.95 | 8.07 | 8.19 | 8.31 | 8.43 | 8.55 | 8.67 | 8.79 | 8.91 | 9.03 | 9.15 | 9.27 | 9.39
21.02 | 4.59 | 4.71 | 4.83 | 4.95 | 5.07 | 5.19 | 5.31 | 5.43 | 5.55 | 5.67 | 5.79 | 5.91 | 6.03 | 6.15 | 6.27 | 6.39 | 6.51 | 6.63 | 6.75 | 6.87 | 6.99 | 7.11 | 7.23 | 7.35 | 7.47 | 7.59 | 7.71 | 7.83 | 7.95 | 8.07 | 8.19 | 8.31 | 8.43 | 8.55 | 8.67 | 8.79 | 8.91 | 9.03 | 9.15 | 9.27 | 9.39
21.03 | 4.58 | 4.7 | 4.82 | 4.95 | 5.07 | 5.19 | 5.31 | 5.43 | 5.55 | 5.67 | 5.79 | 5.91 | 6.03 | 6.15 | 6.27 | 6.39 | 6.51 | 6.63 | 6.75 | 6.87 | 6.99 | 7.11 | 7.23 | 7.35 | 7.47 | 7.59 | 7.71 | 7.83 | 7.95 | 8.07 | 8.19 | 8.31 | 8.43 | 8.55 | 8.67 | 8.79 | 8.91 | 9.03 | 9.15 | 9.27 | 9.39
21.04 | 4.58 | 4.7 | 4.82 | 4.94 | 5.06 | 5.18 | 5.3 | 5.42 | 5.54 | 5.66 | 5.79 | 5.91 | 6.03 | 6.15 | 6.27 | 6.39 | 6.51 | 6.63 | 6.75 | 6.87 | 6.99 | 7.11 | 7.23 | 7.35 | 7.47 | 7.59 | 7.71 | 7.83 | 7.95 | 8.07 | 8.19 | 8.31 | 8.43 | 8.55 | 8.67 | 8.79 | 8.91 | 9.03 | 9.15 | 9.27 | 9.39
21.05 | 4.58 | 4.7 | 4.82 | 4.94 | 5.06 | 5.18 | 5.3 | 5.42 | 5.54 | 5.66 | 5.78 | 5.9 | 6.02 | 6.14 | 6.26 | 6.38 | 6.5 | 6.62 | 6.74 | 6.86 | 6.99 | 7.11 | 7.23 | 7.35 | 7.47 | 7.59 | 7.71 | 7.83 | 7.95 | 8.07 | 8.19 | 8.31 | 8.43 | 8.55 | 8.67 | 8.79 | 8.91 | 9.03 | 9.15 | 9.27 | 9.39
21.06 | 4.58 | 4.7 | 4.82 | 4.94 | 5.06 | 5.18 | 5.3 | 5.42 | 5.54 | 5.66 | 5.78 | 5.9 | 6.02 | 6.14 | 6.26 | 6.38 | 6.5 | 6.62 | 6.74 | 6.86 | 6.98 | 7.1 | 7.22 | 7.34 | 7.46 | 7.58 | 7.7 | 7.82 | 7.94 | 8.06 | 8.18 | 8.3 | 8.42 | 8.54 | 8.66 | 8.78 | 8.9 | 9.02 | 9.14 | 9.26 | 9.38
21.07 | 4.58 | 4.7 | 4.82 | 4.94 | 5.06 | 5.18 | 5.3 | 5.42 | 5.54 | 5.66 | 5.78 | 5.9 | 6.02 | 6.14 | 6.26 | 6.38 | 6.5 | 6.62 | 6.74 | 6.86 | 6.98 | 7.1 | 7.22 | 7.34 | 7.46 | 7.58 | 7.7 | 7.82 | 7.94 | 8.06 | 8.18 | 8.3 | 8.42 | 8.54 | 8.66 | 8.78 | 8.9 | 9.02 | 9.14 | 9.26 | 9.38
21.08 | 4.58 | 4.7 | 4.82 | 4.94 | 5.06 | 5.18 | 5.3 | 5.42 | 5.54 | 5.66 | 5.78 | 5.9 | 6.02 | 6.14 | 6.26 | 6.38 | 6.5 | 6.62 | 6.74 | 6.86 | 6.98 | 7.1 | 7.22 | 7.34 | 7.46 | 7.58 | 7.7 | 7.82 | 7.94 | 8.06 | 8.18 | 8.3 | 8.42 | 8.54 | 8.66 | 8.78 | 8.9 | 9.02 | 9.14 | 9.26 | 9.38
21.09 | 4.58 | 4.7 | 4.82 | 4.94 | 5.06 | 5.18 | 5.3 | 5.42 | 5.54 | 5.66 | 5.78 | 5.9 | 6.02 | 6.14 | 6.26 | 6.38 | 6.5 | 6.62 | 6.74 | 6.86 | 6.98 | 7.1 | 7.22 | 7.34 | 7.46 | 7.58 | 7.7 | 7.82 | 7.94 | 8.06 | 8.18 | 8.3 | 8.42 | 8.54 | 8.66 | 8.78 | 8.9 | 9.02 | 9.14 | 9.26 | 9.38
21.1 | 4.58 | 4.7 | 4.82 | 4.94 | 5.06 | 5.18 | 5.3 | 5.42 | 5.54 | 5.66 | 5.78 | 5.9 | 6.02 | 6.14 | 6.26 | 6.38 | 6.5 | 6.62 | 6.74 | 6.86 | 6.98 | 7.1 | 7.22 | 7.34 | 7.46 | 7.58 | 7.7 | 7.82 | 7.94 | 8.06 | 8.18 | 8.3 | 8.42 | 8.54 | 8.66 | 8.78 | 8.9 | 9.02 | 9.14 | 9.26 | 9.38
21.11 | 4.58 | 4.7 | 4.82 | 4.94 | 5.06 | 5.18 | 5.3 | 5.42 | 5.54 | 5.66 | 5.78 | 5.9 | 6.02 | 6.14 | 6.26 | 6.38 | 6.5 | 6.62 | 6.74 | 6.86 | 6.98 | 7.1 | 7.22 | 7.34 | 7.46 | 7.58 | 7.7 | 7.82 | 7.94 | 8.06 | 8.18 | 8.3 | 8.42 | 8.54 | 8.66 | 8.78 | 8.9 | 9.02 | 9.14 | 9.26 | 9.38
21.12 | 4.58 | 4.7 | 4.82 | 4.94 | 5.06 | 5.18 | 5.3 | 5.42 | 5.54 | 5.66 | 5.78 | 5.9 | 6.02 | 6.14 | 6.26 | 6.38 | 6.5 | 6.62 | 6.74 | 6.86 | 6.98 | 7.1 | 7.22 | 7.34 | 7.46 | 7.57 | 7.69 | 7.81 | 7.93 | 8.05 | 8.17 | 8.29 | 8.41 | 8.53 | 8.65 | 8.77 | 8.89 | 9.01 | 9.13 | 9.25 | 9.37
21.13 | 4.57 | 4.69 | 4.81 | 4.93 | 5.05 | 5.17 | 5.29 | 5.41 | 5.53 | 5.65 | 5.77 | 5.89 | 6.01 | 6.13 | 6.25 | 6.37 | 6.49 | 6.61 | 6.73 | 6.85 | 6.97 | 7.09 | 7.21 | 7.33 | 7.45 | 7.57 | 7.69 | 7.81 | 7.93 | 8.05 | 8.17 | 8.29 | 8.41 | 8.53 | 8.65 | 8.77 | 8.89 | 9.01 | 9.13 | 9.25 | 9.37
21.14 | 4.57 | 4.69 | 4.81 | 4.93 | 5.05 | 5.17 | 5.29 | 5.41 | 5.53 | 5.65 | 5.77 | 5.89 | 6.01 | 6.13 | 6.25 | 6.37 | 6.49 | 6.61 | 6.73 | 6.85 | 6.97 | 7.09 | 7.21 | 7.33 | 7.45 | 7.57 | 7.69 | 7.81 | 7.93 | 8.05 | 8.17 | 8.29 | 8.41 | 8.53 | 8.65 | 8.77 | 8.89 | 9.01 | 9.13 | 9.25 | 9.37
21.15 | 4.57 | 4.69 | 4.81 | 4.93 | 5.05 | 5.17 | 5.29 | 5.41 | 5.53 | 5.65 | 5.77 | 5.89 | 6.01 | 6.13 | 6.25 | 6.37 | 6.49 | 6.61 | 6.73 | 6.85 | 6.97 | 7.09 | 7.21 | 7.33 | 7.45 | 7.57 | 7.69 | 7.81 | 7.93 | 8.05 | 8.17 | 8.29 | 8.41 | 8.53 | 8.65 | 8.77 | 8.89 | 9.01 | 9.13 | 9.25 | 9.37
21.16 | 4.57 | 4.69 | 4.81 | 4.93 | 5.05 | 5.17 | 5.29 | 5.41 | 5.53 | 5.65 | 5.77 | 5.89 | 6.01 | 6.13 | 6.25 | 6.37 | 6.49 | 6.61 | 6.73 | 6.85 | 6.97 | 7.09 | 7.21 | 7.33 | 7.45 | 7.57 | 7.69 | 7.81 | 7.93 | 8.05 | 8.17 | 8.29 | 8.41 | 8.53 | 8.65 | 8.77 | 8.89 | 9.01 | 9.13 | 9.25 | 9.37
21.17 | 4.57 | 4.69 | 4.81 | 4.93 | 5.05 | 5.17 | 5.29 | 5.41 | 5.53 | 5.65 | 5.77 | 5.89 | 6.01 | 6.13 | 6.25 | 6.37 | 6.49 | 6.61 | 6.73 | 6.85 | 6.97 | 7.09 | 7.21 | 7.33 | 7.45 | 7.57 | 7.69 | 7.81 | 7.93 | 8.05 | 8.17 | 8.29 | 8.41 | 8.53 | 8.65 | 8.77 | 8.89 | 9.01 | 9.13 | 9.24 | 9.36
21.18 | 4.57 | 4.69 | 4.81 | 4.93 | 5.05 | 5.17 | 5.29 | 5.41 | 5.53 | 5.65 | 5.77 | 5.89 | 6.01 | 6.13 | 6.25 | 6.37 | 6.49 | 6.61 | 6.73 | 6.85 | 6.97 | 7.09 | 7.21 | 7.33 | 7.45 | 7.57 | 7.69 | 7.81 | 7.93 | 8.05 | 8.16 | 8.28 | 8.4 | 8.52 | 8.64 | 8.76 | 8.88 | 9 | 9.12 | 9.24 | 9.36
21.19 | 4.57 | 4.69 | 4.81 | 4.93 | 5.05 | 5.17 | 5.29 | 5.41 | 5.53 | 5.65 | 5.77 | 5.89 | 6.01 | 6.13 | 6.25 | 6.37 | 6.49 | 6.61 | 6.73 | 6.85 | 6.97 | 7.09 | 7.21 | 7.32 | 7.44 | 7.56 | 7.68 | 7.8 | 7.92 | 8.04 | 8.16 | 8.28 | 8.4 | 8.52 | 8.64 | 8.76 | 8.88 | 9 | 9.12 | 9.24 | 9.36
21.2 | 4.57 | 4.69 | 4.81 | 4.93 | 5.05 | 5.17 | 5.29 | 5.41 | 5.53 | 5.65 | 5.77 | 5.89 | 6.01 | 6.13 | 6.25 | 6.37 | 6.49 | 6.6 | 6.72 | 6.84 | 6.96 | 7.08 | 7.2 | 7.32 | 7.44 | 7.56 | 7.68 | 7.8 | 7.92 | 8.04 | 8.16 | 8.28 | 8.4 | 8.52 | 8.64 | 8.76 | 8.88 | 9 | 9.12 | 9.24 | 9.36
21.21 | 4.57 | 4.69 | 4.81 | 4.93 | 5.05 | 5.17 | 5.29 | 5.41 | 5.53 | 5.65 | 5.77 | 5.88 | 6 | 6.12 | 6.24 | 6.36 | 6.48 | 6.6 | 6.72 | 6.84 | 6.96 | 7.08 | 7.2 | 7.32 | 7.44 | 7.56 | 7.68 | 7.8 | 7.92 | 8.04 | 8.16 | 8.28 | 8.4 | 8.52 | 8.64 | 8.76 | 8.88 | 9 | 9.12 | 9.24 | 9.36
21.22 | 4.57 | 4.69 | 4.81 | 4.93 | 5.05 | 5.17 | 5.28 | 5.4 | 5.52 | 5.64 | 5.76 | 5.88 | 6 | 6.12 | 6.24 | 6.36 | 6.48 | 6.6 | 6.72 | 6.84 | 6.96 | 7.08 | 7.2 | 7.32 | 7.44 | 7.56 | 7.68 | 7.8 | 7.92 | 8.04 | 8.16 | 8.28 | 8.4 | 8.52 | 8.64 | 8.76 | 8.88 | 9 | 9.12 | 9.24 | 9.36
21.23 | 4.57 | 4.69 | 4.8 | 4.92 | 5.04 | 5.16 | 5.28 | 5.4 | 5.52 | 5.64 | 5.76 | 5.88 | 6 | 6.12 | 6.24 | 6.36 | 6.48 | 6.6 | 6.72 | 6.84 | 6.96 | 7.08 | 7.2 | 7.32 | 7.44 | 7.56 | 7.68 | 7.8 | 7.92 | 8.04 | 8.16 | 8.28 | 8.4 | 8.52 | 8.64 | 8.76 | 8.88 | 8.99 | 9.11 | 9.23 | 9.35
21.24 | 4.56 | 4.68 | 4.8 | 4.92 | 5.04 | 5.16 | 5.28 | 5.4 | 5.52 | 5.64 | 5.76 | 5.88 | 6 | 6.12 | 6.24 | 6.36 | 6.48 | 6.6 | 6.72 | 6.84 | 6.96 | 7.08 | 7.2 | 7.32 | 7.44 | 7.56 | 7.68 | 7.8 | 7.92 | 8.04 | 8.16 | 8.28 | 8.39 | 8.51 | 8.63 | 8.75 | 8.87 | 8.99 | 9.11 | 9.23 | 9.35
21.25 | 4.56 | 4.68 | 4.8 | 4.92 | 5.04 | 5.16 | 5.28 | 5.4 | 5.52 | 5.64 | 5.76 | 5.88 | 6 | 6.12 | 6.24 | 6.36 | 6.48 | 6.6 | 6.72 | 6.84 | 6.96 | 7.08 | 7.2 | 7.32 | 7.44 | 7.56 | 7.68 | 7.79 | 7.91 | 8.03 | 8.15 | 8.27 | 8.39 | 8.51 | 8.63 | 8.75 | 8.87 | 8.99 | 9.11 | 9.23 | 9.35
21.26 | 4.56 | 4.68 | 4.8 | 4.92 | 5.04 | 5.16 | 5.28 | 5.4 | 5.52 | 5.64 | 5.76 | 5.88 | 6 | 6.12 | 6.24 | 6.36 | 6.48 | 6.6 | 6.72 | 6.84 | 6.96 | 7.08 | 7.2 | 7.31 | 7.43 | 7.55 | 7.67 | 7.79 | 7.91 | 8.03 | 8.15 | 8.27 | 8.39 | 8.51 | 8.63 | 8.75 | 8.87 | 8.99 | 9.11 | 9.23 | 9.35
21.27 | 4.56 | 4.68 | 4.8 | 4.92 | 5.04 | 5.16 | 5.28 | 5.4 | 5.52 | 5.64 | 5.76 | 5.88 | 6 | 6.12 | 6.24 | 6.36 | 6.48 | 6.6 | 6.72 | 6.83 | 6.95 | 7.07 | 7.19 | 7.31 | 7.43 | 7.55 | 7.67 | 7.79 | 7.91 | 8.03 | 8.15 | 8.27 | 8.39 | 8.51 | 8.63 | 8.75 | 8.87 | 8.99 | 9.11 | 9.23 | 9.35
21.28 | 4.56 | 4.68 | 4.8 | 4.92 | 5.04 | 5.16 | 5.28 | 5.4 | 5.52 | 5.64 | 5.76 | 5.88 | 6 | 6.12 | 6.24 | 6.35 | 6.47 | 6.59 | 6.71 | 6.83 | 6.95 | 7.07 | 7.19 | 7.31 | 7.43 | 7.55 | 7.67 | 7.79 | 7.91 | 8.03 | 8.15 | 8.27 | 8.39 | 8.51 | 8.63 | 8.75 | 8.87 | 8.99 | 9.11 | 9.23 | 9.34
21.29 | 4.56 | 4.68 | 4.8 | 4.92 | 5.04 | 5.16 | 5.28 | 5.4 | 5.52 | 5.64 | 5.76 | 5.88 | 5.99 | 6.11 | 6.23 | 6.35 | 6.47 | 6.59 | 6.71 | 6.83 | 6.95 | 7.07 | 7.19 | 7.31 | 7.43 | 7.55 | 7.67 | 7.79 | 7.91 | 8.03 | 8.15 | 8.27 | 8.39 | 8.51 | 8.63 | 8.75 | 8.86 | 8.98 | 9.1 | 9.22 | 9.34
21.3 | 4.56 | 4.68 | 4.8 | 4.92 | 5.04 | 5.16 | 5.28 | 5.4 | 5.52 | 5.63 | 5.75 | 5.87 | 5.99 | 6.11 | 6.23 | 6.35 | 6.47 | 6.59 | 6.71 | 6.83 | 6.95 | 7.07 | 7.19 | 7.31 | 7.43 | 7.55 | 7.67 | 7.79 | 7.91 | 8.03 | 8.15 | 8.27 | 8.38 | 8.5 | 8.62 | 8.74 | 8.86 | 8.98 | 9.1 | 9.22 | 9.34
21.31 | 4.56 | 4.68 | 4.8 | 4.92 | 5.04 | 5.16 | 5.28 | 5.39 | 5.51 | 5.63 | 5.75 | 5.87 | 5.99 | 6.11 | 6.23 | 6.35 | 6.47 | 6.59 | 6.71 | 6.83 | 6.95 | 7.07 | 7.19 | 7.31 | 7.43 | 7.55 | 7.67 | 7.79 | 7.91 | 8.02 | 8.14 | 8.26 | 8.38 | 8.5 | 8.62 | 8.74 | 8.86 | 8.98 | 9.1 | 9.22 | 9.34
21.32 | 4.56 | 4.68 | 4.8 | 4.92 | 5.03 | 5.15 | 5.27 | 5.39 | 5.51 | 5.63 | 5.75 | 5.87 | 5.99 | 6.11 | 6.23 | 6.35 | 6.47 | 6.59 | 6.71 | 6.83 | 6.95 | 7.07 | 7.19 | 7.31 | 7.43 | 7.55 | 7.66 | 7.78 | 7.9 | 8.02 | 8.14 | 8.26 | 8.38 | 8.5 | 8.62 | 8.74 | 8.86 | 8.98 | 9.1 | 9.22 | 9.34
21.33 | 4.56 | 4.68 | 4.79 | 4.91 | 5.03 | 5.15 | 5.27 | 5.39 | 5.51 | 5.63 | 5.75 | 5.87 | 5.99 | 6.11 | 6.23 | 6.35 | 6.47 | 6.59 | 6.71 | 6.83 | 6.95 | 7.07 | 7.19 | 7.3 | 7.42 | 7.54 | 7.66 | 7.78 | 7.9 | 8.02 | 8.14 | 8.26 | 8.38 | 8.5 | 8.62 | 8.74 | 8.86 | 8.98 | 9.1 | 9.22 | 9.34
21.34 | 4.55 | 4.67 | 4.79 | 4.91 | 5.03 | 5.15 | 5.27 | 5.39 | 5.51 | 5.63 | 5.75 | 5.87 | 5.99 | 6.11 | 6.23 | 6.35 | 6.47 | 6.59 | 6.71 | 6.83 | 6.94 | 7.06 | 7.18 | 7.3 | 7.42 | 7.54 | 7.66 | 7.78 | 7.9 | 8.02 | 8.14 | 8.26 | 8.38 | 8.5 | 8.62 | 8.74 | 8.86 | 8.98 | 9.09 | 9.21 | 9.33
21.35 | 4.55 | 4.67 | 4.79 | 4.91 | 5.03 | 5.15 | 5.27 | 5.39 | 5.51 | 5.63 | 5.75 | 5.87 | 5.99 | 6.11 | 6.23 | 6.35 | 6.47 | 6.59 | 6.7 | 6.82 | 6.94 | 7.06 | 7.18 | 7.3 | 7.42 | 7.54 | 7.66 | 7.78 | 7.9 | 8.02 | 8.14 | 8.26 | 8.38 | 8.5 | 8.62 | 8.73 | 8.85 | 8.97 | 9.09 | 9.21 | 9.33
21.36 | 4.55 | 4.67 | 4.79 | 4.91 | 5.03 | 5.15 | 5.27 | 5.39 | 5.51 | 5.63 | 5.75 | 5.87 | 5.99 | 6.11 | 6.23 | 6.34 | 6.46 | 6.58 | 6.7 | 6.82 | 6.94 | 7.06 | 7.18 | 7.3 | 7.42 | 7.54 | 7.66 | 7.78 | 7.9 | 8.02 | 8.14 | 8.26 | 8.38 | 8.49 | 8.61 | 8.73 | 8.85 | 8.97 | 9.09 | 9.21 | 9.33
21.37 | 4.55 | 4.67 | 4.79 | 4.91 | 5.03 | 5.15 | 5.27 | 5.39 | 5.51 | 5.63 | 5.75 | 5.87 | 5.99 | 6.1 | 6.22 | 6.34 | 6.46 | 6.58 | 6.7 | 6.82 | 6.94 | 7.06 | 7.18 | 7.3 | 7.42 | 7.54 | 7.66 | 7.78 | 7.9 | 8.02 | 8.13 | 8.25 | 8.37 | 8.49 | 8.61 | 8.73 | 8.85 | 8.97 | 9.09 | 9.21 | 9.33
21.38 | 4.55 | 4.67 | 4.79 | 4.91 | 5.03 | 5.15 | 5.27 | 5.39 | 5.51 | 5.63 | 5.75 | 5.86 | 5.98 | 6.1 | 6.22 | 6.34 | 6.46 | 6.58 | 6.7 | 6.82 | 6.94 | 7.06 | 7.18 | 7.3 | 7.42 | 7.54 | 7.66 | 7.77 | 7.89 | 8.01 | 8.13 | 8.25 | 8.37 | 8.49 | 8.61 | 8.73 | 8.85 | 8.97 | 9.09 | 9.21 | 9.33
21.39 | 4.55 | 4.67 | 4.79 | 4.91 | 5.03 | 5.15 | 5.27 | 5.39 | 5.51 | 5.62 | 5.74 | 5.86 | 5.98 | 6.1 | 6.22 | 6.34 | 6.46 | 6.58 | 6.7 | 6.82 | 6.94 | 7.06 | 7.18 | 7.3 | 7.42 | 7.53 | 7.65 | 7.77 | 7.89 | 8.01 | 8.13 | 8.25 | 8.37 | 8.49 | 8.61 | 8.73 | 8.85 | 8.97 | 9.09 | 9.21 | 9.32
21.4 | 4.55 | 4.67 | 4.79 | 4.91 | 5.03 | 5.15 | 5.27 | 5.38 | 5.5 | 5.62 | 5.74 | 5.86 | 5.98 | 6.1 | 6.22 | 6.34 | 6.46 | 6.58 | 6.7 | 6.82 | 6.94 | 7.06 | 7.18 | 7.29 | 7.41 | 7.53 | 7.65 | 7.77 | 7.89 | 8.01 | 8.13 | 8.25 | 8.37 | 8.49 | 8.61 | 8.73 | 8.85 | 8.97 | 9.08 | 9.2 | 9.32
21.41 | 4.55 | 4.67 | 4.79 | 4.91 | 5.03 | 5.14 | 5.26 | 5.38 | 5.5 | 5.62 | 5.74 | 5.86 | 5.98 | 6.1 | 6.22 | 6.34 | 6.46 | 6.58 | 6.7 | 6.82 | 6.94 | 7.05 | 7.17 | 7.29 | 7.41 | 7.53 | 7.65 | 7.77 | 7.89 | 8.01 | 8.13 | 8.25 | 8.37 | 8.49 | 8.61 | 8.72 | 8.84 | 8.96 | 9.08 | 9.2 | 9.32
21.42 | 4.55 | 4.67 | 4.79 | 4.91 | 5.02 | 5.14 | 5.26 | 5.38 | 5.5 | 5.62 | 5.74 | 5.86 | 5.98 | 6.1 | 6.22 | 6.34 | 6.46 | 6.58 | 6.7 | 6.81 | 6.93 | 7.05 | 7.17 | 7.29 | 7.41 | 7.53 | 7.65 | 7.77 | 7.89 | 8.01 | 8.13 | 8.25 | 8.37 | 8.48 | 8.6 | 8.72 | 8.84 | 8.96 | 9.08 | 9.2 | 9.32
21.43 | 4.55 | 4.67 | 4.78 | 4.9 | 5.02 | 5.14 | 5.26 | 5.38 | 5.5 | 5.62 | 5.74 | 5.86 | 5.98 | 6.1 | 6.22 | 6.34 | 6.46 | 6.57 | 6.69 | 6.81 | 6.93 | 7.05 | 7.17 | 7.29 | 7.41 | 7.53 | 7.65 | 7.77 | 7.89 | 8.01 | 8.13 | 8.24 | 8.36 | 8.48 | 8.6 | 8.72 | 8.84 | 8.96 | 9.08 | 9.2 | 9.32
21.44 | 4.55 | 4.66 | 4.78 | 4.9 | 5.02 | 5.14 | 5.26 | 5.38 | 5.5 | 5.62 | 5.74 | 5.86 | 5.98 | 6.1 | 6.22 | 6.33 | 6.45 | 6.57 | 6.69 | 6.81 | 6.93 | 7.05 | 7.17 | 7.29 | 7.41 | 7.53 | 7.65 | 7.77 | 7.89 | 8 | 8.12 | 8.24 | 8.36 | 8.48 | 8.6 | 8.72 | 8.84 | 8.96 | 9.08 | 9.2 | 9.32
21.45 | 4.54 | 4.66 | 4.78 | 4.9 | 5.02 | 5.14 | 5.26 | 5.38 | 5.5 | 5.62 | 5.74 | 5.86 | 5.98 | 6.09 | 6.21 | 6.33 | 6.45 | 6.57 | 6.69 | 6.81 | 6.93 | 7.05 | 7.17 | 7.29 | 7.41 | 7.53 | 7.65 | 7.76 | 7.88 | 8 | 8.12 | 8.24 | 8.36 | 8.48 | 8.6 | 8.72 | 8.84 | 8.96 | 9.08 | 9.19 | 9.31
21.46 | 4.54 | 4.66 | 4.78 | 4.9 | 5.02 | 5.14 | 5.26 | 5.38 | 5.5 | 5.62 | 5.74 | 5.86 | 5.97 | 6.09 | 6.21 | 6.33 | 6.45 | 6.57 | 6.69 | 6.81 | 6.93 | 7.05 | 7.17 | 7.29 | 7.41 | 7.52 | 7.64 | 7.76 | 7.88 | 8 | 8.12 | 8.24 | 8.36 | 8.48 | 8.6 | 8.72 | 8.84 | 8.95 | 9.07 | 9.19 | 9.31
21.47 | 4.54 | 4.66 | 4.78 | 4.9 | 5.02 | 5.14 | 5.26 | 5.38 | 5.5 | 5.62 | 5.73 | 5.85 | 5.97 | 6.09 | 6.21 | 6.33 | 6.45 | 6.57 | 6.69 | 6.81 | 6.93 | 7.05 | 7.17 | 7.28 | 7.4 | 7.52 | 7.64 | 7.76 | 7.88 | 8 | 8.12 | 8.24 | 8.36 | 8.48 | 8.6 | 8.71 | 8.83 | 8.95 | 9.07 | 9.19 | 9.31
21.48 | 4.54 | 4.66 | 4.78 | 4.9 | 5.02 | 5.14 | 5.26 | 5.38 | 5.5 | 5.61 | 5.73 | 5.85 | 5.97 | 6.09 | 6.21 | 6.33 | 6.45 | 6.57 | 6.69 | 6.81 | 6.93 | 7.04 | 7.16 | 7.28 | 7.4 | 7.52 | 7.64 | 7.76 | 7.88 | 8 | 8.12 | 8.24 | 8.36 | 8.47 | 8.59 | 8.71 | 8.83 | 8.95 | 9.07 | 9.19 | 9.31
21.49 | 4.54 | 4.66 | 4.78 | 4.9 | 5.02 | 5.14 | 5.26 | 5.37 | 5.49 | 5.61 | 5.73 | 5.85 | 5.97 | 6.09 | 6.21 | 6.33 | 6.45 | 6.57 | 6.69 | 6.8 | 6.92 | 7.04 | 7.16 | 7.28 | 7.4 | 7.52 | 7.64 | 7.76 | 7.88 | 8 | 8.12 | 8.23 | 8.35 | 8.47 | 8.59 | 8.71 | 8.83 | 8.95 | 9.07 | 9.19 | 9.31
21.5 | 4.54 | 4.66 | 4.78 | 4.9 | 5.02 | 5.14 | 5.25 | 5.37 | 5.49 | 5.61 | 5.73 | 5.85 | 5.97 | 6.09 | 6.21 | 6.33 | 6.45 | 6.57 | 6.68 | 6.8 | 6.92 | 7.04 | 7.16 | 7.28 | 7.4 | 7.52 | 7.64 | 7.76 | 7.88 | 7.99 | 8.11 | 8.23 | 8.35 | 8.47 | 8.59 | 8.71 | 8.83 | 8.95 | 9.07 | 9.19 | 9.31
21.51 | 4.54 | 4.66 | 4.78 | 4.9 | 5.02 | 5.13 | 5.25 | 5.37 | 5.49 | 5.61 | 5.73 | 5.85 | 5.97 | 6.09 | 6.21 | 6.33 | 6.44 | 6.56 | 6.68 | 6.8 | 6.92 | 7.04 | 7.16 | 7.28 | 7.4 | 7.52 | 7.64 | 7.76 | 7.87 | 7.99 | 8.11 | 8.23 | 8.35 | 8.47 | 8.59 | 8.71 | 8.83 | 8.95 | 9.07 | 9.18 | 9.3
21.52 | 4.54 | 4.66 | 4.78 | 4.9 | 5.01 | 5.13 | 5.25 | 5.37 | 5.49 | 5.61 | 5.73 | 5.85 | 5.97 | 6.09 | 6.21 | 6.32 | 6.44 | 6.56 | 6.68 | 6.8 | 6.92 | 7.04 | 7.16 | 7.28 | 7.4 | 7.52 | 7.63 | 7.75 | 7.87 | 7.99 | 8.11 | 8.23 | 8.35 | 8.47 | 8.59 | 8.71 | 8.83 | 8.94 | 9.06 | 9.18 | 9.3
21.53 | 4.54 | 4.66 | 4.77 | 4.89 | 5.01 | 5.13 | 5.25 | 5.37 | 5.49 | 5.61 | 5.73 | 5.85 | 5.97 | 6.09 | 6.2 | 6.32 | 6.44 | 6.56 | 6.68 | 6.8 | 6.92 | 7.04 | 7.16 | 7.28 | 7.39 | 7.51 | 7.63 | 7.75 | 7.87 | 7.99 | 8.11 | 8.23 | 8.35 | 8.47 | 8.59 | 8.7 | 8.82 | 8.94 | 9.06 | 9.18 | 9.3
21.54 | 4.54 | 4.65 | 4.77 | 4.89 | 5.01 | 5.13 | 5.25 | 5.37 | 5.49 | 5.61 | 5.73 | 5.85 | 5.96 | 6.08 | 6.2 | 6.32 | 6.44 | 6.56 | 6.68 | 6.8 | 6.92 | 7.04 | 7.16 | 7.27 | 7.39 | 7.51 | 7.63 | 7.75 | 7.87 | 7.99 | 8.11 | 8.23 | 8.35 | 8.46 | 8.58 | 8.7 | 8.82 | 8.94 | 9.06 | 9.18 | 9.3
21.55 | 4.53 | 4.65 | 4.77 | 4.89 | 5.01 | 5.13 | 5.25 | 5.37 | 5.49 | 5.61 | 5.73 | 5.84 | 5.96 | 6.08 | 6.2 | 6.32 | 6.44 | 6.56 | 6.68 | 6.8 | 6.92 | 7.03 | 7.15 | 7.27 | 7.39 | 7.51 | 7.63 | 7.75 | 7.87 | 7.99 | 8.11 | 8.23 | 8.34 | 8.46 | 8.58 | 8.7 | 8.82 | 8.94 | 9.06 | 9.18 | 9.3
21.56 | 4.53 | 4.65 | 4.77 | 4.89 | 5.01 | 5.13 | 5.25 | 5.37 | 5.49 | 5.61 | 5.72 | 5.84 | 5.96 | 6.08 | 6.2 | 6.32 | 6.44 | 6.56 | 6.68 | 6.8 | 6.91 | 7.03 | 7.15 | 7.27 | 7.39 | 7.51 | 7.63 | 7.75 | 7.87 | 7.99 | 8.1 | 8.22 | 8.34 | 8.46 | 8.58 | 8.7 | 8.82 | 8.94 | 9.06 | 9.18 | 9.29
21.57 | 4.53 | 4.65 | 4.77 | 4.89 | 5.01 | 5.13 | 5.25 | 5.37 | 5.49 | 5.6 | 5.72 | 5.84 | 5.96 | 6.08 | 6.2 | 6.32 | 6.44 | 6.56 | 6.68 | 6.79 | 6.91 | 7.03 | 7.15 | 7.27 | 7.39 | 7.51 | 7.63 | 7.75 | 7.86 | 7.98 | 8.1 | 8.22 | 8.34 | 8.46 | 8.58 | 8.7 | 8.82 | 8.94 | 9.05 | 9.17 | 9.29
21.58 | 4.53 | 4.65 | 4.77 | 4.89 | 5.01 | 5.13 | 5.25 | 5.36 | 5.48 | 5.6 | 5.72 | 5.84 | 5.96 | 6.08 | 6.2 | 6.32 | 6.44 | 6.55 | 6.67 | 6.79 | 6.91 | 7.03 | 7.15 | 7.27 | 7.39 | 7.51 | 7.63 | 7.74 | 7.86 | 7.98 | 8.1 | 8.22 | 8.34 | 8.46 | 8.58 | 8.7 | 8.81 | 8.93 | 9.05 | 9.17 | 9.29
21.59 | 4.53 | 4.65 | 4.77 | 4.89 | 5.01 | 5.13 | 5.24 | 5.36 | 5.48 | 5.6 | 5.72 | 5.84 | 5.96 | 6.08 | 6.2 | 6.32 | 6.43 | 6.55 | 6.67 | 6.79 | 6.91 | 7.03 | 7.15 | 7.27 | 7.39 | 7.51 | 7.62 | 7.74 | 7.86 | 7.98 | 8.1 | 8.22 | 8.34 | 8.46 | 8.58 | 8.69 | 8.81 | 8.93 | 9.05 | 9.17 | 9.29
21.6 | 4.53 | 4.65 | 4.77 | 4.89 | 5.01 | 5.12 | 5.24 | 5.36 | 5.48 | 5.6 | 5.72 | 5.84 | 5.96 | 6.08 | 6.2 | 6.31 | 6.43 | 6.55 | 6.67 | 6.79 | 6.91 | 7.03 | 7.15 | 7.27 | 7.38 | 7.5 | 7.62 | 7.74 | 7.86 | 7.98 | 8.1 | 8.22 | 8.34 | 8.45 | 8.57 | 8.69 | 8.81 | 8.93 | 9.05 | 9.17 | 9.29
21.61 | 4.53 | 4.65 | 4.77 | 4.89 | 5 | 5.12 | 5.24 | 5.36 | 5.48 | 5.6 | 5.72 | 5.84 | 5.96 | 6.08 | 6.19 | 6.31 | 6.43 | 6.55 | 6.67 | 6.79 | 6.91 | 7.03 | 7.15 | 7.26 | 7.38 | 7.5 | 7.62 | 7.74 | 7.86 | 7.98 | 8.1 | 8.22 | 8.33 | 8.45 | 8.57 | 8.69 | 8.81 | 8.93 | 9.05 | 9.17 | 9.29
21.62 | 4.53 | 4.65 | 4.77 | 4.88 | 5 | 5.12 | 5.24 | 5.36 | 5.48 | 5.6 | 5.72 | 5.84 | 5.96 | 6.07 | 6.19 | 6.31 | 6.43 | 6.55 | 6.67 | 6.79 | 6.91 | 7.03 | 7.14 | 7.26 | 7.38 | 7.5 | 7.62 | 7.74 | 7.86 | 7.98 | 8.09 | 8.21 | 8.33 | 8.45 | 8.57 | 8.69 | 8.81 | 8.93 | 9.05 | 9.16 | 9.28
21.63 | 4.53 | 4.65 | 4.77 | 4.88 | 5 | 5.12 | 5.24 | 5.36 | 5.48 | 5.6 | 5.72 | 5.84 | 5.95 | 6.07 | 6.19 | 6.31 | 6.43 | 6.55 | 6.67 | 6.79 | 6.9 | 7.02 | 7.14 | 7.26 | 7.38 | 7.5 | 7.62 | 7.74 | 7.86 | 7.97 | 8.09 | 8.21 | 8.33 | 8.45 | 8.57 | 8.69 | 8.81 | 8.93 | 9.04 | 9.16 | 9.28
21.64 | 4.53 | 4.65 | 4.76 | 4.88 | 5 | 5.12 | 5.24 | 5.36 | 5.48 | 5.6 | 5.72 | 5.83 | 5.95 | 6.07 | 6.19 | 6.31 | 6.43 | 6.55 | 6.67 | 6.78 | 6.9 | 7.02 | 7.14 | 7.26 | 7.38 | 7.5 | 7.62 | 7.74 | 7.85 | 7.97 | 8.09 | 8.21 | 8.33 | 8.45 | 8.57 | 8.69 | 8.8 | 8.92 | 9.04 | 9.16 | 9.28
21.65 | 4.53 | 4.64 | 4.76 | 4.88 | 5 | 5.12 | 5.24 | 5.36 | 5.48 | 5.6 | 5.71 | 5.83 | 5.95 | 6.07 | 6.19 | 6.31 | 6.43 | 6.55 | 6.66 | 6.78 | 6.9 | 7.02 | 7.14 | 7.26 | 7.38 | 7.5 | 7.62 | 7.73 | 7.85 | 7.97 | 8.09 | 8.21 | 8.33 | 8.45 | 8.57 | 8.68 | 8.8 | 8.92 | 9.04 | 9.16 | 9.28
21.66 | 4.52 | 4.64 | 4.76 | 4.88 | 5 | 5.12 | 5.24 | 5.36 | 5.48 | 5.59 | 5.71 | 5.83 | 5.95 | 6.07 | 6.19 | 6.31 | 6.43 | 6.54 | 6.66 | 6.78 | 6.9 | 7.02 | 7.14 | 7.26 | 7.38 | 7.49 | 7.61 | 7.73 | 7.85 | 7.97 | 8.09 | 8.21 | 8.33 | 8.44 | 8.56 | 8.68 | 8.8 | 8.92 | 9.04 | 9.16 | 9.28
21.67 | 4.52 | 4.64 | 4.76 | 4.88 | 5 | 5.12 | 5.24 | 5.36 | 5.47 | 5.59 | 5.71 | 5.83 | 5.95 | 6.07 | 6.19 | 6.31 | 6.42 | 6.54 | 6.66 | 6.78 | 6.9 | 7.02 | 7.14 | 7.26 | 7.37 | 7.49 | 7.61 | 7.73 | 7.85 | 7.97 | 8.09 | 8.21 | 8.32 | 8.44 | 8.56 | 8.68 | 8.8 | 8.92 | 9.04 | 9.16 | 9.27
21.68 | 4.52 | 4.64 | 4.76 | 4.88 | 5 | 5.12 | 5.24 | 5.35 | 5.47 | 5.59 | 5.71 | 5.83 | 5.95 | 6.07 | 6.19 | 6.3 | 6.42 | 6.54 | 6.66 | 6.78 | 6.9 | 7.02 | 7.14 | 7.25 | 7.37 | 7.49 | 7.61 | 7.73 | 7.85 | 7.97 | 8.09 | 8.2 | 8.32 | 8.44 | 8.56 | 8.68 | 8.8 | 8.92 | 9.04 | 9.15 | 9.27
21.69 | 4.52 | 4.64 | 4.76 | 4.88 | 5 | 5.12 | 5.23 | 5.35 | 5.47 | 5.59 | 5.71 | 5.83 | 5.95 | 6.07 | 6.18 | 6.3 | 6.42 | 6.54 | 6.66 | 6.78 | 6.9 | 7.02 | 7.13 | 7.25 | 7.37 | 7.49 | 7.61 | 7.73 | 7.85 | 7.97 | 8.08 | 8.2 | 8.32 | 8.44 | 8.56 | 8.68 | 8.8 | 8.91 | 9.03 | 9.15 | 9.27
21.7 | 4.52 | 4.64 | 4.76 | 4.88 | 5 | 5.11 | 5.23 | 5.35 | 5.47 | 5.59 | 5.71 | 5.83 | 5.95 | 6.06 | 6.18 | 6.3 | 6.42 | 6.54 | 6.66 | 6.78 | 6.9 | 7.01 | 7.13 | 7.25 | 7.37 | 7.49 | 7.61 | 7.73 | 7.84 | 7.96 | 8.08 | 8.2 | 8.32 | 8.44 | 8.56 | 8.68 | 8.79 | 8.91 | 9.03 | 9.15 | 9.27
21.71 | 4.52 | 4.64 | 4.76 | 4.88 | 4.99 | 5.11 | 5.23 | 5.35 | 5.47 | 5.59 | 5.71 | 5.83 | 5.94 | 6.06 | 6.18 | 6.3 | 6.42 | 6.54 | 6.66 | 6.78 | 6.89 | 7.01 | 7.13 | 7.25 | 7.37 | 7.49 | 7.61 | 7.72 | 7.84 | 7.96 | 8.08 | 8.2 | 8.32 | 8.44 | 8.56 | 8.67 | 8.79 | 8.91 | 9.03 | 9.15 | 9.27
21.72 | 4.52 | 4.64 | 4.76 | 4.87 | 4.99 | 5.11 | 5.23 | 5.35 | 5.47 | 5.59 | 5.71 | 5.82 | 5.94 | 6.06 | 6.18 | 6.3 | 6.42 | 6.54 | 6.66 | 6.77 | 6.89 | 7.01 | 7.13 | 7.25 | 7.37 | 7.49 | 7.6 | 7.72 | 7.84 | 7.96 | 8.08 | 8.2 | 8.32 | 8.44 | 8.55 | 8.67 | 8.79 | 8.91 | 9.03 | 9.15 | 9.27
21.73 | 4.52 | 4.64 | 4.76 | 4.87 | 4.99 | 5.11 | 5.23 | 5.35 | 5.47 | 5.59 | 5.7 | 5.82 | 5.94 | 6.06 | 6.18 | 6.3 | 6.42 | 6.54 | 6.65 | 6.77 | 6.89 | 7.01 | 7.13 | 7.25 | 7.37 | 7.48 | 7.6 | 7.72 | 7.84 | 7.96 | 8.08 | 8.2 | 8.31 | 8.43 | 8.55 | 8.67 | 8.79 | 8.91 | 9.03 | 9.15 | 9.26
21.74 | 4.52 | 4.64 | 4.75 | 4.87 | 4.99 | 5.11 | 5.23 | 5.35 | 5.47 | 5.58 | 5.7 | 5.82 | 5.94 | 6.06 | 6.18 | 6.3 | 6.42 | 6.53 | 6.65 | 6.77 | 6.89 | 7.01 | 7.13 | 7.25 | 7.36 | 7.48 | 7.6 | 7.72 | 7.84 | 7.96 | 8.08 | 8.19 | 8.31 | 8.43 | 8.55 | 8.67 | 8.79 | 8.91 | 9.02 | 9.14 | 9.26
21.75 | 4.52 | 4.63 | 4.75 | 4.87 | 4.99 | 5.11 | 5.23 | 5.35 | 5.47 | 5.58 | 5.7 | 5.82 | 5.94 | 6.06 | 6.18 | 6.3 | 6.41 | 6.53 | 6.65 | 6.77 | 6.89 | 7.01 | 7.13 | 7.24 | 7.36 | 7.48 | 7.6 | 7.72 | 7.84 | 7.96 | 8.07 | 8.19 | 8.31 | 8.43 | 8.55 | 8.67 | 8.79 | 8.9 | 9.02 | 9.14 | 9.26
21.76 | 4.51 | 4.63 | 4.75 | 4.87 | 4.99 | 5.11 | 5.23 | 5.35 | 5.46 | 5.58 | 5.7 | 5.82 | 5.94 | 6.06 | 6.18 | 6.29 | 6.41 | 6.53 | 6.65 | 6.77 | 6.89 | 7.01 | 7.12 | 7.24 | 7.36 | 7.48 | 7.6 | 7.72 | 7.84 | 7.95 | 8.07 | 8.19 | 8.31 | 8.43 | 8.55 | 8.67 | 8.78 | 8.9 | 9.02 | 9.14 | 9.26
21.77 | 4.51 | 4.63 | 4.75 | 4.87 | 4.99 | 5.11 | 5.23 | 5.34 | 5.46 | 5.58 | 5.7 | 5.82 | 5.94 | 6.06 | 6.17 | 6.29 | 6.41 | 6.53 | 6.65 | 6.77 | 6.89 | 7 | 7.12 | 7.24 | 7.36 | 7.48 | 7.6 | 7.72 | 7.83 | 7.95 | 8.07 | 8.19 | 8.31 | 8.43 | 8.55 | 8.66 | 8.78 | 8.9 | 9.02 | 9.14 | 9.26
21.78 | 4.51 | 4.63 | 4.75 | 4.87 | 4.99 | 5.11 | 5.22 | 5.34 | 5.46 | 5.58 | 5.7 | 5.82 | 5.94 | 6.05 | 6.17 | 6.29 | 6.41 | 6.53 | 6.65 | 6.77 | 6.88 | 7 | 7.12 | 7.24 | 7.36 | 7.48 | 7.6 | 7.71 | 7.83 | 7.95 | 8.07 | 8.19 | 8.31 | 8.43 | 8.54 | 8.66 | 8.78 | 8.9 | 9.02 | 9.14 | 9.25
21.79 | 4.51 | 4.63 | 4.75 | 4.87 | 4.99 | 5.1 | 5.22 | 5.34 | 5.46 | 5.58 | 5.7 | 5.82 | 5.93 | 6.05 | 6.17 | 6.29 | 6.41 | 6.53 | 6.65 | 6.76 | 6.88 | 7 | 7.12 | 7.24 | 7.36 | 7.48 | 7.59 | 7.71 | 7.83 | 7.95 | 8.07 | 8.19 | 8.31 | 8.42 | 8.54 | 8.66 | 8.78 | 8.9 | 9.02 | 9.13 | 9.25
21.8 | 4.51 | 4.63 | 4.75 | 4.87 | 4.99 | 5.1 | 5.22 | 5.34 | 5.46 | 5.58 | 5.7 | 5.82 | 5.93 | 6.05 | 6.17 | 6.29 | 6.41 | 6.53 | 6.64 | 6.76 | 6.88 | 7 | 7.12 | 7.24 | 7.36 | 7.47 | 7.59 | 7.71 | 7.83 | 7.95 | 8.07 | 8.19 | 8.3 | 8.42 | 8.54 | 8.66 | 8.78 | 8.9 | 9.01 | 9.13 | 9.25
21.81 | 4.51 | 4.63 | 4.75 | 4.87 | 4.98 | 5.1 | 5.22 | 5.34 | 5.46 | 5.58 | 5.7 | 5.81 | 5.93 | 6.05 | 6.17 | 6.29 | 6.41 | 6.52 | 6.64 | 6.76 | 6.88 | 7 | 7.12 | 7.24 | 7.35 | 7.47 | 7.59 | 7.71 | 7.83 | 7.95 | 8.07 | 8.18 | 8.3 | 8.42 | 8.54 | 8.66 | 8.78 | 8.89 | 9.01 | 9.13 | 9.25
21.82 | 4.51 | 4.63 | 4.75 | 4.86 | 4.98 | 5.1 | 5.22 | 5.34 | 5.46 | 5.58 | 5.69 | 5.81 | 5.93 | 6.05 | 6.17 | 6.29 | 6.41 | 6.52 | 6.64 | 6.76 | 6.88 | 7 | 7.12 | 7.23 | 7.35 | 7.47 | 7.59 | 7.71 | 7.83 | 7.95 | 8.06 | 8.18 | 8.3 | 8.42 | 8.54 | 8.66 | 8.77 | 8.89 | 9.01 | 9.13 | 9.25
21.83 | 4.51 | 4.63 | 4.75 | 4.86 | 4.98 | 5.1 | 5.22 | 5.34 | 5.46 | 5.57 | 5.69 | 5.81 | 5.93 | 6.05 | 6.17 | 6.29 | 6.4 | 6.52 | 6.64 | 6.76 | 6.88 | 7 | 7.11 | 7.23 | 7.35 | 7.47 | 7.59 | 7.71 | 7.83 | 7.94 | 8.06 | 8.18 | 8.3 | 8.42 | 8.54 | 8.65 | 8.77 | 8.89 | 9.01 | 9.13 | 9.25
21.84 | 4.51 | 4.63 | 4.74 | 4.86 | 4.98 | 5.1 | 5.22 | 5.34 | 5.46 | 5.57 | 5.69 | 5.81 | 5.93 | 6.05 | 6.17 | 6.28 | 6.4 | 6.52 | 6.64 | 6.76 | 6.88 | 6.99 | 7.11 | 7.23 | 7.35 | 7.47 | 7.59 | 7.71 | 7.82 | 7.94 | 8.06 | 8.18 | 8.3 | 8.42 | 8.53 | 8.65 | 8.77 | 8.89 | 9.01 | 9.13 | 9.24
21.85 | 4.51 | 4.62 | 4.74 | 4.86 | 4.98 | 5.1 | 5.22 | 5.34 | 5.45 | 5.57 | 5.69 | 5.81 | 5.93 | 6.05 | 6.16 | 6.28 | 6.4 | 6.52 | 6.64 | 6.76 | 6.87 | 6.99 | 7.11 | 7.23 | 7.35 | 7.47 | 7.59 | 7.7 | 7.82 | 7.94 | 8.06 | 8.18 | 8.3 | 8.41 | 8.53 | 8.65 | 8.77 | 8.89 | 9.01 | 9.12 | 9.24
21.86 | 4.51 | 4.62 | 4.74 | 4.86 | 4.98 | 5.1 | 5.22 | 5.33 | 5.45 | 5.57 | 5.69 | 5.81 | 5.93 | 6.04 | 6.16 | 6.28 | 6.4 | 6.52 | 6.64 | 6.76 | 6.87 | 6.99 | 7.11 | 7.23 | 7.35 | 7.47 | 7.58 | 7.7 | 7.82 | 7.94 | 8.06 | 8.18 | 8.29 | 8.41 | 8.53 | 8.65 | 8.77 | 8.89 | 9 | 9.12 | 9.24
21.87 | 4.5 | 4.62 | 4.74 | 4.86 | 4.98 | 5.1 | 5.22 | 5.33 | 5.45 | 5.57 | 5.69 | 5.81 | 5.93 | 6.04 | 6.16 | 6.28 | 6.4 | 6.52 | 6.64 | 6.75 | 6.87 | 6.99 | 7.11 | 7.23 | 7.35 | 7.46 | 7.58 | 7.7 | 7.82 | 7.94 | 8.06 | 8.17 | 8.29 | 8.41 | 8.53 | 8.65 | 8.77 | 8.88 | 9 | 9.12 | 9.24
21.88 | 4.5 | 4.62 | 4.74 | 4.86 | 4.98 | 5.1 | 5.21 | 5.33 | 5.45 | 5.57 | 5.69 | 5.81 | 5.92 | 6.04 | 6.16 | 6.28 | 6.4 | 6.52 | 6.63 | 6.75 | 6.87 | 6.99 | 7.11 | 7.23 | 7.34 | 7.46 | 7.58 | 7.7 | 7.82 | 7.94 | 8.05 | 8.17 | 8.29 | 8.41 | 8.53 | 8.65 | 8.76 | 8.88 | 9 | 9.12 | 9.24
21.89 | 4.5 | 4.62 | 4.74 | 4.86 | 4.98 | 5.09 | 5.21 | 5.33 | 5.45 | 5.57 | 5.69 | 5.8 | 5.92 | 6.04 | 6.16 | 6.28 | 6.4 | 6.51 | 6.63 | 6.75 | 6.87 | 6.99 | 7.11 | 7.22 | 7.34 | 7.46 | 7.58 | 7.7 | 7.82 | 7.93 | 8.05 | 8.17 | 8.29 | 8.41 | 8.53 | 8.64 | 8.76 | 8.88 | 9 | 9.12 | 9.24
21.9 | 4.5 | 4.62 | 4.74 | 4.86 | 4.98 | 5.09 | 5.21 | 5.33 | 5.45 | 5.57 | 5.69 | 5.8 | 5.92 | 6.04 | 6.16 | 6.28 | 6.39 | 6.51 | 6.63 | 6.75 | 6.87 | 6.99 | 7.1 | 7.22 | 7.34 | 7.46 | 7.58 | 7.7 | 7.81 | 7.93 | 8.05 | 8.17 | 8.29 | 8.41 | 8.52 | 8.64 | 8.76 | 8.88 | 9 | 9.12 | 9.23
21.91 | 4.5 | 4.62 | 4.74 | 4.86 | 4.97 | 5.09 | 5.21 | 5.33 | 5.45 | 5.57 | 5.68 | 5.8 | 5.92 | 6.04 | 6.16 | 6.28 | 6.39 | 6.51 | 6.63 | 6.75 | 6.87 | 6.99 | 7.1 | 7.22 | 7.34 | 7.46 | 7.58 | 7.69 | 7.81 | 7.93 | 8.05 | 8.17 | 8.29 | 8.4 | 8.52 | 8.64 | 8.76 | 8.88 | 9 | 9.11 | 9.23
21.92 | 4.5 | 4.62 | 4.74 | 4.85 | 4.97 | 5.09 | 5.21 | 5.33 | 5.45 | 5.56 | 5.68 | 5.8 | 5.92 | 6.04 | 6.16 | 6.27 | 6.39 | 6.51 | 6.63 | 6.75 | 6.87 | 6.98 | 7.1 | 7.22 | 7.34 | 7.46 | 7.57 | 7.69 | 7.81 | 7.93 | 8.05 | 8.17 | 8.28 | 8.4 | 8.52 | 8.64 | 8.76 | 8.88 | 8.99 | 9.11 | 9.23
21.93 | 4.5 | 4.62 | 4.74 | 4.85 | 4.97 | 5.09 | 5.21 | 5.33 | 5.45 | 5.56 | 5.68 | 5.8 | 5.92 | 6.04 | 6.15 | 6.27 | 6.39 | 6.51 | 6.63 | 6.75 | 6.86 | 6.98 | 7.1 | 7.22 | 7.34 | 7.46 | 7.57 | 7.69 | 7.81 | 7.93 | 8.05 | 8.16 | 8.28 | 8.4 | 8.52 | 8.64 | 8.76 | 8.87 | 8.99 | 9.11 | 9.23
21.94 | 4.5 | 4.62 | 4.73 | 4.85 | 4.97 | 5.09 | 5.21 | 5.33 | 5.44 | 5.56 | 5.68 | 5.8 | 5.92 | 6.04 | 6.15 | 6.27 | 6.39 | 6.51 | 6.63 | 6.74 | 6.86 | 6.98 | 7.1 | 7.22 | 7.34 | 7.45 | 7.57 | 7.69 | 7.81 | 7.93 | 8.04 | 8.16 | 8.28 | 8.4 | 8.52 | 8.64 | 8.75 | 8.87 | 8.99 | 9.11 | 9.23
21.95 | 4.5 | 4.62 | 4.73 | 4.85 | 4.97 | 5.09 | 5.21 | 5.32 | 5.44 | 5.56 | 5.68 | 5.8 | 5.92 | 6.03 | 6.15 | 6.27 | 6.39 | 6.51 | 6.62 | 6.74 | 6.86 | 6.98 | 7.1 | 7.22 | 7.33 | 7.45 | 7.57 | 7.69 | 7.81 | 7.92 | 8.04 | 8.16 | 8.28 | 8.4 | 8.52 | 8.63 | 8.75 | 8.87 | 8.99 | 9.11 | 9.22
21.96 | 4.5 | 4.61 | 4.73 | 4.85 | 4.97 | 5.09 | 5.21 | 5.32 | 5.44 | 5.56 | 5.68 | 5.8 | 5.91 | 6.03 | 6.15 | 6.27 | 6.39 | 6.51 | 6.62 | 6.74 | 6.86 | 6.98 | 7.1 | 7.21 | 7.33 | 7.45 | 7.57 | 7.69 | 7.81 | 7.92 | 8.04 | 8.16 | 8.28 | 8.4 | 8.51 | 8.63 | 8.75 | 8.87 | 8.99 | 9.1 | 9.22
21.97 | 4.5 | 4.61 | 4.73 | 4.85 | 4.97 | 5.09 | 5.2 | 5.32 | 5.44 | 5.56 | 5.68 | 5.8 | 5.91 | 6.03 | 6.15 | 6.27 | 6.39 | 6.5 | 6.62 | 6.74 | 6.86 | 6.98 | 7.09 | 7.21 | 7.33 | 7.45 | 7.57 | 7.69 | 7.8 | 7.92 | 8.04 | 8.16 | 8.28 | 8.39 | 8.51 | 8.63 | 8.75 | 8.87 | 8.98 | 9.1 | 9.22
21.98 | 4.49 | 4.61 | 4.73 | 4.85 | 4.97 | 5.09 | 5.2 | 5.32 | 5.44 | 5.56 | 5.68 | 5.79 | 5.91 | 6.03 | 6.15 | 6.27 | 6.38 | 6.5 | 6.62 | 6.74 | 6.86 | 6.98 | 7.09 | 7.21 | 7.33 | 7.45 | 7.57 | 7.68 | 7.8 | 7.92 | 8.04 | 8.16 | 8.27 | 8.39 | 8.51 | 8.63 | 8.75 | 8.87 | 8.98 | 9.1 | 9.22
21.99 | 4.49 | 4.61 | 4.73 | 4.85 | 4.97 | 5.08 | 5.2 | 5.32 | 5.44 | 5.56 | 5.67 | 5.79 | 5.91 | 6.03 | 6.15 | 6.27 | 6.38 | 6.5 | 6.62 | 6.74 | 6.86 | 6.97 | 7.09 | 7.21 | 7.33 | 7.45 | 7.56 | 7.68 | 7.8 | 7.92 | 8.04 | 8.16 | 8.27 | 8.39 | 8.51 | 8.63 | 8.75 | 8.86 | 8.98 | 9.1 | 9.22
22 | 4.49 | 4.61 | 4.73 | 4.85 | 4.96 | 5.08 | 5.2 | 5.32 | 5.44 | 5.56 | 5.67 | 5.79 | 5.91 | 6.03 | 6.15 | 6.26 | 6.38 | 6.5 | 6.62 | 6.74 | 6.85 | 6.97 | 7.09 | 7.21 | 7.33 | 7.45 | 7.56 | 7.68 | 7.8 | 7.92 | 8.04 | 8.15 | 8.27 | 8.39 | 8.51 | 8.63 | 8.74 | 8.86 | 8.98 | 9.1 | 9.22
22.01 | 4.49 | 4.61 | 4.73 | 4.85 | 4.96 | 5.08 | 5.2 | 5.32 | 5.44 | 5.55 | 5.67 | 5.79 | 5.91 | 6.03 | 6.14 | 6.26 | 6.38 | 6.5 | 6.62 | 6.74 | 6.85 | 6.97 | 7.09 | 7.21 | 7.33 | 7.44 | 7.56 | 7.68 | 7.8 | 7.92 | 8.03 | 8.15 | 8.27 | 8.39 | 8.51 | 8.62 | 8.74 | 8.86 | 8.98 | 9.1 | 9.21
22.02 | 4.49 | 4.61 | 4.73 | 4.84 | 4.96 | 5.08 | 5.2 | 5.32 | 5.44 | 5.55 | 5.67 | 5.79 | 5.91 | 6.03 | 6.14 | 6.26 | 6.38 | 6.5 | 6.62 | 6.73 | 6.85 | 6.97 | 7.09 | 7.21 | 7.32 | 7.44 | 7.56 | 7.68 | 7.8 | 7.91 | 8.03 | 8.15 | 8.27 | 8.39 | 8.5 | 8.62 | 8.74 | 8.86 | 8.98 | 9.09 | 9.21
22.03 | 4.49 | 4.61 | 4.73 | 4.84 | 4.96 | 5.08 | 5.2 | 5.32 | 5.43 | 5.55 | 5.67 | 5.79 | 5.91 | 6.02 | 6.14 | 6.26 | 6.38 | 6.5 | 6.61 | 6.73 | 6.85 | 6.97 | 7.09 | 7.2 | 7.32 | 7.44 | 7.56 | 7.68 | 7.79 | 7.91 | 8.03 | 8.15 | 8.27 | 8.38 | 8.5 | 8.62 | 8.74 | 8.86 | 8.97 | 9.09 | 9.21
22.04 | 4.49 | 4.61 | 4.72 | 4.84 | 4.96 | 5.08 | 5.2 | 5.32 | 5.43 | 5.55 | 5.67 | 5.79 | 5.91 | 6.02 | 6.14 | 6.26 | 6.38 | 6.5 | 6.61 | 6.73 | 6.85 | 6.97 | 7.09 | 7.2 | 7.32 | 7.44 | 7.56 | 7.68 | 7.79 | 7.91 | 8.03 | 8.15 | 8.27 | 8.38 | 8.5 | 8.62 | 8.74 | 8.85 | 8.97 | 9.09 | 9.21
22.05 | 4.49 | 4.61 | 4.72 | 4.84 | 4.96 | 5.08 | 5.2 | 5.31 | 5.43 | 5.55 | 5.67 | 5.79 | 5.9 | 6.02 | 6.14 | 6.26 | 6.38 | 6.49 | 6.61 | 6.73 | 6.85 | 6.97 | 7.08 | 7.2 | 7.32 | 7.44 | 7.56 | 7.67 | 7.79 | 7.91 | 8.03 | 8.15 | 8.26 | 8.38 | 8.5 | 8.62 | 8.74 | 8.85 | 8.97 | 9.09 | 9.21
22.06 | 4.49 | 4.6 | 4.72 | 4.84 | 4.96 | 5.08 | 5.19 | 5.31 | 5.43 | 5.55 | 5.67 | 5.78 | 5.9 | 6.02 | 6.14 | 6.26 | 6.37 | 6.49 | 6.61 | 6.73 | 6.85 | 6.96 | 7.08 | 7.2 | 7.32 | 7.44 | 7.55 | 7.67 | 7.79 | 7.91 | 8.03 | 8.14 | 8.26 | 8.38 | 8.5 | 8.62 | 8.73 | 8.85 | 8.97 | 9.09 | 9.21
22.07 | 4.49 | 4.6 | 4.72 | 4.84 | 4.96 | 5.08 | 5.19 | 5.31 | 5.43 | 5.55 | 5.67 | 5.78 | 5.9 | 6.02 | 6.14 | 6.26 | 6.37 | 6.49 | 6.61 | 6.73 | 6.85 | 6.96 | 7.08 | 7.2 | 7.32 | 7.43 | 7.55 | 7.67 | 7.79 | 7.91 | 8.02 | 8.14 | 8.26 | 8.38 | 8.5 | 8.61 | 8.73 | 8.85 | 8.97 | 9.09 | 9.2
22.08 | 4.49 | 4.6 | 4.72 | 4.84 | 4.96 | 5.07 | 5.19 | 5.31 | 5.43 | 5.55 | 5.66 | 5.78 | 5.9 | 6.02 | 6.14 | 6.25 | 6.37 | 6.49 | 6.61 | 6.73 | 6.84 | 6.96 | 7.08 | 7.2 | 7.32 | 7.43 | 7.55 | 7.67 | 7.79 | 7.91 | 8.02 | 8.14 | 8.26 | 8.38 | 8.49 | 8.61 | 8.73 | 8.85 | 8.97 | 9.08 | 9.2
22.09 | 4.48 | 4.6 | 4.72 | 4.84 | 4.96 | 5.07 | 5.19 | 5.31 | 5.43 | 5.55 | 5.66 | 5.78 | 5.9 | 6.02 | 6.14 | 6.25 | 6.37 | 6.49 | 6.61 | 6.72 | 6.84 | 6.96 | 7.08 | 7.2 | 7.31 | 7.43 | 7.55 | 7.67 | 7.79 | 7.9 | 8.02 | 8.14 | 8.26 | 8.37 | 8.49 | 8.61 | 8.73 | 8.85 | 8.96 | 9.08 | 9.2
22.1 | 4.48 | 4.6 | 4.72 | 4.84 | 4.95 | 5.07 | 5.19 | 5.31 | 5.43 | 5.54 | 5.66 | 5.78 | 5.9 | 6.02 | 6.13 | 6.25 | 6.37 | 6.49 | 6.61 | 6.72 | 6.84 | 6.96 | 7.08 | 7.19 | 7.31 | 7.43 | 7.55 | 7.67 | 7.78 | 7.9 | 8.02 | 8.14 | 8.26 | 8.37 | 8.49 | 8.61 | 8.73 | 8.84 | 8.96 | 9.08 | 9.2
22.11 | 4.48 | 4.6 | 4.72 | 4.84 | 4.95 | 5.07 | 5.19 | 5.31 | 5.43 | 5.54 | 5.66 | 5.78 | 5.9 | 6.01 | 6.13 | 6.25 | 6.37 | 6.49 | 6.6 | 6.72 | 6.84 | 6.96 | 7.08 | 7.19 | 7.31 | 7.43 | 7.55 | 7.66 | 7.78 | 7.9 | 8.02 | 8.14 | 8.25 | 8.37 | 8.49 | 8.61 | 8.73 | 8.84 | 8.96 | 9.08 | 9.2
22.12 | 4.48 | 4.6 | 4.72 | 4.83 | 4.95 | 5.07 | 5.19 | 5.31 | 5.42 | 5.54 | 5.66 | 5.78 | 5.9 | 6.01 | 6.13 | 6.25 | 6.37 | 6.48 | 6.6 | 6.72 | 6.84 | 6.96 | 7.07 | 7.19 | 7.31 | 7.43 | 7.55 | 7.66 | 7.78 | 7.9 | 8.02 | 8.13 | 8.25 | 8.37 | 8.49 | 8.61 | 8.72 | 8.84 | 8.96 | 9.08 | 9.19
22.13 | 4.48 | 4.6 | 4.72 | 4.83 | 4.95 | 5.07 | 5.19 | 5.31 | 5.42 | 5.54 | 5.66 | 5.78 | 5.89 | 6.01 | 6.13 | 6.25 | 6.37 | 6.48 | 6.6 | 6.72 | 6.84 | 6.95 | 7.07 | 7.19 | 7.31 | 7.43 | 7.54 | 7.66 | 7.78 | 7.9 | 8.02 | 8.13 | 8.25 | 8.37 | 8.49 | 8.6 | 8.72 | 8.84 | 8.96 | 9.08 | 9.19
22.14 | 4.48 | 4.6 | 4.72 | 4.83 | 4.95 | 5.07 | 5.19 | 5.3 | 5.42 | 5.54 | 5.66 | 5.78 | 5.89 | 6.01 | 6.13 | 6.25 | 6.36 | 6.48 | 6.6 | 6.72 | 6.84 | 6.95 | 7.07 | 7.19 | 7.31 | 7.42 | 7.54 | 7.66 | 7.78 | 7.9 | 8.01 | 8.13 | 8.25 | 8.37 | 8.48 | 8.6 | 8.72 | 8.84 | 8.96 | 9.07 | 9.19
22.15 | 4.48 | 4.6 | 4.71 | 4.83 | 4.95 | 5.07 | 5.19 | 5.3 | 5.42 | 5.54 | 5.66 | 5.77 | 5.89 | 6.01 | 6.13 | 6.25 | 6.36 | 6.48 | 6.6 | 6.72 | 6.83 | 6.95 | 7.07 | 7.19 | 7.31 | 7.42 | 7.54 | 7.66 | 7.78 | 7.89 | 8.01 | 8.13 | 8.25 | 8.37 | 8.48 | 8.6 | 8.72 | 8.84 | 8.95 | 9.07 | 9.19
22.16 | 4.48 | 4.6 | 4.71 | 4.83 | 4.95 | 5.07 | 5.18 | 5.3 | 5.42 | 5.54 | 5.66 | 5.77 | 5.89 | 6.01 | 6.13 | 6.24 | 6.36 | 6.48 | 6.6 | 6.72 | 6.83 | 6.95 | 7.07 | 7.19 | 7.3 | 7.42 | 7.54 | 7.66 | 7.78 | 7.89 | 8.01 | 8.13 | 8.25 | 8.36 | 8.48 | 8.6 | 8.72 | 8.83 | 8.95 | 9.07 | 9.19
22.17 | 4.48 | 4.59 | 4.71 | 4.83 | 4.95 | 5.07 | 5.18 | 5.3 | 5.42 | 5.54 | 5.65 | 5.77 | 5.89 | 6.01 | 6.13 | 6.24 | 6.36 | 6.48 | 6.6 | 6.71 | 6.83 | 6.95 | 7.07 | 7.18 | 7.3 | 7.42 | 7.54 | 7.66 | 7.77 | 7.89 | 8.01 | 8.13 | 8.24 | 8.36 | 8.48 | 8.6 | 8.72 | 8.83 | 8.95 | 9.07 | 9.19
22.18 | 4.48 | 4.59 | 4.71 | 4.83 | 4.95 | 5.06 | 5.18 | 5.3 | 5.42 | 5.54 | 5.65 | 5.77 | 5.89 | 6.01 | 6.12 | 6.24 | 6.36 | 6.48 | 6.59 | 6.71 | 6.83 | 6.95 | 7.07 | 7.18 | 7.3 | 7.42 | 7.54 | 7.65 | 7.77 | 7.89 | 8.01 | 8.13 | 8.24 | 8.36 | 8.48 | 8.6 | 8.71 | 8.83 | 8.95 | 9.07 | 9.18
22.19 | 4.47 | 4.59 | 4.71 | 4.83 | 4.95 | 5.06 | 5.18 | 5.3 | 5.42 | 5.53 | 5.65 | 5.77 | 5.89 | 6.01 | 6.12 | 6.24 | 6.36 | 6.48 | 6.59 | 6.71 | 6.83 | 6.95 | 7.06 | 7.18 | 7.3 | 7.42 | 7.54 | 7.65 | 7.77 | 7.89 | 8.01 | 8.12 | 8.24 | 8.36 | 8.48 | 8.59 | 8.71 | 8.83 | 8.95 | 9.06 | 9.18
22.2 | 4.47 | 4.59 | 4.71 | 4.83 | 4.94 | 5.06 | 5.18 | 5.3 | 5.42 | 5.53 | 5.65 | 5.77 | 5.89 | 6 | 6.12 | 6.24 | 6.36 | 6.47 | 6.59 | 6.71 | 6.83 | 6.95 | 7.06 | 7.18 | 7.3 | 7.42 | 7.53 | 7.65 | 7.77 | 7.89 | 8 | 8.12 | 8.24 | 8.36 | 8.47 | 8.59 | 8.71 | 8.83 | 8.95 | 9.06 | 9.18
22.21 | 4.47 | 4.59 | 4.71 | 4.83 | 4.94 | 5.06 | 5.18 | 5.3 | 5.41 | 5.53 | 5.65 | 5.77 | 5.89 | 6 | 6.12 | 6.24 | 6.36 | 6.47 | 6.59 | 6.71 | 6.83 | 6.94 | 7.06 | 7.18 | 7.3 | 7.41 | 7.53 | 7.65 | 7.77 | 7.89 | 8 | 8.12 | 8.24 | 8.36 | 8.47 | 8.59 | 8.71 | 8.83 | 8.94 | 9.06 | 9.18
22.22 | 4.47 | 4.59 | 4.71 | 4.82 | 4.94 | 5.06 | 5.18 | 5.3 | 5.41 | 5.53 | 5.65 | 5.77 | 5.88 | 6 | 6.12 | 6.24 | 6.35 | 6.47 | 6.59 | 6.71 | 6.83 | 6.94 | 7.06 | 7.18 | 7.3 | 7.41 | 7.53 | 7.65 | 7.77 | 7.88 | 8 | 8.12 | 8.24 | 8.35 | 8.47 | 8.59 | 8.71 | 8.82 | 8.94 | 9.06 | 9.18
22.23 | 4.47 | 4.59 | 4.71 | 4.82 | 4.94 | 5.06 | 5.18 | 5.29 | 5.41 | 5.53 | 5.65 | 5.77 | 5.88 | 6 | 6.12 | 6.24 | 6.35 | 6.47 | 6.59 | 6.71 | 6.82 | 6.94 | 7.06 | 7.18 | 7.29 | 7.41 | 7.53 | 7.65 | 7.76 | 7.88 | 8 | 8.12 | 8.23 | 8.35 | 8.47 | 8.59 | 8.71 | 8.82 | 8.94 | 9.06 | 9.18
22.24 | 4.47 | 4.59 | 4.71 | 4.82 | 4.94 | 5.06 | 5.18 | 5.29 | 5.41 | 5.53 | 5.65 | 5.76 | 5.88 | 6 | 6.12 | 6.23 | 6.35 | 6.47 | 6.59 | 6.7 | 6.82 | 6.94 | 7.06 | 7.18 | 7.29 | 7.41 | 7.53 | 7.65 | 7.76 | 7.88 | 8 | 8.12 | 8.23 | 8.35 | 8.47 | 8.59 | 8.7 | 8.82 | 8.94 | 9.06 | 9.17
22.25 | 4.47 | 4.59 | 4.7 | 4.82 | 4.94 | 5.06 | 5.17 | 5.29 | 5.41 | 5.53 | 5.65 | 5.76 | 5.88 | 6 | 6.12 | 6.23 | 6.35 | 6.47 | 6.59 | 6.7 | 6.82 | 6.94 | 7.06 | 7.17 | 7.29 | 7.41 | 7.53 | 7.64 | 7.76 | 7.88 | 8 | 8.11 | 8.23 | 8.35 | 8.47 | 8.58 | 8.7 | 8.82 | 8.94 | 9.05 | 9.17
22.26 | 4.47 | 4.59 | 4.7 | 4.82 | 4.94 | 5.06 | 5.17 | 5.29 | 5.41 | 5.53 | 5.64 | 5.76 | 5.88 | 6 | 6.11 | 6.23 | 6.35 | 6.47 | 6.58 | 6.7 | 6.82 | 6.94 | 7.05 | 7.17 | 7.29 | 7.41 | 7.52 | 7.64 | 7.76 | 7.88 | 8 | 8.11 | 8.23 | 8.35 | 8.47 | 8.58 | 8.7 | 8.82 | 8.94 | 9.05 | 9.17
22.27 | 4.47 | 4.58 | 4.7 | 4.82 | 4.94 | 5.06 | 5.17 | 5.29 | 5.41 | 5.53 | 5.64 | 5.76 | 5.88 | 6 | 6.11 | 6.23 | 6.35 | 6.47 | 6.58 | 6.7 | 6.82 | 6.94 | 7.05 | 7.17 | 7.29 | 7.41 | 7.52 | 7.64 | 7.76 | 7.88 | 7.99 | 8.11 | 8.23 | 8.35 | 8.46 | 8.58 | 8.7 | 8.82 | 8.93 | 9.05 | 9.17
22.28 | 4.47 | 4.58 | 4.7 | 4.82 | 4.94 | 5.05 | 5.17 | 5.29 | 5.41 | 5.52 | 5.64 | 5.76 | 5.88 | 5.99 | 6.11 | 6.23 | 6.35 | 6.46 | 6.58 | 6.7 | 6.82 | 6.93 | 7.05 | 7.17 | 7.29 | 7.4 | 7.52 | 7.64 | 7.76 | 7.87 | 7.99 | 8.11 | 8.23 | 8.34 | 8.46 | 8.58 | 8.7 | 8.81 | 8.93 | 9.05 | 9.17
22.29 | 4.47 | 4.58 | 4.7 | 4.82 | 4.94 | 5.05 | 5.17 | 5.29 | 5.41 | 5.52 | 5.64 | 5.76 | 5.88 | 5.99 | 6.11 | 6.23 | 6.35 | 6.46 | 6.58 | 6.7 | 6.82 | 6.93 | 7.05 | 7.17 | 7.29 | 7.4 | 7.52 | 7.64 | 7.76 | 7.87 | 7.99 | 8.11 | 8.23 | 8.34 | 8.46 | 8.58 | 8.7 | 8.81 | 8.93 | 9.05 | 9.16
22.3 | 4.46 | 4.58 | 4.7 | 4.82 | 4.93 | 5.05 | 5.17 | 5.29 | 5.4 | 5.52 | 5.64 | 5.76 | 5.87 | 5.99 | 6.11 | 6.23 | 6.34 | 6.46 | 6.58 | 6.7 | 6.81 | 6.93 | 7.05 | 7.17 | 7.28 | 7.4 | 7.52 | 7.64 | 7.75 | 7.87 | 7.99 | 8.11 | 8.22 | 8.34 | 8.46 | 8.58 | 8.69 | 8.81 | 8.93 | 9.05 | 9.16
22.31 | 4.46 | 4.58 | 4.7 | 4.82 | 4.93 | 5.05 | 5.17 | 5.29 | 5.4 | 5.52 | 5.64 | 5.76 | 5.87 | 5.99 | 6.11 | 6.23 | 6.34 | 6.46 | 6.58 | 6.7 | 6.81 | 6.93 | 7.05 | 7.17 | 7.28 | 7.4 | 7.52 | 7.64 | 7.75 | 7.87 | 7.99 | 8.1 | 8.22 | 8.34 | 8.46 | 8.57 | 8.69 | 8.81 | 8.93 | 9.04 | 9.16
22.32 | 4.46 | 4.58 | 4.7 | 4.82 | 4.93 | 5.05 | 5.17 | 5.28 | 5.4 | 5.52 | 5.64 | 5.75 | 5.87 | 5.99 | 6.11 | 6.22 | 6.34 | 6.46 | 6.58 | 6.69 | 6.81 | 6.93 | 7.05 | 7.16 | 7.28 | 7.4 | 7.52 | 7.63 | 7.75 | 7.87 | 7.99 | 8.1 | 8.22 | 8.34 | 8.46 | 8.57 | 8.69 | 8.81 | 8.92 | 9.04 | 9.16
22.33 | 4.46 | 4.58 | 4.7 | 4.81 | 4.93 | 5.05 | 5.17 | 5.28 | 5.4 | 5.52 | 5.64 | 5.75 | 5.87 | 5.99 | 6.11 | 6.22 | 6.34 | 6.46 | 6.58 | 6.69 | 6.81 | 6.93 | 7.05 | 7.16 | 7.28 | 7.4 | 7.51 | 7.63 | 7.75 | 7.87 | 7.98 | 8.1 | 8.22 | 8.34 | 8.45 | 8.57 | 8.69 | 8.81 | 8.92 | 9.04 | 9.16
22.34 | 4.46 | 4.58 | 4.7 | 4.81 | 4.93 | 5.05 | 5.17 | 5.28 | 5.4 | 5.52 | 5.64 | 5.75 | 5.87 | 5.99 | 6.1 | 6.22 | 6.34 | 6.46 | 6.57 | 6.69 | 6.81 | 6.93 | 7.04 | 7.16 | 7.28 | 7.4 | 7.51 | 7.63 | 7.75 | 7.87 | 7.98 | 8.1 | 8.22 | 8.33 | 8.45 | 8.57 | 8.69 | 8.8 | 8.92 | 9.04 | 9.16
22.35 | 4.46 | 4.58 | 4.69 | 4.81 | 4.93 | 5.05 | 5.16 | 5.28 | 5.4 | 5.52 | 5.63 | 5.75 | 5.87 | 5.99 | 6.1 | 6.22 | 6.34 | 6.46 | 6.57 | 6.69 | 6.81 | 6.92 | 7.04 | 7.16 | 7.28 | 7.39 | 7.51 | 7.63 | 7.75 | 7.86 | 7.98 | 8.1 | 8.22 | 8.33 | 8.45 | 8.57 | 8.69 | 8.8 | 8.92 | 9.04 | 9.15
22.36 | 4.46 | 4.58 | 4.69 | 4.81 | 4.93 | 5.05 | 5.16 | 5.28 | 5.4 | 5.52 | 5.63 | 5.75 | 5.87 | 5.98 | 6.1 | 6.22 | 6.34 | 6.45 | 6.57 | 6.69 | 6.81 | 6.92 | 7.04 | 7.16 | 7.28 | 7.39 | 7.51 | 7.63 | 7.74 | 7.86 | 7.98 | 8.1 | 8.21 | 8.33 | 8.45 | 8.57 | 8.68 | 8.8 | 8.92 | 9.04 | 9.15
22.37 | 4.46 | 4.58 | 4.69 | 4.81 | 4.93 | 5.04 | 5.16 | 5.28 | 5.4 | 5.51 | 5.63 | 5.75 | 5.87 | 5.98 | 6.1 | 6.22 | 6.34 | 6.45 | 6.57 | 6.69 | 6.8 | 6.92 | 7.04 | 7.16 | 7.27 | 7.39 | 7.51 | 7.63 | 7.74 | 7.86 | 7.98 | 8.1 | 8.21 | 8.33 | 8.45 | 8.56 | 8.68 | 8.8 | 8.92 | 9.03 | 9.15
22.38 | 4.46 | 4.57 | 4.69 | 4.81 | 4.93 | 5.04 | 5.16 | 5.28 | 5.4 | 5.51 | 5.63 | 5.75 | 5.87 | 5.98 | 6.1 | 6.22 | 6.33 | 6.45 | 6.57 | 6.69 | 6.8 | 6.92 | 7.04 | 7.16 | 7.27 | 7.39 | 7.51 | 7.62 | 7.74 | 7.86 | 7.98 | 8.09 | 8.21 | 8.33 | 8.45 | 8.56 | 8.68 | 8.8 | 8.91 | 9.03 | 9.15
22.39 | 4.46 | 4.57 | 4.69 | 4.81 | 4.93 | 5.04 | 5.16 | 5.28 | 5.39 | 5.51 | 5.63 | 5.75 | 5.86 | 5.98 | 6.1 | 6.22 | 6.33 | 6.45 | 6.57 | 6.68 | 6.8 | 6.92 | 7.04 | 7.15 | 7.27 | 7.39 | 7.51 | 7.62 | 7.74 | 7.86 | 7.97 | 8.09 | 8.21 | 8.33 | 8.44 | 8.56 | 8.68 | 8.8 | 8.91 | 9.03 | 9.15
22.4 | 4.46 | 4.57 | 4.69 | 4.81 | 4.92 | 5.04 | 5.16 | 5.28 | 5.39 | 5.51 | 5.63 | 5.75 | 5.86 | 5.98 | 6.1 | 6.21 | 6.33 | 6.45 | 6.57 | 6.68 | 6.8 | 6.92 | 7.04 | 7.15 | 7.27 | 7.39 | 7.5 | 7.62 | 7.74 | 7.86 | 7.97 | 8.09 | 8.21 | 8.33 | 8.44 | 8.56 | 8.68 | 8.79 | 8.91 | 9.03 | 9.15
22.41 | 4.45 | 4.57 | 4.69 | 4.81 | 4.92 | 5.04 | 5.16 | 5.28 | 5.39 | 5.51 | 5.63 | 5.74 | 5.86 | 5.98 | 6.1 | 6.21 | 6.33 | 6.45 | 6.57 | 6.68 | 6.8 | 6.92 | 7.03 | 7.15 | 7.27 | 7.39 | 7.5 | 7.62 | 7.74 | 7.85 | 7.97 | 8.09 | 8.21 | 8.32 | 8.44 | 8.56 | 8.68 | 8.79 | 8.91 | 9.03 | 9.14
22.42 | 4.45 | 4.57 | 4.69 | 4.81 | 4.92 | 5.04 | 5.16 | 5.27 | 5.39 | 5.51 | 5.63 | 5.74 | 5.86 | 5.98 | 6.09 | 6.21 | 6.33 | 6.45 | 6.56 | 6.68 | 6.8 | 6.92 | 7.03 | 7.15 | 7.27 | 7.38 | 7.5 | 7.62 | 7.74 | 7.85 | 7.97 | 8.09 | 8.2 | 8.32 | 8.44 | 8.56 | 8.67 | 8.79 | 8.91 | 9.03 | 9.14
22.43 | 4.45 | 4.57 | 4.69 | 4.8 | 4.92 | 5.04 | 5.16 | 5.27 | 5.39 | 5.51 | 5.62 | 5.74 | 5.86 | 5.98 | 6.09 | 6.21 | 6.33 | 6.45 | 6.56 | 6.68 | 6.8 | 6.91 | 7.03 | 7.15 | 7.27 | 7.38 | 7.5 | 7.62 | 7.73 | 7.85 | 7.97 | 8.09 | 8.2 | 8.32 | 8.44 | 8.55 | 8.67 | 8.79 | 8.91 | 9.02 | 9.14
22.44 | 4.45 | 4.57 | 4.69 | 4.8 | 4.92 | 5.04 | 5.15 | 5.27 | 5.39 | 5.51 | 5.62 | 5.74 | 5.86 | 5.98 | 6.09 | 6.21 | 6.33 | 6.44 | 6.56 | 6.68 | 6.8 | 6.91 | 7.03 | 7.15 | 7.26 | 7.38 | 7.5 | 7.62 | 7.73 | 7.85 | 7.97 | 8.08 | 8.2 | 8.32 | 8.44 | 8.55 | 8.67 | 8.79 | 8.9 | 9.02 | 9.14
22.45 | 4.45 | 4.57 | 4.69 | 4.8 | 4.92 | 5.04 | 5.15 | 5.27 | 5.39 | 5.51 | 5.62 | 5.74 | 5.86 | 5.97 | 6.09 | 6.21 | 6.33 | 6.44 | 6.56 | 6.68 | 6.79 | 6.91 | 7.03 | 7.15 | 7.26 | 7.38 | 7.5 | 7.61 | 7.73 | 7.85 | 7.97 | 8.08 | 8.2 | 8.32 | 8.43 | 8.55 | 8.67 | 8.79 | 8.9 | 9.02 | 9.14
22.46 | 4.45 | 4.57 | 4.68 | 4.8 | 4.92 | 5.04 | 5.15 | 5.27 | 5.39 | 5.5 | 5.62 | 5.74 | 5.86 | 5.97 | 6.09 | 6.21 | 6.32 | 6.44 | 6.56 | 6.68 | 6.79 | 6.91 | 7.03 | 7.14 | 7.26 | 7.38 | 7.5 | 7.61 | 7.73 | 7.85 | 7.96 | 8.08 | 8.2 | 8.32 | 8.43 | 8.55 | 8.67 | 8.78 | 8.9 | 9.02 | 9.14
22.47 | 4.45 | 4.57 | 4.68 | 4.8 | 4.92 | 5.03 | 5.15 | 5.27 | 5.39 | 5.5 | 5.62 | 5.74 | 5.85 | 5.97 | 6.09 | 6.21 | 6.32 | 6.44 | 6.56 | 6.67 | 6.79 | 6.91 | 7.03 | 7.14 | 7.26 | 7.38 | 7.49 | 7.61 | 7.73 | 7.85 | 7.96 | 8.08 | 8.2 | 8.31 | 8.43 | 8.55 | 8.67 | 8.78 | 8.9 | 9.02 | 9.13
22.48 | 4.45 | 4.57 | 4.68 | 4.8 | 4.92 | 5.03 | 5.15 | 5.27 | 5.38 | 5.5 | 5.62 | 5.74 | 5.85 | 5.97 | 6.09 | 6.2 | 6.32 | 6.44 | 6.56 | 6.67 | 6.79 | 6.91 | 7.02 | 7.14 | 7.26 | 7.38 | 7.49 | 7.61 | 7.73 | 7.84 | 7.96 | 8.08 | 8.2 | 8.31 | 8.43 | 8.55 | 8.66 | 8.78 | 8.9 | 9.01 | 9.13
22.49 | 4.45 | 4.56 | 4.68 | 4.8 | 4.92 | 5.03 | 5.15 | 5.27 | 5.38 | 5.5 | 5.62 | 5.74 | 5.85 | 5.97 | 6.09 | 6.2 | 6.32 | 6.44 | 6.55 | 6.67 | 6.79 | 6.91 | 7.02 | 7.14 | 7.26 | 7.37 | 7.49 | 7.61 | 7.73 | 7.84 | 7.96 | 8.08 | 8.19 | 8.31 | 8.43 | 8.54 | 8.66 | 8.78 | 8.9 | 9.01 | 9.13
22.5 | 4.45 | 4.56 | 4.68 | 4.8 | 4.91 | 5.03 | 5.15 | 5.27 | 5.38 | 5.5 | 5.62 | 5.73 | 5.85 | 5.97 | 6.09 | 6.2 | 6.32 | 6.44 | 6.55 | 6.67 | 6.79 | 6.9 | 7.02 | 7.14 | 7.26 | 7.37 | 7.49 | 7.61 | 7.72 | 7.84 | 7.96 | 8.08 | 8.19 | 8.31 | 8.43 | 8.54 | 8.66 | 8.78 | 8.89 | 9.01 | 9.13
22.51 | 4.45 | 4.56 | 4.68 | 4.8 | 4.91 | 5.03 | 5.15 | 5.26 | 5.38 | 5.5 | 5.62 | 5.73 | 5.85 | 5.97 | 6.08 | 6.2 | 6.32 | 6.44 | 6.55 | 6.67 | 6.79 | 6.9 | 7.02 | 7.14 | 7.25 | 7.37 | 7.49 | 7.61 | 7.72 | 7.84 | 7.96 | 8.07 | 8.19 | 8.31 | 8.42 | 8.54 | 8.66 | 8.78 | 8.89 | 9.01 | 9.13
22.52 | 4.44 | 4.56 | 4.68 | 4.8 | 4.91 | 5.03 | 5.15 | 5.26 | 5.38 | 5.5 | 5.61 | 5.73 | 5.85 | 5.97 | 6.08 | 6.2 | 6.32 | 6.43 | 6.55 | 6.67 | 6.78 | 6.9 | 7.02 | 7.14 | 7.25 | 7.37 | 7.49 | 7.6 | 7.72 | 7.84 | 7.96 | 8.07 | 8.19 | 8.31 | 8.42 | 8.54 | 8.66 | 8.77 | 8.89 | 9.01 | 9.12
22.53 | 4.44 | 4.56 | 4.68 | 4.79 | 4.91 | 5.03 | 5.15 | 5.26 | 5.38 | 5.5 | 5.61 | 5.73 | 5.85 | 5.96 | 6.08 | 6.2 | 6.32 | 6.43 | 6.55 | 6.67 | 6.78 | 6.9 | 7.02 | 7.13 | 7.25 | 7.37 | 7.49 | 7.6 | 7.72 | 7.84 | 7.95 | 8.07 | 8.19 | 8.3 | 8.42 | 8.54 | 8.66 | 8.77 | 8.89 | 9.01 | 9.12
22.54 | 4.44 | 4.56 | 4.68 | 4.79 | 4.91 | 5.03 | 5.14 | 5.26 | 5.38 | 5.5 | 5.61 | 5.73 | 5.85 | 5.96 | 6.08 | 6.2 | 6.31 | 6.43 | 6.55 | 6.67 | 6.78 | 6.9 | 7.02 | 7.13 | 7.25 | 7.37 | 7.48 | 7.6 | 7.72 | 7.84 | 7.95 | 8.07 | 8.19 | 8.3 | 8.42 | 8.54 | 8.65 | 8.77 | 8.89 | 9 | 9.12
22.55 | 4.44 | 4.56 | 4.68 | 4.79 | 4.91 | 5.03 | 5.14 | 5.26 | 5.38 | 5.49 | 5.61 | 5.73 | 5.85 | 5.96 | 6.08 | 6.2 | 6.31 | 6.43 | 6.55 | 6.66 | 6.78 | 6.9 | 7.01 | 7.13 | 7.25 | 7.37 | 7.48 | 7.6 | 7.72 | 7.83 | 7.95 | 8.07 | 8.18 | 8.3 | 8.42 | 8.54 | 8.65 | 8.77 | 8.89 | 9 | 9.12
22.56 | 4.44 | 4.56 | 4.67 | 4.79 | 4.91 | 5.03 | 5.14 | 5.26 | 5.38 | 5.49 | 5.61 | 5.73 | 5.84 | 5.96 | 6.08 | 6.19 | 6.31 | 6.43 | 6.55 | 6.66 | 6.78 | 6.9 | 7.01 | 7.13 | 7.25 | 7.36 | 7.48 | 7.6 | 7.72 | 7.83 | 7.95 | 8.07 | 8.18 | 8.3 | 8.42 | 8.53 | 8.65 | 8.77 | 8.88 | 9 | 9.12
22.57 | 4.44 | 4.56 | 4.67 | 4.79 | 4.91 | 5.02 | 5.14 | 5.26 | 5.38 | 5.49 | 5.61 | 5.73 | 5.84 | 5.96 | 6.08 | 6.19 | 6.31 | 6.43 | 6.54 | 6.66 | 6.78 | 6.9 | 7.01 | 7.13 | 7.25 | 7.36 | 7.48 | 7.6 | 7.71 | 7.83 | 7.95 | 8.06 | 8.18 | 8.3 | 8.41 | 8.53 | 8.65 | 8.77 | 8.88 | 9 | 9.12
22.58 | 4.44 | 4.56 | 4.67 | 4.79 | 4.91 | 5.02 | 5.14 | 5.26 | 5.37 | 5.49 | 5.61 | 5.72 | 5.84 | 5.96 | 6.08 | 6.19 | 6.31 | 6.43 | 6.54 | 6.66 | 6.78 | 6.89 | 7.01 | 7.13 | 7.24 | 7.36 | 7.48 | 7.6 | 7.71 | 7.83 | 7.95 | 8.06 | 8.18 | 8.3 | 8.41 | 8.53 | 8.65 | 8.76 | 8.88 | 9 | 9.11
22.59 | 4.44 | 4.55 | 4.67 | 4.79 | 4.91 | 5.02 | 5.14 | 5.26 | 5.37 | 5.49 | 5.61 | 5.72 | 5.84 | 5.96 | 6.07 | 6.19 | 6.31 | 6.43 | 6.54 | 6.66 | 6.78 | 6.89 | 7.01 | 7.13 | 7.24 | 7.36 | 7.48 | 7.59 | 7.71 | 7.83 | 7.94 | 8.06 | 8.18 | 8.29 | 8.41 | 8.53 | 8.65 | 8.76 | 8.88 | 9 | 9.11
22.6 | 4.44 | 4.55 | 4.67 | 4.79 | 4.9 | 5.02 | 5.14 | 5.26 | 5.37 | 5.49 | 5.61 | 5.72 | 5.84 | 5.96 | 6.07 | 6.19 | 6.31 | 6.42 | 6.54 | 6.66 | 6.77 | 6.89 | 7.01 | 7.12 | 7.24 | 7.36 | 7.48 | 7.59 | 7.71 | 7.83 | 7.94 | 8.06 | 8.18 | 8.29 | 8.41 | 8.53 | 8.64 | 8.76 | 8.88 | 8.99 | 9.11
22.61 | 4.44 | 4.55 | 4.67 | 4.79 | 4.9 | 5.02 | 5.14 | 5.25 | 5.37 | 5.49 | 5.6 | 5.72 | 5.84 | 5.96 | 6.07 | 6.19 | 6.31 | 6.42 | 6.54 | 6.66 | 6.77 | 6.89 | 7.01 | 7.12 | 7.24 | 7.36 | 7.47 | 7.59 | 7.71 | 7.82 | 7.94 | 8.06 | 8.17 | 8.29 | 8.41 | 8.53 | 8.64 | 8.76 | 8.88 | 8.99 | 9.11
22.62 | 4.44 | 4.55 | 4.67 | 4.79 | 4.9 | 5.02 | 5.14 | 5.25 | 5.37 | 5.49 | 5.6 | 5.72 | 5.84 | 5.95 | 6.07 | 6.19 | 6.3 | 6.42 | 6.54 | 6.65 | 6.77 | 6.89 | 7.01 | 7.12 | 7.24 | 7.36 | 7.47 | 7.59 | 7.71 | 7.82 | 7.94 | 8.06 | 8.17 | 8.29 | 8.41 | 8.52 | 8.64 | 8.76 | 8.87 | 8.99 | 9.11
22.63 | 4.43 | 4.55 | 4.67 | 4.78 | 4.9 | 5.02 | 5.14 | 5.25 | 5.37 | 5.49 | 5.6 | 5.72 | 5.84 | 5.95 | 6.07 | 6.19 | 6.3 | 6.42 | 6.54 | 6.65 | 6.77 | 6.89 | 7 | 7.12 | 7.24 | 7.35 | 7.47 | 7.59 | 7.7 | 7.82 | 7.94 | 8.05 | 8.17 | 8.29 | 8.41 | 8.52 | 8.64 | 8.76 | 8.87 | 8.99 | 9.11
22.64 | 4.43 | 4.55 | 4.67 | 4.78 | 4.9 | 5.02 | 5.13 | 5.25 | 5.37 | 5.48 | 5.6 | 5.72 | 5.83 | 5.95 | 6.07 | 6.19 | 6.3 | 6.42 | 6.54 | 6.65 | 6.77 | 6.89 | 7 | 7.12 | 7.24 | 7.35 | 7.47 | 7.59 | 7.7 | 7.82 | 7.94 | 8.05 | 8.17 | 8.29 | 8.4 | 8.52 | 8.64 | 8.75 | 8.87 | 8.99 | 9.1
22.65 | 4.43 | 4.55 | 4.67 | 4.78 | 4.9 | 5.02 | 5.13 | 5.25 | 5.37 | 5.48 | 5.6 | 5.72 | 5.83 | 5.95 | 6.07 | 6.18 | 6.3 | 6.42 | 6.53 | 6.65 | 6.77 | 6.88 | 7 | 7.12 | 7.23 | 7.35 | 7.47 | 7.58 | 7.7 | 7.82 | 7.94 | 8.05 | 8.17 | 8.29 | 8.4 | 8.52 | 8.64 | 8.75 | 8.87 | 8.99 | 9.1
22.66 | 4.43 | 4.55 | 4.66 | 4.78 | 4.9 | 5.02 | 5.13 | 5.25 | 5.37 | 5.48 | 5.6 | 5.72 | 5.83 | 5.95 | 6.07 | 6.18 | 6.3 | 6.42 | 6.53 | 6.65 | 6.77 | 6.88 | 7 | 7.12 | 7.23 | 7.35 | 7.47 | 7.58 | 7.7 | 7.82 | 7.93 | 8.05 | 8.17 | 8.28 | 8.4 | 8.52 | 8.63 | 8.75 | 8.87 | 8.98 | 9.1
22.67 | 4.43 | 4.55 | 4.66 | 4.78 | 4.9 | 5.01 | 5.13 | 5.25 | 5.36 | 5.48 | 5.6 | 5.71 | 5.83 | 5.95 | 6.06 | 6.18 | 6.3 | 6.41 | 6.53 | 6.65 | 6.77 | 6.88 | 7 | 7.12 | 7.23 | 7.35 | 7.47 | 7.58 | 7.7 | 7.82 | 7.93 | 8.05 | 8.17 | 8.28 | 8.4 | 8.52 | 8.63 | 8.75 | 8.87 | 8.98 | 9.1
22.68 | 4.43 | 4.55 | 4.66 | 4.78 | 4.9 | 5.01 | 5.13 | 5.25 | 5.36 | 5.48 | 5.6 | 5.71 | 5.83 | 5.95 | 6.06 | 6.18 | 6.3 | 6.41 | 6.53 | 6.65 | 6.76 | 6.88 | 7 | 7.11 | 7.23 | 7.35 | 7.46 | 7.58 | 7.7 | 7.81 | 7.93 | 8.05 | 8.16 | 8.28 | 8.4 | 8.51 | 8.63 | 8.75 | 8.86 | 8.98 | 9.1
22.69 | 4.43 | 4.55 | 4.66 | 4.78 | 4.9 | 5.01 | 5.13 | 5.25 | 5.36 | 5.48 | 5.6 | 5.71 | 5.83 | 5.95 | 6.06 | 6.18 | 6.3 | 6.41 | 6.53 | 6.65 | 6.76 | 6.88 | 7 | 7.11 | 7.23 | 7.35 | 7.46 | 7.58 | 7.7 | 7.81 | 7.93 | 8.05 | 8.16 | 8.28 | 8.4 | 8.51 | 8.63 | 8.75 | 8.86 | 8.98 | 9.1
22.7 | 4.43 | 4.54 | 4.66 | 4.78 | 4.89 | 5.01 | 5.13 | 5.24 | 5.36 | 5.48 | 5.59 | 5.71 | 5.83 | 5.94 | 6.06 | 6.18 | 6.29 | 6.41 | 6.53 | 6.64 | 6.76 | 6.88 | 6.99 | 7.11 | 7.23 | 7.34 | 7.46 | 7.58 | 7.69 | 7.81 | 7.93 | 8.04 | 8.16 | 8.28 | 8.39 | 8.51 | 8.63 | 8.74 | 8.86 | 8.98 | 9.09
22.71 | 4.43 | 4.54 | 4.66 | 4.78 | 4.89 | 5.01 | 5.13 | 5.24 | 5.36 | 5.48 | 5.59 | 5.71 | 5.83 | 5.94 | 6.06 | 6.18 | 6.29 | 6.41 | 6.53 | 6.64 | 6.76 | 6.88 | 6.99 | 7.11 | 7.23 | 7.34 | 7.46 | 7.58 | 7.69 | 7.81 | 7.93 | 8.04 | 8.16 | 8.28 | 8.39 | 8.51 | 8.63 | 8.74 | 8.86 | 8.98 | 9.09
22.72 | 4.43 | 4.54 | 4.66 | 4.78 | 4.89 | 5.01 | 5.13 | 5.24 | 5.36 | 5.48 | 5.59 | 5.71 | 5.83 | 5.94 | 6.06 | 6.18 | 6.29 | 6.41 | 6.53 | 6.64 | 6.76 | 6.88 | 6.99 | 7.11 | 7.22 | 7.34 | 7.46 | 7.57 | 7.69 | 7.81 | 7.92 | 8.04 | 8.16 | 8.27 | 8.39 | 8.51 | 8.62 | 8.74 | 8.86 | 8.97 | 9.09
22.73 | 4.42 | 4.54 | 4.66 | 4.77 | 4.89 | 5.01 | 5.12 | 5.24 | 5.36 | 5.47 | 5.59 | 5.71 | 5.82 | 5.94 | 6.06 | 6.17 | 6.29 | 6.41 | 6.52 | 6.64 | 6.76 | 6.87 | 6.99 | 7.11 | 7.22 | 7.34 | 7.46 | 7.57 | 7.69 | 7.81 | 7.92 | 8.04 | 8.16 | 8.27 | 8.39 | 8.51 | 8.62 | 8.74 | 8.86 | 8.97 | 9.09
22.74 | 4.42 | 4.54 | 4.66 | 4.77 | 4.89 | 5.01 | 5.12 | 5.24 | 5.36 | 5.47 | 5.59 | 5.71 | 5.82 | 5.94 | 6.06 | 6.17 | 6.29 | 6.41 | 6.52 | 6.64 | 6.76 | 6.87 | 6.99 | 7.11 | 7.22 | 7.34 | 7.46 | 7.57 | 7.69 | 7.8 | 7.92 | 8.04 | 8.15 | 8.27 | 8.39 | 8.5 | 8.62 | 8.74 | 8.85 | 8.97 | 9.09
22.75 | 4.42 | 4.54 | 4.66 | 4.77 | 4.89 | 5.01 | 5.12 | 5.24 | 5.36 | 5.47 | 5.59 | 5.71 | 5.82 | 5.94 | 6.06 | 6.17 | 6.29 | 6.4 | 6.52 | 6.64 | 6.75 | 6.87 | 6.99 | 7.1 | 7.22 | 7.34 | 7.45 | 7.57 | 7.69 | 7.8 | 7.92 | 8.04 | 8.15 | 8.27 | 8.39 | 8.5 | 8.62 | 8.74 | 8.85 | 8.97 | 9.09
22.76 | 4.42 | 4.54 | 4.66 | 4.77 | 4.89 | 5.01 | 5.12 | 5.24 | 5.35 | 5.47 | 5.59 | 5.7 | 5.82 | 5.94 | 6.05 | 6.17 | 6.29 | 6.4 | 6.52 | 6.64 | 6.75 | 6.87 | 6.99 | 7.1 | 7.22 | 7.34 | 7.45 | 7.57 | 7.69 | 7.8 | 7.92 | 8.03 | 8.15 | 8.27 | 8.38 | 8.5 | 8.62 | 8.73 | 8.85 | 8.97 | 9.08
22.77 | 4.42 | 4.54 | 4.65 | 4.77 | 4.89 | 5 | 5.12 | 5.24 | 5.35 | 5.47 | 5.59 | 5.7 | 5.82 | 5.94 | 6.05 | 6.17 | 6.29 | 6.4 | 6.52 | 6.64 | 6.75 | 6.87 | 6.98 | 7.1 | 7.22 | 7.33 | 7.45 | 7.57 | 7.68 | 7.8 | 7.92 | 8.03 | 8.15 | 8.27 | 8.38 | 8.5 | 8.62 | 8.73 | 8.85 | 8.97 | 9.08
22.78 | 4.42 | 4.54 | 4.65 | 4.77 | 4.89 | 5 | 5.12 | 5.24 | 5.35 | 5.47 | 5.59 | 5.7 | 5.82 | 5.94 | 6.05 | 6.17 | 6.28 | 6.4 | 6.52 | 6.63 | 6.75 | 6.87 | 6.98 | 7.1 | 7.22 | 7.33 | 7.45 | 7.57 | 7.68 | 7.8 | 7.92 | 8.03 | 8.15 | 8.26 | 8.38 | 8.5 | 8.61 | 8.73 | 8.85 | 8.96 | 9.08
22.79 | 4.42 | 4.54 | 4.65 | 4.77 | 4.89 | 5 | 5.12 | 5.24 | 5.35 | 5.47 | 5.58 | 5.7 | 5.82 | 5.93 | 6.05 | 6.17 | 6.28 | 6.4 | 6.52 | 6.63 | 6.75 | 6.87 | 6.98 | 7.1 | 7.22 | 7.33 | 7.45 | 7.56 | 7.68 | 7.8 | 7.91 | 8.03 | 8.15 | 8.26 | 8.38 | 8.5 | 8.61 | 8.73 | 8.85 | 8.96 | 9.08
22.8 | 4.42 | 4.54 | 4.65 | 4.77 | 4.88 | 5 | 5.12 | 5.23 | 5.35 | 5.47 | 5.58 | 5.7 | 5.82 | 5.93 | 6.05 | 6.17 | 6.28 | 6.4 | 6.52 | 6.63 | 6.75 | 6.86 | 6.98 | 7.1 | 7.21 | 7.33 | 7.45 | 7.56 | 7.68 | 7.8 | 7.91 | 8.03 | 8.15 | 8.26 | 8.38 | 8.49 | 8.61 | 8.73 | 8.84 | 8.96 | 9.08
22.81 | 4.42 | 4.53 | 4.65 | 4.77 | 4.88 | 5 | 5.12 | 5.23 | 5.35 | 5.47 | 5.58 | 5.7 | 5.82 | 5.93 | 6.05 | 6.16 | 6.28 | 6.4 | 6.51 | 6.63 | 6.75 | 6.86 | 6.98 | 7.1 | 7.21 | 7.33 | 7.45 | 7.56 | 7.68 | 7.79 | 7.91 | 8.03 | 8.14 | 8.26 | 8.38 | 8.49 | 8.61 | 8.73 | 8.84 | 8.96 | 9.07
22.82 | 4.42 | 4.53 | 4.65 | 4.77 | 4.88 | 5 | 5.12 | 5.23 | 5.35 | 5.46 | 5.58 | 5.7 | 5.81 | 5.93 | 6.05 | 6.16 | 6.28 | 6.4 | 6.51 | 6.63 | 6.75 | 6.86 | 6.98 | 7.09 | 7.21 | 7.33 | 7.44 | 7.56 | 7.68 | 7.79 | 7.91 | 8.03 | 8.14 | 8.26 | 8.37 | 8.49 | 8.61 | 8.72 | 8.84 | 8.96 | 9.07
22.83 | 4.42 | 4.53 | 4.65 | 4.77 | 4.88 | 5 | 5.11 | 5.23 | 5.35 | 5.46 | 5.58 | 5.7 | 5.81 | 5.93 | 6.05 | 6.16 | 6.28 | 6.39 | 6.51 | 6.63 | 6.74 | 6.86 | 6.98 | 7.09 | 7.21 | 7.33 | 7.44 | 7.56 | 7.68 | 7.79 | 7.91 | 8.02 | 8.14 | 8.26 | 8.37 | 8.49 | 8.61 | 8.72 | 8.84 | 8.96 | 9.07
22.84 | 4.41 | 4.53 | 4.65 | 4.76 | 4.88 | 5 | 5.11 | 5.23 | 5.35 | 5.46 | 5.58 | 5.7 | 5.81 | 5.93 | 6.04 | 6.16 | 6.28 | 6.39 | 6.51 | 6.63 | 6.74 | 6.86 | 6.98 | 7.09 | 7.21 | 7.32 | 7.44 | 7.56 | 7.67 | 7.79 | 7.91 | 8.02 | 8.14 | 8.26 | 8.37 | 8.49 | 8.6 | 8.72 | 8.84 | 8.95 | 9.07
22.85 | 4.41 | 4.53 | 4.65 | 4.76 | 4.88 | 5 | 5.11 | 5.23 | 5.35 | 5.46 | 5.58 | 5.69 | 5.81 | 5.93 | 6.04 | 6.16 | 6.28 | 6.39 | 6.51 | 6.63 | 6.74 | 6.86 | 6.97 | 7.09 | 7.21 | 7.32 | 7.44 | 7.56 | 7.67 | 7.79 | 7.9 | 8.02 | 8.14 | 8.25 | 8.37 | 8.49 | 8.6 | 8.72 | 8.84 | 8.95 | 9.07
22.86 | 4.41 | 4.53 | 4.65 | 4.76 | 4.88 | 4.99 | 5.11 | 5.23 | 5.34 | 5.46 | 5.58 | 5.69 | 5.81 | 5.93 | 6.04 | 6.16 | 6.27 | 6.39 | 6.51 | 6.62 | 6.74 | 6.86 | 6.97 | 7.09 | 7.21 | 7.32 | 7.44 | 7.55 | 7.67 | 7.79 | 7.9 | 8.02 | 8.14 | 8.25 | 8.37 | 8.48 | 8.6 | 8.72 | 8.83 | 8.95 | 9.07
22.87 | 4.41 | 4.53 | 4.64 | 4.76 | 4.88 | 4.99 | 5.11 | 5.23 | 5.34 | 5.46 | 5.58 | 5.69 | 5.81 | 5.92 | 6.04 | 6.16 | 6.27 | 6.39 | 6.51 | 6.62 | 6.74 | 6.86 | 6.97 | 7.09 | 7.2 | 7.32 | 7.44 | 7.55 | 7.67 | 7.79 | 7.9 | 8.02 | 8.13 | 8.25 | 8.37 | 8.48 | 8.6 | 8.72 | 8.83 | 8.95 | 9.06
22.88 | 4.41 | 4.53 | 4.64 | 4.76 | 4.88 | 4.99 | 5.11 | 5.23 | 5.34 | 5.46 | 5.57 | 5.69 | 5.81 | 5.92 | 6.04 | 6.16 | 6.27 | 6.39 | 6.5 | 6.62 | 6.74 | 6.85 | 6.97 | 7.09 | 7.2 | 7.32 | 7.44 | 7.55 | 7.67 | 7.78 | 7.9 | 8.02 | 8.13 | 8.25 | 8.37 | 8.48 | 8.6 | 8.71 | 8.83 | 8.95 | 9.06
22.89 | 4.41 | 4.53 | 4.64 | 4.76 | 4.88 | 4.99 | 5.11 | 5.22 | 5.34 | 5.46 | 5.57 | 5.69 | 5.81 | 5.92 | 6.04 | 6.15 | 6.27 | 6.39 | 6.5 | 6.62 | 6.74 | 6.85 | 6.97 | 7.08 | 7.2 | 7.32 | 7.43 | 7.55 | 7.67 | 7.78 | 7.9 | 8.01 | 8.13 | 8.25 | 8.36 | 8.48 | 8.6 | 8.71 | 8.83 | 8.94 | 9.06
22.9 | 4.41 | 4.53 | 4.64 | 4.76 | 4.87 | 4.99 | 5.11 | 5.22 | 5.34 | 5.46 | 5.57 | 5.69 | 5.8 | 5.92 | 6.04 | 6.15 | 6.27 | 6.39 | 6.5 | 6.62 | 6.73 | 6.85 | 6.97 | 7.08 | 7.2 | 7.32 | 7.43 | 7.55 | 7.66 | 7.78 | 7.9 | 8.01 | 8.13 | 8.25 | 8.36 | 8.48 | 8.59 | 8.71 | 8.83 | 8.94 | 9.06
22.91 | 4.41 | 4.52 | 4.64 | 4.76 | 4.87 | 4.99 | 5.11 | 5.22 | 5.34 | 5.45 | 5.57 | 5.69 | 5.8 | 5.92 | 6.04 | 6.15 | 6.27 | 6.38 | 6.5 | 6.62 | 6.73 | 6.85 | 6.97 | 7.08 | 7.2 | 7.31 | 7.43 | 7.55 | 7.66 | 7.78 | 7.9 | 8.01 | 8.13 | 8.24 | 8.36 | 8.48 | 8.59 | 8.71 | 8.83 | 8.94 | 9.06
22.92 | 4.41 | 4.52 | 4.64 | 4.76 | 4.87 | 4.99 | 5.11 | 5.22 | 5.34 | 5.45 | 5.57 | 5.69 | 5.8 | 5.92 | 6.03 | 6.15 | 6.27 | 6.38 | 6.5 | 6.62 | 6.73 | 6.85 | 6.96 | 7.08 | 7.2 | 7.31 | 7.43 | 7.55 | 7.66 | 7.78 | 7.89 | 8.01 | 8.13 | 8.24 | 8.36 | 8.48 | 8.59 | 8.71 | 8.82 | 8.94 | 9.06
22.93 | 4.41 | 4.52 | 4.64 | 4.76 | 4.87 | 4.99 | 5.1 | 5.22 | 5.34 | 5.45 | 5.57 | 5.69 | 5.8 | 5.92 | 6.03 | 6.15 | 6.27 | 6.38 | 6.5 | 6.61 | 6.73 | 6.85 | 6.96 | 7.08 | 7.2 | 7.31 | 7.43 | 7.54 | 7.66 | 7.78 | 7.89 | 8.01 | 8.12 | 8.24 | 8.36 | 8.47 | 8.59 | 8.71 | 8.82 | 8.94 | 9.05
22.94 | 4.41 | 4.52 | 4.64 | 4.75 | 4.87 | 4.99 | 5.1 | 5.22 | 5.34 | 5.45 | 5.57 | 5.68 | 5.8 | 5.92 | 6.03 | 6.15 | 6.26 | 6.38 | 6.5 | 6.61 | 6.73 | 6.85 | 6.96 | 7.08 | 7.19 | 7.31 | 7.43 | 7.54 | 7.66 | 7.78 | 7.89 | 8.01 | 8.12 | 8.24 | 8.36 | 8.47 | 8.59 | 8.7 | 8.82 | 8.94 | 9.05
22.95 | 4.4 | 4.52 | 4.64 | 4.75 | 4.87 | 4.99 | 5.1 | 5.22 | 5.33 | 5.45 | 5.57 | 5.68 | 5.8 | 5.92 | 6.03 | 6.15 | 6.26 | 6.38 | 6.5 | 6.61 | 6.73 | 6.84 | 6.96 | 7.08 | 7.19 | 7.31 | 7.43 | 7.54 | 7.66 | 7.77 | 7.89 | 8.01 | 8.12 | 8.24 | 8.35 | 8.47 | 8.59 | 8.7 | 8.82 | 8.93 | 9.05
22.96 | 4.4 | 4.52 | 4.64 | 4.75 | 4.87 | 4.98 | 5.1 | 5.22 | 5.33 | 5.45 | 5.57 | 5.68 | 5.8 | 5.91 | 6.03 | 6.15 | 6.26 | 6.38 | 6.49 | 6.61 | 6.73 | 6.84 | 6.96 | 7.08 | 7.19 | 7.31 | 7.42 | 7.54 | 7.66 | 7.77 | 7.89 | 8 | 8.12 | 8.24 | 8.35 | 8.47 | 8.58 | 8.7 | 8.82 | 8.93 | 9.05
22.97 | 4.4 | 4.52 | 4.64 | 4.75 | 4.87 | 4.98 | 5.1 | 5.22 | 5.33 | 5.45 | 5.56 | 5.68 | 5.8 | 5.91 | 6.03 | 6.15 | 6.26 | 6.38 | 6.49 | 6.61 | 6.73 | 6.84 | 6.96 | 7.07 | 7.19 | 7.31 | 7.42 | 7.54 | 7.65 | 7.77 | 7.89 | 8 | 8.12 | 8.23 | 8.35 | 8.47 | 8.58 | 8.7 | 8.82 | 8.93 | 9.05
22.98 | 4.4 | 4.52 | 4.63 | 4.75 | 4.87 | 4.98 | 5.1 | 5.22 | 5.33 | 5.45 | 5.56 | 5.68 | 5.8 | 5.91 | 6.03 | 6.14 | 6.26 | 6.38 | 6.49 | 6.61 | 6.72 | 6.84 | 6.96 | 7.07 | 7.19 | 7.3 | 7.42 | 7.54 | 7.65 | 7.77 | 7.89 | 8 | 8.12 | 8.23 | 8.35 | 8.47 | 8.58 | 8.7 | 8.81 | 8.93 | 9.05
22.99 | 4.4 | 4.52 | 4.63 | 4.75 | 4.87 | 4.98 | 5.1 | 5.21 | 5.33 | 5.45 | 5.56 | 5.68 | 5.79 | 5.91 | 6.03 | 6.14 | 6.26 | 6.37 | 6.49 | 6.61 | 6.72 | 6.84 | 6.96 | 7.07 | 7.19 | 7.3 | 7.42 | 7.54 | 7.65 | 7.77 | 7.88 | 8 | 8.12 | 8.23 | 8.35 | 8.46 | 8.58 | 8.7 | 8.81 | 8.93 | 9.04
23 | 4.4 | 4.52 | 4.63 | 4.75 | 4.86 | 4.98 | 5.1 | 5.21 | 5.33 | 5.45 | 5.56 | 5.68 | 5.79 | 5.91 | 6.03 | 6.14 | 6.26 | 6.37 | 6.49 | 6.61 | 6.72 | 6.84 | 6.95 | 7.07 | 7.19 | 7.3 | 7.42 | 7.53 | 7.65 | 7.77 | 7.88 | 8 | 8.11 | 8.23 | 8.35 | 8.46 | 8.58 | 8.69 | 8.81 | 8.93 | 9.04
23.01 | 4.4 | 4.52 | 4.63 | 4.75 | 4.86 | 4.98 | 5.1 | 5.21 | 5.33 | 5.44 | 5.56 | 5.68 | 5.79 | 5.91 | 6.02 | 6.14 | 6.26 | 6.37 | 6.49 | 6.6 | 6.72 | 6.84 | 6.95 | 7.07 | 7.18 | 7.3 | 7.42 | 7.53 | 7.65 | 7.76 | 7.88 | 8 | 8.11 | 8.23 | 8.34 | 8.46 | 8.58 | 8.69 | 8.81 | 8.92 | 9.04
23.02 | 4.4 | 4.51 | 4.63 | 4.75 | 4.86 | 4.98 | 5.09 | 5.21 | 5.33 | 5.44 | 5.56 | 5.67 | 5.79 | 5.91 | 6.02 | 6.14 | 6.26 | 6.37 | 6.49 | 6.6 | 6.72 | 6.84 | 6.95 | 7.07 | 7.18 | 7.3 | 7.42 | 7.53 | 7.65 | 7.76 | 7.88 | 8 | 8.11 | 8.23 | 8.34 | 8.46 | 8.57 | 8.69 | 8.81 | 8.92 | 9.04
23.03 | 4.4 | 4.51 | 4.63 | 4.75 | 4.86 | 4.98 | 5.09 | 5.21 | 5.33 | 5.44 | 5.56 | 5.67 | 5.79 | 5.91 | 6.02 | 6.14 | 6.25 | 6.37 | 6.49 | 6.6 | 6.72 | 6.83 | 6.95 | 7.07 | 7.18 | 7.3 | 7.41 | 7.53 | 7.65 | 7.76 | 7.88 | 7.99 | 8.11 | 8.23 | 8.34 | 8.46 | 8.57 | 8.69 | 8.81 | 8.92 | 9.04
23.04 | 4.4 | 4.51 | 4.63 | 4.74 | 4.86 | 4.98 | 5.09 | 5.21 | 5.32 | 5.44 | 5.56 | 5.67 | 5.79 | 5.9 | 6.02 | 6.14 | 6.25 | 6.37 | 6.48 | 6.6 | 6.72 | 6.83 | 6.95 | 7.06 | 7.18 | 7.3 | 7.41 | 7.53 | 7.64 | 7.76 | 7.88 | 7.99 | 8.11 | 8.22 | 8.34 | 8.46 | 8.57 | 8.69 | 8.8 | 8.92 | 9.04
23.05 | 4.4 | 4.51 | 4.63 | 4.74 | 4.86 | 4.98 | 5.09 | 5.21 | 5.32 | 5.44 | 5.56 | 5.67 | 5.79 | 5.9 | 6.02 | 6.14 | 6.25 | 6.37 | 6.48 | 6.6 | 6.72 | 6.83 | 6.95 | 7.06 | 7.18 | 7.29 | 7.41 | 7.53 | 7.64 | 7.76 | 7.87 | 7.99 | 8.11 | 8.22 | 8.34 | 8.45 | 8.57 | 8.69 | 8.8 | 8.92 | 9.03
23.06 | 4.39 | 4.51 | 4.63 | 4.74 | 4.86 | 4.97 | 5.09 | 5.21 | 5.32 | 5.44 | 5.55 | 5.67 | 5.79 | 5.9 | 6.02 | 6.13 | 6.25 | 6.37 | 6.48 | 6.6 | 6.71 | 6.83 | 6.95 | 7.06 | 7.18 | 7.29 | 7.41 | 7.53 | 7.64 | 7.76 | 7.87 | 7.99 | 8.1 | 8.22 | 8.34 | 8.45 | 8.57 | 8.68 | 8.8 | 8.92 | 9.03
23.07 | 4.39 | 4.51 | 4.63 | 4.74 | 4.86 | 4.97 | 5.09 | 5.21 | 5.32 | 5.44 | 5.55 | 5.67 | 5.79 | 5.9 | 6.02 | 6.13 | 6.25 | 6.36 | 6.48 | 6.6 | 6.71 | 6.83 | 6.94 | 7.06 | 7.18 | 7.29 | 7.41 | 7.52 | 7.64 | 7.76 | 7.87 | 7.99 | 8.1 | 8.22 | 8.34 | 8.45 | 8.57 | 8.68 | 8.8 | 8.91 | 9.03
23.08 | 4.39 | 4.51 | 4.63 | 4.74 | 4.86 | 4.97 | 5.09 | 5.2 | 5.32 | 5.44 | 5.55 | 5.67 | 5.78 | 5.9 | 6.02 | 6.13 | 6.25 | 6.36 | 6.48 | 6.6 | 6.71 | 6.83 | 6.94 | 7.06 | 7.17 | 7.29 | 7.41 | 7.52 | 7.64 | 7.75 | 7.87 | 7.99 | 8.1 | 8.22 | 8.33 | 8.45 | 8.57 | 8.68 | 8.8 | 8.91 | 9.03
23.09 | 4.39 | 4.51 | 4.62 | 4.74 | 4.86 | 4.97 | 5.09 | 5.2 | 5.32 | 5.44 | 5.55 | 5.67 | 5.78 | 5.9 | 6.01 | 6.13 | 6.25 | 6.36 | 6.48 | 6.59 | 6.71 | 6.83 | 6.94 | 7.06 | 7.17 | 7.29 | 7.41 | 7.52 | 7.64 | 7.75 | 7.87 | 7.98 | 8.1 | 8.22 | 8.33 | 8.45 | 8.56 | 8.68 | 8.8 | 8.91 | 9.03
23.1 | 4.39 | 4.51 | 4.62 | 4.74 | 4.85 | 4.97 | 5.09 | 5.2 | 5.32 | 5.43 | 5.55 | 5.67 | 5.78 | 5.9 | 6.01 | 6.13 | 6.25 | 6.36 | 6.48 | 6.59 | 6.71 | 6.82 | 6.94 | 7.06 | 7.17 | 7.29 | 7.4 | 7.52 | 7.64 | 7.75 | 7.87 | 7.98 | 8.1 | 8.21 | 8.33 | 8.45 | 8.56 | 8.68 | 8.79 | 8.91 | 9.03
23.11 | 4.39 | 4.51 | 4.62 | 4.74 | 4.85 | 4.97 | 5.09 | 5.2 | 5.32 | 5.43 | 5.55 | 5.66 | 5.78 | 5.9 | 6.01 | 6.13 | 6.24 | 6.36 | 6.48 | 6.59 | 6.71 | 6.82 | 6.94 | 7.05 | 7.17 | 7.29 | 7.4 | 7.52 | 7.63 | 7.75 | 7.87 | 7.98 | 8.1 | 8.21 | 8.33 | 8.44 | 8.56 | 8.68 | 8.79 | 8.91 | 9.02
23.12 | 4.39 | 4.51 | 4.62 | 4.74 | 4.85 | 4.97 | 5.08 | 5.2 | 5.32 | 5.43 | 5.55 | 5.66 | 5.78 | 5.9 | 6.01 | 6.13 | 6.24 | 6.36 | 6.47 | 6.59 | 6.71 | 6.82 | 6.94 | 7.05 | 7.17 | 7.29 | 7.4 | 7.52 | 7.63 | 7.75 | 7.86 | 7.98 | 8.1 | 8.21 | 8.33 | 8.44 | 8.56 | 8.67 | 8.79 | 8.91 | 9.02
23.13 | 4.39 | 4.5 | 4.62 | 4.74 | 4.85 | 4.97 | 5.08 | 5.2 | 5.32 | 5.43 | 5.55 | 5.66 | 5.78 | 5.89 | 6.01 | 6.13 | 6.24 | 6.36 | 6.47 | 6.59 | 6.7 | 6.82 | 6.94 | 7.05 | 7.17 | 7.28 | 7.4 | 7.52 | 7.63 | 7.75 | 7.86 | 7.98 | 8.09 | 8.21 | 8.33 | 8.44 | 8.56 | 8.67 | 8.79 | 8.9 | 9.02
23.14 | 4.39 | 4.5 | 4.62 | 4.74 | 4.85 | 4.97 | 5.08 | 5.2 | 5.31 | 5.43 | 5.55 | 5.66 | 5.78 | 5.89 | 6.01 | 6.12 | 6.24 | 6.36 | 6.47 | 6.59 | 6.7 | 6.82 | 6.94 | 7.05 | 7.17 | 7.28 | 7.4 | 7.51 | 7.63 | 7.75 | 7.86 | 7.98 | 8.09 | 8.21 | 8.32 | 8.44 | 8.56 | 8.67 | 8.79 | 8.9 | 9.02
23.15 | 4.39 | 4.5 | 4.62 | 4.73 | 4.85 | 4.97 | 5.08 | 5.2 | 5.31 | 5.43 | 5.54 | 5.66 | 5.78 | 5.89 | 6.01 | 6.12 | 6.24 | 6.35 | 6.47 | 6.59 | 6.7 | 6.82 | 6.93 | 7.05 | 7.17 | 7.28 | 7.4 | 7.51 | 7.63 | 7.74 | 7.86 | 7.98 | 8.09 | 8.21 | 8.32 | 8.44 | 8.55 | 8.67 | 8.79 | 8.9 | 9.02
23.16 | 4.39 | 4.5 | 4.62 | 4.73 | 4.85 | 4.96 | 5.08 | 5.2 | 5.31 | 5.43 | 5.54 | 5.66 | 5.78 | 5.89 | 6.01 | 6.12 | 6.24 | 6.35 | 6.47 | 6.59 | 6.7 | 6.82 | 6.93 | 7.05 | 7.16 | 7.28 | 7.4 | 7.51 | 7.63 | 7.74 | 7.86 | 7.97 | 8.09 | 8.21 | 8.32 | 8.44 | 8.55 | 8.67 | 8.78 | 8.9 | 9.02
23.17 | 4.39 | 4.5 | 4.62 | 4.73 | 4.85 | 4.96 | 5.08 | 5.2 | 5.31 | 5.43 | 5.54 | 5.66 | 5.77 | 5.89 | 6.01 | 6.12 | 6.24 | 6.35 | 6.47 | 6.58 | 6.7 | 6.82 | 6.93 | 7.05 | 7.16 | 7.28 | 7.39 | 7.51 | 7.63 | 7.74 | 7.86 | 7.97 | 8.09 | 8.2 | 8.32 | 8.43 | 8.55 | 8.67 | 8.78 | 8.9 | 9.01
23.18 | 4.38 | 4.5 | 4.62 | 4.73 | 4.85 | 4.96 | 5.08 | 5.19 | 5.31 | 5.43 | 5.54 | 5.66 | 5.77 | 5.89 | 6 | 6.12 | 6.24 | 6.35 | 6.47 | 6.58 | 6.7 | 6.81 | 6.93 | 7.05 | 7.16 | 7.28 | 7.39 | 7.51 | 7.62 | 7.74 | 7.86 | 7.97 | 8.09 | 8.2 | 8.32 | 8.43 | 8.55 | 8.66 | 8.78 | 8.9 | 9.01
23.19 | 4.38 | 4.5 | 4.61 | 4.73 | 4.85 | 4.96 | 5.08 | 5.19 | 5.31 | 5.42 | 5.54 | 5.66 | 5.77 | 5.89 | 6 | 6.12 | 6.23 | 6.35 | 6.47 | 6.58 | 6.7 | 6.81 | 6.93 | 7.04 | 7.16 | 7.28 | 7.39 | 7.51 | 7.62 | 7.74 | 7.85 | 7.97 | 8.08 | 8.2 | 8.32 | 8.43 | 8.55 | 8.66 | 8.78 | 8.89 | 9.01
23.2 | 4.38 | 4.5 | 4.61 | 4.73 | 4.85 | 4.96 | 5.08 | 5.19 | 5.31 | 5.42 | 5.54 | 5.65 | 5.77 | 5.89 | 6 | 6.12 | 6.23 | 6.35 | 6.46 | 6.58 | 6.7 | 6.81 | 6.93 | 7.04 | 7.16 | 7.27 | 7.39 | 7.51 | 7.62 | 7.74 | 7.85 | 7.97 | 8.08 | 8.2 | 8.31 | 8.43 | 8.55 | 8.66 | 8.78 | 8.89 | 9.01
23.21 | 4.38 | 4.5 | 4.61 | 4.73 | 4.84 | 4.96 | 5.08 | 5.19 | 5.31 | 5.42 | 5.54 | 5.65 | 5.77 | 5.88 | 6 | 6.12 | 6.23 | 6.35 | 6.46 | 6.58 | 6.69 | 6.81 | 6.93 | 7.04 | 7.16 | 7.27 | 7.39 | 7.5 | 7.62 | 7.73 | 7.85 | 7.97 | 8.08 | 8.2 | 8.31 | 8.43 | 8.54 | 8.66 | 8.78 | 8.89 | 9.01
23.22 | 4.38 | 4.5 | 4.61 | 4.73 | 4.84 | 4.96 | 5.07 | 5.19 | 5.31 | 5.42 | 5.54 | 5.65 | 5.77 | 5.88 | 6 | 6.12 | 6.23 | 6.35 | 6.46 | 6.58 | 6.69 | 6.81 | 6.92 | 7.04 | 7.16 | 7.27 | 7.39 | 7.5 | 7.62 | 7.73 | 7.85 | 7.96 | 8.08 | 8.2 | 8.31 | 8.43 | 8.54 | 8.66 | 8.77 | 8.89 | 9
23.23 | 4.38 | 4.5 | 4.61 | 4.73 | 4.84 | 4.96 | 5.07 | 5.19 | 5.3 | 5.42 | 5.54 | 5.65 | 5.77 | 5.88 | 6 | 6.11 | 6.23 | 6.35 | 6.46 | 6.58 | 6.69 | 6.81 | 6.92 | 7.04 | 7.15 | 7.27 | 7.39 | 7.5 | 7.62 | 7.73 | 7.85 | 7.96 | 8.08 | 8.19 | 8.31 | 8.43 | 8.54 | 8.66 | 8.77 | 8.89 | 9
23.24 | 4.38 | 4.49 | 4.61 | 4.73 | 4.84 | 4.96 | 5.07 | 5.19 | 5.3 | 5.42 | 5.53 | 5.65 | 5.77 | 5.88 | 6 | 6.11 | 6.23 | 6.34 | 6.46 | 6.57 | 6.69 | 6.81 | 6.92 | 7.04 | 7.15 | 7.27 | 7.38 | 7.5 | 7.61 | 7.73 | 7.85 | 7.96 | 8.08 | 8.19 | 8.31 | 8.42 | 8.54 | 8.65 | 8.77 | 8.89 | 9
23.25 | 4.38 | 4.49 | 4.61 | 4.72 | 4.84 | 4.96 | 5.07 | 5.19 | 5.3 | 5.42 | 5.53 | 5.65 | 5.76 | 5.88 | 6 | 6.11 | 6.23 | 6.34 | 6.46 | 6.57 | 6.69 | 6.8 | 6.92 | 7.04 | 7.15 | 7.27 | 7.38 | 7.5 | 7.61 | 7.73 | 7.84 | 7.96 | 8.08 | 8.19 | 8.31 | 8.42 | 8.54 | 8.65 | 8.77 | 8.88 | 9
23.26 | 4.38 | 4.49 | 4.61 | 4.72 | 4.84 | 4.95 | 5.07 | 5.19 | 5.3 | 5.42 | 5.53 | 5.65 | 5.76 | 5.88 | 5.99 | 6.11 | 6.23 | 6.34 | 6.46 | 6.57 | 6.69 | 6.8 | 6.92 | 7.03 | 7.15 | 7.27 | 7.38 | 7.5 | 7.61 | 7.73 | 7.84 | 7.96 | 8.07 | 8.19 | 8.31 | 8.42 | 8.54 | 8.65 | 8.77 | 8.88 | 9
23.27 | 4.38 | 4.49 | 4.61 | 4.72 | 4.84 | 4.95 | 5.07 | 5.18 | 5.3 | 5.42 | 5.53 | 5.65 | 5.76 | 5.88 | 5.99 | 6.11 | 6.22 | 6.34 | 6.46 | 6.57 | 6.69 | 6.8 | 6.92 | 7.03 | 7.15 | 7.26 | 7.38 | 7.5 | 7.61 | 7.73 | 7.84 | 7.96 | 8.07 | 8.19 | 8.3 | 8.42 | 8.53 | 8.65 | 8.77 | 8.88 | 9
23.28 | 4.38 | 4.49 | 4.61 | 4.72 | 4.84 | 4.95 | 5.07 | 5.18 | 5.3 | 5.41 | 5.53 | 5.65 | 5.76 | 5.88 | 5.99 | 6.11 | 6.22 | 6.34 | 6.45 | 6.57 | 6.69 | 6.8 | 6.92 | 7.03 | 7.15 | 7.26 | 7.38 | 7.49 | 7.61 | 7.72 | 7.84 | 7.96 | 8.07 | 8.19 | 8.3 | 8.42 | 8.53 | 8.65 | 8.76 | 8.88 | 8.99
23.29 | 4.37 | 4.49 | 4.61 | 4.72 | 4.84 | 4.95 | 5.07 | 5.18 | 5.3 | 5.41 | 5.53 | 5.64 | 5.76 | 5.88 | 5.99 | 6.11 | 6.22 | 6.34 | 6.45 | 6.57 | 6.68 | 6.8 | 6.91 | 7.03 | 7.15 | 7.26 | 7.38 | 7.49 | 7.61 | 7.72 | 7.84 | 7.95 | 8.07 | 8.18 | 8.3 | 8.42 | 8.53 | 8.65 | 8.76 | 8.88 | 8.99
23.3 | 4.37 | 4.49 | 4.6 | 4.72 | 4.84 | 4.95 | 5.07 | 5.18 | 5.3 | 5.41 | 5.53 | 5.64 | 5.76 | 5.87 | 5.99 | 6.11 | 6.22 | 6.34 | 6.45 | 6.57 | 6.68 | 6.8 | 6.91 | 7.03 | 7.14 | 7.26 | 7.38 | 7.49 | 7.61 | 7.72 | 7.84 | 7.95 | 8.07 | 8.18 | 8.3 | 8.41 | 8.53 | 8.65 | 8.76 | 8.88 | 8.99
23.31 | 4.37 | 4.49 | 4.6 | 4.72 | 4.83 | 4.95 | 5.07 | 5.18 | 5.3 | 5.41 | 5.53 | 5.64 | 5.76 | 5.87 | 5.99 | 6.1 | 6.22 | 6.34 | 6.45 | 6.57 | 6.68 | 6.8 | 6.91 | 7.03 | 7.14 | 7.26 | 7.37 | 7.49 | 7.6 | 7.72 | 7.84 | 7.95 | 8.07 | 8.18 | 8.3 | 8.41 | 8.53 | 8.64 | 8.76 | 8.87 | 8.99
23.32 | 4.37 | 4.49 | 4.6 | 4.72 | 4.83 | 4.95 | 5.06 | 5.18 | 5.3 | 5.41 | 5.53 | 5.64 | 5.76 | 5.87 | 5.99 | 6.1 | 6.22 | 6.33 | 6.45 | 6.56 | 6.68 | 6.8 | 6.91 | 7.03 | 7.14 | 7.26 | 7.37 | 7.49 | 7.6 | 7.72 | 7.83 | 7.95 | 8.06 | 8.18 | 8.3 | 8.41 | 8.53 | 8.64 | 8.76 | 8.87 | 8.99
23.33 | 4.37 | 4.49 | 4.6 | 4.72 | 4.83 | 4.95 | 5.06 | 5.18 | 5.29 | 5.41 | 5.52 | 5.64 | 5.76 | 5.87 | 5.99 | 6.1 | 6.22 | 6.33 | 6.45 | 6.56 | 6.68 | 6.79 | 6.91 | 7.03 | 7.14 | 7.26 | 7.37 | 7.49 | 7.6 | 7.72 | 7.83 | 7.95 | 8.06 | 8.18 | 8.29 | 8.41 | 8.52 | 8.64 | 8.76 | 8.87 | 8.99
23.34 | 4.37 | 4.49 | 4.6 | 4.72 | 4.83 | 4.95 | 5.06 | 5.18 | 5.29 | 5.41 | 5.52 | 5.64 | 5.75 | 5.87 | 5.99 | 6.1 | 6.22 | 6.33 | 6.45 | 6.56 | 6.68 | 6.79 | 6.91 | 7.02 | 7.14 | 7.25 | 7.37 | 7.49 | 7.6 | 7.72 | 7.83 | 7.95 | 8.06 | 8.18 | 8.29 | 8.41 | 8.52 | 8.64 | 8.75 | 8.87 | 8.98
23.35 | 4.37 | 4.48 | 4.6 | 4.71 | 4.83 | 4.95 | 5.06 | 5.18 | 5.29 | 5.41 | 5.52 | 5.64 | 5.75 | 5.87 | 5.98 | 6.1 | 6.21 | 6.33 | 6.45 | 6.56 | 6.68 | 6.79 | 6.91 | 7.02 | 7.14 | 7.25 | 7.37 | 7.48 | 7.6 | 7.71 | 7.83 | 7.94 | 8.06 | 8.18 | 8.29 | 8.41 | 8.52 | 8.64 | 8.75 | 8.87 | 8.98
23.36 | 4.37 | 4.48 | 4.6 | 4.71 | 4.83 | 4.94 | 5.06 | 5.18 | 5.29 | 5.41 | 5.52 | 5.64 | 5.75 | 5.87 | 5.98 | 6.1 | 6.21 | 6.33 | 6.44 | 6.56 | 6.67 | 6.79 | 6.91 | 7.02 | 7.14 | 7.25 | 7.37 | 7.48 | 7.6 | 7.71 | 7.83 | 7.94 | 8.06 | 8.17 | 8.29 | 8.4 | 8.52 | 8.64 | 8.75 | 8.87 | 8.98
23.37 | 4.37 | 4.48 | 4.6 | 4.71 | 4.83 | 4.94 | 5.06 | 5.17 | 5.29 | 5.41 | 5.52 | 5.64 | 5.75 | 5.87 | 5.98 | 6.1 | 6.21 | 6.33 | 6.44 | 6.56 | 6.67 | 6.79 | 6.9 | 7.02 | 7.13 | 7.25 | 7.37 | 7.48 | 7.6 | 7.71 | 7.83 | 7.94 | 8.06 | 8.17 | 8.29 | 8.4 | 8.52 | 8.63 | 8.75 | 8.86 | 8.98
23.38 | 4.37 | 4.48 | 4.6 | 4.71 | 4.83 | 4.94 | 5.06 | 5.17 | 5.29 | 5.4 | 5.52 | 5.63 | 5.75 | 5.87 | 5.98 | 6.1 | 6.21 | 6.33 | 6.44 | 6.56 | 6.67 | 6.79 | 6.9 | 7.02 | 7.13 | 7.25 | 7.36 | 7.48 | 7.59 | 7.71 | 7.83 | 7.94 | 8.06 | 8.17 | 8.29 | 8.4 | 8.52 | 8.63 | 8.75 | 8.86 | 8.98
23.39 | 4.37 | 4.48 | 4.6 | 4.71 | 4.83 | 4.94 | 5.06 | 5.17 | 5.29 | 5.4 | 5.52 | 5.63 | 5.75 | 5.86 | 5.98 | 6.09 | 6.21 | 6.33 | 6.44 | 6.56 | 6.67 | 6.79 | 6.9 | 7.02 | 7.13 | 7.25 | 7.36 | 7.48 | 7.59 | 7.71 | 7.82 | 7.94 | 8.05 | 8.17 | 8.28 | 8.4 | 8.52 | 8.63 | 8.75 | 8.86 | 8.98
23.4 | 4.36 | 4.48 | 4.59 | 4.71 | 4.83 | 4.94 | 5.06 | 5.17 | 5.29 | 5.4 | 5.52 | 5.63 | 5.75 | 5.86 | 5.98 | 6.09 | 6.21 | 6.32 | 6.44 | 6.55 | 6.67 | 6.79 | 6.9 | 7.02 | 7.13 | 7.25 | 7.36 | 7.48 | 7.59 | 7.71 | 7.82 | 7.94 | 8.05 | 8.17 | 8.28 | 8.4 | 8.51 | 8.63 | 8.74 | 8.86 | 8.97
23.41 | 4.36 | 4.48 | 4.59 | 4.71 | 4.82 | 4.94 | 5.06 | 5.17 | 5.29 | 5.4 | 5.52 | 5.63 | 5.75 | 5.86 | 5.98 | 6.09 | 6.21 | 6.32 | 6.44 | 6.55 | 6.67 | 6.78 | 6.9 | 7.01 | 7.13 | 7.24 | 7.36 | 7.48 | 7.59 | 7.71 | 7.82 | 7.94 | 8.05 | 8.17 | 8.28 | 8.4 | 8.51 | 8.63 | 8.74 | 8.86 | 8.97
23.42 | 4.36 | 4.48 | 4.59 | 4.71 | 4.82 | 4.94 | 5.05 | 5.17 | 5.28 | 5.4 | 5.51 | 5.63 | 5.75 | 5.86 | 5.98 | 6.09 | 6.21 | 6.32 | 6.44 | 6.55 | 6.67 | 6.78 | 6.9 | 7.01 | 7.13 | 7.24 | 7.36 | 7.47 | 7.59 | 7.7 | 7.82 | 7.93 | 8.05 | 8.16 | 8.28 | 8.4 | 8.51 | 8.63 | 8.74 | 8.86 | 8.97
23.43 | 4.36 | 4.48 | 4.59 | 4.71 | 4.82 | 4.94 | 5.05 | 5.17 | 5.28 | 5.4 | 5.51 | 5.63 | 5.74 | 5.86 | 5.97 | 6.09 | 6.21 | 6.32 | 6.44 | 6.55 | 6.67 | 6.78 | 6.9 | 7.01 | 7.13 | 7.24 | 7.36 | 7.47 | 7.59 | 7.7 | 7.82 | 7.93 | 8.05 | 8.16 | 8.28 | 8.39 | 8.51 | 8.62 | 8.74 | 8.85 | 8.97
23.44 | 4.36 | 4.48 | 4.59 | 4.71 | 4.82 | 4.94 | 5.05 | 5.17 | 5.28 | 5.4 | 5.51 | 5.63 | 5.74 | 5.86 | 5.97 | 6.09 | 6.2 | 6.32 | 6.43 | 6.55 | 6.66 | 6.78 | 6.89 | 7.01 | 7.13 | 7.24 | 7.36 | 7.47 | 7.59 | 7.7 | 7.82 | 7.93 | 8.05 | 8.16 | 8.28 | 8.39 | 8.51 | 8.62 | 8.74 | 8.85 | 8.97
23.45 | 4.36 | 4.48 | 4.59 | 4.71 | 4.82 | 4.94 | 5.05 | 5.17 | 5.28 | 5.4 | 5.51 | 5.63 | 5.74 | 5.86 | 5.97 | 6.09 | 6.2 | 6.32 | 6.43 | 6.55 | 6.66 | 6.78 | 6.89 | 7.01 | 7.12 | 7.24 | 7.35 | 7.47 | 7.58 | 7.7 | 7.81 | 7.93 | 8.05 | 8.16 | 8.28 | 8.39 | 8.51 | 8.62 | 8.74 | 8.85 | 8.97
23.46 | 4.36 | 4.47 | 4.59 | 4.7 | 4.82 | 4.93 | 5.05 | 5.17 | 5.28 | 5.4 | 5.51 | 5.63 | 5.74 | 5.86 | 5.97 | 6.09 | 6.2 | 6.32 | 6.43 | 6.55 | 6.66 | 6.78 | 6.89 | 7.01 | 7.12 | 7.24 | 7.35 | 7.47 | 7.58 | 7.7 | 7.81 | 7.93 | 8.04 | 8.16 | 8.27 | 8.39 | 8.5 | 8.62 | 8.73 | 8.85 | 8.96
23.47 | 4.36 | 4.47 | 4.59 | 4.7 | 4.82 | 4.93 | 5.05 | 5.16 | 5.28 | 5.39 | 5.51 | 5.62 | 5.74 | 5.85 | 5.97 | 6.09 | 6.2 | 6.32 | 6.43 | 6.55 | 6.66 | 6.78 | 6.89 | 7.01 | 7.12 | 7.24 | 7.35 | 7.47 | 7.58 | 7.7 | 7.81 | 7.93 | 8.04 | 8.16 | 8.27 | 8.39 | 8.5 | 8.62 | 8.73 | 8.85 | 8.96
23.48 | 4.36 | 4.47 | 4.59 | 4.7 | 4.82 | 4.93 | 5.05 | 5.16 | 5.28 | 5.39 | 5.51 | 5.62 | 5.74 | 5.85 | 5.97 | 6.08 | 6.2 | 6.31 | 6.43 | 6.54 | 6.66 | 6.77 | 6.89 | 7 | 7.12 | 7.23 | 7.35 | 7.47 | 7.58 | 7.7 | 7.81 | 7.93 | 8.04 | 8.16 | 8.27 | 8.39 | 8.5 | 8.62 | 8.73 | 8.85 | 8.96
23.49 | 4.36 | 4.47 | 4.59 | 4.7 | 4.82 | 4.93 | 5.05 | 5.16 | 5.28 | 5.39 | 5.51 | 5.62 | 5.74 | 5.85 | 5.97 | 6.08 | 6.2 | 6.31 | 6.43 | 6.54 | 6.66 | 6.77 | 6.89 | 7 | 7.12 | 7.23 | 7.35 | 7.46 | 7.58 | 7.69 | 7.81 | 7.92 | 8.04 | 8.15 | 8.27 | 8.38 | 8.5 | 8.61 | 8.73 | 8.84 | 8.96
23.5 | 4.36 | 4.47 | 4.59 | 4.7 | 4.82 | 4.93 | 5.05 | 5.16 | 5.28 | 5.39 | 5.51 | 5.62 | 5.74 | 5.85 | 5.97 | 6.08 | 6.2 | 6.31 | 6.43 | 6.54 | 6.66 | 6.77 | 6.89 | 7 | 7.12 | 7.23 | 7.35 | 7.46 | 7.58 | 7.69 | 7.81 | 7.92 | 8.04 | 8.15 | 8.27 | 8.38 | 8.5 | 8.61 | 8.73 | 8.84 | 8.96
23.51 | 4.35 | 4.47 | 4.58 | 4.7 | 4.81 | 4.93 | 5.04 | 5.16 | 5.28 | 5.39 | 5.51 | 5.62 | 5.74 | 5.85 | 5.97 | 6.08 | 6.2 | 6.31 | 6.43 | 6.54 | 6.66 | 6.77 | 6.89 | 7 | 7.12 | 7.23 | 7.35 | 7.46 | 7.58 | 7.69 | 7.81 | 7.92 | 8.04 | 8.15 | 8.27 | 8.38 | 8.5 | 8.61 | 8.73 | 8.84 | 8.96
23.52 | 4.35 | 4.47 | 4.58 | 4.7 | 4.81 | 4.93 | 5.04 | 5.16 | 5.27 | 5.39 | 5.5 | 5.62 | 5.73 | 5.85 | 5.96 | 6.08 | 6.19 | 6.31 | 6.42 | 6.54 | 6.65 | 6.77 | 6.88 | 7 | 7.11 | 7.23 | 7.34 | 7.46 | 7.57 | 7.69 | 7.8 | 7.92 | 8.03 | 8.15 | 8.26 | 8.38 | 8.49 | 8.61 | 8.72 | 8.84 | 8.95
23.53 | 4.35 | 4.47 | 4.58 | 4.7 | 4.81 | 4.93 | 5.04 | 5.16 | 5.27 | 5.39 | 5.5 | 5.62 | 5.73 | 5.85 | 5.96 | 6.08 | 6.19 | 6.31 | 6.42 | 6.54 | 6.65 | 6.77 | 6.88 | 7 | 7.11 | 7.23 | 7.34 | 7.46 | 7.57 | 7.69 | 7.8 | 7.92 | 8.03 | 8.15 | 8.26 | 8.38 | 8.49 | 8.61 | 8.72 | 8.84 | 8.95
23.54 | 4.35 | 4.47 | 4.58 | 4.7 | 4.81 | 4.93 | 5.04 | 5.16 | 5.27 | 5.39 | 5.5 | 5.62 | 5.73 | 5.85 | 5.96 | 6.08 | 6.19 | 6.31 | 6.42 | 6.54 | 6.65 | 6.77 | 6.88 | 7 | 7.11 | 7.23 | 7.34 | 7.46 | 7.57 | 7.69 | 7.8 | 7.92 | 8.03 | 8.15 | 8.26 | 8.38 | 8.49 | 8.61 | 8.72 | 8.84 | 8.95
23.55 | 4.35 | 4.47 | 4.58 | 4.7 | 4.81 | 4.93 | 5.04 | 5.16 | 5.27 | 5.39 | 5.5 | 5.62 | 5.73 | 5.85 | 5.96 | 6.08 | 6.19 | 6.31 | 6.42 | 6.54 | 6.65 | 6.77 | 6.88 | 7 | 7.11 | 7.23 | 7.34 | 7.46 | 7.57 | 7.69 | 7.8 | 7.91 | 8.03 | 8.14 | 8.26 | 8.37 | 8.49 | 8.6 | 8.72 | 8.83 | 8.95
23.56 | 4.35 | 4.47 | 4.58 | 4.69 | 4.81 | 4.92 | 5.04 | 5.15 | 5.27 | 5.38 | 5.5 | 5.61 | 5.73 | 5.84 | 5.96 | 6.07 | 6.19 | 6.3 | 6.42 | 6.53 | 6.65 | 6.76 | 6.88 | 6.99 | 7.11 | 7.22 | 7.34 | 7.45 | 7.57 | 7.68 | 7.8 | 7.91 | 8.03 | 8.14 | 8.26 | 8.37 | 8.49 | 8.6 | 8.72 | 8.83 | 8.95
23.57 | 4.35 | 4.46 | 4.58 | 4.69 | 4.81 | 4.92 | 5.04 | 5.15 | 5.27 | 5.38 | 5.5 | 5.61 | 5.73 | 5.84 | 5.96 | 6.07 | 6.19 | 6.3 | 6.42 | 6.53 | 6.65 | 6.76 | 6.88 | 6.99 | 7.11 | 7.22 | 7.34 | 7.45 | 7.57 | 7.68 | 7.8 | 7.91 | 8.03 | 8.14 | 8.26 | 8.37 | 8.49 | 8.6 | 8.72 | 8.83 | 8.95
23.58 | 4.35 | 4.46 | 4.58 | 4.69 | 4.81 | 4.92 | 5.04 | 5.15 | 5.27 | 5.38 | 5.5 | 5.61 | 5.73 | 5.84 | 5.96 | 6.07 | 6.19 | 6.3 | 6.42 | 6.53 | 6.65 | 6.76 | 6.88 | 6.99 | 7.11 | 7.22 | 7.34 | 7.45 | 7.57 | 7.68 | 7.8 | 7.91 | 8.03 | 8.14 | 8.26 | 8.37 | 8.48 | 8.6 | 8.71 | 8.83 | 8.94
23.59 | 4.35 | 4.46 | 4.58 | 4.69 | 4.81 | 4.92 | 5.04 | 5.15 | 5.27 | 5.38 | 5.5 | 5.61 | 5.73 | 5.84 | 5.96 | 6.07 | 6.19 | 6.3 | 6.42 | 6.53 | 6.65 | 6.76 | 6.88 | 6.99 | 7.1 | 7.22 | 7.33 | 7.45 | 7.56 | 7.68 | 7.79 | 7.91 | 8.02 | 8.14 | 8.25 | 8.37 | 8.48 | 8.6 | 8.71 | 8.83 | 8.94
23.6 | 4.35 | 4.46 | 4.58 | 4.69 | 4.81 | 4.92 | 5.04 | 5.15 | 5.27 | 5.38 | 5.5 | 5.61 | 5.73 | 5.84 | 5.95 | 6.07 | 6.18 | 6.3 | 6.41 | 6.53 | 6.64 | 6.76 | 6.87 | 6.99 | 7.1 | 7.22 | 7.33 | 7.45 | 7.56 | 7.68 | 7.79 | 7.91 | 8.02 | 8.14 | 8.25 | 8.37 | 8.48 | 8.6 | 8.71 | 8.83 | 8.94
23.61 | 4.35 | 4.46 | 4.58 | 4.69 | 4.81 | 4.92 | 5.03 | 5.15 | 5.26 | 5.38 | 5.49 | 5.61 | 5.72 | 5.84 | 5.95 | 6.07 | 6.18 | 6.3 | 6.41 | 6.53 | 6.64 | 6.76 | 6.87 | 6.99 | 7.1 | 7.22 | 7.33 | 7.45 | 7.56 | 7.68 | 7.79 | 7.91 | 8.02 | 8.14 | 8.25 | 8.37 | 8.48 | 8.59 | 8.71 | 8.82 | 8.94
23.62 | 4.34 | 4.46 | 4.57 | 4.69 | 4.8 | 4.92 | 5.03 | 5.15 | 5.26 | 5.38 | 5.49 | 5.61 | 5.72 | 5.84 | 5.95 | 6.07 | 6.18 | 6.3 | 6.41 | 6.53 | 6.64 | 6.76 | 6.87 | 6.99 | 7.1 | 7.22 | 7.33 | 7.45 | 7.56 | 7.67 | 7.79 | 7.9 | 8.02 | 8.13 | 8.25 | 8.36 | 8.48 | 8.59 | 8.71 | 8.82 | 8.94
23.63 | 4.34 | 4.46 | 4.57 | 4.69 | 4.8 | 4.92 | 5.03 | 5.15 | 5.26 | 5.38 | 5.49 | 5.61 | 5.72 | 5.84 | 5.95 | 6.07 | 6.18 | 6.3 | 6.41 | 6.53 | 6.64 | 6.76 | 6.87 | 6.98 | 7.1 | 7.21 | 7.33 | 7.44 | 7.56 | 7.67 | 7.79 | 7.9 | 8.02 | 8.13 | 8.25 | 8.36 | 8.48 | 8.59 | 8.71 | 8.82 | 8.94
23.64 | 4.34 | 4.46 | 4.57 | 4.69 | 4.8 | 4.92 | 5.03 | 5.15 | 5.26 | 5.38 | 5.49 | 5.61 | 5.72 | 5.84 | 5.95 | 6.07 | 6.18 | 6.29 | 6.41 | 6.52 | 6.64 | 6.75 | 6.87 | 6.98 | 7.1 | 7.21 | 7.33 | 7.44 | 7.56 | 7.67 | 7.79 | 7.9 | 8.02 | 8.13 | 8.25 | 8.36 | 8.48 | 8.59 | 8.7 | 8.82 | 8.93
23.65 | 4.34 | 4.46 | 4.57 | 4.69 | 4.8 | 4.92 | 5.03 | 5.15 | 5.26 | 5.38 | 5.49 | 5.6 | 5.72 | 5.83 | 5.95 | 6.06 | 6.18 | 6.29 | 6.41 | 6.52 | 6.64 | 6.75 | 6.87 | 6.98 | 7.1 | 7.21 | 7.33 | 7.44 | 7.56 | 7.67 | 7.79 | 7.9 | 8.01 | 8.13 | 8.24 | 8.36 | 8.47 | 8.59 | 8.7 | 8.82 | 8.93
23.66 | 4.34 | 4.46 | 4.57 | 4.69 | 4.8 | 4.92 | 5.03 | 5.14 | 5.26 | 5.37 | 5.49 | 5.6 | 5.72 | 5.83 | 5.95 | 6.06 | 6.18 | 6.29 | 6.41 | 6.52 | 6.64 | 6.75 | 6.87 | 6.98 | 7.1 | 7.21 | 7.32 | 7.44 | 7.55 | 7.67 | 7.78 | 7.9 | 8.01 | 8.13 | 8.24 | 8.36 | 8.47 | 8.59 | 8.7 | 8.82 | 8.93
23.67 | 4.34 | 4.46 | 4.57 | 4.68 | 4.8 | 4.91 | 5.03 | 5.14 | 5.26 | 5.37 | 5.49 | 5.6 | 5.72 | 5.83 | 5.95 | 6.06 | 6.18 | 6.29 | 6.41 | 6.52 | 6.64 | 6.75 | 6.86 | 6.98 | 7.09 | 7.21 | 7.32 | 7.44 | 7.55 | 7.67 | 7.78 | 7.9 | 8.01 | 8.13 | 8.24 | 8.36 | 8.47 | 8.59 | 8.7 | 8.81 | 8.93
23.68 | 4.34 | 4.45 | 4.57 | 4.68 | 4.8 | 4.91 | 5.03 | 5.14 | 5.26 | 5.37 | 5.49 | 5.6 | 5.72 | 5.83 | 5.95 | 6.06 | 6.18 | 6.29 | 6.4 | 6.52 | 6.63 | 6.75 | 6.86 | 6.98 | 7.09 | 7.21 | 7.32 | 7.44 | 7.55 | 7.67 | 7.78 | 7.9 | 8.01 | 8.12 | 8.24 | 8.35 | 8.47 | 8.58 | 8.7 | 8.81 | 8.93
23.69 | 4.34 | 4.45 | 4.57 | 4.68 | 4.8 | 4.91 | 5.03 | 5.14 | 5.26 | 5.37 | 5.49 | 5.6 | 5.72 | 5.83 | 5.94 | 6.06 | 6.17 | 6.29 | 6.4 | 6.52 | 6.63 | 6.75 | 6.86 | 6.98 | 7.09 | 7.21 | 7.32 | 7.44 | 7.55 | 7.66 | 7.78 | 7.89 | 8.01 | 8.12 | 8.24 | 8.35 | 8.47 | 8.58 | 8.7 | 8.81 | 8.93
23.7 | 4.34 | 4.45 | 4.57 | 4.68 | 4.8 | 4.91 | 5.03 | 5.14 | 5.26 | 5.37 | 5.48 | 5.6 | 5.71 | 5.83 | 5.94 | 6.06 | 6.17 | 6.29 | 6.4 | 6.52 | 6.63 | 6.75 | 6.86 | 6.98 | 7.09 | 7.2 | 7.32 | 7.43 | 7.55 | 7.66 | 7.78 | 7.89 | 8.01 | 8.12 | 8.24 | 8.35 | 8.47 | 8.58 | 8.69 | 8.81 | 8.92
23.71 | 4.34 | 4.45 | 4.57 | 4.68 | 4.8 | 4.91 | 5.02 | 5.14 | 5.25 | 5.37 | 5.48 | 5.6 | 5.71 | 5.83 | 5.94 | 6.06 | 6.17 | 6.29 | 6.4 | 6.52 | 6.63 | 6.74 | 6.86 | 6.97 | 7.09 | 7.2 | 7.32 | 7.43 | 7.55 | 7.66 | 7.78 | 7.89 | 8.01 | 8.12 | 8.23 | 8.35 | 8.46 | 8.58 | 8.69 | 8.81 | 8.92
23.72 | 4.34 | 4.45 | 4.57 | 4.68 | 4.79 | 4.91 | 5.02 | 5.14 | 5.25 | 5.37 | 5.48 | 5.6 | 5.71 | 5.83 | 5.94 | 6.06 | 6.17 | 6.28 | 6.4 | 6.51 | 6.63 | 6.74 | 6.86 | 6.97 | 7.09 | 7.2 | 7.32 | 7.43 | 7.55 | 7.66 | 7.77 | 7.89 | 8 | 8.12 | 8.23 | 8.35 | 8.46 | 8.58 | 8.69 | 8.81 | 8.92
23.73 | 4.33 | 4.45 | 4.56 | 4.68 | 4.79 | 4.91 | 5.02 | 5.14 | 5.25 | 5.37 | 5.48 | 5.6 | 5.71 | 5.83 | 5.94 | 6.05 | 6.17 | 6.28 | 6.4 | 6.51 | 6.63 | 6.74 | 6.86 | 6.97 | 7.09 | 7.2 | 7.32 | 7.43 | 7.54 | 7.66 | 7.77 | 7.89 | 8 | 8.12 | 8.23 | 8.35 | 8.46 | 8.58 | 8.69 | 8.8 | 8.92
23.74 | 4.33 | 4.45 | 4.56 | 4.68 | 4.79 | 4.91 | 5.02 | 5.14 | 5.25 | 5.37 | 5.48 | 5.59 | 5.71 | 5.82 | 5.94 | 6.05 | 6.17 | 6.28 | 6.4 | 6.51 | 6.63 | 6.74 | 6.86 | 6.97 | 7.08 | 7.2 | 7.31 | 7.43 | 7.54 | 7.66 | 7.77 | 7.89 | 8 | 8.12 | 8.23 | 8.34 | 8.46 | 8.57 | 8.69 | 8.8 | 8.92
23.75 | 4.33 | 4.45 | 4.56 | 4.68 | 4.79 | 4.91 | 5.02 | 5.14 | 5.25 | 5.36 | 5.48 | 5.59 | 5.71 | 5.82 | 5.94 | 6.05 | 6.17 | 6.28 | 6.4 | 6.51 | 6.62 | 6.74 | 6.85 | 6.97 | 7.08 | 7.2 | 7.31 | 7.43 | 7.54 | 7.66 | 7.77 | 7.89 | 8 | 8.11 | 8.23 | 8.34 | 8.46 | 8.57 | 8.69 | 8.8 | 8.92
23.76 | 4.33 | 4.45 | 4.56 | 4.68 | 4.79 | 4.91 | 5.02 | 5.13 | 5.25 | 5.36 | 5.48 | 5.59 | 5.71 | 5.82 | 5.94 | 6.05 | 6.17 | 6.28 | 6.39 | 6.51 | 6.62 | 6.74 | 6.85 | 6.97 | 7.08 | 7.2 | 7.31 | 7.43 | 7.54 | 7.65 | 7.77 | 7.88 | 8 | 8.11 | 8.23 | 8.34 | 8.46 | 8.57 | 8.69 | 8.8 | 8.91
23.77 | 4.33 | 4.45 | 4.56 | 4.68 | 4.79 | 4.9 | 5.02 | 5.13 | 5.25 | 5.36 | 5.48 | 5.59 | 5.71 | 5.82 | 5.94 | 6.05 | 6.16 | 6.28 | 6.39 | 6.51 | 6.62 | 6.74 | 6.85 | 6.97 | 7.08 | 7.19 | 7.31 | 7.42 | 7.54 | 7.65 | 7.77 | 7.88 | 8 | 8.11 | 8.23 | 8.34 | 8.45 | 8.57 | 8.68 | 8.8 | 8.91
23.78 | 4.33 | 4.45 | 4.56 | 4.67 | 4.79 | 4.9 | 5.02 | 5.13 | 5.25 | 5.36 | 5.48 | 5.59 | 5.7 | 5.82 | 5.93 | 6.05 | 6.16 | 6.28 | 6.39 | 6.51 | 6.62 | 6.74 | 6.85 | 6.96 | 7.08 | 7.19 | 7.31 | 7.42 | 7.54 | 7.65 | 7.77 | 7.88 | 8 | 8.11 | 8.22 | 8.34 | 8.45 | 8.57 | 8.68 | 8.8 | 8.91
23.79 | 4.33 | 4.44 | 4.56 | 4.67 | 4.79 | 4.9 | 5.02 | 5.13 | 5.25 | 5.36 | 5.47 | 5.59 | 5.7 | 5.82 | 5.93 | 6.05 | 6.16 | 6.28 | 6.39 | 6.51 | 6.62 | 6.73 | 6.85 | 6.96 | 7.08 | 7.19 | 7.31 | 7.42 | 7.54 | 7.65 | 7.76 | 7.88 | 7.99 | 8.11 | 8.22 | 8.34 | 8.45 | 8.57 | 8.68 | 8.79 | 8.91
23.8 | 4.33 | 4.44 | 4.56 | 4.67 | 4.79 | 4.9 | 5.02 | 5.13 | 5.24 | 5.36 | 5.47 | 5.59 | 5.7 | 5.82 | 5.93 | 6.05 | 6.16 | 6.28 | 6.39 | 6.5 | 6.62 | 6.73 | 6.85 | 6.96 | 7.08 | 7.19 | 7.31 | 7.42 | 7.53 | 7.65 | 7.76 | 7.88 | 7.99 | 8.11 | 8.22 | 8.34 | 8.45 | 8.56 | 8.68 | 8.79 | 8.91
23.81 | 4.33 | 4.44 | 4.56 | 4.67 | 4.79 | 4.9 | 5.01 | 5.13 | 5.24 | 5.36 | 5.47 | 5.59 | 5.7 | 5.82 | 5.93 | 6.04 | 6.16 | 6.27 | 6.39 | 6.5 | 6.62 | 6.73 | 6.85 | 6.96 | 7.08 | 7.19 | 7.3 | 7.42 | 7.53 | 7.65 | 7.76 | 7.88 | 7.99 | 8.1 | 8.22 | 8.33 | 8.45 | 8.56 | 8.68 | 8.79 | 8.91
23.82 | 4.33 | 4.44 | 4.56 | 4.67 | 4.78 | 4.9 | 5.01 | 5.13 | 5.24 | 5.36 | 5.47 | 5.59 | 5.7 | 5.81 | 5.93 | 6.04 | 6.16 | 6.27 | 6.39 | 6.5 | 6.62 | 6.73 | 6.84 | 6.96 | 7.07 | 7.19 | 7.3 | 7.42 | 7.53 | 7.65 | 7.76 | 7.87 | 7.99 | 8.1 | 8.22 | 8.33 | 8.45 | 8.56 | 8.68 | 8.79 | 8.9
23.83 | 4.33 | 4.44 | 4.55 | 4.67 | 4.78 | 4.9 | 5.01 | 5.13 | 5.24 | 5.36 | 5.47 | 5.58 | 5.7 | 5.81 | 5.93 | 6.04 | 6.16 | 6.27 | 6.39 | 6.5 | 6.61 | 6.73 | 6.84 | 6.96 | 7.07 | 7.19 | 7.3 | 7.42 | 7.53 | 7.64 | 7.76 | 7.87 | 7.99 | 8.1 | 8.22 | 8.33 | 8.45 | 8.56 | 8.67 | 8.79 | 8.9
23.84 | 4.33 | 4.44 | 4.55 | 4.67 | 4.78 | 4.9 | 5.01 | 5.13 | 5.24 | 5.35 | 5.47 | 5.58 | 5.7 | 5.81 | 5.93 | 6.04 | 6.16 | 6.27 | 6.38 | 6.5 | 6.61 | 6.73 | 6.84 | 6.96 | 7.07 | 7.19 | 7.3 | 7.41 | 7.53 | 7.64 | 7.76 | 7.87 | 7.99 | 8.1 | 8.21 | 8.33 | 8.44 | 8.56 | 8.67 | 8.79 | 8.9
23.85 | 4.32 | 4.44 | 4.55 | 4.67 | 4.78 | 4.9 | 5.01 | 5.13 | 5.24 | 5.35 | 5.47 | 5.58 | 5.7 | 5.81 | 5.93 | 6.04 | 6.15 | 6.27 | 6.38 | 6.5 | 6.61 | 6.73 | 6.84 | 6.96 | 7.07 | 7.18 | 7.3 | 7.41 | 7.53 | 7.64 | 7.76 | 7.87 | 7.98 | 8.1 | 8.21 | 8.33 | 8.44 | 8.56 | 8.67 | 8.78 | 8.9
23.86 | 4.32 | 4.44 | 4.55 | 4.67 | 4.78 | 4.9 | 5.01 | 5.12 | 5.24 | 5.35 | 5.47 | 5.58 | 5.7 | 5.81 | 5.92 | 6.04 | 6.15 | 6.27 | 6.38 | 6.5 | 6.61 | 6.73 | 6.84 | 6.95 | 7.07 | 7.18 | 7.3 | 7.41 | 7.53 | 7.64 | 7.75 | 7.87 | 7.98 | 8.1 | 8.21 | 8.33 | 8.44 | 8.55 | 8.67 | 8.78 | 8.9
23.87 | 4.32 | 4.44 | 4.55 | 4.67 | 4.78 | 4.89 | 5.01 | 5.12 | 5.24 | 5.35 | 5.47 | 5.58 | 5.69 | 5.81 | 5.92 | 6.04 | 6.15 | 6.27 | 6.38 | 6.5 | 6.61 | 6.72 | 6.84 | 6.95 | 7.07 | 7.18 | 7.3 | 7.41 | 7.52 | 7.64 | 7.75 | 7.87 | 7.98 | 8.1 | 8.21 | 8.32 | 8.44 | 8.55 | 8.67 | 8.78 | 8.9
23.88 | 4.32 | 4.44 | 4.55 | 4.66 | 4.78 | 4.89 | 5.01 | 5.12 | 5.24 | 5.35 | 5.47 | 5.58 | 5.69 | 5.81 | 5.92 | 6.04 | 6.15 | 6.27 | 6.38 | 6.49 | 6.61 | 6.72 | 6.84 | 6.95 | 7.07 | 7.18 | 7.29 | 7.41 | 7.52 | 7.64 | 7.75 | 7.87 | 7.98 | 8.09 | 8.21 | 8.32 | 8.44 | 8.55 | 8.67 | 8.78 | 8.89
23.89 | 4.32 | 4.44 | 4.55 | 4.66 | 4.78 | 4.89 | 5.01 | 5.12 | 5.24 | 5.35 | 5.46 | 5.58 | 5.69 | 5.81 | 5.92 | 6.04 | 6.15 | 6.26 | 6.38 | 6.49 | 6.61 | 6.72 | 6.84 | 6.95 | 7.06 | 7.18 | 7.29 | 7.41 | 7.52 | 7.64 | 7.75 | 7.86 | 7.98 | 8.09 | 8.21 | 8.32 | 8.44 | 8.55 | 8.66 | 8.78 | 8.89
23.9 | 4.32 | 4.43 | 4.55 | 4.66 | 4.78 | 4.89 | 5.01 | 5.12 | 5.23 | 5.35 | 5.46 | 5.58 | 5.69 | 5.81 | 5.92 | 6.03 | 6.15 | 6.26 | 6.38 | 6.49 | 6.61 | 6.72 | 6.83 | 6.95 | 7.06 | 7.18 | 7.29 | 7.41 | 7.52 | 7.63 | 7.75 | 7.86 | 7.98 | 8.09 | 8.21 | 8.32 | 8.43 | 8.55 | 8.66 | 8.78 | 8.89
23.91 | 4.32 | 4.43 | 4.55 | 4.66 | 4.78 | 4.89 | 5 | 5.12 | 5.23 | 5.35 | 5.46 | 5.58 | 5.69 | 5.8 | 5.92 | 6.03 | 6.15 | 6.26 | 6.38 | 6.49 | 6.6 | 6.72 | 6.83 | 6.95 | 7.06 | 7.18 | 7.29 | 7.4 | 7.52 | 7.63 | 7.75 | 7.86 | 7.98 | 8.09 | 8.2 | 8.32 | 8.43 | 8.55 | 8.66 | 8.78 | 8.89
23.92 | 4.32 | 4.43 | 4.55 | 4.66 | 4.78 | 4.89 | 5 | 5.12 | 5.23 | 5.35 | 5.46 | 5.58 | 5.69 | 5.8 | 5.92 | 6.03 | 6.15 | 6.26 | 6.37 | 6.49 | 6.6 | 6.72 | 6.83 | 6.95 | 7.06 | 7.17 | 7.29 | 7.4 | 7.52 | 7.63 | 7.75 | 7.86 | 7.97 | 8.09 | 8.2 | 8.32 | 8.43 | 8.55 | 8.66 | 8.77 | 8.89
23.93 | 4.32 | 4.43 | 4.55 | 4.66 | 4.77 | 4.89 | 5 | 5.12 | 5.23 | 5.35 | 5.46 | 5.57 | 5.69 | 5.8 | 5.92 | 6.03 | 6.15 | 6.26 | 6.37 | 6.49 | 6.6 | 6.72 | 6.83 | 6.94 | 7.06 | 7.17 | 7.29 | 7.4 | 7.52 | 7.63 | 7.74 | 7.86 | 7.97 | 8.09 | 8.2 | 8.32 | 8.43 | 8.54 | 8.66 | 8.77 | 8.89
23.94 | 4.32 | 4.43 | 4.54 | 4.66 | 4.77 | 4.89 | 5 | 5.12 | 5.23 | 5.34 | 5.46 | 5.57 | 5.69 | 5.8 | 5.92 | 6.03 | 6.14 | 6.26 | 6.37 | 6.49 | 6.6 | 6.71 | 6.83 | 6.94 | 7.06 | 7.17 | 7.29 | 7.4 | 7.51 | 7.63 | 7.74 | 7.86 | 7.97 | 8.09 | 8.2 | 8.31 | 8.43 | 8.54 | 8.66 | 8.77 | 8.88
23.95 | 4.32 | 4.43 | 4.54 | 4.66 | 4.77 | 4.89 | 5 | 5.11 | 5.23 | 5.34 | 5.46 | 5.57 | 5.69 | 5.8 | 5.91 | 6.03 | 6.14 | 6.26 | 6.37 | 6.49 | 6.6 | 6.71 | 6.83 | 6.94 | 7.06 | 7.17 | 7.28 | 7.4 | 7.51 | 7.63 | 7.74 | 7.86 | 7.97 | 8.08 | 8.2 | 8.31 | 8.43 | 8.54 | 8.65 | 8.77 | 8.88
23.96 | 4.31 | 4.43 | 4.54 | 4.66 | 4.77 | 4.89 | 5 | 5.11 | 5.23 | 5.34 | 5.46 | 5.57 | 5.68 | 5.8 | 5.91 | 6.03 | 6.14 | 6.26 | 6.37 | 6.48 | 6.6 | 6.71 | 6.83 | 6.94 | 7.05 | 7.17 | 7.28 | 7.4 | 7.51 | 7.63 | 7.74 | 7.85 | 7.97 | 8.08 | 8.2 | 8.31 | 8.42 | 8.54 | 8.65 | 8.77 | 8.88
23.97 | 4.31 | 4.43 | 4.54 | 4.66 | 4.77 | 4.88 | 5 | 5.11 | 5.23 | 5.34 | 5.46 | 5.57 | 5.68 | 5.8 | 5.91 | 6.03 | 6.14 | 6.25 | 6.37 | 6.48 | 6.6 | 6.71 | 6.83 | 6.94 | 7.05 | 7.17 | 7.28 | 7.4 | 7.51 | 7.62 | 7.74 | 7.85 | 7.97 | 8.08 | 8.19 | 8.31 | 8.42 | 8.54 | 8.65 | 8.77 | 8.88
23.98 | 4.31 | 4.43 | 4.54 | 4.66 | 4.77 | 4.88 | 5 | 5.11 | 5.23 | 5.34 | 5.45 | 5.57 | 5.68 | 5.8 | 5.91 | 6.03 | 6.14 | 6.25 | 6.37 | 6.48 | 6.6 | 6.71 | 6.82 | 6.94 | 7.05 | 7.17 | 7.28 | 7.39 | 7.51 | 7.62 | 7.74 | 7.85 | 7.96 | 8.08 | 8.19 | 8.31 | 8.42 | 8.54 | 8.65 | 8.76 | 8.88
23.99 | 4.31 | 4.43 | 4.54 | 4.65 | 4.77 | 4.88 | 5 | 5.11 | 5.23 | 5.34 | 5.45 | 5.57 | 5.68 | 5.8 | 5.91 | 6.02 | 6.14 | 6.25 | 6.37 | 6.48 | 6.59 | 6.71 | 6.82 | 6.94 | 7.05 | 7.16 | 7.28 | 7.39 | 7.51 | 7.62 | 7.74 | 7.85 | 7.96 | 8.08 | 8.19 | 8.31 | 8.42 | 8.53 | 8.65 | 8.76 | 8.88
24 | 4.31 | 4.43 | 4.54 | 4.65 | 4.77 | 4.88 | 5 | 5.11 | 5.22 | 5.34 | 5.45 | 5.57 | 5.68 | 5.79 | 5.91 | 6.02 | 6.14 | 6.25 | 6.36 | 6.48 | 6.59 | 6.71 | 6.82 | 6.94 | 7.05 | 7.16 | 7.28 | 7.39 | 7.51 | 7.62 | 7.73 | 7.85 | 7.96 | 8.08 | 8.19 | 8.3 | 8.42 | 8.53 | 8.65 | 8.76 | 8.87
24.01 | 4.31 | 4.42 | 4.54 | 4.65 | 4.77 | 4.88 | 4.99 | 5.11 | 5.22 | 5.34 | 5.45 | 5.57 | 5.68 | 5.79 | 5.91 | 6.02 | 6.14 | 6.25 | 6.36 | 6.48 | 6.59 | 6.71 | 6.82 | 6.93 | 7.05 | 7.16 | 7.28 | 7.39 | 7.5 | 7.62 | 7.73 | 7.85 | 7.96 | 8.07 | 8.19 | 8.3 | 8.42 | 8.53 | 8.64 | 8.76 | 8.87
24.02 | 4.31 | 4.42 | 4.54 | 4.65 | 4.77 | 4.88 | 4.99 | 5.11 | 5.22 | 5.34 | 5.45 | 5.56 | 5.68 | 5.79 | 5.91 | 6.02 | 6.13 | 6.25 | 6.36 | 6.48 | 6.59 | 6.7 | 6.82 | 6.93 | 7.05 | 7.16 | 7.27 | 7.39 | 7.5 | 7.62 | 7.73 | 7.84 | 7.96 | 8.07 | 8.19 | 8.3 | 8.42 | 8.53 | 8.64 | 8.76 | 8.87
24.03 | 4.31 | 4.42 | 4.54 | 4.65 | 4.76 | 4.88 | 4.99 | 5.11 | 5.22 | 5.33 | 5.45 | 5.56 | 5.68 | 5.79 | 5.91 | 6.02 | 6.13 | 6.25 | 6.36 | 6.48 | 6.59 | 6.7 | 6.82 | 6.93 | 7.05 | 7.16 | 7.27 | 7.39 | 7.5 | 7.62 | 7.73 | 7.84 | 7.96 | 8.07 | 8.19 | 8.3 | 8.41 | 8.53 | 8.64 | 8.76 | 8.87
24.04 | 4.31 | 4.42 | 4.54 | 4.65 | 4.76 | 4.88 | 4.99 | 5.11 | 5.22 | 5.33 | 5.45 | 5.56 | 5.68 | 5.79 | 5.9 | 6.02 | 6.13 | 6.25 | 6.36 | 6.47 | 6.59 | 6.7 | 6.82 | 6.93 | 7.04 | 7.16 | 7.27 | 7.39 | 7.5 | 7.61 | 7.73 | 7.84 | 7.96 | 8.07 | 8.18 | 8.3 | 8.41 | 8.53 | 8.64 | 8.75 | 8.87
24.05 | 4.31 | 4.42 | 4.53 | 4.65 | 4.76 | 4.88 | 4.99 | 5.1 | 5.22 | 5.33 | 5.45 | 5.56 | 5.67 | 5.79 | 5.9 | 6.02 | 6.13 | 6.24 | 6.36 | 6.47 | 6.59 | 6.7 | 6.81 | 6.93 | 7.04 | 7.16 | 7.27 | 7.38 | 7.5 | 7.61 | 7.73 | 7.84 | 7.95 | 8.07 | 8.18 | 8.3 | 8.41 | 8.52 | 8.64 | 8.75 | 8.87
24.06 | 4.31 | 4.42 | 4.53 | 4.65 | 4.76 | 4.88 | 4.99 | 5.1 | 5.22 | 5.33 | 5.45 | 5.56 | 5.67 | 5.79 | 5.9 | 6.02 | 6.13 | 6.24 | 6.36 | 6.47 | 6.59 | 6.7 | 6.81 | 6.93 | 7.04 | 7.16 | 7.27 | 7.38 | 7.5 | 7.61 | 7.73 | 7.84 | 7.95 | 8.07 | 8.18 | 8.29 | 8.41 | 8.52 | 8.64 | 8.75 | 8.86
24.07 | 4.3 | 4.42 | 4.53 | 4.65 | 4.76 | 4.87 | 4.99 | 5.1 | 5.22 | 5.33 | 5.44 | 5.56 | 5.67 | 5.79 | 5.9 | 6.01 | 6.13 | 6.24 | 6.36 | 6.47 | 6.58 | 6.7 | 6.81 | 6.93 | 7.04 | 7.15 | 7.27 | 7.38 | 7.5 | 7.61 | 7.72 | 7.84 | 7.95 | 8.07 | 8.18 | 8.29 | 8.41 | 8.52 | 8.64 | 8.75 | 8.86
24.08 | 4.3 | 4.42 | 4.53 | 4.65 | 4.76 | 4.87 | 4.99 | 5.1 | 5.22 | 5.33 | 5.44 | 5.56 | 5.67 | 5.79 | 5.9 | 6.01 | 6.13 | 6.24 | 6.36 | 6.47 | 6.58 | 6.7 | 6.81 | 6.92 | 7.04 | 7.15 | 7.27 | 7.38 | 7.49 | 7.61 | 7.72 | 7.84 | 7.95 | 8.06 | 8.18 | 8.29 | 8.41 | 8.52 | 8.63 | 8.75 | 8.86
24.09 | 4.3 | 4.42 | 4.53 | 4.64 | 4.76 | 4.87 | 4.99 | 5.1 | 5.21 | 5.33 | 5.44 | 5.56 | 5.67 | 5.78 | 5.9 | 6.01 | 6.13 | 6.24 | 6.35 | 6.47 | 6.58 | 6.7 | 6.81 | 6.92 | 7.04 | 7.15 | 7.27 | 7.38 | 7.49 | 7.61 | 7.72 | 7.83 | 7.95 | 8.06 | 8.18 | 8.29 | 8.4 | 8.52 | 8.63 | 8.75 | 8.86
24.1 | 4.3 | 4.42 | 4.53 | 4.64 | 4.76 | 4.87 | 4.99 | 5.1 | 5.21 | 5.33 | 5.44 | 5.56 | 5.67 | 5.78 | 5.9 | 6.01 | 6.12 | 6.24 | 6.35 | 6.47 | 6.58 | 6.69 | 6.81 | 6.92 | 7.04 | 7.15 | 7.26 | 7.38 | 7.49 | 7.61 | 7.72 | 7.83 | 7.95 | 8.06 | 8.17 | 8.29 | 8.4 | 8.52 | 8.63 | 8.74 | 8.86
24.11 | 4.3 | 4.42 | 4.53 | 4.64 | 4.76 | 4.87 | 4.98 | 5.1 | 5.21 | 5.33 | 5.44 | 5.55 | 5.67 | 5.78 | 5.9 | 6.01 | 6.12 | 6.24 | 6.35 | 6.47 | 6.58 | 6.69 | 6.81 | 6.92 | 7.03 | 7.15 | 7.26 | 7.38 | 7.49 | 7.6 | 7.72 | 7.83 | 7.95 | 8.06 | 8.17 | 8.29 | 8.4 | 8.51 | 8.63 | 8.74 | 8.86
24.12 | 4.3 | 4.41 | 4.53 | 4.64 | 4.76 | 4.87 | 4.98 | 5.1 | 5.21 | 5.33 | 5.44 | 5.55 | 5.67 | 5.78 | 5.89 | 6.01 | 6.12 | 6.24 | 6.35 | 6.46 | 6.58 | 6.69 | 6.81 | 6.92 | 7.03 | 7.15 | 7.26 | 7.37 | 7.49 | 7.6 | 7.72 | 7.83 | 7.94 | 8.06 | 8.17 | 8.29 | 8.4 | 8.51 | 8.63 | 8.74 | 8.85
24.13 | 4.3 | 4.41 | 4.53 | 4.64 | 4.76 | 4.87 | 4.98 | 5.1 | 5.21 | 5.32 | 5.44 | 5.55 | 5.67 | 5.78 | 5.89 | 6.01 | 6.12 | 6.24 | 6.35 | 6.46 | 6.58 | 6.69 | 6.8 | 6.92 | 7.03 | 7.15 | 7.26 | 7.37 | 7.49 | 7.6 | 7.71 | 7.83 | 7.94 | 8.06 | 8.17 | 8.28 | 8.4 | 8.51 | 8.63 | 8.74 | 8.85
24.14 | 4.3 | 4.41 | 4.53 | 4.64 | 4.75 | 4.87 | 4.98 | 5.1 | 5.21 | 5.32 | 5.44 | 5.55 | 5.66 | 5.78 | 5.89 | 6.01 | 6.12 | 6.23 | 6.35 | 6.46 | 6.58 | 6.69 | 6.8 | 6.92 | 7.03 | 7.14 | 7.26 | 7.37 | 7.49 | 7.6 | 7.71 | 7.83 | 7.94 | 8.05 | 8.17 | 8.28 | 8.4 | 8.51 | 8.62 | 8.74 | 8.85
24.15 | 4.3 | 4.41 | 4.53 | 4.64 | 4.75 | 4.87 | 4.98 | 5.09 | 5.21 | 5.32 | 5.44 | 5.55 | 5.66 | 5.78 | 5.89 | 6.01 | 6.12 | 6.23 | 6.35 | 6.46 | 6.57 | 6.69 | 6.8 | 6.92 | 7.03 | 7.14 | 7.26 | 7.37 | 7.48 | 7.6 | 7.71 | 7.83 | 7.94 | 8.05 | 8.17 | 8.28 | 8.39 | 8.51 | 8.62 | 8.74 | 8.85
24.16 | 4.3 | 4.41 | 4.52 | 4.64 | 4.75 | 4.87 | 4.98 | 5.09 | 5.21 | 5.32 | 5.44 | 5.55 | 5.66 | 5.78 | 5.89 | 6 | 6.12 | 6.23 | 6.35 | 6.46 | 6.57 | 6.69 | 6.8 | 6.91 | 7.03 | 7.14 | 7.26 | 7.37 | 7.48 | 7.6 | 7.71 | 7.82 | 7.94 | 8.05 | 8.17 | 8.28 | 8.39 | 8.51 | 8.62 | 8.73 | 8.85
24.17 | 4.3 | 4.41 | 4.52 | 4.64 | 4.75 | 4.87 | 4.98 | 5.09 | 5.21 | 5.32 | 5.43 | 5.55 | 5.66 | 5.78 | 5.89 | 6 | 6.12 | 6.23 | 6.34 | 6.46 | 6.57 | 6.69 | 6.8 | 6.91 | 7.03 | 7.14 | 7.25 | 7.37 | 7.48 | 7.6 | 7.71 | 7.82 | 7.94 | 8.05 | 8.16 | 8.28 | 8.39 | 8.51 | 8.62 | 8.73 | 8.85
24.18 | 4.3 | 4.41 | 4.52 | 4.64 | 4.75 | 4.86 | 4.98 | 5.09 | 5.21 | 5.32 | 5.43 | 5.55 | 5.66 | 5.77 | 5.89 | 6 | 6.12 | 6.23 | 6.34 | 6.46 | 6.57 | 6.68 | 6.8 | 6.91 | 7.03 | 7.14 | 7.25 | 7.37 | 7.48 | 7.59 | 7.71 | 7.82 | 7.94 | 8.05 | 8.16 | 8.28 | 8.39 | 8.5 | 8.62 | 8.73 | 8.84
24.19 | 4.29 | 4.41 | 4.52 | 4.64 | 4.75 | 4.86 | 4.98 | 5.09 | 5.2 | 5.32 | 5.43 | 5.55 | 5.66 | 5.77 | 5.89 | 6 | 6.11 | 6.23 | 6.34 | 6.46 | 6.57 | 6.68 | 6.8 | 6.91 | 7.02 | 7.14 | 7.25 | 7.37 | 7.48 | 7.59 | 7.71 | 7.82 | 7.93 | 8.05 | 8.16 | 8.27 | 8.39 | 8.5 | 8.62 | 8.73 | 8.84
24.2 | 4.29 | 4.41 | 4.52 | 4.63 | 4.75 | 4.86 | 4.98 | 5.09 | 5.2 | 5.32 | 5.43 | 5.54 | 5.66 | 5.77 | 5.89 | 6 | 6.11 | 6.23 | 6.34 | 6.45 | 6.57 | 6.68 | 6.8 | 6.91 | 7.02 | 7.14 | 7.25 | 7.36 | 7.48 | 7.59 | 7.7 | 7.82 | 7.93 | 8.05 | 8.16 | 8.27 | 8.39 | 8.5 | 8.61 | 8.73 | 8.84
24.21 | 4.29 | 4.41 | 4.52 | 4.63 | 4.75 | 4.86 | 4.97 | 5.09 | 5.2 | 5.32 | 5.43 | 5.54 | 5.66 | 5.77 | 5.88 | 6 | 6.11 | 6.23 | 6.34 | 6.45 | 6.57 | 6.68 | 6.79 | 6.91 | 7.02 | 7.14 | 7.25 | 7.36 | 7.48 | 7.59 | 7.7 | 7.82 | 7.93 | 8.04 | 8.16 | 8.27 | 8.39 | 8.5 | 8.61 | 8.73 | 8.84
24.22 | 4.29 | 4.41 | 4.52 | 4.63 | 4.75 | 4.86 | 4.97 | 5.09 | 5.2 | 5.31 | 5.43 | 5.54 | 5.66 | 5.77 | 5.88 | 6 | 6.11 | 6.22 | 6.34 | 6.45 | 6.57 | 6.68 | 6.79 | 6.91 | 7.02 | 7.13 | 7.25 | 7.36 | 7.47 | 7.59 | 7.7 | 7.82 | 7.93 | 8.04 | 8.16 | 8.27 | 8.38 | 8.5 | 8.61 | 8.72 | 8.84
24.23 | 4.29 | 4.4 | 4.52 | 4.63 | 4.75 | 4.86 | 4.97 | 5.09 | 5.2 | 5.31 | 5.43 | 5.54 | 5.65 | 5.77 | 5.88 | 6 | 6.11 | 6.22 | 6.34 | 6.45 | 6.56 | 6.68 | 6.79 | 6.91 | 7.02 | 7.13 | 7.25 | 7.36 | 7.47 | 7.59 | 7.7 | 7.81 | 7.93 | 8.04 | 8.15 | 8.27 | 8.38 | 8.5 | 8.61 | 8.72 | 8.84
24.24 | 4.29 | 4.4 | 4.52 | 4.63 | 4.74 | 4.86 | 4.97 | 5.09 | 5.2 | 5.31 | 5.43 | 5.54 | 5.65 | 5.77 | 5.88 | 5.99 | 6.11 | 6.22 | 6.34 | 6.45 | 6.56 | 6.68 | 6.79 | 6.9 | 7.02 | 7.13 | 7.24 | 7.36 | 7.47 | 7.59 | 7.7 | 7.81 | 7.93 | 8.04 | 8.15 | 8.27 | 8.38 | 8.49 | 8.61 | 8.72 | 8.83
24.25 | 4.29 | 4.4 | 4.52 | 4.63 | 4.74 | 4.86 | 4.97 | 5.08 | 5.2 | 5.31 | 5.43 | 5.54 | 5.65 | 5.77 | 5.88 | 5.99 | 6.11 | 6.22 | 6.33 | 6.45 | 6.56 | 6.68 | 6.79 | 6.9 | 7.02 | 7.13 | 7.24 | 7.36 | 7.47 | 7.58 | 7.7 | 7.81 | 7.92 | 8.04 | 8.15 | 8.27 | 8.38 | 8.49 | 8.61 | 8.72 | 8.83
24.26 | 4.29 | 4.4 | 4.52 | 4.63 | 4.74 | 4.86 | 4.97 | 5.08 | 5.2 | 5.31 | 5.42 | 5.54 | 5.65 | 5.77 | 5.88 | 5.99 | 6.11 | 6.22 | 6.33 | 6.45 | 6.56 | 6.67 | 6.79 | 6.9 | 7.01 | 7.13 | 7.24 | 7.36 | 7.47 | 7.58 | 7.7 | 7.81 | 7.92 | 8.04 | 8.15 | 8.26 | 8.38 | 8.49 | 8.6 | 8.72 | 8.83
24.27 | 4.29 | 4.4 | 4.51 | 4.63 | 4.74 | 4.86 | 4.97 | 5.08 | 5.2 | 5.31 | 5.42 | 5.54 | 5.65 | 5.76 | 5.88 | 5.99 | 6.1 | 6.22 | 6.33 | 6.45 | 6.56 | 6.67 | 6.79 | 6.9 | 7.01 | 7.13 | 7.24 | 7.35 | 7.47 | 7.58 | 7.69 | 7.81 | 7.92 | 8.04 | 8.15 | 8.26 | 8.38 | 8.49 | 8.6 | 8.72 | 8.83
24.28 | 4.29 | 4.4 | 4.51 | 4.63 | 4.74 | 4.85 | 4.97 | 5.08 | 5.2 | 5.31 | 5.42 | 5.54 | 5.65 | 5.76 | 5.88 | 5.99 | 6.1 | 6.22 | 6.33 | 6.44 | 6.56 | 6.67 | 6.78 | 6.9 | 7.01 | 7.13 | 7.24 | 7.35 | 7.47 | 7.58 | 7.69 | 7.81 | 7.92 | 8.03 | 8.15 | 8.26 | 8.37 | 8.49 | 8.6 | 8.71 | 8.83
24.29 | 4.29 | 4.4 | 4.51 | 4.63 | 4.74 | 4.85 | 4.97 | 5.08 | 5.19 | 5.31 | 5.42 | 5.53 | 5.65 | 5.76 | 5.88 | 5.99 | 6.1 | 6.22 | 6.33 | 6.44 | 6.56 | 6.67 | 6.78 | 6.9 | 7.01 | 7.12 | 7.24 | 7.35 | 7.46 | 7.58 | 7.69 | 7.81 | 7.92 | 8.03 | 8.15 | 8.26 | 8.37 | 8.49 | 8.6 | 8.71 | 8.83
24.3 | 4.28 | 4.4 | 4.51 | 4.63 | 4.74 | 4.85 | 4.97 | 5.08 | 5.19 | 5.31 | 5.42 | 5.53 | 5.65 | 5.76 | 5.87 | 5.99 | 6.1 | 6.21 | 6.33 | 6.44 | 6.56 | 6.67 | 6.78 | 6.9 | 7.01 | 7.12 | 7.24 | 7.35 | 7.46 | 7.58 | 7.69 | 7.8 | 7.92 | 8.03 | 8.14 | 8.26 | 8.37 | 8.48 | 8.6 | 8.71 | 8.83
24.31 | 4.28 | 4.4 | 4.51 | 4.62 | 4.74 | 4.85 | 4.97 | 5.08 | 5.19 | 5.31 | 5.42 | 5.53 | 5.65 | 5.76 | 5.87 | 5.99 | 6.1 | 6.21 | 6.33 | 6.44 | 6.55 | 6.67 | 6.78 | 6.89 | 7.01 | 7.12 | 7.24 | 7.35 | 7.46 | 7.58 | 7.69 | 7.8 | 7.92 | 8.03 | 8.14 | 8.26 | 8.37 | 8.48 | 8.6 | 8.71 | 8.82
24.32 | 4.28 | 4.4 | 4.51 | 4.62 | 4.74 | 4.85 | 4.96 | 5.08 | 5.19 | 5.3 | 5.42 | 5.53 | 5.64 | 5.76 | 5.87 | 5.99 | 6.1 | 6.21 | 6.33 | 6.44 | 6.55 | 6.67 | 6.78 | 6.89 | 7.01 | 7.12 | 7.23 | 7.35 | 7.46 | 7.57 | 7.69 | 7.8 | 7.91 | 8.03 | 8.14 | 8.25 | 8.37 | 8.48 | 8.6 | 8.71 | 8.82
24.33 | 4.28 | 4.4 | 4.51 | 4.62 | 4.74 | 4.85 | 4.96 | 5.08 | 5.19 | 5.3 | 5.42 | 5.53 | 5.64 | 5.76 | 5.87 | 5.98 | 6.1 | 6.21 | 6.32 | 6.44 | 6.55 | 6.67 | 6.78 | 6.89 | 7.01 | 7.12 | 7.23 | 7.35 | 7.46 | 7.57 | 7.69 | 7.8 | 7.91 | 8.03 | 8.14 | 8.25 | 8.37 | 8.48 | 8.59 | 8.71 | 8.82
24.34 | 4.28 | 4.39 | 4.51 | 4.62 | 4.74 | 4.85 | 4.96 | 5.08 | 5.19 | 5.3 | 5.42 | 5.53 | 5.64 | 5.76 | 5.87 | 5.98 | 6.1 | 6.21 | 6.32 | 6.44 | 6.55 | 6.66 | 6.78 | 6.89 | 7 | 7.12 | 7.23 | 7.34 | 7.46 | 7.57 | 7.68 | 7.8 | 7.91 | 8.02 | 8.14 | 8.25 | 8.37 | 8.48 | 8.59 | 8.71 | 8.82
24.35 | 4.28 | 4.39 | 4.51 | 4.62 | 4.73 | 4.85 | 4.96 | 5.07 | 5.19 | 5.3 | 5.41 | 5.53 | 5.64 | 5.76 | 5.87 | 5.98 | 6.1 | 6.21 | 6.32 | 6.44 | 6.55 | 6.66 | 6.78 | 6.89 | 7 | 7.12 | 7.23 | 7.34 | 7.46 | 7.57 | 7.68 | 7.8 | 7.91 | 8.02 | 8.14 | 8.25 | 8.36 | 8.48 | 8.59 | 8.7 | 8.82
24.36 | 4.28 | 4.39 | 4.51 | 4.62 | 4.73 | 4.85 | 4.96 | 5.07 | 5.19 | 5.3 | 5.41 | 5.53 | 5.64 | 5.75 | 5.87 | 5.98 | 6.09 | 6.21 | 6.32 | 6.43 | 6.55 | 6.66 | 6.77 | 6.89 | 7 | 7.11 | 7.23 | 7.34 | 7.45 | 7.57 | 7.68 | 7.8 | 7.91 | 8.02 | 8.14 | 8.25 | 8.36 | 8.48 | 8.59 | 8.7 | 8.82
24.37 | 4.28 | 4.39 | 4.51 | 4.62 | 4.73 | 4.85 | 4.96 | 5.07 | 5.19 | 5.3 | 5.41 | 5.53 | 5.64 | 5.75 | 5.87 | 5.98 | 6.09 | 6.21 | 6.32 | 6.43 | 6.55 | 6.66 | 6.77 | 6.89 | 7 | 7.11 | 7.23 | 7.34 | 7.45 | 7.57 | 7.68 | 7.79 | 7.91 | 8.02 | 8.13 | 8.25 | 8.36 | 8.47 | 8.59 | 8.7 | 8.81
24.38 | 4.28 | 4.39 | 4.5 | 4.62 | 4.73 | 4.84 | 4.96 | 5.07 | 5.18 | 5.3 | 5.41 | 5.53 | 5.64 | 5.75 | 5.87 | 5.98 | 6.09 | 6.21 | 6.32 | 6.43 | 6.55 | 6.66 | 6.77 | 6.89 | 7 | 7.11 | 7.23 | 7.34 | 7.45 | 7.57 | 7.68 | 7.79 | 7.91 | 8.02 | 8.13 | 8.25 | 8.36 | 8.47 | 8.59 | 8.7 | 8.81
24.39 | 4.28 | 4.39 | 4.5 | 4.62 | 4.73 | 4.84 | 4.96 | 5.07 | 5.18 | 5.3 | 5.41 | 5.52 | 5.64 | 5.75 | 5.86 | 5.98 | 6.09 | 6.2 | 6.32 | 6.43 | 6.54 | 6.66 | 6.77 | 6.88 | 7 | 7.11 | 7.22 | 7.34 | 7.45 | 7.56 | 7.68 | 7.79 | 7.9 | 8.02 | 8.13 | 8.24 | 8.36 | 8.47 | 8.58 | 8.7 | 8.81
24.4 | 4.28 | 4.39 | 4.5 | 4.62 | 4.73 | 4.84 | 4.96 | 5.07 | 5.18 | 5.3 | 5.41 | 5.52 | 5.64 | 5.75 | 5.86 | 5.98 | 6.09 | 6.2 | 6.32 | 6.43 | 6.54 | 6.66 | 6.77 | 6.88 | 7 | 7.11 | 7.22 | 7.34 | 7.45 | 7.56 | 7.68 | 7.79 | 7.9 | 8.02 | 8.13 | 8.24 | 8.36 | 8.47 | 8.58 | 8.7 | 8.81
24.41 | 4.28 | 4.39 | 4.5 | 4.62 | 4.73 | 4.84 | 4.96 | 5.07 | 5.18 | 5.3 | 5.41 | 5.52 | 5.64 | 5.75 | 5.86 | 5.98 | 6.09 | 6.2 | 6.31 | 6.43 | 6.54 | 6.65 | 6.77 | 6.88 | 6.99 | 7.11 | 7.22 | 7.33 | 7.45 | 7.56 | 7.67 | 7.79 | 7.9 | 8.01 | 8.13 | 8.24 | 8.35 | 8.47 | 8.58 | 8.69 | 8.81
24.42 | 4.27 | 4.39 | 4.5 | 4.61 | 4.73 | 4.84 | 4.95 | 5.07 | 5.18 | 5.29 | 5.41 | 5.52 | 5.63 | 5.75 | 5.86 | 5.97 | 6.09 | 6.2 | 6.31 | 6.43 | 6.54 | 6.65 | 6.77 | 6.88 | 6.99 | 7.11 | 7.22 | 7.33 | 7.45 | 7.56 | 7.67 | 7.79 | 7.9 | 8.01 | 8.13 | 8.24 | 8.35 | 8.47 | 8.58 | 8.69 | 8.81
24.43 | 4.27 | 4.39 | 4.5 | 4.61 | 4.73 | 4.84 | 4.95 | 5.07 | 5.18 | 5.29 | 5.41 | 5.52 | 5.63 | 5.75 | 5.86 | 5.97 | 6.09 | 6.2 | 6.31 | 6.43 | 6.54 | 6.65 | 6.77 | 6.88 | 6.99 | 7.11 | 7.22 | 7.33 | 7.45 | 7.56 | 7.67 | 7.78 | 7.9 | 8.01 | 8.12 | 8.24 | 8.35 | 8.46 | 8.58 | 8.69 | 8.8
24.44 | 4.27 | 4.39 | 4.5 | 4.61 | 4.73 | 4.84 | 4.95 | 5.07 | 5.18 | 5.29 | 5.41 | 5.52 | 5.63 | 5.75 | 5.86 | 5.97 | 6.08 | 6.2 | 6.31 | 6.42 | 6.54 | 6.65 | 6.76 | 6.88 | 6.99 | 7.1 | 7.22 | 7.33 | 7.44 | 7.56 | 7.67 | 7.78 | 7.9 | 8.01 | 8.12 | 8.24 | 8.35 | 8.46 | 8.58 | 8.69 | 8.8
24.45 | 4.27 | 4.38 | 4.5 | 4.61 | 4.72 | 4.84 | 4.95 | 5.06 | 5.18 | 5.29 | 5.4 | 5.52 | 5.63 | 5.74 | 5.86 | 5.97 | 6.08 | 6.2 | 6.31 | 6.42 | 6.54 | 6.65 | 6.76 | 6.88 | 6.99 | 7.1 | 7.22 | 7.33 | 7.44 | 7.56 | 7.67 | 7.78 | 7.9 | 8.01 | 8.12 | 8.23 | 8.35 | 8.46 | 8.57 | 8.69 | 8.8
24.46 | 4.27 | 4.38 | 4.5 | 4.61 | 4.72 | 4.84 | 4.95 | 5.06 | 5.18 | 5.29 | 5.4 | 5.52 | 5.63 | 5.74 | 5.86 | 5.97 | 6.08 | 6.2 | 6.31 | 6.42 | 6.54 | 6.65 | 6.76 | 6.87 | 6.99 | 7.1 | 7.21 | 7.33 | 7.44 | 7.55 | 7.67 | 7.78 | 7.89 | 8.01 | 8.12 | 8.23 | 8.35 | 8.46 | 8.57 | 8.69 | 8.8
24.47 | 4.27 | 4.38 | 4.5 | 4.61 | 4.72 | 4.84 | 4.95 | 5.06 | 5.18 | 5.29 | 5.4 | 5.52 | 5.63 | 5.74 | 5.85 | 5.97 | 6.08 | 6.19 | 6.31 | 6.42 | 6.53 | 6.65 | 6.76 | 6.87 | 6.99 | 7.1 | 7.21 | 7.33 | 7.44 | 7.55 | 7.67 | 7.78 | 7.89 | 8.01 | 8.12 | 8.23 | 8.34 | 8.46 | 8.57 | 8.68 | 8.8
24.48 | 4.27 | 4.38 | 4.5 | 4.61 | 4.72 | 4.84 | 4.95 | 5.06 | 5.17 | 5.29 | 5.4 | 5.51 | 5.63 | 5.74 | 5.85 | 5.97 | 6.08 | 6.19 | 6.31 | 6.42 | 6.53 | 6.65 | 6.76 | 6.87 | 6.99 | 7.1 | 7.21 | 7.32 | 7.44 | 7.55 | 7.66 | 7.78 | 7.89 | 8 | 8.12 | 8.23 | 8.34 | 8.46 | 8.57 | 8.68 | 8.8
24.49 | 4.27 | 4.38 | 4.49 | 4.61 | 4.72 | 4.83 | 4.95 | 5.06 | 5.17 | 5.29 | 5.4 | 5.51 | 5.63 | 5.74 | 5.85 | 5.97 | 6.08 | 6.19 | 6.31 | 6.42 | 6.53 | 6.64 | 6.76 | 6.87 | 6.98 | 7.1 | 7.21 | 7.32 | 7.44 | 7.55 | 7.66 | 7.78 | 7.89 | 8 | 8.12 | 8.23 | 8.34 | 8.45 | 8.57 | 8.68 | 8.79
24.5 | 4.27 | 4.38 | 4.49 | 4.61 | 4.72 | 4.83 | 4.95 | 5.06 | 5.17 | 5.29 | 5.4 | 5.51 | 5.63 | 5.74 | 5.85 | 5.96 | 6.08 | 6.19 | 6.3 | 6.42 | 6.53 | 6.64 | 6.76 | 6.87 | 6.98 | 7.1 | 7.21 | 7.32 | 7.44 | 7.55 | 7.66 | 7.77 | 7.89 | 8 | 8.11 | 8.23 | 8.34 | 8.45 | 8.57 | 8.68 | 8.79
24.51 | 4.27 | 4.38 | 4.49 | 4.61 | 4.72 | 4.83 | 4.95 | 5.06 | 5.17 | 5.28 | 5.4 | 5.51 | 5.62 | 5.74 | 5.85 | 5.96 | 6.08 | 6.19 | 6.3 | 6.42 | 6.53 | 6.64 | 6.76 | 6.87 | 6.98 | 7.09 | 7.21 | 7.32 | 7.43 | 7.55 | 7.66 | 7.77 | 7.89 | 8 | 8.11 | 8.23 | 8.34 | 8.45 | 8.56 | 8.68 | 8.79
24.52 | 4.27 | 4.38 | 4.49 | 4.6 | 4.72 | 4.83 | 4.94 | 5.06 | 5.17 | 5.28 | 5.4 | 5.51 | 5.62 | 5.74 | 5.85 | 5.96 | 6.08 | 6.19 | 6.3 | 6.41 | 6.53 | 6.64 | 6.75 | 6.87 | 6.98 | 7.09 | 7.21 | 7.32 | 7.43 | 7.55 | 7.66 | 7.77 | 7.88 | 8 | 8.11 | 8.22 | 8.34 | 8.45 | 8.56 | 8.68 | 8.79
24.53 | 4.26 | 4.38 | 4.49 | 4.6 | 4.72 | 4.83 | 4.94 | 5.06 | 5.17 | 5.28 | 5.4 | 5.51 | 5.62 | 5.74 | 5.85 | 5.96 | 6.07 | 6.19 | 6.3 | 6.41 | 6.53 | 6.64 | 6.75 | 6.87 | 6.98 | 7.09 | 7.21 | 7.32 | 7.43 | 7.54 | 7.66 | 7.77 | 7.88 | 8 | 8.11 | 8.22 | 8.34 | 8.45 | 8.56 | 8.67 | 8.79
24.54 | 4.26 | 4.38 | 4.49 | 4.6 | 4.72 | 4.83 | 4.94 | 5.06 | 5.17 | 5.28 | 5.39 | 5.51 | 5.62 | 5.73 | 5.85 | 5.96 | 6.07 | 6.19 | 6.3 | 6.41 | 6.53 | 6.64 | 6.75 | 6.86 | 6.98 | 7.09 | 7.2 | 7.32 | 7.43 | 7.54 | 7.66 | 7.77 | 7.88 | 7.99 | 8.11 | 8.22 | 8.33 | 8.45 | 8.56 | 8.67 | 8.79
24.55 | 4.26 | 4.38 | 4.49 | 4.6 | 4.72 | 4.83 | 4.94 | 5.05 | 5.17 | 5.28 | 5.39 | 5.51 | 5.62 | 5.73 | 5.85 | 5.96 | 6.07 | 6.18 | 6.3 | 6.41 | 6.52 | 6.64 | 6.75 | 6.86 | 6.98 | 7.09 | 7.2 | 7.32 | 7.43 | 7.54 | 7.65 | 7.77 | 7.88 | 7.99 | 8.11 | 8.22 | 8.33 | 8.45 | 8.56 | 8.67 | 8.78
24.56 | 4.26 | 4.38 | 4.49 | 4.6 | 4.71 | 4.83 | 4.94 | 5.05 | 5.17 | 5.28 | 5.39 | 5.51 | 5.62 | 5.73 | 5.84 | 5.96 | 6.07 | 6.18 | 6.3 | 6.41 | 6.52 | 6.64 | 6.75 | 6.86 | 6.97 | 7.09 | 7.2 | 7.31 | 7.43 | 7.54 | 7.65 | 7.77 | 7.88 | 7.99 | 8.1 | 8.22 | 8.33 | 8.44 | 8.56 | 8.67 | 8.78
24.57 | 4.26 | 4.37 | 4.49 | 4.6 | 4.71 | 4.83 | 4.94 | 5.05 | 5.17 | 5.28 | 5.39 | 5.5 | 5.62 | 5.73 | 5.84 | 5.96 | 6.07 | 6.18 | 6.3 | 6.41 | 6.52 | 6.63 | 6.75 | 6.86 | 6.97 | 7.09 | 7.2 | 7.31 | 7.43 | 7.54 | 7.65 | 7.76 | 7.88 | 7.99 | 8.1 | 8.22 | 8.33 | 8.44 | 8.56 | 8.67 | 8.78
24.58 | 4.26 | 4.37 | 4.49 | 4.6 | 4.71 | 4.83 | 4.94 | 5.05 | 5.16 | 5.28 | 5.39 | 5.5 | 5.62 | 5.73 | 5.84 | 5.96 | 6.07 | 6.18 | 6.29 | 6.41 | 6.52 | 6.63 | 6.75 | 6.86 | 6.97 | 7.09 | 7.2 | 7.31 | 7.42 | 7.54 | 7.65 | 7.76 | 7.88 | 7.99 | 8.1 | 8.21 | 8.33 | 8.44 | 8.55 | 8.67 | 8.78
24.59 | 4.26 | 4.37 | 4.49 | 4.6 | 4.71 | 4.82 | 4.94 | 5.05 | 5.16 | 5.28 | 5.39 | 5.5 | 5.62 | 5.73 | 5.84 | 5.95 | 6.07 | 6.18 | 6.29 | 6.41 | 6.52 | 6.63 | 6.75 | 6.86 | 6.97 | 7.08 | 7.2 | 7.31 | 7.42 | 7.54 | 7.65 | 7.76 | 7.87 | 7.99 | 8.1 | 8.21 | 8.33 | 8.44 | 8.55 | 8.67 | 8.78
24.6 | 4.26 | 4.37 | 4.48 | 4.6 | 4.71 | 4.82 | 4.94 | 5.05 | 5.16 | 5.28 | 5.39 | 5.5 | 5.61 | 5.73 | 5.84 | 5.95 | 6.07 | 6.18 | 6.29 | 6.4 | 6.52 | 6.63 | 6.74 | 6.86 | 6.97 | 7.08 | 7.2 | 7.31 | 7.42 | 7.53 | 7.65 | 7.76 | 7.87 | 7.99 | 8.1 | 8.21 | 8.32 | 8.44 | 8.55 | 8.66 | 8.78
24.61 | 4.26 | 4.37 | 4.48 | 4.6 | 4.71 | 4.82 | 4.94 | 5.05 | 5.16 | 5.27 | 5.39 | 5.5 | 5.61 | 5.73 | 5.84 | 5.95 | 6.06 | 6.18 | 6.29 | 6.4 | 6.52 | 6.63 | 6.74 | 6.86 | 6.97 | 7.08 | 7.19 | 7.31 | 7.42 | 7.53 | 7.65 | 7.76 | 7.87 | 7.98 | 8.1 | 8.21 | 8.32 | 8.44 | 8.55 | 8.66 | 8.77
24.62 | 4.26 | 4.37 | 4.48 | 4.6 | 4.71 | 4.82 | 4.93 | 5.05 | 5.16 | 5.27 | 5.39 | 5.5 | 5.61 | 5.73 | 5.84 | 5.95 | 6.06 | 6.18 | 6.29 | 6.4 | 6.52 | 6.63 | 6.74 | 6.85 | 6.97 | 7.08 | 7.19 | 7.31 | 7.42 | 7.53 | 7.64 | 7.76 | 7.87 | 7.98 | 8.1 | 8.21 | 8.32 | 8.43 | 8.55 | 8.66 | 8.77
24.63 | 4.26 | 4.37 | 4.48 | 4.59 | 4.71 | 4.82 | 4.93 | 5.05 | 5.16 | 5.27 | 5.39 | 5.5 | 5.61 | 5.72 | 5.84 | 5.95 | 6.06 | 6.18 | 6.29 | 6.4 | 6.51 | 6.63 | 6.74 | 6.85 | 6.97 | 7.08 | 7.19 | 7.3 | 7.42 | 7.53 | 7.64 | 7.76 | 7.87 | 7.98 | 8.09 | 8.21 | 8.32 | 8.43 | 8.55 | 8.66 | 8.77
24.64 | 4.26 | 4.37 | 4.48 | 4.59 | 4.71 | 4.82 | 4.93 | 5.05 | 5.16 | 5.27 | 5.38 | 5.5 | 5.61 | 5.72 | 5.84 | 5.95 | 6.06 | 6.17 | 6.29 | 6.4 | 6.51 | 6.63 | 6.74 | 6.85 | 6.96 | 7.08 | 7.19 | 7.3 | 7.42 | 7.53 | 7.64 | 7.75 | 7.87 | 7.98 | 8.09 | 8.21 | 8.32 | 8.43 | 8.54 | 8.66 | 8.77
24.65 | 4.25 | 4.37 | 4.48 | 4.59 | 4.71 | 4.82 | 4.93 | 5.04 | 5.16 | 5.27 | 5.38 | 5.5 | 5.61 | 5.72 | 5.83 | 5.95 | 6.06 | 6.17 | 6.29 | 6.4 | 6.51 | 6.62 | 6.74 | 6.85 | 6.96 | 7.08 | 7.19 | 7.3 | 7.41 | 7.53 | 7.64 | 7.75 | 7.87 | 7.98 | 8.09 | 8.2 | 8.32 | 8.43 | 8.54 | 8.66 | 8.77
24.66 | 4.25 | 4.37 | 4.48 | 4.59 | 4.7 | 4.82 | 4.93 | 5.04 | 5.16 | 5.27 | 5.38 | 5.49 | 5.61 | 5.72 | 5.83 | 5.95 | 6.06 | 6.17 | 6.28 | 6.4 | 6.51 | 6.62 | 6.74 | 6.85 | 6.96 | 7.07 | 7.19 | 7.3 | 7.41 | 7.53 | 7.64 | 7.75 | 7.86 | 7.98 | 8.09 | 8.2 | 8.32 | 8.43 | 8.54 | 8.65 | 8.77
24.67 | 4.25 | 4.37 | 4.48 | 4.59 | 4.7 | 4.82 | 4.93 | 5.04 | 5.16 | 5.27 | 5.38 | 5.49 | 5.61 | 5.72 | 5.83 | 5.95 | 6.06 | 6.17 | 6.28 | 6.4 | 6.51 | 6.62 | 6.73 | 6.85 | 6.96 | 7.07 | 7.19 | 7.3 | 7.41 | 7.52 | 7.64 | 7.75 | 7.86 | 7.98 | 8.09 | 8.2 | 8.31 | 8.43 | 8.54 | 8.65 | 8.77
24.68 | 4.25 | 4.36 | 4.48 | 4.59 | 4.7 | 4.82 | 4.93 | 5.04 | 5.15 | 5.27 | 5.38 | 5.49 | 5.61 | 5.72 | 5.83 | 5.94 | 6.06 | 6.17 | 6.28 | 6.4 | 6.51 | 6.62 | 6.73 | 6.85 | 6.96 | 7.07 | 7.18 | 7.3 | 7.41 | 7.52 | 7.64 | 7.75 | 7.86 | 7.97 | 8.09 | 8.2 | 8.31 | 8.43 | 8.54 | 8.65 | 8.76
24.69 | 4.25 | 4.36 | 4.48 | 4.59 | 4.7 | 4.81 | 4.93 | 5.04 | 5.15 | 5.27 | 5.38 | 5.49 | 5.6 | 5.72 | 5.83 | 5.94 | 6.06 | 6.17 | 6.28 | 6.39 | 6.51 | 6.62 | 6.73 | 6.85 | 6.96 | 7.07 | 7.18 | 7.3 | 7.41 | 7.52 | 7.63 | 7.75 | 7.86 | 7.97 | 8.09 | 8.2 | 8.31 | 8.42 | 8.54 | 8.65 | 8.76
24.7 | 4.25 | 4.36 | 4.48 | 4.59 | 4.7 | 4.81 | 4.93 | 5.04 | 5.15 | 5.27 | 5.38 | 5.49 | 5.6 | 5.72 | 5.83 | 5.94 | 6.05 | 6.17 | 6.28 | 6.39 | 6.51 | 6.62 | 6.73 | 6.84 | 6.96 | 7.07 | 7.18 | 7.29 | 7.41 | 7.52 | 7.63 | 7.75 | 7.86 | 7.97 | 8.08 | 8.2 | 8.31 | 8.42 | 8.53 | 8.65 | 8.76
24.71 | 4.25 | 4.36 | 4.47 | 4.59 | 4.7 | 4.81 | 4.93 | 5.04 | 5.15 | 5.26 | 5.38 | 5.49 | 5.6 | 5.72 | 5.83 | 5.94 | 6.05 | 6.17 | 6.28 | 6.39 | 6.5 | 6.62 | 6.73 | 6.84 | 6.96 | 7.07 | 7.18 | 7.29 | 7.41 | 7.52 | 7.63 | 7.74 | 7.86 | 7.97 | 8.08 | 8.2 | 8.31 | 8.42 | 8.53 | 8.65 | 8.76
24.72 | 4.25 | 4.36 | 4.47 | 4.59 | 4.7 | 4.81 | 4.92 | 5.04 | 5.15 | 5.26 | 5.38 | 5.49 | 5.6 | 5.71 | 5.83 | 5.94 | 6.05 | 6.16 | 6.28 | 6.39 | 6.5 | 6.62 | 6.73 | 6.84 | 6.95 | 7.07 | 7.18 | 7.29 | 7.4 | 7.52 | 7.63 | 7.74 | 7.86 | 7.97 | 8.08 | 8.19 | 8.31 | 8.42 | 8.53 | 8.64 | 8.76
24.73 | 4.25 | 4.36 | 4.47 | 4.59 | 4.7 | 4.81 | 4.92 | 5.04 | 5.15 | 5.26 | 5.37 | 5.49 | 5.6 | 5.71 | 5.83 | 5.94 | 6.05 | 6.16 | 6.28 | 6.39 | 6.5 | 6.61 | 6.73 | 6.84 | 6.95 | 7.07 | 7.18 | 7.29 | 7.4 | 7.52 | 7.63 | 7.74 | 7.85 | 7.97 | 8.08 | 8.19 | 8.3 | 8.42 | 8.53 | 8.64 | 8.76
24.74 | 4.25 | 4.36 | 4.47 | 4.58 | 4.7 | 4.81 | 4.92 | 5.04 | 5.15 | 5.26 | 5.37 | 5.49 | 5.6 | 5.71 | 5.82 | 5.94 | 6.05 | 6.16 | 6.28 | 6.39 | 6.5 | 6.61 | 6.73 | 6.84 | 6.95 | 7.06 | 7.18 | 7.29 | 7.4 | 7.51 | 7.63 | 7.74 | 7.85 | 7.97 | 8.08 | 8.19 | 8.3 | 8.42 | 8.53 | 8.64 | 8.75
24.75 | 4.25 | 4.36 | 4.47 | 4.58 | 4.7 | 4.81 | 4.92 | 5.03 | 5.15 | 5.26 | 5.37 | 5.49 | 5.6 | 5.71 | 5.82 | 5.94 | 6.05 | 6.16 | 6.27 | 6.39 | 6.5 | 6.61 | 6.72 | 6.84 | 6.95 | 7.06 | 7.18 | 7.29 | 7.4 | 7.51 | 7.63 | 7.74 | 7.85 | 7.96 | 8.08 | 8.19 | 8.3 | 8.41 | 8.53 | 8.64 | 8.75
24.76 | 4.24 | 4.36 | 4.47 | 4.58 | 4.7 | 4.81 | 4.92 | 5.03 | 5.15 | 5.26 | 5.37 | 5.48 | 5.6 | 5.71 | 5.82 | 5.93 | 6.05 | 6.16 | 6.27 | 6.39 | 6.5 | 6.61 | 6.72 | 6.84 | 6.95 | 7.06 | 7.17 | 7.29 | 7.4 | 7.51 | 7.62 | 7.74 | 7.85 | 7.96 | 8.07 | 8.19 | 8.3 | 8.41 | 8.53 | 8.64 | 8.75
24.77 | 4.24 | 4.36 | 4.47 | 4.58 | 4.69 | 4.81 | 4.92 | 5.03 | 5.15 | 5.26 | 5.37 | 5.48 | 5.6 | 5.71 | 5.82 | 5.93 | 6.05 | 6.16 | 6.27 | 6.38 | 6.5 | 6.61 | 6.72 | 6.83 | 6.95 | 7.06 | 7.17 | 7.29 | 7.4 | 7.51 | 7.62 | 7.74 | 7.85 | 7.96 | 8.07 | 8.19 | 8.3 | 8.41 | 8.52 | 8.64 | 8.75
24.78 | 4.24 | 4.36 | 4.47 | 4.58 | 4.69 | 4.81 | 4.92 | 5.03 | 5.14 | 5.26 | 5.37 | 5.48 | 5.59 | 5.71 | 5.82 | 5.93 | 6.05 | 6.16 | 6.27 | 6.38 | 6.5 | 6.61 | 6.72 | 6.83 | 6.95 | 7.06 | 7.17 | 7.28 | 7.4 | 7.51 | 7.62 | 7.73 | 7.85 | 7.96 | 8.07 | 8.18 | 8.3 | 8.41 | 8.52 | 8.63 | 8.75
24.79 | 4.24 | 4.35 | 4.47 | 4.58 | 4.69 | 4.81 | 4.92 | 5.03 | 5.14 | 5.26 | 5.37 | 5.48 | 5.59 | 5.71 | 5.82 | 5.93 | 6.04 | 6.16 | 6.27 | 6.38 | 6.49 | 6.61 | 6.72 | 6.83 | 6.94 | 7.06 | 7.17 | 7.28 | 7.4 | 7.51 | 7.62 | 7.73 | 7.85 | 7.96 | 8.07 | 8.18 | 8.3 | 8.41 | 8.52 | 8.63 | 8.75
24.8 | 4.24 | 4.35 | 4.47 | 4.58 | 4.69 | 4.8 | 4.92 | 5.03 | 5.14 | 5.25 | 5.37 | 5.48 | 5.59 | 5.71 | 5.82 | 5.93 | 6.04 | 6.16 | 6.27 | 6.38 | 6.49 | 6.61 | 6.72 | 6.83 | 6.94 | 7.06 | 7.17 | 7.28 | 7.39 | 7.51 | 7.62 | 7.73 | 7.84 | 7.96 | 8.07 | 8.18 | 8.29 | 8.41 | 8.52 | 8.63 | 8.74
24.81 | 4.24 | 4.35 | 4.47 | 4.58 | 4.69 | 4.8 | 4.92 | 5.03 | 5.14 | 5.25 | 5.37 | 5.48 | 5.59 | 5.7 | 5.82 | 5.93 | 6.04 | 6.15 | 6.27 | 6.38 | 6.49 | 6.6 | 6.72 | 6.83 | 6.94 | 7.05 | 7.17 | 7.28 | 7.39 | 7.5 | 7.62 | 7.73 | 7.84 | 7.95 | 8.07 | 8.18 | 8.29 | 8.4 | 8.52 | 8.63 | 8.74
24.82 | 4.24 | 4.35 | 4.46 | 4.58 | 4.69 | 4.8 | 4.92 | 5.03 | 5.14 | 5.25 | 5.37 | 5.48 | 5.59 | 5.7 | 5.82 | 5.93 | 6.04 | 6.15 | 6.27 | 6.38 | 6.49 | 6.6 | 6.72 | 6.83 | 6.94 | 7.05 | 7.17 | 7.28 | 7.39 | 7.5 | 7.62 | 7.73 | 7.84 | 7.95 | 8.07 | 8.18 | 8.29 | 8.4 | 8.52 | 8.63 | 8.74
24.83 | 4.24 | 4.35 | 4.46 | 4.58 | 4.69 | 4.8 | 4.91 | 5.03 | 5.14 | 5.25 | 5.36 | 5.48 | 5.59 | 5.7 | 5.81 | 5.93 | 6.04 | 6.15 | 6.26 | 6.38 | 6.49 | 6.6 | 6.71 | 6.83 | 6.94 | 7.05 | 7.16 | 7.28 | 7.39 | 7.5 | 7.61 | 7.73 | 7.84 | 7.95 | 8.06 | 8.18 | 8.29 | 8.4 | 8.51 | 8.63 | 8.74
24.84 | 4.24 | 4.35 | 4.46 | 4.58 | 4.69 | 4.8 | 4.91 | 5.03 | 5.14 | 5.25 | 5.36 | 5.48 | 5.59 | 5.7 | 5.81 | 5.93 | 6.04 | 6.15 | 6.26 | 6.38 | 6.49 | 6.6 | 6.71 | 6.83 | 6.94 | 7.05 | 7.16 | 7.28 | 7.39 | 7.5 | 7.61 | 7.73 | 7.84 | 7.95 | 8.06 | 8.18 | 8.29 | 8.4 | 8.51 | 8.63 | 8.74
24.85 | 4.24 | 4.35 | 4.46 | 4.57 | 4.69 | 4.8 | 4.91 | 5.02 | 5.14 | 5.25 | 5.36 | 5.47 | 5.59 | 5.7 | 5.81 | 5.92 | 6.04 | 6.15 | 6.26 | 6.37 | 6.49 | 6.6 | 6.71 | 6.82 | 6.94 | 7.05 | 7.16 | 7.27 | 7.39 | 7.5 | 7.61 | 7.72 | 7.84 | 7.95 | 8.06 | 8.17 | 8.29 | 8.4 | 8.51 | 8.62 | 8.74
24.86 | 4.24 | 4.35 | 4.46 | 4.57 | 4.69 | 4.8 | 4.91 | 5.02 | 5.14 | 5.25 | 5.36 | 5.47 | 5.59 | 5.7 | 5.81 | 5.92 | 6.04 | 6.15 | 6.26 | 6.37 | 6.49 | 6.6 | 6.71 | 6.82 | 6.94 | 7.05 | 7.16 | 7.27 | 7.39 | 7.5 | 7.61 | 7.72 | 7.84 | 7.95 | 8.06 | 8.17 | 8.28 | 8.4 | 8.51 | 8.62 | 8.73
24.87 | 4.24 | 4.35 | 4.46 | 4.57 | 4.69 | 4.8 | 4.91 | 5.02 | 5.14 | 5.25 | 5.36 | 5.47 | 5.58 | 5.7 | 5.81 | 5.92 | 6.03 | 6.15 | 6.26 | 6.37 | 6.48 | 6.6 | 6.71 | 6.82 | 6.93 | 7.05 | 7.16 | 7.27 | 7.38 | 7.5 | 7.61 | 7.72 | 7.83 | 7.95 | 8.06 | 8.17 | 8.28 | 8.4 | 8.51 | 8.62 | 8.73
24.88 | 4.23 | 4.35 | 4.46 | 4.57 | 4.68 | 4.8 | 4.91 | 5.02 | 5.13 | 5.25 | 5.36 | 5.47 | 5.58 | 5.7 | 5.81 | 5.92 | 6.03 | 6.15 | 6.26 | 6.37 | 6.48 | 6.6 | 6.71 | 6.82 | 6.93 | 7.05 | 7.16 | 7.27 | 7.38 | 7.49 | 7.61 | 7.72 | 7.83 | 7.94 | 8.06 | 8.17 | 8.28 | 8.39 | 8.51 | 8.62 | 8.73
24.89 | 4.23 | 4.35 | 4.46 | 4.57 | 4.68 | 4.8 | 4.91 | 5.02 | 5.13 | 5.25 | 5.36 | 5.47 | 5.58 | 5.7 | 5.81 | 5.92 | 6.03 | 6.14 | 6.26 | 6.37 | 6.48 | 6.59 | 6.71 | 6.82 | 6.93 | 7.04 | 7.16 | 7.27 | 7.38 | 7.49 | 7.61 | 7.72 | 7.83 | 7.94 | 8.06 | 8.17 | 8.28 | 8.39 | 8.5 | 8.62 | 8.73
24.9 | 4.23 | 4.35 | 4.46 | 4.57 | 4.68 | 4.79 | 4.91 | 5.02 | 5.13 | 5.24 | 5.36 | 5.47 | 5.58 | 5.69 | 5.81 | 5.92 | 6.03 | 6.14 | 6.26 | 6.37 | 6.48 | 6.59 | 6.71 | 6.82 | 6.93 | 7.04 | 7.16 | 7.27 | 7.38 | 7.49 | 7.6 | 7.72 | 7.83 | 7.94 | 8.05 | 8.17 | 8.28 | 8.39 | 8.5 | 8.62 | 8.73
24.91 | 4.23 | 4.34 | 4.46 | 4.57 | 4.68 | 4.79 | 4.91 | 5.02 | 5.13 | 5.24 | 5.36 | 5.47 | 5.58 | 5.69 | 5.81 | 5.92 | 6.03 | 6.14 | 6.25 | 6.37 | 6.48 | 6.59 | 6.7 | 6.82 | 6.93 | 7.04 | 7.15 | 7.27 | 7.38 | 7.49 | 7.6 | 7.72 | 7.83 | 7.94 | 8.05 | 8.16 | 8.28 | 8.39 | 8.5 | 8.61 | 8.73
24.92 | 4.23 | 4.34 | 4.46 | 4.57 | 4.68 | 4.79 | 4.91 | 5.02 | 5.13 | 5.24 | 5.35 | 5.47 | 5.58 | 5.69 | 5.8 | 5.92 | 6.03 | 6.14 | 6.25 | 6.37 | 6.48 | 6.59 | 6.7 | 6.82 | 6.93 | 7.04 | 7.15 | 7.26 | 7.38 | 7.49 | 7.6 | 7.71 | 7.83 | 7.94 | 8.05 | 8.16 | 8.28 | 8.39 | 8.5 | 8.61 | 8.72
24.93 | 4.23 | 4.34 | 4.45 | 4.57 | 4.68 | 4.79 | 4.9 | 5.02 | 5.13 | 5.24 | 5.35 | 5.47 | 5.58 | 5.69 | 5.8 | 5.92 | 6.03 | 6.14 | 6.25 | 6.36 | 6.48 | 6.59 | 6.7 | 6.81 | 6.93 | 7.04 | 7.15 | 7.26 | 7.38 | 7.49 | 7.6 | 7.71 | 7.82 | 7.94 | 8.05 | 8.16 | 8.27 | 8.39 | 8.5 | 8.61 | 8.72
24.94 | 4.23 | 4.34 | 4.45 | 4.57 | 4.68 | 4.79 | 4.9 | 5.02 | 5.13 | 5.24 | 5.35 | 5.47 | 5.58 | 5.69 | 5.8 | 5.91 | 6.03 | 6.14 | 6.25 | 6.36 | 6.48 | 6.59 | 6.7 | 6.81 | 6.93 | 7.04 | 7.15 | 7.26 | 7.37 | 7.49 | 7.6 | 7.71 | 7.82 | 7.94 | 8.05 | 8.16 | 8.27 | 8.38 | 8.5 | 8.61 | 8.72
24.95 | 4.23 | 4.34 | 4.45 | 4.57 | 4.68 | 4.79 | 4.9 | 5.01 | 5.13 | 5.24 | 5.35 | 5.46 | 5.58 | 5.69 | 5.8 | 5.91 | 6.03 | 6.14 | 6.25 | 6.36 | 6.47 | 6.59 | 6.7 | 6.81 | 6.92 | 7.04 | 7.15 | 7.26 | 7.37 | 7.49 | 7.6 | 7.71 | 7.82 | 7.93 | 8.05 | 8.16 | 8.27 | 8.38 | 8.5 | 8.61 | 8.72
24.96 | 4.23 | 4.34 | 4.45 | 4.56 | 4.68 | 4.79 | 4.9 | 5.01 | 5.13 | 5.24 | 5.35 | 5.46 | 5.58 | 5.69 | 5.8 | 5.91 | 6.02 | 6.14 | 6.25 | 6.36 | 6.47 | 6.59 | 6.7 | 6.81 | 6.92 | 7.03 | 7.15 | 7.26 | 7.37 | 7.48 | 7.6 | 7.71 | 7.82 | 7.93 | 8.05 | 8.16 | 8.27 | 8.38 | 8.49 | 8.61 | 8.72
24.97 | 4.23 | 4.34 | 4.45 | 4.56 | 4.68 | 4.79 | 4.9 | 5.01 | 5.13 | 5.24 | 5.35 | 5.46 | 5.57 | 5.69 | 5.8 | 5.91 | 6.02 | 6.14 | 6.25 | 6.36 | 6.47 | 6.58 | 6.7 | 6.81 | 6.92 | 7.03 | 7.15 | 7.26 | 7.37 | 7.48 | 7.59 | 7.71 | 7.82 | 7.93 | 8.04 | 8.16 | 8.27 | 8.38 | 8.49 | 8.6 | 8.72
24.98 | 4.23 | 4.34 | 4.45 | 4.56 | 4.67 | 4.79 | 4.9 | 5.01 | 5.12 | 5.24 | 5.35 | 5.46 | 5.57 | 5.69 | 5.8 | 5.91 | 6.02 | 6.13 | 6.25 | 6.36 | 6.47 | 6.58 | 6.7 | 6.81 | 6.92 | 7.03 | 7.14 | 7.26 | 7.37 | 7.48 | 7.59 | 7.71 | 7.82 | 7.93 | 8.04 | 8.15 | 8.27 | 8.38 | 8.49 | 8.6 | 8.72
24.99 | 4.23 | 4.34 | 4.45 | 4.56 | 4.67 | 4.79 | 4.9 | 5.01 | 5.12 | 5.24 | 5.35 | 5.46 | 5.57 | 5.68 | 5.8 | 5.91 | 6.02 | 6.13 | 6.25 | 6.36 | 6.47 | 6.58 | 6.69 | 6.81 | 6.92 | 7.03 | 7.14 | 7.26 | 7.37 | 7.48 | 7.59 | 7.7 | 7.82 | 7.93 | 8.04 | 8.15 | 8.26 | 8.38 | 8.49 | 8.6 | 8.71
25 | 4.22 | 4.34 | 4.45 | 4.56 | 4.67 | 4.79 | 4.9 | 5.01 | 5.12 | 5.23 | 5.35 | 5.46 | 5.57 | 5.68 | 5.8 | 5.91 | 6.02 | 6.13 | 6.24 | 6.36 | 6.47 | 6.58 | 6.69 | 6.81 | 6.92 | 7.03 | 7.14 | 7.25 | 7.37 | 7.48 | 7.59 | 7.7 | 7.81 | 7.93 | 8.04 | 8.15 | 8.26 | 8.38 | 8.49 | 8.6 | 8.71
25.01 | 4.22 | 4.34 | 4.45 | 4.56 | 4.67 | 4.78 | 4.9 | 5.01 | 5.12 | 5.23 | 5.35 | 5.46 | 5.57 | 5.68 | 5.79 | 5.91 | 6.02 | 6.13 | 6.24 | 6.36 | 6.47 | 6.58 | 6.69 | 6.8 | 6.92 | 7.03 | 7.14 | 7.25 | 7.36 | 7.48 | 7.59 | 7.7 | 7.81 | 7.93 | 8.04 | 8.15 | 8.26 | 8.37 | 8.49 | 8.6 | 8.71
25.02 | 4.22 | 4.33 | 4.45 | 4.56 | 4.67 | 4.78 | 4.9 | 5.01 | 5.12 | 5.23 | 5.34 | 5.46 | 5.57 | 5.68 | 5.79 | 5.91 | 6.02 | 6.13 | 6.24 | 6.35 | 6.47 | 6.58 | 6.69 | 6.8 | 6.91 | 7.03 | 7.14 | 7.25 | 7.36 | 7.48 | 7.59 | 7.7 | 7.81 | 7.92 | 8.04 | 8.15 | 8.26 | 8.37 | 8.48 | 8.6 | 8.71
25.03 | 4.22 | 4.33 | 4.45 | 4.56 | 4.67 | 4.78 | 4.89 | 5.01 | 5.12 | 5.23 | 5.34 | 5.46 | 5.57 | 5.68 | 5.79 | 5.9 | 6.02 | 6.13 | 6.24 | 6.35 | 6.46 | 6.58 | 6.69 | 6.8 | 6.91 | 7.03 | 7.14 | 7.25 | 7.36 | 7.47 | 7.59 | 7.7 | 7.81 | 7.92 | 8.03 | 8.15 | 8.26 | 8.37 | 8.48 | 8.6 | 8.71
25.04 | 4.22 | 4.33 | 4.45 | 4.56 | 4.67 | 4.78 | 4.89 | 5.01 | 5.12 | 5.23 | 5.34 | 5.45 | 5.57 | 5.68 | 5.79 | 5.9 | 6.02 | 6.13 | 6.24 | 6.35 | 6.46 | 6.58 | 6.69 | 6.8 | 6.91 | 7.02 | 7.14 | 7.25 | 7.36 | 7.47 | 7.58 | 7.7 | 7.81 | 7.92 | 8.03 | 8.15 | 8.26 | 8.37 | 8.48 | 8.59 | 8.71
25.05 | 4.22 | 4.33 | 4.44 | 4.56 | 4.67 | 4.78 | 4.89 | 5 | 5.12 | 5.23 | 5.34 | 5.45 | 5.57 | 5.68 | 5.79 | 5.9 | 6.01 | 6.13 | 6.24 | 6.35 | 6.46 | 6.57 | 6.69 | 6.8 | 6.91 | 7.02 | 7.14 | 7.25 | 7.36 | 7.47 | 7.58 | 7.7 | 7.81 | 7.92 | 8.03 | 8.14 | 8.26 | 8.37 | 8.48 | 8.59 | 8.7
25.06 | 4.22 | 4.33 | 4.44 | 4.56 | 4.67 | 4.78 | 4.89 | 5 | 5.12 | 5.23 | 5.34 | 5.45 | 5.56 | 5.68 | 5.79 | 5.9 | 6.01 | 6.12 | 6.24 | 6.35 | 6.46 | 6.57 | 6.69 | 6.8 | 6.91 | 7.02 | 7.13 | 7.25 | 7.36 | 7.47 | 7.58 | 7.69 | 7.81 | 7.92 | 8.03 | 8.14 | 8.25 | 8.37 | 8.48 | 8.59 | 8.7
25.07 | 4.22 | 4.33 | 4.44 | 4.55 | 4.67 | 4.78 | 4.89 | 5 | 5.11 | 5.23 | 5.34 | 5.45 | 5.56 | 5.68 | 5.79 | 5.9 | 6.01 | 6.12 | 6.24 | 6.35 | 6.46 | 6.57 | 6.68 | 6.8 | 6.91 | 7.02 | 7.13 | 7.24 | 7.36 | 7.47 | 7.58 | 7.69 | 7.8 | 7.92 | 8.03 | 8.14 | 8.25 | 8.36 | 8.48 | 8.59 | 8.7
25.08 | 4.22 | 4.33 | 4.44 | 4.55 | 4.67 | 4.78 | 4.89 | 5 | 5.11 | 5.23 | 5.34 | 5.45 | 5.56 | 5.67 | 5.79 | 5.9 | 6.01 | 6.12 | 6.23 | 6.35 | 6.46 | 6.57 | 6.68 | 6.79 | 6.91 | 7.02 | 7.13 | 7.24 | 7.36 | 7.47 | 7.58 | 7.69 | 7.8 | 7.92 | 8.03 | 8.14 | 8.25 | 8.36 | 8.48 | 8.59 | 8.7
25.09 | 4.22 | 4.33 | 4.44 | 4.55 | 4.66 | 4.78 | 4.89 | 5 | 5.11 | 5.23 | 5.34 | 5.45 | 5.56 | 5.67 | 5.79 | 5.9 | 6.01 | 6.12 | 6.23 | 6.35 | 6.46 | 6.57 | 6.68 | 6.79 | 6.91 | 7.02 | 7.13 | 7.24 | 7.35 | 7.47 | 7.58 | 7.69 | 7.8 | 7.91 | 8.03 | 8.14 | 8.25 | 8.36 | 8.47 | 8.59 | 8.7
25.1 | 4.22 | 4.33 | 4.44 | 4.55 | 4.66 | 4.78 | 4.89 | 5 | 5.11 | 5.22 | 5.34 | 5.45 | 5.56 | 5.67 | 5.78 | 5.9 | 6.01 | 6.12 | 6.23 | 6.34 | 6.46 | 6.57 | 6.68 | 6.79 | 6.9 | 7.02 | 7.13 | 7.24 | 7.35 | 7.46 | 7.58 | 7.69 | 7.8 | 7.91 | 8.02 | 8.14 | 8.25 | 8.36 | 8.47 | 8.58 | 8.7
25.11 | 4.21 | 4.33 | 4.44 | 4.55 | 4.66 | 4.77 | 4.89 | 5 | 5.11 | 5.22 | 5.34 | 5.45 | 5.56 | 5.67 | 5.78 | 5.9 | 6.01 | 6.12 | 6.23 | 6.34 | 6.46 | 6.57 | 6.68 | 6.79 | 6.9 | 7.02 | 7.13 | 7.24 | 7.35 | 7.46 | 7.57 | 7.69 | 7.8 | 7.91 | 8.02 | 8.13 | 8.25 | 8.36 | 8.47 | 8.58 | 8.69
25.12 | 4.21 | 4.33 | 4.44 | 4.55 | 4.66 | 4.77 | 4.89 | 5 | 5.11 | 5.22 | 5.33 | 5.45 | 5.56 | 5.67 | 5.78 | 5.89 | 6.01 | 6.12 | 6.23 | 6.34 | 6.45 | 6.57 | 6.68 | 6.79 | 6.9 | 7.01 | 7.13 | 7.24 | 7.35 | 7.46 | 7.57 | 7.69 | 7.8 | 7.91 | 8.02 | 8.13 | 8.25 | 8.36 | 8.47 | 8.58 | 8.69
25.13 | 4.21 | 4.33 | 4.44 | 4.55 | 4.66 | 4.77 | 4.89 | 5 | 5.11 | 5.22 | 5.33 | 5.44 | 5.56 | 5.67 | 5.78 | 5.89 | 6 | 6.12 | 6.23 | 6.34 | 6.45 | 6.56 | 6.68 | 6.79 | 6.9 | 7.01 | 7.12 | 7.24 | 7.35 | 7.46 | 7.57 | 7.68 | 7.8 | 7.91 | 8.02 | 8.13 | 8.24 | 8.36 | 8.47 | 8.58 | 8.69
25.14 | 4.21 | 4.32 | 4.44 | 4.55 | 4.66 | 4.77 | 4.88 | 5 | 5.11 | 5.22 | 5.33 | 5.44 | 5.56 | 5.67 | 5.78 | 5.89 | 6 | 6.12 | 6.23 | 6.34 | 6.45 | 6.56 | 6.68 | 6.79 | 6.9 | 7.01 | 7.12 | 7.23 | 7.35 | 7.46 | 7.57 | 7.68 | 7.79 | 7.91 | 8.02 | 8.13 | 8.24 | 8.35 | 8.47 | 8.58 | 8.69
25.15 | 4.21 | 4.32 | 4.44 | 4.55 | 4.66 | 4.77 | 4.88 | 5 | 5.11 | 5.22 | 5.33 | 5.44 | 5.55 | 5.67 | 5.78 | 5.89 | 6 | 6.11 | 6.23 | 6.34 | 6.45 | 6.56 | 6.67 | 6.79 | 6.9 | 7.01 | 7.12 | 7.23 | 7.35 | 7.46 | 7.57 | 7.68 | 7.79 | 7.9 | 8.02 | 8.13 | 8.24 | 8.35 | 8.46 | 8.58 | 8.69
25.16 | 4.21 | 4.32 | 4.43 | 4.55 | 4.66 | 4.77 | 4.88 | 4.99 | 5.11 | 5.22 | 5.33 | 5.44 | 5.55 | 5.67 | 5.78 | 5.89 | 6 | 6.11 | 6.23 | 6.34 | 6.45 | 6.56 | 6.67 | 6.78 | 6.9 | 7.01 | 7.12 | 7.23 | 7.34 | 7.46 | 7.57 | 7.68 | 7.79 | 7.9 | 8.02 | 8.13 | 8.24 | 8.35 | 8.46 | 8.57 | 8.69
25.17 | 4.21 | 4.32 | 4.43 | 4.55 | 4.66 | 4.77 | 4.88 | 4.99 | 5.1 | 5.22 | 5.33 | 5.44 | 5.55 | 5.66 | 5.78 | 5.89 | 6 | 6.11 | 6.22 | 6.34 | 6.45 | 6.56 | 6.67 | 6.78 | 6.9 | 7.01 | 7.12 | 7.23 | 7.34 | 7.45 | 7.57 | 7.68 | 7.79 | 7.9 | 8.01 | 8.13 | 8.24 | 8.35 | 8.46 | 8.57 | 8.69
25.18 | 4.21 | 4.32 | 4.43 | 4.54 | 4.66 | 4.77 | 4.88 | 4.99 | 5.1 | 5.22 | 5.33 | 5.44 | 5.55 | 5.66 | 5.78 | 5.89 | 6 | 6.11 | 6.22 | 6.33 | 6.45 | 6.56 | 6.67 | 6.78 | 6.89 | 7.01 | 7.12 | 7.23 | 7.34 | 7.45 | 7.57 | 7.68 | 7.79 | 7.9 | 8.01 | 8.12 | 8.24 | 8.35 | 8.46 | 8.57 | 8.68
25.19 | 4.21 | 4.32 | 4.43 | 4.54 | 4.66 | 4.77 | 4.88 | 4.99 | 5.1 | 5.21 | 5.33 | 5.44 | 5.55 | 5.66 | 5.77 | 5.89 | 6 | 6.11 | 6.22 | 6.33 | 6.45 | 6.56 | 6.67 | 6.78 | 6.89 | 7 | 7.12 | 7.23 | 7.34 | 7.45 | 7.56 | 7.68 | 7.79 | 7.9 | 8.01 | 8.12 | 8.23 | 8.35 | 8.46 | 8.57 | 8.68
25.2 | 4.21 | 4.32 | 4.43 | 4.54 | 4.65 | 4.77 | 4.88 | 4.99 | 5.1 | 5.21 | 5.33 | 5.44 | 5.55 | 5.66 | 5.77 | 5.88 | 6 | 6.11 | 6.22 | 6.33 | 6.44 | 6.56 | 6.67 | 6.78 | 6.89 | 7 | 7.12 | 7.23 | 7.34 | 7.45 | 7.56 | 7.67 | 7.79 | 7.9 | 8.01 | 8.12 | 8.23 | 8.34 | 8.46 | 8.57 | 8.68
25.21 | 4.21 | 4.32 | 4.43 | 4.54 | 4.65 | 4.77 | 4.88 | 4.99 | 5.1 | 5.21 | 5.32 | 5.44 | 5.55 | 5.66 | 5.77 | 5.88 | 6 | 6.11 | 6.22 | 6.33 | 6.44 | 6.55 | 6.67 | 6.78 | 6.89 | 7 | 7.11 | 7.23 | 7.34 | 7.45 | 7.56 | 7.67 | 7.78 | 7.9 | 8.01 | 8.12 | 8.23 | 8.34 | 8.46 | 8.57 | 8.68
25.22 | 4.21 | 4.32 | 4.43 | 4.54 | 4.65 | 4.76 | 4.88 | 4.99 | 5.1 | 5.21 | 5.32 | 5.44 | 5.55 | 5.66 | 5.77 | 5.88 | 5.99 | 6.11 | 6.22 | 6.33 | 6.44 | 6.55 | 6.67 | 6.78 | 6.89 | 7 | 7.11 | 7.22 | 7.34 | 7.45 | 7.56 | 7.67 | 7.78 | 7.89 | 8.01 | 8.12 | 8.23 | 8.34 | 8.45 | 8.57 | 8.68
25.23 | 4.2 | 4.32 | 4.43 | 4.54 | 4.65 | 4.76 | 4.88 | 4.99 | 5.1 | 5.21 | 5.32 | 5.43 | 5.55 | 5.66 | 5.77 | 5.88 | 5.99 | 6.11 | 6.22 | 6.33 | 6.44 | 6.55 | 6.66 | 6.78 | 6.89 | 7 | 7.11 | 7.22 | 7.33 | 7.45 | 7.56 | 7.67 | 7.78 | 7.89 | 8.01 | 8.12 | 8.23 | 8.34 | 8.45 | 8.56 | 8.68
25.24 | 4.2 | 4.32 | 4.43 | 4.54 | 4.65 | 4.76 | 4.87 | 4.99 | 5.1 | 5.21 | 5.32 | 5.43 | 5.55 | 5.66 | 5.77 | 5.88 | 5.99 | 6.1 | 6.22 | 6.33 | 6.44 | 6.55 | 6.66 | 6.77 | 6.89 | 7 | 7.11 | 7.22 | 7.33 | 7.44 | 7.56 | 7.67 | 7.78 | 7.89 | 8 | 8.12 | 8.23 | 8.34 | 8.45 | 8.56 | 8.67
25.25 | 4.2 | 4.31 | 4.43 | 4.54 | 4.65 | 4.76 | 4.87 | 4.99 | 5.1 | 5.21 | 5.32 | 5.43 | 5.54 | 5.66 | 5.77 | 5.88 | 5.99 | 6.1 | 6.21 | 6.33 | 6.44 | 6.55 | 6.66 | 6.77 | 6.88 | 7 | 7.11 | 7.22 | 7.33 | 7.44 | 7.56 | 7.67 | 7.78 | 7.89 | 8 | 8.11 | 8.23 | 8.34 | 8.45 | 8.56 | 8.67
25.26 | 4.2 | 4.31 | 4.43 | 4.54 | 4.65 | 4.76 | 4.87 | 4.98 | 5.1 | 5.21 | 5.32 | 5.43 | 5.54 | 5.65 | 5.77 | 5.88 | 5.99 | 6.1 | 6.21 | 6.33 | 6.44 | 6.55 | 6.66 | 6.77 | 6.88 | 7 | 7.11 | 7.22 | 7.33 | 7.44 | 7.55 | 7.67 | 7.78 | 7.89 | 8 | 8.11 | 8.22 | 8.34 | 8.45 | 8.56 | 8.67
25.27 | 4.2 | 4.31 | 4.42 | 4.54 | 4.65 | 4.76 | 4.87 | 4.98 | 5.1 | 5.21 | 5.32 | 5.43 | 5.54 | 5.65 | 5.77 | 5.88 | 5.99 | 6.1 | 6.21 | 6.32 | 6.44 | 6.55 | 6.66 | 6.77 | 6.88 | 6.99 | 7.11 | 7.22 | 7.33 | 7.44 | 7.55 | 7.66 | 7.78 | 7.89 | 8 | 8.11 | 8.22 | 8.33 | 8.45 | 8.56 | 8.67
25.28 | 4.2 | 4.31 | 4.42 | 4.54 | 4.65 | 4.76 | 4.87 | 4.98 | 5.09 | 5.21 | 5.32 | 5.43 | 5.54 | 5.65 | 5.76 | 5.88 | 5.99 | 6.1 | 6.21 | 6.32 | 6.43 | 6.55 | 6.66 | 6.77 | 6.88 | 6.99 | 7.1 | 7.22 | 7.33 | 7.44 | 7.55 | 7.66 | 7.77 | 7.89 | 8 | 8.11 | 8.22 | 8.33 | 8.44 | 8.56 | 8.67
25.29 | 4.2 | 4.31 | 4.42 | 4.53 | 4.65 | 4.76 | 4.87 | 4.98 | 5.09 | 5.2 | 5.32 | 5.43 | 5.54 | 5.65 | 5.76 | 5.87 | 5.99 | 6.1 | 6.21 | 6.32 | 6.43 | 6.54 | 6.66 | 6.77 | 6.88 | 6.99 | 7.1 | 7.21 | 7.33 | 7.44 | 7.55 | 7.66 | 7.77 | 7.88 | 8 | 8.11 | 8.22 | 8.33 | 8.44 | 8.55 | 8.67
25.3 | 4.2 | 4.31 | 4.42 | 4.53 | 4.65 | 4.76 | 4.87 | 4.98 | 5.09 | 5.2 | 5.32 | 5.43 | 5.54 | 5.65 | 5.76 | 5.87 | 5.99 | 6.1 | 6.21 | 6.32 | 6.43 | 6.54 | 6.66 | 6.77 | 6.88 | 6.99 | 7.1 | 7.21 | 7.33 | 7.44 | 7.55 | 7.66 | 7.77 | 7.88 | 7.99 | 8.11 | 8.22 | 8.33 | 8.44 | 8.55 | 8.66
25.31 | 4.2 | 4.31 | 4.42 | 4.53 | 4.64 | 4.76 | 4.87 | 4.98 | 5.09 | 5.2 | 5.31 | 5.43 | 5.54 | 5.65 | 5.76 | 5.87 | 5.98 | 6.1 | 6.21 | 6.32 | 6.43 | 6.54 | 6.65 | 6.77 | 6.88 | 6.99 | 7.1 | 7.21 | 7.32 | 7.44 | 7.55 | 7.66 | 7.77 | 7.88 | 7.99 | 8.1 | 8.22 | 8.33 | 8.44 | 8.55 | 8.66
25.32 | 4.2 | 4.31 | 4.42 | 4.53 | 4.64 | 4.76 | 4.87 | 4.98 | 5.09 | 5.2 | 5.31 | 5.42 | 5.54 | 5.65 | 5.76 | 5.87 | 5.98 | 6.09 | 6.21 | 6.32 | 6.43 | 6.54 | 6.65 | 6.76 | 6.88 | 6.99 | 7.1 | 7.21 | 7.32 | 7.43 | 7.55 | 7.66 | 7.77 | 7.88 | 7.99 | 8.1 | 8.21 | 8.33 | 8.44 | 8.55 | 8.66
25.33 | 4.2 | 4.31 | 4.42 | 4.53 | 4.64 | 4.75 | 4.87 | 4.98 | 5.09 | 5.2 | 5.31 | 5.42 | 5.54 | 5.65 | 5.76 | 5.87 | 5.98 | 6.09 | 6.21 | 6.32 | 6.43 | 6.54 | 6.65 | 6.76 | 6.87 | 6.99 | 7.1 | 7.21 | 7.32 | 7.43 | 7.54 | 7.66 | 7.77 | 7.88 | 7.99 | 8.1 | 8.21 | 8.33 | 8.44 | 8.55 | 8.66
25.34 | 4.2 | 4.31 | 4.42 | 4.53 | 4.64 | 4.75 | 4.86 | 4.98 | 5.09 | 5.2 | 5.31 | 5.42 | 5.53 | 5.65 | 5.76 | 5.87 | 5.98 | 6.09 | 6.2 | 6.32 | 6.43 | 6.54 | 6.65 | 6.76 | 6.87 | 6.98 | 7.1 | 7.21 | 7.32 | 7.43 | 7.54 | 7.65 | 7.77 | 7.88 | 7.99 | 8.1 | 8.21 | 8.32 | 8.44 | 8.55 | 8.66
25.35 | 4.19 | 4.31 | 4.42 | 4.53 | 4.64 | 4.75 | 4.86 | 4.98 | 5.09 | 5.2 | 5.31 | 5.42 | 5.53 | 5.64 | 5.76 | 5.87 | 5.98 | 6.09 | 6.2 | 6.31 | 6.43 | 6.54 | 6.65 | 6.76 | 6.87 | 6.98 | 7.1 | 7.21 | 7.32 | 7.43 | 7.54 | 7.65 | 7.76 | 7.88 | 7.99 | 8.1 | 8.21 | 8.32 | 8.43 | 8.55 | 8.66
25.36 | 4.19 | 4.31 | 4.42 | 4.53 | 4.64 | 4.75 | 4.86 | 4.97 | 5.09 | 5.2 | 5.31 | 5.42 | 5.53 | 5.64 | 5.76 | 5.87 | 5.98 | 6.09 | 6.2 | 6.31 | 6.42 | 6.54 | 6.65 | 6.76 | 6.87 | 6.98 | 7.09 | 7.21 | 7.32 | 7.43 | 7.54 | 7.65 | 7.76 | 7.87 | 7.99 | 8.1 | 8.21 | 8.32 | 8.43 | 8.54 | 8.65
25.37 | 4.19 | 4.3 | 4.42 | 4.53 | 4.64 | 4.75 | 4.86 | 4.97 | 5.09 | 5.2 | 5.31 | 5.42 | 5.53 | 5.64 | 5.75 | 5.87 | 5.98 | 6.09 | 6.2 | 6.31 | 6.42 | 6.53 | 6.65 | 6.76 | 6.87 | 6.98 | 7.09 | 7.2 | 7.32 | 7.43 | 7.54 | 7.65 | 7.76 | 7.87 | 7.98 | 8.1 | 8.21 | 8.32 | 8.43 | 8.54 | 8.65
25.38 | 4.19 | 4.3 | 4.41 | 4.53 | 4.64 | 4.75 | 4.86 | 4.97 | 5.08 | 5.2 | 5.31 | 5.42 | 5.53 | 5.64 | 5.75 | 5.86 | 5.98 | 6.09 | 6.2 | 6.31 | 6.42 | 6.53 | 6.65 | 6.76 | 6.87 | 6.98 | 7.09 | 7.2 | 7.31 | 7.43 | 7.54 | 7.65 | 7.76 | 7.87 | 7.98 | 8.09 | 8.21 | 8.32 | 8.43 | 8.54 | 8.65
25.39 | 4.19 | 4.3 | 4.41 | 4.53 | 4.64 | 4.75 | 4.86 | 4.97 | 5.08 | 5.19 | 5.31 | 5.42 | 5.53 | 5.64 | 5.75 | 5.86 | 5.98 | 6.09 | 6.2 | 6.31 | 6.42 | 6.53 | 6.64 | 6.76 | 6.87 | 6.98 | 7.09 | 7.2 | 7.31 | 7.42 | 7.54 | 7.65 | 7.76 | 7.87 | 7.98 | 8.09 | 8.2 | 8.32 | 8.43 | 8.54 | 8.65
25.4 | 4.19 | 4.3 | 4.41 | 4.52 | 4.64 | 4.75 | 4.86 | 4.97 | 5.08 | 5.19 | 5.31 | 5.42 | 5.53 | 5.64 | 5.75 | 5.86 | 5.97 | 6.09 | 6.2 | 6.31 | 6.42 | 6.53 | 6.64 | 6.75 | 6.87 | 6.98 | 7.09 | 7.2 | 7.31 | 7.42 | 7.53 | 7.65 | 7.76 | 7.87 | 7.98 | 8.09 | 8.2 | 8.31 | 8.43 | 8.54 | 8.65
25.41 | 4.19 | 4.3 | 4.41 | 4.52 | 4.64 | 4.75 | 4.86 | 4.97 | 5.08 | 5.19 | 5.3 | 5.42 | 5.53 | 5.64 | 5.75 | 5.86 | 5.97 | 6.08 | 6.2 | 6.31 | 6.42 | 6.53 | 6.64 | 6.75 | 6.86 | 6.98 | 7.09 | 7.2 | 7.31 | 7.42 | 7.53 | 7.64 | 7.76 | 7.87 | 7.98 | 8.09 | 8.2 | 8.31 | 8.42 | 8.54 | 8.65
25.42 | 4.19 | 4.3 | 4.41 | 4.52 | 4.63 | 4.75 | 4.86 | 4.97 | 5.08 | 5.19 | 5.3 | 5.41 | 5.53 | 5.64 | 5.75 | 5.86 | 5.97 | 6.08 | 6.19 | 6.31 | 6.42 | 6.53 | 6.64 | 6.75 | 6.86 | 6.97 | 7.09 | 7.2 | 7.31 | 7.42 | 7.53 | 7.64 | 7.75 | 7.87 | 7.98 | 8.09 | 8.2 | 8.31 | 8.42 | 8.53 | 8.65
25.43 | 4.19 | 4.3 | 4.41 | 4.52 | 4.63 | 4.74 | 4.86 | 4.97 | 5.08 | 5.19 | 5.3 | 5.41 | 5.52 | 5.64 | 5.75 | 5.86 | 5.97 | 6.08 | 6.19 | 6.3 | 6.42 | 6.53 | 6.64 | 6.75 | 6.86 | 6.97 | 7.08 | 7.2 | 7.31 | 7.42 | 7.53 | 7.64 | 7.75 | 7.86 | 7.98 | 8.09 | 8.2 | 8.31 | 8.42 | 8.53 | 8.64
25.44 | 4.19 | 4.3 | 4.41 | 4.52 | 4.63 | 4.74 | 4.86 | 4.97 | 5.08 | 5.19 | 5.3 | 5.41 | 5.52 | 5.64 | 5.75 | 5.86 | 5.97 | 6.08 | 6.19 | 6.3 | 6.41 | 6.53 | 6.64 | 6.75 | 6.86 | 6.97 | 7.08 | 7.19 | 7.31 | 7.42 | 7.53 | 7.64 | 7.75 | 7.86 | 7.97 | 8.09 | 8.2 | 8.31 | 8.42 | 8.53 | 8.64
25.45 | 4.19 | 4.3 | 4.41 | 4.52 | 4.63 | 4.74 | 4.85 | 4.97 | 5.08 | 5.19 | 5.3 | 5.41 | 5.52 | 5.63 | 5.75 | 5.86 | 5.97 | 6.08 | 6.19 | 6.3 | 6.41 | 6.53 | 6.64 | 6.75 | 6.86 | 6.97 | 7.08 | 7.19 | 7.3 | 7.42 | 7.53 | 7.64 | 7.75 | 7.86 | 7.97 | 8.08 | 8.2 | 8.31 | 8.42 | 8.53 | 8.64
25.46 | 4.19 | 4.3 | 4.41 | 4.52 | 4.63 | 4.74 | 4.85 | 4.96 | 5.08 | 5.19 | 5.3 | 5.41 | 5.52 | 5.63 | 5.74 | 5.86 | 5.97 | 6.08 | 6.19 | 6.3 | 6.41 | 6.52 | 6.64 | 6.75 | 6.86 | 6.97 | 7.08 | 7.19 | 7.3 | 7.41 | 7.53 | 7.64 | 7.75 | 7.86 | 7.97 | 8.08 | 8.19 | 8.31 | 8.42 | 8.53 | 8.64
25.47 | 4.18 | 4.3 | 4.41 | 4.52 | 4.63 | 4.74 | 4.85 | 4.96 | 5.08 | 5.19 | 5.3 | 5.41 | 5.52 | 5.63 | 5.74 | 5.85 | 5.97 | 6.08 | 6.19 | 6.3 | 6.41 | 6.52 | 6.63 | 6.75 | 6.86 | 6.97 | 7.08 | 7.19 | 7.3 | 7.41 | 7.52 | 7.64 | 7.75 | 7.86 | 7.97 | 8.08 | 8.19 | 8.3 | 8.42 | 8.53 | 8.64
25.48 | 4.18 | 4.29 | 4.41 | 4.52 | 4.63 | 4.74 | 4.85 | 4.96 | 5.07 | 5.19 | 5.3 | 5.41 | 5.52 | 5.63 | 5.74 | 5.85 | 5.96 | 6.08 | 6.19 | 6.3 | 6.41 | 6.52 | 6.63 | 6.74 | 6.86 | 6.97 | 7.08 | 7.19 | 7.3 | 7.41 | 7.52 | 7.63 | 7.75 | 7.86 | 7.97 | 8.08 | 8.19 | 8.3 | 8.41 | 8.52 | 8.64
25.49 | 4.18 | 4.29 | 4.41 | 4.52 | 4.63 | 4.74 | 4.85 | 4.96 | 5.07 | 5.18 | 5.3 | 5.41 | 5.52 | 5.63 | 5.74 | 5.85 | 5.96 | 6.07 | 6.19 | 6.3 | 6.41 | 6.52 | 6.63 | 6.74 | 6.85 | 6.97 | 7.08 | 7.19 | 7.3 | 7.41 | 7.52 | 7.63 | 7.74 | 7.86 | 7.97 | 8.08 | 8.19 | 8.3 | 8.41 | 8.52 | 8.63
25.5 | 4.18 | 4.29 | 4.4 | 4.52 | 4.63 | 4.74 | 4.85 | 4.96 | 5.07 | 5.18 | 5.29 | 5.41 | 5.52 | 5.63 | 5.74 | 5.85 | 5.96 | 6.07 | 6.19 | 6.3 | 6.41 | 6.52 | 6.63 | 6.74 | 6.85 | 6.96 | 7.08 | 7.19 | 7.3 | 7.41 | 7.52 | 7.63 | 7.74 | 7.85 | 7.97 | 8.08 | 8.19 | 8.3 | 8.41 | 8.52 | 8.63
25.51 | 4.18 | 4.29 | 4.4 | 4.51 | 4.63 | 4.74 | 4.85 | 4.96 | 5.07 | 5.18 | 5.29 | 5.41 | 5.52 | 5.63 | 5.74 | 5.85 | 5.96 | 6.07 | 6.18 | 6.3 | 6.41 | 6.52 | 6.63 | 6.74 | 6.85 | 6.96 | 7.07 | 7.19 | 7.3 | 7.41 | 7.52 | 7.63 | 7.74 | 7.85 | 7.96 | 8.08 | 8.19 | 8.3 | 8.41 | 8.52 | 8.63
25.52 | 4.18 | 4.29 | 4.4 | 4.51 | 4.63 | 4.74 | 4.85 | 4.96 | 5.07 | 5.18 | 5.29 | 5.4 | 5.52 | 5.63 | 5.74 | 5.85 | 5.96 | 6.07 | 6.18 | 6.29 | 6.41 | 6.52 | 6.63 | 6.74 | 6.85 | 6.96 | 7.07 | 7.18 | 7.3 | 7.41 | 7.52 | 7.63 | 7.74 | 7.85 | 7.96 | 8.07 | 8.18 | 8.3 | 8.41 | 8.52 | 8.63
25.53 | 4.18 | 4.29 | 4.4 | 4.51 | 4.62 | 4.74 | 4.85 | 4.96 | 5.07 | 5.18 | 5.29 | 5.4 | 5.51 | 5.63 | 5.74 | 5.85 | 5.96 | 6.07 | 6.18 | 6.29 | 6.4 | 6.52 | 6.63 | 6.74 | 6.85 | 6.96 | 7.07 | 7.18 | 7.29 | 7.41 | 7.52 | 7.63 | 7.74 | 7.85 | 7.96 | 8.07 | 8.18 | 8.29 | 8.41 | 8.52 | 8.63
25.54 | 4.18 | 4.29 | 4.4 | 4.51 | 4.62 | 4.73 | 4.85 | 4.96 | 5.07 | 5.18 | 5.29 | 5.4 | 5.51 | 5.62 | 5.74 | 5.85 | 5.96 | 6.07 | 6.18 | 6.29 | 6.4 | 6.51 | 6.63 | 6.74 | 6.85 | 6.96 | 7.07 | 7.18 | 7.29 | 7.4 | 7.51 | 7.63 | 7.74 | 7.85 | 7.96 | 8.07 | 8.18 | 8.29 | 8.4 | 8.52 | 8.63
25.55 | 4.18 | 4.29 | 4.4 | 4.51 | 4.62 | 4.73 | 4.84 | 4.96 | 5.07 | 5.18 | 5.29 | 5.4 | 5.51 | 5.62 | 5.73 | 5.85 | 5.96 | 6.07 | 6.18 | 6.29 | 6.4 | 6.51 | 6.62 | 6.74 | 6.85 | 6.96 | 7.07 | 7.18 | 7.29 | 7.4 | 7.51 | 7.62 | 7.74 | 7.85 | 7.96 | 8.07 | 8.18 | 8.29 | 8.4 | 8.51 | 8.63
25.56 | 4.18 | 4.29 | 4.4 | 4.51 | 4.62 | 4.73 | 4.84 | 4.96 | 5.07 | 5.18 | 5.29 | 5.4 | 5.51 | 5.62 | 5.73 | 5.84 | 5.96 | 6.07 | 6.18 | 6.29 | 6.4 | 6.51 | 6.62 | 6.73 | 6.85 | 6.96 | 7.07 | 7.18 | 7.29 | 7.4 | 7.51 | 7.62 | 7.73 | 7.85 | 7.96 | 8.07 | 8.18 | 8.29 | 8.4 | 8.51 | 8.62
25.57 | 4.18 | 4.29 | 4.4 | 4.51 | 4.62 | 4.73 | 4.84 | 4.95 | 5.07 | 5.18 | 5.29 | 5.4 | 5.51 | 5.62 | 5.73 | 5.84 | 5.95 | 6.07 | 6.18 | 6.29 | 6.4 | 6.51 | 6.62 | 6.73 | 6.84 | 6.95 | 7.07 | 7.18 | 7.29 | 7.4 | 7.51 | 7.62 | 7.73 | 7.84 | 7.96 | 8.07 | 8.18 | 8.29 | 8.4 | 8.51 | 8.62
25.58 | 4.17 | 4.29 | 4.4 | 4.51 | 4.62 | 4.73 | 4.84 | 4.95 | 5.06 | 5.18 | 5.29 | 5.4 | 5.51 | 5.62 | 5.73 | 5.84 | 5.95 | 6.06 | 6.18 | 6.29 | 6.4 | 6.51 | 6.62 | 6.73 | 6.84 | 6.95 | 7.06 | 7.18 | 7.29 | 7.4 | 7.51 | 7.62 | 7.73 | 7.84 | 7.95 | 8.06 | 8.18 | 8.29 | 8.4 | 8.51 | 8.62
25.59 | 4.17 | 4.29 | 4.4 | 4.51 | 4.62 | 4.73 | 4.84 | 4.95 | 5.06 | 5.17 | 5.29 | 5.4 | 5.51 | 5.62 | 5.73 | 5.84 | 5.95 | 6.06 | 6.17 | 6.29 | 6.4 | 6.51 | 6.62 | 6.73 | 6.84 | 6.95 | 7.06 | 7.17 | 7.29 | 7.4 | 7.51 | 7.62 | 7.73 | 7.84 | 7.95 | 8.06 | 8.17 | 8.29 | 8.4 | 8.51 | 8.62
25.6 | 4.17 | 4.28 | 4.4 | 4.51 | 4.62 | 4.73 | 4.84 | 4.95 | 5.06 | 5.17 | 5.28 | 5.4 | 5.51 | 5.62 | 5.73 | 5.84 | 5.95 | 6.06 | 6.17 | 6.28 | 6.4 | 6.51 | 6.62 | 6.73 | 6.84 | 6.95 | 7.06 | 7.17 | 7.28 | 7.4 | 7.51 | 7.62 | 7.73 | 7.84 | 7.95 | 8.06 | 8.17 | 8.28 | 8.4 | 8.51 | 8.62
25.61 | 4.17 | 4.28 | 4.39 | 4.51 | 4.62 | 4.73 | 4.84 | 4.95 | 5.06 | 5.17 | 5.28 | 5.39 | 5.51 | 5.62 | 5.73 | 5.84 | 5.95 | 6.06 | 6.17 | 6.28 | 6.39 | 6.51 | 6.62 | 6.73 | 6.84 | 6.95 | 7.06 | 7.17 | 7.28 | 7.39 | 7.51 | 7.62 | 7.73 | 7.84 | 7.95 | 8.06 | 8.17 | 8.28 | 8.39 | 8.5 | 8.62
25.62 | 4.17 | 4.28 | 4.39 | 4.5 | 4.62 | 4.73 | 4.84 | 4.95 | 5.06 | 5.17 | 5.28 | 5.39 | 5.5 | 5.62 | 5.73 | 5.84 | 5.95 | 6.06 | 6.17 | 6.28 | 6.39 | 6.5 | 6.62 | 6.73 | 6.84 | 6.95 | 7.06 | 7.17 | 7.28 | 7.39 | 7.5 | 7.61 | 7.73 | 7.84 | 7.95 | 8.06 | 8.17 | 8.28 | 8.39 | 8.5 | 8.61
25.63 | 4.17 | 4.28 | 4.39 | 4.5 | 4.62 | 4.73 | 4.84 | 4.95 | 5.06 | 5.17 | 5.28 | 5.39 | 5.5 | 5.61 | 5.73 | 5.84 | 5.95 | 6.06 | 6.17 | 6.28 | 6.39 | 6.5 | 6.61 | 6.73 | 6.84 | 6.95 | 7.06 | 7.17 | 7.28 | 7.39 | 7.5 | 7.61 | 7.72 | 7.84 | 7.95 | 8.06 | 8.17 | 8.28 | 8.39 | 8.5 | 8.61
25.64 | 4.17 | 4.28 | 4.39 | 4.5 | 4.61 | 4.73 | 4.84 | 4.95 | 5.06 | 5.17 | 5.28 | 5.39 | 5.5 | 5.61 | 5.72 | 5.84 | 5.95 | 6.06 | 6.17 | 6.28 | 6.39 | 6.5 | 6.61 | 6.72 | 6.83 | 6.95 | 7.06 | 7.17 | 7.28 | 7.39 | 7.5 | 7.61 | 7.72 | 7.83 | 7.95 | 8.06 | 8.17 | 8.28 | 8.39 | 8.5 | 8.61
25.65 | 4.17 | 4.28 | 4.39 | 4.5 | 4.61 | 4.72 | 4.84 | 4.95 | 5.06 | 5.17 | 5.28 | 5.39 | 5.5 | 5.61 | 5.72 | 5.83 | 5.95 | 6.06 | 6.17 | 6.28 | 6.39 | 6.5 | 6.61 | 6.72 | 6.83 | 6.94 | 7.06 | 7.17 | 7.28 | 7.39 | 7.5 | 7.61 | 7.72 | 7.83 | 7.94 | 8.05 | 8.17 | 8.28 | 8.39 | 8.5 | 8.61
25.66 | 4.17 | 4.28 | 4.39 | 4.5 | 4.61 | 4.72 | 4.83 | 4.95 | 5.06 | 5.17 | 5.28 | 5.39 | 5.5 | 5.61 | 5.72 | 5.83 | 5.94 | 6.06 | 6.17 | 6.28 | 6.39 | 6.5 | 6.61 | 6.72 | 6.83 | 6.94 | 7.05 | 7.17 | 7.28 | 7.39 | 7.5 | 7.61 | 7.72 | 7.83 | 7.94 | 8.05 | 8.16 | 8.27 | 8.39 | 8.5 | 8.61
25.67 | 4.17 | 4.28 | 4.39 | 4.5 | 4.61 | 4.72 | 4.83 | 4.94 | 5.06 | 5.17 | 5.28 | 5.39 | 5.5 | 5.61 | 5.72 | 5.83 | 5.94 | 6.05 | 6.17 | 6.28 | 6.39 | 6.5 | 6.61 | 6.72 | 6.83 | 6.94 | 7.05 | 7.16 | 7.27 | 7.39 | 7.5 | 7.61 | 7.72 | 7.83 | 7.94 | 8.05 | 8.16 | 8.27 | 8.38 | 8.5 | 8.61
25.68 | 4.17 | 4.28 | 4.39 | 4.5 | 4.61 | 4.72 | 4.83 | 4.94 | 5.05 | 5.17 | 5.28 | 5.39 | 5.5 | 5.61 | 5.72 | 5.83 | 5.94 | 6.05 | 6.16 | 6.28 | 6.39 | 6.5 | 6.61 | 6.72 | 6.83 | 6.94 | 7.05 | 7.16 | 7.27 | 7.38 | 7.5 | 7.61 | 7.72 | 7.83 | 7.94 | 8.05 | 8.16 | 8.27 | 8.38 | 8.49 | 8.6
25.69 | 4.17 | 4.28 | 4.39 | 4.5 | 4.61 | 4.72 | 4.83 | 4.94 | 5.05 | 5.16 | 5.28 | 5.39 | 5.5 | 5.61 | 5.72 | 5.83 | 5.94 | 6.05 | 6.16 | 6.27 | 6.38 | 6.5 | 6.61 | 6.72 | 6.83 | 6.94 | 7.05 | 7.16 | 7.27 | 7.38 | 7.49 | 7.61 | 7.72 | 7.83 | 7.94 | 8.05 | 8.16 | 8.27 | 8.38 | 8.49 | 8.6
25.7 | 4.16 | 4.28 | 4.39 | 4.5 | 4.61 | 4.72 | 4.83 | 4.94 | 5.05 | 5.16 | 5.27 | 5.39 | 5.5 | 5.61 | 5.72 | 5.83 | 5.94 | 6.05 | 6.16 | 6.27 | 6.38 | 6.49 | 6.61 | 6.72 | 6.83 | 6.94 | 7.05 | 7.16 | 7.27 | 7.38 | 7.49 | 7.6 | 7.71 | 7.83 | 7.94 | 8.05 | 8.16 | 8.27 | 8.38 | 8.49 | 8.6
25.71 | 4.16 | 4.27 | 4.39 | 4.5 | 4.61 | 4.72 | 4.83 | 4.94 | 5.05 | 5.16 | 5.27 | 5.38 | 5.5 | 5.61 | 5.72 | 5.83 | 5.94 | 6.05 | 6.16 | 6.27 | 6.38 | 6.49 | 6.6 | 6.72 | 6.83 | 6.94 | 7.05 | 7.16 | 7.27 | 7.38 | 7.49 | 7.6 | 7.71 | 7.82 | 7.93 | 8.05 | 8.16 | 8.27 | 8.38 | 8.49 | 8.6
25.72 | 4.16 | 4.27 | 4.38 | 4.5 | 4.61 | 4.72 | 4.83 | 4.94 | 5.05 | 5.16 | 5.27 | 5.38 | 5.49 | 5.6 | 5.72 | 5.83 | 5.94 | 6.05 | 6.16 | 6.27 | 6.38 | 6.49 | 6.6 | 6.71 | 6.82 | 6.94 | 7.05 | 7.16 | 7.27 | 7.38 | 7.49 | 7.6 | 7.71 | 7.82 | 7.93 | 8.04 | 8.16 | 8.27 | 8.38 | 8.49 | 8.6
25.73 | 4.16 | 4.27 | 4.38 | 4.5 | 4.61 | 4.72 | 4.83 | 4.94 | 5.05 | 5.16 | 5.27 | 5.38 | 5.49 | 5.6 | 5.71 | 5.83 | 5.94 | 6.05 | 6.16 | 6.27 | 6.38 | 6.49 | 6.6 | 6.71 | 6.82 | 6.93 | 7.05 | 7.16 | 7.27 | 7.38 | 7.49 | 7.6 | 7.71 | 7.82 | 7.93 | 8.04 | 8.15 | 8.26 | 8.38 | 8.49 | 8.6
25.74 | 4.16 | 4.27 | 4.38 | 4.49 | 4.6 | 4.72 | 4.83 | 4.94 | 5.05 | 5.16 | 5.27 | 5.38 | 5.49 | 5.6 | 5.71 | 5.82 | 5.94 | 6.05 | 6.16 | 6.27 | 6.38 | 6.49 | 6.6 | 6.71 | 6.82 | 6.93 | 7.04 | 7.15 | 7.27 | 7.38 | 7.49 | 7.6 | 7.71 | 7.82 | 7.93 | 8.04 | 8.15 | 8.26 | 8.37 | 8.48 | 8.6
25.75 | 4.16 | 4.27 | 4.38 | 4.49 | 4.6 | 4.71 | 4.83 | 4.94 | 5.05 | 5.16 | 5.27 | 5.38 | 5.49 | 5.6 | 5.71 | 5.82 | 5.93 | 6.05 | 6.16 | 6.27 | 6.38 | 6.49 | 6.6 | 6.71 | 6.82 | 6.93 | 7.04 | 7.15 | 7.26 | 7.37 | 7.49 | 7.6 | 7.71 | 7.82 | 7.93 | 8.04 | 8.15 | 8.26 | 8.37 | 8.48 | 8.59
25.76 | 4.16 | 4.27 | 4.38 | 4.49 | 4.6 | 4.71 | 4.82 | 4.94 | 5.05 | 5.16 | 5.27 | 5.38 | 5.49 | 5.6 | 5.71 | 5.82 | 5.93 | 6.04 | 6.15 | 6.27 | 6.38 | 6.49 | 6.6 | 6.71 | 6.82 | 6.93 | 7.04 | 7.15 | 7.26 | 7.37 | 7.48 | 7.6 | 7.71 | 7.82 | 7.93 | 8.04 | 8.15 | 8.26 | 8.37 | 8.48 | 8.59
25.77 | 4.16 | 4.27 | 4.38 | 4.49 | 4.6 | 4.71 | 4.82 | 4.93 | 5.05 | 5.16 | 5.27 | 5.38 | 5.49 | 5.6 | 5.71 | 5.82 | 5.93 | 6.04 | 6.15 | 6.26 | 6.38 | 6.49 | 6.6 | 6.71 | 6.82 | 6.93 | 7.04 | 7.15 | 7.26 | 7.37 | 7.48 | 7.59 | 7.7 | 7.82 | 7.93 | 8.04 | 8.15 | 8.26 | 8.37 | 8.48 | 8.59
25.78 | 4.16 | 4.27 | 4.38 | 4.49 | 4.6 | 4.71 | 4.82 | 4.93 | 5.04 | 5.16 | 5.27 | 5.38 | 5.49 | 5.6 | 5.71 | 5.82 | 5.93 | 6.04 | 6.15 | 6.26 | 6.37 | 6.48 | 6.6 | 6.71 | 6.82 | 6.93 | 7.04 | 7.15 | 7.26 | 7.37 | 7.48 | 7.59 | 7.7 | 7.81 | 7.92 | 8.04 | 8.15 | 8.26 | 8.37 | 8.48 | 8.59
25.79 | 4.16 | 4.27 | 4.38 | 4.49 | 4.6 | 4.71 | 4.82 | 4.93 | 5.04 | 5.15 | 5.27 | 5.38 | 5.49 | 5.6 | 5.71 | 5.82 | 5.93 | 6.04 | 6.15 | 6.26 | 6.37 | 6.48 | 6.59 | 6.71 | 6.82 | 6.93 | 7.04 | 7.15 | 7.26 | 7.37 | 7.48 | 7.59 | 7.7 | 7.81 | 7.92 | 8.03 | 8.14 | 8.26 | 8.37 | 8.48 | 8.59
25.8 | 4.16 | 4.27 | 4.38 | 4.49 | 4.6 | 4.71 | 4.82 | 4.93 | 5.04 | 5.15 | 5.26 | 5.37 | 5.49 | 5.6 | 5.71 | 5.82 | 5.93 | 6.04 | 6.15 | 6.26 | 6.37 | 6.48 | 6.59 | 6.7 | 6.81 | 6.93 | 7.04 | 7.15 | 7.26 | 7.37 | 7.48 | 7.59 | 7.7 | 7.81 | 7.92 | 8.03 | 8.14 | 8.25 | 8.36 | 8.48 | 8.59
25.81 | 4.16 | 4.27 | 4.38 | 4.49 | 4.6 | 4.71 | 4.82 | 4.93 | 5.04 | 5.15 | 5.26 | 5.37 | 5.48 | 5.6 | 5.71 | 5.82 | 5.93 | 6.04 | 6.15 | 6.26 | 6.37 | 6.48 | 6.59 | 6.7 | 6.81 | 6.92 | 7.03 | 7.15 | 7.26 | 7.37 | 7.48 | 7.59 | 7.7 | 7.81 | 7.92 | 8.03 | 8.14 | 8.25 | 8.36 | 8.47 | 8.58
25.82 | 4.15 | 4.27 | 4.38 | 4.49 | 4.6 | 4.71 | 4.82 | 4.93 | 5.04 | 5.15 | 5.26 | 5.37 | 5.48 | 5.59 | 5.7 | 5.82 | 5.93 | 6.04 | 6.15 | 6.26 | 6.37 | 6.48 | 6.59 | 6.7 | 6.81 | 6.92 | 7.03 | 7.14 | 7.25 | 7.37 | 7.48 | 7.59 | 7.7 | 7.81 | 7.92 | 8.03 | 8.14 | 8.25 | 8.36 | 8.47 | 8.58
25.83 | 4.15 | 4.26 | 4.38 | 4.49 | 4.6 | 4.71 | 4.82 | 4.93 | 5.04 | 5.15 | 5.26 | 5.37 | 5.48 | 5.59 | 5.7 | 5.81 | 5.93 | 6.04 | 6.15 | 6.26 | 6.37 | 6.48 | 6.59 | 6.7 | 6.81 | 6.92 | 7.03 | 7.14 | 7.25 | 7.36 | 7.47 | 7.59 | 7.7 | 7.81 | 7.92 | 8.03 | 8.14 | 8.25 | 8.36 | 8.47 | 8.58
25.84 | 4.15 | 4.26 | 4.37 | 4.49 | 4.6 | 4.71 | 4.82 | 4.93 | 5.04 | 5.15 | 5.26 | 5.37 | 5.48 | 5.59 | 5.7 | 5.81 | 5.92 | 6.03 | 6.15 | 6.26 | 6.37 | 6.48 | 6.59 | 6.7 | 6.81 | 6.92 | 7.03 | 7.14 | 7.25 | 7.36 | 7.47 | 7.58 | 7.69 | 7.81 | 7.92 | 8.03 | 8.14 | 8.25 | 8.36 | 8.47 | 8.58
25.85 | 4.15 | 4.26 | 4.37 | 4.48 | 4.59 | 4.71 | 4.82 | 4.93 | 5.04 | 5.15 | 5.26 | 5.37 | 5.48 | 5.59 | 5.7 | 5.81 | 5.92 | 6.03 | 6.14 | 6.25 | 6.37 | 6.48 | 6.59 | 6.7 | 6.81 | 6.92 | 7.03 | 7.14 | 7.25 | 7.36 | 7.47 | 7.58 | 7.69 | 7.8 | 7.91 | 8.03 | 8.14 | 8.25 | 8.36 | 8.47 | 8.58
25.86 | 4.15 | 4.26 | 4.37 | 4.48 | 4.59 | 4.7 | 4.82 | 4.93 | 5.04 | 5.15 | 5.26 | 5.37 | 5.48 | 5.59 | 5.7 | 5.81 | 5.92 | 6.03 | 6.14 | 6.25 | 6.36 | 6.48 | 6.59 | 6.7 | 6.81 | 6.92 | 7.03 | 7.14 | 7.25 | 7.36 | 7.47 | 7.58 | 7.69 | 7.8 | 7.91 | 8.02 | 8.13 | 8.24 | 8.36 | 8.47 | 8.58
25.87 | 4.15 | 4.26 | 4.37 | 4.48 | 4.59 | 4.7 | 4.81 | 4.93 | 5.04 | 5.15 | 5.26 | 5.37 | 5.48 | 5.59 | 5.7 | 5.81 | 5.92 | 6.03 | 6.14 | 6.25 | 6.36 | 6.47 | 6.58 | 6.7 | 6.81 | 6.92 | 7.03 | 7.14 | 7.25 | 7.36 | 7.47 | 7.58 | 7.69 | 7.8 | 7.91 | 8.02 | 8.13 | 8.24 | 8.35 | 8.46 | 8.58
25.88 | 4.15 | 4.26 | 4.37 | 4.48 | 4.59 | 4.7 | 4.81 | 4.92 | 5.03 | 5.15 | 5.26 | 5.37 | 5.48 | 5.59 | 5.7 | 5.81 | 5.92 | 6.03 | 6.14 | 6.25 | 6.36 | 6.47 | 6.58 | 6.69 | 6.8 | 6.92 | 7.03 | 7.14 | 7.25 | 7.36 | 7.47 | 7.58 | 7.69 | 7.8 | 7.91 | 8.02 | 8.13 | 8.24 | 8.35 | 8.46 | 8.57
25.89 | 4.15 | 4.26 | 4.37 | 4.48 | 4.59 | 4.7 | 4.81 | 4.92 | 5.03 | 5.14 | 5.25 | 5.37 | 5.48 | 5.59 | 5.7 | 5.81 | 5.92 | 6.03 | 6.14 | 6.25 | 6.36 | 6.47 | 6.58 | 6.69 | 6.8 | 6.91 | 7.02 | 7.13 | 7.25 | 7.36 | 7.47 | 7.58 | 7.69 | 7.8 | 7.91 | 8.02 | 8.13 | 8.24 | 8.35 | 8.46 | 8.57
25.9 | 4.15 | 4.26 | 4.37 | 4.48 | 4.59 | 4.7 | 4.81 | 4.92 | 5.03 | 5.14 | 5.25 | 5.36 | 5.48 | 5.59 | 5.7 | 5.81 | 5.92 | 6.03 | 6.14 | 6.25 | 6.36 | 6.47 | 6.58 | 6.69 | 6.8 | 6.91 | 7.02 | 7.13 | 7.24 | 7.35 | 7.47 | 7.58 | 7.69 | 7.8 | 7.91 | 8.02 | 8.13 | 8.24 | 8.35 | 8.46 | 8.57
25.91 | 4.15 | 4.26 | 4.37 | 4.48 | 4.59 | 4.7 | 4.81 | 4.92 | 5.03 | 5.14 | 5.25 | 5.36 | 5.47 | 5.58 | 5.7 | 5.81 | 5.92 | 6.03 | 6.14 | 6.25 | 6.36 | 6.47 | 6.58 | 6.69 | 6.8 | 6.91 | 7.02 | 7.13 | 7.24 | 7.35 | 7.46 | 7.57 | 7.68 | 7.8 | 7.91 | 8.02 | 8.13 | 8.24 | 8.35 | 8.46 | 8.57
25.92 | 4.15 | 4.26 | 4.37 | 4.48 | 4.59 | 4.7 | 4.81 | 4.92 | 5.03 | 5.14 | 5.25 | 5.36 | 5.47 | 5.58 | 5.69 | 5.8 | 5.92 | 6.03 | 6.14 | 6.25 | 6.36 | 6.47 | 6.58 | 6.69 | 6.8 | 6.91 | 7.02 | 7.13 | 7.24 | 7.35 | 7.46 | 7.57 | 7.68 | 7.79 | 7.9 | 8.01 | 8.13 | 8.24 | 8.35 | 8.46 | 8.57
25.93 | 4.15 | 4.26 | 4.37 | 4.48 | 4.59 | 4.7 | 4.81 | 4.92 | 5.03 | 5.14 | 5.25 | 5.36 | 5.47 | 5.58 | 5.69 | 5.8 | 5.91 | 6.02 | 6.14 | 6.25 | 6.36 | 6.47 | 6.58 | 6.69 | 6.8 | 6.91 | 7.02 | 7.13 | 7.24 | 7.35 | 7.46 | 7.57 | 7.68 | 7.79 | 7.9 | 8.01 | 8.12 | 8.23 | 8.34 | 8.46 | 8.57
25.94 | 4.14 | 4.26 | 4.37 | 4.48 | 4.59 | 4.7 | 4.81 | 4.92 | 5.03 | 5.14 | 5.25 | 5.36 | 5.47 | 5.58 | 5.69 | 5.8 | 5.91 | 6.02 | 6.13 | 6.24 | 6.35 | 6.47 | 6.58 | 6.69 | 6.8 | 6.91 | 7.02 | 7.13 | 7.24 | 7.35 | 7.46 | 7.57 | 7.68 | 7.79 | 7.9 | 8.01 | 8.12 | 8.23 | 8.34 | 8.45 | 8.56
25.95 | 4.14 | 4.25 | 4.36 | 4.48 | 4.59 | 4.7 | 4.81 | 4.92 | 5.03 | 5.14 | 5.25 | 5.36 | 5.47 | 5.58 | 5.69 | 5.8 | 5.91 | 6.02 | 6.13 | 6.24 | 6.35 | 6.46 | 6.57 | 6.69 | 6.8 | 6.91 | 7.02 | 7.13 | 7.24 | 7.35 | 7.46 | 7.57 | 7.68 | 7.79 | 7.9 | 8.01 | 8.12 | 8.23 | 8.34 | 8.45 | 8.56
25.96 | 4.14 | 4.25 | 4.36 | 4.47 | 4.59 | 4.7 | 4.81 | 4.92 | 5.03 | 5.14 | 5.25 | 5.36 | 5.47 | 5.58 | 5.69 | 5.8 | 5.91 | 6.02 | 6.13 | 6.24 | 6.35 | 6.46 | 6.57 | 6.68 | 6.79 | 6.9 | 7.02 | 7.13 | 7.24 | 7.35 | 7.46 | 7.57 | 7.68 | 7.79 | 7.9 | 8.01 | 8.12 | 8.23 | 8.34 | 8.45 | 8.56
25.97 | 4.14 | 4.25 | 4.36 | 4.47 | 4.58 | 4.69 | 4.81 | 4.92 | 5.03 | 5.14 | 5.25 | 5.36 | 5.47 | 5.58 | 5.69 | 5.8 | 5.91 | 6.02 | 6.13 | 6.24 | 6.35 | 6.46 | 6.57 | 6.68 | 6.79 | 6.9 | 7.01 | 7.12 | 7.23 | 7.35 | 7.46 | 7.57 | 7.68 | 7.79 | 7.9 | 8.01 | 8.12 | 8.23 | 8.34 | 8.45 | 8.56
25.98 | 4.14 | 4.25 | 4.36 | 4.47 | 4.58 | 4.69 | 4.8 | 4.91 | 5.02 | 5.14 | 5.25 | 5.36 | 5.47 | 5.58 | 5.69 | 5.8 | 5.91 | 6.02 | 6.13 | 6.24 | 6.35 | 6.46 | 6.57 | 6.68 | 6.79 | 6.9 | 7.01 | 7.12 | 7.23 | 7.34 | 7.45 | 7.56 | 7.68 | 7.79 | 7.9 | 8.01 | 8.12 | 8.23 | 8.34 | 8.45 | 8.56
25.99 | 4.14 | 4.25 | 4.36 | 4.47 | 4.58 | 4.69 | 4.8 | 4.91 | 5.02 | 5.13 | 5.24 | 5.36 | 5.47 | 5.58 | 5.69 | 5.8 | 5.91 | 6.02 | 6.13 | 6.24 | 6.35 | 6.46 | 6.57 | 6.68 | 6.79 | 6.9 | 7.01 | 7.12 | 7.23 | 7.34 | 7.45 | 7.56 | 7.67 | 7.78 | 7.89 | 8 | 8.12 | 8.23 | 8.34 | 8.45 | 8.56
26 | 4.14 | 4.25 | 4.36 | 4.47 | 4.58 | 4.69 | 4.8 | 4.91 | 5.02 | 5.13 | 5.24 | 5.35 | 5.46 | 5.58 | 5.69 | 5.8 | 5.91 | 6.02 | 6.13 | 6.24 | 6.35 | 6.46 | 6.57 | 6.68 | 6.79 | 6.9 | 7.01 | 7.12 | 7.23 | 7.34 | 7.45 | 7.56 | 7.67 | 7.78 | 7.89 | 8 | 8.11 | 8.22 | 8.33 | 8.44 | 8.56
26.01 | 4.14 | 4.25 | 4.36 | 4.47 | 4.58 | 4.69 | 4.8 | 4.91 | 5.02 | 5.13 | 5.24 | 5.35 | 5.46 | 5.57 | 5.68 | 5.79 | 5.91 | 6.02 | 6.13 | 6.24 | 6.35 | 6.46 | 6.57 | 6.68 | 6.79 | 6.9 | 7.01 | 7.12 | 7.23 | 7.34 | 7.45 | 7.56 | 7.67 | 7.78 | 7.89 | 8 | 8.11 | 8.22 | 8.33 | 8.44 | 8.55
26.02 | 4.14 | 4.25 | 4.36 | 4.47 | 4.58 | 4.69 | 4.8 | 4.91 | 5.02 | 5.13 | 5.24 | 5.35 | 5.46 | 5.57 | 5.68 | 5.79 | 5.9 | 6.01 | 6.12 | 6.24 | 6.35 | 6.46 | 6.57 | 6.68 | 6.79 | 6.9 | 7.01 | 7.12 | 7.23 | 7.34 | 7.45 | 7.56 | 7.67 | 7.78 | 7.89 | 8 | 8.11 | 8.22 | 8.33 | 8.44 | 8.55
26.03 | 4.14 | 4.25 | 4.36 | 4.47 | 4.58 | 4.69 | 4.8 | 4.91 | 5.02 | 5.13 | 5.24 | 5.35 | 5.46 | 5.57 | 5.68 | 5.79 | 5.9 | 6.01 | 6.12 | 6.23 | 6.34 | 6.45 | 6.56 | 6.68 | 6.79 | 6.9 | 7.01 | 7.12 | 7.23 | 7.34 | 7.45 | 7.56 | 7.67 | 7.78 | 7.89 | 8 | 8.11 | 8.22 | 8.33 | 8.44 | 8.55
26.04 | 4.14 | 4.25 | 4.36 | 4.47 | 4.58 | 4.69 | 4.8 | 4.91 | 5.02 | 5.13 | 5.24 | 5.35 | 5.46 | 5.57 | 5.68 | 5.79 | 5.9 | 6.01 | 6.12 | 6.23 | 6.34 | 6.45 | 6.56 | 6.67 | 6.78 | 6.89 | 7 | 7.12 | 7.23 | 7.34 | 7.45 | 7.56 | 7.67 | 7.78 | 7.89 | 8 | 8.11 | 8.22 | 8.33 | 8.44 | 8.55
26.05 | 4.14 | 4.25 | 4.36 | 4.47 | 4.58 | 4.69 | 4.8 | 4.91 | 5.02 | 5.13 | 5.24 | 5.35 | 5.46 | 5.57 | 5.68 | 5.79 | 5.9 | 6.01 | 6.12 | 6.23 | 6.34 | 6.45 | 6.56 | 6.67 | 6.78 | 6.89 | 7 | 7.11 | 7.22 | 7.33 | 7.44 | 7.55 | 7.67 | 7.78 | 7.89 | 8 | 8.11 | 8.22 | 8.33 | 8.44 | 8.55
26.06 | 4.13 | 4.25 | 4.36 | 4.47 | 4.58 | 4.69 | 4.8 | 4.91 | 5.02 | 5.13 | 5.24 | 5.35 | 5.46 | 5.57 | 5.68 | 5.79 | 5.9 | 6.01 | 6.12 | 6.23 | 6.34 | 6.45 | 6.56 | 6.67 | 6.78 | 6.89 | 7 | 7.11 | 7.22 | 7.33 | 7.44 | 7.55 | 7.66 | 7.77 | 7.88 | 7.99 | 8.1 | 8.22 | 8.33 | 8.44 | 8.55
26.07 | 4.13 | 4.24 | 4.35 | 4.46 | 4.58 | 4.69 | 4.8 | 4.91 | 5.02 | 5.13 | 5.24 | 5.35 | 5.46 | 5.57 | 5.68 | 5.79 | 5.9 | 6.01 | 6.12 | 6.23 | 6.34 | 6.45 | 6.56 | 6.67 | 6.78 | 6.89 | 7 | 7.11 | 7.22 | 7.33 | 7.44 | 7.55 | 7.66 | 7.77 | 7.88 | 7.99 | 8.1 | 8.21 | 8.32 | 8.43 | 8.54
26.08 | 4.13 | 4.24 | 4.35 | 4.46 | 4.57 | 4.68 | 4.79 | 4.9 | 5.02 | 5.13 | 5.24 | 5.35 | 5.46 | 5.57 | 5.68 | 5.79 | 5.9 | 6.01 | 6.12 | 6.23 | 6.34 | 6.45 | 6.56 | 6.67 | 6.78 | 6.89 | 7 | 7.11 | 7.22 | 7.33 | 7.44 | 7.55 | 7.66 | 7.77 | 7.88 | 7.99 | 8.1 | 8.21 | 8.32 | 8.43 | 8.54
26.09 | 4.13 | 4.24 | 4.35 | 4.46 | 4.57 | 4.68 | 4.79 | 4.9 | 5.01 | 5.12 | 5.23 | 5.34 | 5.46 | 5.57 | 5.68 | 5.79 | 5.9 | 6.01 | 6.12 | 6.23 | 6.34 | 6.45 | 6.56 | 6.67 | 6.78 | 6.89 | 7 | 7.11 | 7.22 | 7.33 | 7.44 | 7.55 | 7.66 | 7.77 | 7.88 | 7.99 | 8.1 | 8.21 | 8.32 | 8.43 | 8.54
26.1 | 4.13 | 4.24 | 4.35 | 4.46 | 4.57 | 4.68 | 4.79 | 4.9 | 5.01 | 5.12 | 5.23 | 5.34 | 5.45 | 5.56 | 5.67 | 5.78 | 5.9 | 6.01 | 6.12 | 6.23 | 6.34 | 6.45 | 6.56 | 6.67 | 6.78 | 6.89 | 7 | 7.11 | 7.22 | 7.33 | 7.44 | 7.55 | 7.66 | 7.77 | 7.88 | 7.99 | 8.1 | 8.21 | 8.32 | 8.43 | 8.54
26.11 | 4.13 | 4.24 | 4.35 | 4.46 | 4.57 | 4.68 | 4.79 | 4.9 | 5.01 | 5.12 | 5.23 | 5.34 | 5.45 | 5.56 | 5.67 | 5.78 | 5.89 | 6 | 6.11 | 6.22 | 6.33 | 6.44 | 6.56 | 6.67 | 6.78 | 6.89 | 7 | 7.11 | 7.22 | 7.33 | 7.44 | 7.55 | 7.66 | 7.77 | 7.88 | 7.99 | 8.1 | 8.21 | 8.32 | 8.43 | 8.54
26.12 | 4.13 | 4.24 | 4.35 | 4.46 | 4.57 | 4.68 | 4.79 | 4.9 | 5.01 | 5.12 | 5.23 | 5.34 | 5.45 | 5.56 | 5.67 | 5.78 | 5.89 | 6 | 6.11 | 6.22 | 6.33 | 6.44 | 6.55 | 6.66 | 6.77 | 6.88 | 6.99 | 7.1 | 7.21 | 7.32 | 7.44 | 7.55 | 7.66 | 7.77 | 7.88 | 7.99 | 8.1 | 8.21 | 8.32 | 8.43 | 8.54
26.13 | 4.13 | 4.24 | 4.35 | 4.46 | 4.57 | 4.68 | 4.79 | 4.9 | 5.01 | 5.12 | 5.23 | 5.34 | 5.45 | 5.56 | 5.67 | 5.78 | 5.89 | 6 | 6.11 | 6.22 | 6.33 | 6.44 | 6.55 | 6.66 | 6.77 | 6.88 | 6.99 | 7.1 | 7.21 | 7.32 | 7.43 | 7.54 | 7.65 | 7.76 | 7.87 | 7.98 | 8.09 | 8.2 | 8.31 | 8.42 | 8.53
26.14 | 4.13 | 4.24 | 4.35 | 4.46 | 4.57 | 4.68 | 4.79 | 4.9 | 5.01 | 5.12 | 5.23 | 5.34 | 5.45 | 5.56 | 5.67 | 5.78 | 5.89 | 6 | 6.11 | 6.22 | 6.33 | 6.44 | 6.55 | 6.66 | 6.77 | 6.88 | 6.99 | 7.1 | 7.21 | 7.32 | 7.43 | 7.54 | 7.65 | 7.76 | 7.87 | 7.98 | 8.09 | 8.2 | 8.31 | 8.42 | 8.53
26.15 | 4.13 | 4.24 | 4.35 | 4.46 | 4.57 | 4.68 | 4.79 | 4.9 | 5.01 | 5.12 | 5.23 | 5.34 | 5.45 | 5.56 | 5.67 | 5.78 | 5.89 | 6 | 6.11 | 6.22 | 6.33 | 6.44 | 6.55 | 6.66 | 6.77 | 6.88 | 6.99 | 7.1 | 7.21 | 7.32 | 7.43 | 7.54 | 7.65 | 7.76 | 7.87 | 7.98 | 8.09 | 8.2 | 8.31 | 8.42 | 8.53
26.16 | 4.13 | 4.24 | 4.35 | 4.46 | 4.57 | 4.68 | 4.79 | 4.9 | 5.01 | 5.12 | 5.23 | 5.34 | 5.45 | 5.56 | 5.67 | 5.78 | 5.89 | 6 | 6.11 | 6.22 | 6.33 | 6.44 | 6.55 | 6.66 | 6.77 | 6.88 | 6.99 | 7.1 | 7.21 | 7.32 | 7.43 | 7.54 | 7.65 | 7.76 | 7.87 | 7.98 | 8.09 | 8.2 | 8.31 | 8.42 | 8.53
26.17 | 4.13 | 4.24 | 4.35 | 4.46 | 4.57 | 4.68 | 4.79 | 4.9 | 5.01 | 5.12 | 5.23 | 5.34 | 5.45 | 5.56 | 5.67 | 5.78 | 5.89 | 6 | 6.11 | 6.22 | 6.33 | 6.44 | 6.55 | 6.66 | 6.77 | 6.88 | 6.99 | 7.1 | 7.21 | 7.32 | 7.43 | 7.54 | 7.65 | 7.76 | 7.87 | 7.98 | 8.09 | 8.2 | 8.31 | 8.42 | 8.53
26.18 | 4.12 | 4.23 | 4.34 | 4.46 | 4.57 | 4.68 | 4.79 | 4.9 | 5.01 | 5.12 | 5.23 | 5.34 | 5.45 | 5.56 | 5.67 | 5.78 | 5.89 | 6 | 6.11 | 6.22 | 6.33 | 6.44 | 6.55 | 6.66 | 6.77 | 6.88 | 6.99 | 7.1 | 7.21 | 7.32 | 7.43 | 7.54 | 7.65 | 7.76 | 7.87 | 7.98 | 8.09 | 8.2 | 8.31 | 8.42 | 8.53
26.19 | 4.12 | 4.23 | 4.34 | 4.45 | 4.56 | 4.67 | 4.78 | 4.89 | 5 | 5.11 | 5.22 | 5.33 | 5.44 | 5.55 | 5.66 | 5.77 | 5.89 | 6 | 6.11 | 6.22 | 6.33 | 6.44 | 6.55 | 6.66 | 6.77 | 6.88 | 6.99 | 7.1 | 7.21 | 7.32 | 7.43 | 7.54 | 7.65 | 7.76 | 7.87 | 7.98 | 8.09 | 8.2 | 8.31 | 8.42 | 8.53
26.2 | 4.12 | 4.23 | 4.34 | 4.45 | 4.56 | 4.67 | 4.78 | 4.89 | 5 | 5.11 | 5.22 | 5.33 | 5.44 | 5.55 | 5.66 | 5.77 | 5.88 | 5.99 | 6.1 | 6.21 | 6.32 | 6.43 | 6.54 | 6.65 | 6.76 | 6.87 | 6.98 | 7.09 | 7.2 | 7.31 | 7.42 | 7.53 | 7.64 | 7.75 | 7.86 | 7.97 | 8.08 | 8.19 | 8.3 | 8.41 | 8.52
26.21 | 4.12 | 4.23 | 4.34 | 4.45 | 4.56 | 4.67 | 4.78 | 4.89 | 5 | 5.11 | 5.22 | 5.33 | 5.44 | 5.55 | 5.66 | 5.77 | 5.88 | 5.99 | 6.1 | 6.21 | 6.32 | 6.43 | 6.54 | 6.65 | 6.76 | 6.87 | 6.98 | 7.09 | 7.2 | 7.31 | 7.42 | 7.53 | 7.64 | 7.75 | 7.86 | 7.97 | 8.08 | 8.19 | 8.3 | 8.41 | 8.52
26.22 | 4.12 | 4.23 | 4.34 | 4.45 | 4.56 | 4.67 | 4.78 | 4.89 | 5 | 5.11 | 5.22 | 5.33 | 5.44 | 5.55 | 5.66 | 5.77 | 5.88 | 5.99 | 6.1 | 6.21 | 6.32 | 6.43 | 6.54 | 6.65 | 6.76 | 6.87 | 6.98 | 7.09 | 7.2 | 7.31 | 7.42 | 7.53 | 7.64 | 7.75 | 7.86 | 7.97 | 8.08 | 8.19 | 8.3 | 8.41 | 8.52
26.23 | 4.12 | 4.23 | 4.34 | 4.45 | 4.56 | 4.67 | 4.78 | 4.89 | 5 | 5.11 | 5.22 | 5.33 | 5.44 | 5.55 | 5.66 | 5.77 | 5.88 | 5.99 | 6.1 | 6.21 | 6.32 | 6.43 | 6.54 | 6.65 | 6.76 | 6.87 | 6.98 | 7.09 | 7.2 | 7.31 | 7.42 | 7.53 | 7.64 | 7.75 | 7.86 | 7.97 | 8.08 | 8.19 | 8.3 | 8.41 | 8.52
26.24 | 4.12 | 4.23 | 4.34 | 4.45 | 4.56 | 4.67 | 4.78 | 4.89 | 5 | 5.11 | 5.22 | 5.33 | 5.44 | 5.55 | 5.66 | 5.77 | 5.88 | 5.99 | 6.1 | 6.21 | 6.32 | 6.43 | 6.54 | 6.65 | 6.76 | 6.87 | 6.98 | 7.09 | 7.2 | 7.31 | 7.42 | 7.53 | 7.64 | 7.75 | 7.86 | 7.97 | 8.08 | 8.19 | 8.3 | 8.41 | 8.52
26.25 | 4.12 | 4.23 | 4.34 | 4.45 | 4.56 | 4.67 | 4.78 | 4.89 | 5 | 5.11 | 5.22 | 5.33 | 5.44 | 5.55 | 5.66 | 5.77 | 5.88 | 5.99 | 6.1 | 6.21 | 6.32 | 6.43 | 6.54 | 6.65 | 6.76 | 6.87 | 6.98 | 7.09 | 7.2 | 7.31 | 7.42 | 7.53 | 7.64 | 7.75 | 7.86 | 7.97 | 8.08 | 8.19 | 8.3 | 8.41 | 8.52
26.26 | 4.12 | 4.23 | 4.34 | 4.45 | 4.56 | 4.67 | 4.78 | 4.89 | 5 | 5.11 | 5.22 | 5.33 | 5.44 | 5.55 | 5.66 | 5.77 | 5.88 | 5.99 | 6.1 | 6.21 | 6.32 | 6.43 | 6.54 | 6.65 | 6.76 | 6.87 | 6.98 | 7.09 | 7.2 | 7.31 | 7.42 | 7.53 | 7.64 | 7.75 | 7.86 | 7.97 | 8.08 | 8.19 | 8.3 | 8.41 | 8.51
26.27 | 4.12 | 4.23 | 4.34 | 4.45 | 4.56 | 4.67 | 4.78 | 4.89 | 5 | 5.11 | 5.22 | 5.33 | 5.44 | 5.55 | 5.66 | 5.77 | 5.88 | 5.99 | 6.1 | 6.21 | 6.32 | 6.43 | 6.54 | 6.65 | 6.76 | 6.87 | 6.98 | 7.09 | 7.19 | 7.3 | 7.41 | 7.52 | 7.63 | 7.74 | 7.85 | 7.96 | 8.07 | 8.18 | 8.29 | 8.4 | 8.51
26.28 | 4.12 | 4.23 | 4.34 | 4.45 | 4.56 | 4.67 | 4.78 | 4.89 | 5 | 5.11 | 5.22 | 5.33 | 5.44 | 5.55 | 5.66 | 5.77 | 5.88 | 5.98 | 6.09 | 6.2 | 6.31 | 6.42 | 6.53 | 6.64 | 6.75 | 6.86 | 6.97 | 7.08 | 7.19 | 7.3 | 7.41 | 7.52 | 7.63 | 7.74 | 7.85 | 7.96 | 8.07 | 8.18 | 8.29 | 8.4 | 8.51
26.29 | 4.12 | 4.23 | 4.34 | 4.45 | 4.56 | 4.67 | 4.78 | 4.88 | 4.99 | 5.1 | 5.21 | 5.32 | 5.43 | 5.54 | 5.65 | 5.76 | 5.87 | 5.98 | 6.09 | 6.2 | 6.31 | 6.42 | 6.53 | 6.64 | 6.75 | 6.86 | 6.97 | 7.08 | 7.19 | 7.3 | 7.41 | 7.52 | 7.63 | 7.74 | 7.85 | 7.96 | 8.07 | 8.18 | 8.29 | 8.4 | 8.51
26.3 | 4.11 | 4.22 | 4.33 | 4.44 | 4.55 | 4.66 | 4.77 | 4.88 | 4.99 | 5.1 | 5.21 | 5.32 | 5.43 | 5.54 | 5.65 | 5.76 | 5.87 | 5.98 | 6.09 | 6.2 | 6.31 | 6.42 | 6.53 | 6.64 | 6.75 | 6.86 | 6.97 | 7.08 | 7.19 | 7.3 | 7.41 | 7.52 | 7.63 | 7.74 | 7.85 | 7.96 | 8.07 | 8.18 | 8.29 | 8.4 | 8.51
26.31 | 4.11 | 4.22 | 4.33 | 4.44 | 4.55 | 4.66 | 4.77 | 4.88 | 4.99 | 5.1 | 5.21 | 5.32 | 5.43 | 5.54 | 5.65 | 5.76 | 5.87 | 5.98 | 6.09 | 6.2 | 6.31 | 6.42 | 6.53 | 6.64 | 6.75 | 6.86 | 6.97 | 7.08 | 7.19 | 7.3 | 7.41 | 7.52 | 7.63 | 7.74 | 7.85 | 7.96 | 8.07 | 8.18 | 8.29 | 8.4 | 8.51
26.32 | 4.11 | 4.22 | 4.33 | 4.44 | 4.55 | 4.66 | 4.77 | 4.88 | 4.99 | 5.1 | 5.21 | 5.32 | 5.43 | 5.54 | 5.65 | 5.76 | 5.87 | 5.98 | 6.09 | 6.2 | 6.31 | 6.42 | 6.53 | 6.64 | 6.75 | 6.86 | 6.97 | 7.08 | 7.19 | 7.3 | 7.41 | 7.52 | 7.63 | 7.74 | 7.85 | 7.96 | 8.07 | 8.18 | 8.29 | 8.4 | 8.51
26.33 | 4.11 | 4.22 | 4.33 | 4.44 | 4.55 | 4.66 | 4.77 | 4.88 | 4.99 | 5.1 | 5.21 | 5.32 | 5.43 | 5.54 | 5.65 | 5.76 | 5.87 | 5.98 | 6.09 | 6.2 | 6.31 | 6.42 | 6.53 | 6.64 | 6.75 | 6.86 | 6.97 | 7.08 | 7.19 | 7.3 | 7.41 | 7.52 | 7.63 | 7.74 | 7.85 | 7.96 | 8.07 | 8.17 | 8.28 | 8.39 | 8.5
26.34 | 4.11 | 4.22 | 4.33 | 4.44 | 4.55 | 4.66 | 4.77 | 4.88 | 4.99 | 5.1 | 5.21 | 5.32 | 5.43 | 5.54 | 5.65 | 5.76 | 5.87 | 5.98 | 6.09 | 6.2 | 6.31 | 6.42 | 6.53 | 6.64 | 6.75 | 6.86 | 6.97 | 7.08 | 7.19 | 7.3 | 7.41 | 7.51 | 7.62 | 7.73 | 7.84 | 7.95 | 8.06 | 8.17 | 8.28 | 8.39 | 8.5
26.35 | 4.11 | 4.22 | 4.33 | 4.44 | 4.55 | 4.66 | 4.77 | 4.88 | 4.99 | 5.1 | 5.21 | 5.32 | 5.43 | 5.54 | 5.65 | 5.76 | 5.87 | 5.98 | 6.09 | 6.2 | 6.31 | 6.42 | 6.53 | 6.64 | 6.75 | 6.86 | 6.96 | 7.07 | 7.18 | 7.29 | 7.4 | 7.51 | 7.62 | 7.73 | 7.84 | 7.95 | 8.06 | 8.17 | 8.28 | 8.39 | 8.5
26.36 | 4.11 | 4.22 | 4.33 | 4.44 | 4.55 | 4.66 | 4.77 | 4.88 | 4.99 | 5.1 | 5.21 | 5.32 | 5.43 | 5.54 | 5.65 | 5.76 | 5.87 | 5.98 | 6.09 | 6.2 | 6.31 | 6.41 | 6.52 | 6.63 | 6.74 | 6.85 | 6.96 | 7.07 | 7.18 | 7.29 | 7.4 | 7.51 | 7.62 | 7.73 | 7.84 | 7.95 | 8.06 | 8.17 | 8.28 | 8.39 | 8.5
26.37 | 4.11 | 4.22 | 4.33 | 4.44 | 4.55 | 4.66 | 4.77 | 4.88 | 4.99 | 5.1 | 5.21 | 5.32 | 5.43 | 5.54 | 5.65 | 5.76 | 5.87 | 5.97 | 6.08 | 6.19 | 6.3 | 6.41 | 6.52 | 6.63 | 6.74 | 6.85 | 6.96 | 7.07 | 7.18 | 7.29 | 7.4 | 7.51 | 7.62 | 7.73 | 7.84 | 7.95 | 8.06 | 8.17 | 8.28 | 8.39 | 8.5
26.38 | 4.11 | 4.22 | 4.33 | 4.44 | 4.55 | 4.66 | 4.77 | 4.88 | 4.99 | 5.1 | 5.21 | 5.32 | 5.43 | 5.53 | 5.64 | 5.75 | 5.86 | 5.97 | 6.08 | 6.19 | 6.3 | 6.41 | 6.52 | 6.63 | 6.74 | 6.85 | 6.96 | 7.07 | 7.18 | 7.29 | 7.4 | 7.51 | 7.62 | 7.73 | 7.84 | 7.95 | 8.06 | 8.17 | 8.28 | 8.39 | 8.5
26.39 | 4.11 | 4.22 | 4.33 | 4.44 | 4.55 | 4.66 | 4.77 | 4.88 | 4.99 | 5.09 | 5.2 | 5.31 | 5.42 | 5.53 | 5.64 | 5.75 | 5.86 | 5.97 | 6.08 | 6.19 | 6.3 | 6.41 | 6.52 | 6.63 | 6.74 | 6.85 | 6.96 | 7.07 | 7.18 | 7.29 | 7.4 | 7.51 | 7.62 | 7.73 | 7.84 | 7.95 | 8.06 | 8.17 | 8.28 | 8.39 | 8.5
26.4 | 4.11 | 4.22 | 4.33 | 4.44 | 4.55 | 4.66 | 4.76 | 4.87 | 4.98 | 5.09 | 5.2 | 5.31 | 5.42 | 5.53 | 5.64 | 5.75 | 5.86 | 5.97 | 6.08 | 6.19 | 6.3 | 6.41 | 6.52 | 6.63 | 6.74 | 6.85 | 6.96 | 7.07 | 7.18 | 7.29 | 7.4 | 7.51 | 7.62 | 7.73 | 7.84 | 7.95 | 8.06 | 8.16 | 8.27 | 8.38 | 8.49
26.41 | 4.11 | 4.22 | 4.33 | 4.43 | 4.54 | 4.65 | 4.76 | 4.87 | 4.98 | 5.09 | 5.2 | 5.31 | 5.42 | 5.53 | 5.64 | 5.75 | 5.86 | 5.97 | 6.08 | 6.19 | 6.3 | 6.41 | 6.52 | 6.63 | 6.74 | 6.85 | 6.96 | 7.07 | 7.18 | 7.29 | 7.4 | 7.51 | 7.62 | 7.72 | 7.83 | 7.94 | 8.05 | 8.16 | 8.27 | 8.38 | 8.49
26.42 | 4.1 | 4.21 | 4.32 | 4.43 | 4.54 | 4.65 | 4.76 | 4.87 | 4.98 | 5.09 | 5.2 | 5.31 | 5.42 | 5.53 | 5.64 | 5.75 | 5.86 | 5.97 | 6.08 | 6.19 | 6.3 | 6.41 | 6.52 | 6.63 | 6.74 | 6.85 | 6.96 | 7.07 | 7.18 | 7.28 | 7.39 | 7.5 | 7.61 | 7.72 | 7.83 | 7.94 | 8.05 | 8.16 | 8.27 | 8.38 | 8.49
26.43 | 4.1 | 4.21 | 4.32 | 4.43 | 4.54 | 4.65 | 4.76 | 4.87 | 4.98 | 5.09 | 5.2 | 5.31 | 5.42 | 5.53 | 5.64 | 5.75 | 5.86 | 5.97 | 6.08 | 6.19 | 6.3 | 6.41 | 6.52 | 6.63 | 6.74 | 6.85 | 6.95 | 7.06 | 7.17 | 7.28 | 7.39 | 7.5 | 7.61 | 7.72 | 7.83 | 7.94 | 8.05 | 8.16 | 8.27 | 8.38 | 8.49
26.44 | 4.1 | 4.21 | 4.32 | 4.43 | 4.54 | 4.65 | 4.76 | 4.87 | 4.98 | 5.09 | 5.2 | 5.31 | 5.42 | 5.53 | 5.64 | 5.75 | 5.86 | 5.97 | 6.08 | 6.19 | 6.3 | 6.41 | 6.51 | 6.62 | 6.73 | 6.84 | 6.95 | 7.06 | 7.17 | 7.28 | 7.39 | 7.5 | 7.61 | 7.72 | 7.83 | 7.94 | 8.05 | 8.16 | 8.27 | 8.38 | 8.49
26.45 | 4.1 | 4.21 | 4.32 | 4.43 | 4.54 | 4.65 | 4.76 | 4.87 | 4.98 | 5.09 | 5.2 | 5.31 | 5.42 | 5.53 | 5.64 | 5.75 | 5.86 | 5.97 | 6.08 | 6.18 | 6.29 | 6.4 | 6.51 | 6.62 | 6.73 | 6.84 | 6.95 | 7.06 | 7.17 | 7.28 | 7.39 | 7.5 | 7.61 | 7.72 | 7.83 | 7.94 | 8.05 | 8.16 | 8.27 | 8.38 | 8.49
26.46 | 4.1 | 4.21 | 4.32 | 4.43 | 4.54 | 4.65 | 4.76 | 4.87 | 4.98 | 5.09 | 5.2 | 5.31 | 5.42 | 5.53 | 5.64 | 5.75 | 5.86 | 5.96 | 6.07 | 6.18 | 6.29 | 6.4 | 6.51 | 6.62 | 6.73 | 6.84 | 6.95 | 7.06 | 7.17 | 7.28 | 7.39 | 7.5 | 7.61 | 7.72 | 7.83 | 7.94 | 8.05 | 8.16 | 8.27 | 8.37 | 8.48
26.47 | 4.1 | 4.21 | 4.32 | 4.43 | 4.54 | 4.65 | 4.76 | 4.87 | 4.98 | 5.09 | 5.2 | 5.31 | 5.42 | 5.53 | 5.63 | 5.74 | 5.85 | 5.96 | 6.07 | 6.18 | 6.29 | 6.4 | 6.51 | 6.62 | 6.73 | 6.84 | 6.95 | 7.06 | 7.17 | 7.28 | 7.39 | 7.5 | 7.61 | 7.72 | 7.83 | 7.94 | 8.04 | 8.15 | 8.26 | 8.37 | 8.48
26.48 | 4.1 | 4.21 | 4.32 | 4.43 | 4.54 | 4.65 | 4.76 | 4.87 | 4.98 | 5.09 | 5.2 | 5.31 | 5.41 | 5.52 | 5.63 | 5.74 | 5.85 | 5.96 | 6.07 | 6.18 | 6.29 | 6.4 | 6.51 | 6.62 | 6.73 | 6.84 | 6.95 | 7.06 | 7.17 | 7.28 | 7.39 | 7.5 | 7.61 | 7.71 | 7.82 | 7.93 | 8.04 | 8.15 | 8.26 | 8.37 | 8.48
26.49 | 4.1 | 4.21 | 4.32 | 4.43 | 4.54 | 4.65 | 4.76 | 4.87 | 4.98 | 5.09 | 5.19 | 5.3 | 5.41 | 5.52 | 5.63 | 5.74 | 5.85 | 5.96 | 6.07 | 6.18 | 6.29 | 6.4 | 6.51 | 6.62 | 6.73 | 6.84 | 6.95 | 7.06 | 7.17 | 7.28 | 7.38 | 7.49 | 7.6 | 7.71 | 7.82 | 7.93 | 8.04 | 8.15 | 8.26 | 8.37 | 8.48
26.5 | 4.1 | 4.21 | 4.32 | 4.43 | 4.54 | 4.65 | 4.76 | 4.87 | 4.97 | 5.08 | 5.19 | 5.3 | 5.41 | 5.52 | 5.63 | 5.74 | 5.85 | 5.96 | 6.07 | 6.18 | 6.29 | 6.4 | 6.51 | 6.62 | 6.73 | 6.84 | 6.95 | 7.06 | 7.16 | 7.27 | 7.38 | 7.49 | 7.6 | 7.71 | 7.82 | 7.93 | 8.04 | 8.15 | 8.26 | 8.37 | 8.48
26.51 | 4.1 | 4.21 | 4.32 | 4.43 | 4.54 | 4.65 | 4.75 | 4.86 | 4.97 | 5.08 | 5.19 | 5.3 | 5.41 | 5.52 | 5.63 | 5.74 | 5.85 | 5.96 | 6.07 | 6.18 | 6.29 | 6.4 | 6.51 | 6.62 | 6.73 | 6.83 | 6.94 | 7.05 | 7.16 | 7.27 | 7.38 | 7.49 | 7.6 | 7.71 | 7.82 | 7.93 | 8.04 | 8.15 | 8.26 | 8.37 | 8.48
26.52 | 4.1 | 4.21 | 4.32 | 4.43 | 4.53 | 4.64 | 4.75 | 4.86 | 4.97 | 5.08 | 5.19 | 5.3 | 5.41 | 5.52 | 5.63 | 5.74 | 5.85 | 5.96 | 6.07 | 6.18 | 6.29 | 6.4 | 6.51 | 6.61 | 6.72 | 6.83 | 6.94 | 7.05 | 7.16 | 7.27 | 7.38 | 7.49 | 7.6 | 7.71 | 7.82 | 7.93 | 8.04 | 8.15 | 8.26 | 8.37 | 8.48
26.53 | 4.1 | 4.21 | 4.31 | 4.42 | 4.53 | 4.64 | 4.75 | 4.86 | 4.97 | 5.08 | 5.19 | 5.3 | 5.41 | 5.52 | 5.63 | 5.74 | 5.85 | 5.96 | 6.07 | 6.18 | 6.29 | 6.39 | 6.5 | 6.61 | 6.72 | 6.83 | 6.94 | 7.05 | 7.16 | 7.27 | 7.38 | 7.49 | 7.6 | 7.71 | 7.82 | 7.93 | 8.04 | 8.15 | 8.25 | 8.36 | 8.47
26.54 | 4.1 | 4.2 | 4.31 | 4.42 | 4.53 | 4.64 | 4.75 | 4.86 | 4.97 | 5.08 | 5.19 | 5.3 | 5.41 | 5.52 | 5.63 | 5.74 | 5.85 | 5.96 | 6.07 | 6.17 | 6.28 | 6.39 | 6.5 | 6.61 | 6.72 | 6.83 | 6.94 | 7.05 | 7.16 | 7.27 | 7.38 | 7.49 | 7.6 | 7.71 | 7.82 | 7.93 | 8.03 | 8.14 | 8.25 | 8.36 | 8.47
26.55 | 4.09 | 4.2 | 4.31 | 4.42 | 4.53 | 4.64 | 4.75 | 4.86 | 4.97 | 5.08 | 5.19 | 5.3 | 5.41 | 5.52 | 5.63 | 5.74 | 5.85 | 5.95 | 6.06 | 6.17 | 6.28 | 6.39 | 6.5 | 6.61 | 6.72 | 6.83 | 6.94 | 7.05 | 7.16 | 7.27 | 7.38 | 7.49 | 7.6 | 7.71 | 7.81 | 7.92 | 8.03 | 8.14 | 8.25 | 8.36 | 8.47
26.56 | 4.09 | 4.2 | 4.31 | 4.42 | 4.53 | 4.64 | 4.75 | 4.86 | 4.97 | 5.08 | 5.19 | 5.3 | 5.41 | 5.52 | 5.63 | 5.73 | 5.84 | 5.95 | 6.06 | 6.17 | 6.28 | 6.39 | 6.5 | 6.61 | 6.72 | 6.83 | 6.94 | 7.05 | 7.16 | 7.27 | 7.38 | 7.48 | 7.59 | 7.7 | 7.81 | 7.92 | 8.03 | 8.14 | 8.25 | 8.36 | 8.47
26.57 | 4.09 | 4.2 | 4.31 | 4.42 | 4.53 | 4.64 | 4.75 | 4.86 | 4.97 | 5.08 | 5.19 | 5.3 | 5.41 | 5.51 | 5.62 | 5.73 | 5.84 | 5.95 | 6.06 | 6.17 | 6.28 | 6.39 | 6.5 | 6.61 | 6.72 | 6.83 | 6.94 | 7.05 | 7.16 | 7.26 | 7.37 | 7.48 | 7.59 | 7.7 | 7.81 | 7.92 | 8.03 | 8.14 | 8.25 | 8.36 | 8.47
26.58 | 4.09 | 4.2 | 4.31 | 4.42 | 4.53 | 4.64 | 4.75 | 4.86 | 4.97 | 5.08 | 5.19 | 5.3 | 5.4 | 5.51 | 5.62 | 5.73 | 5.84 | 5.95 | 6.06 | 6.17 | 6.28 | 6.39 | 6.5 | 6.61 | 6.72 | 6.83 | 6.94 | 7.04 | 7.15 | 7.26 | 7.37 | 7.48 | 7.59 | 7.7 | 7.81 | 7.92 | 8.03 | 8.14 | 8.25 | 8.36 | 8.47
26.59 | 4.09 | 4.2 | 4.31 | 4.42 | 4.53 | 4.64 | 4.75 | 4.86 | 4.97 | 5.08 | 5.18 | 5.29 | 5.4 | 5.51 | 5.62 | 5.73 | 5.84 | 5.95 | 6.06 | 6.17 | 6.28 | 6.39 | 6.5 | 6.61 | 6.72 | 6.82 | 6.93 | 7.04 | 7.15 | 7.26 | 7.37 | 7.48 | 7.59 | 7.7 | 7.81 | 7.92 | 8.03 | 8.14 | 8.25 | 8.36 | 8.46
26.6 | 4.09 | 4.2 | 4.31 | 4.42 | 4.53 | 4.64 | 4.75 | 4.86 | 4.96 | 5.07 | 5.18 | 5.29 | 5.4 | 5.51 | 5.62 | 5.73 | 5.84 | 5.95 | 6.06 | 6.17 | 6.28 | 6.39 | 6.5 | 6.6 | 6.71 | 6.82 | 6.93 | 7.04 | 7.15 | 7.26 | 7.37 | 7.48 | 7.59 | 7.7 | 7.81 | 7.92 | 8.03 | 8.14 | 8.24 | 8.35 | 8.46
26.61 | 4.09 | 4.2 | 4.31 | 4.42 | 4.53 | 4.64 | 4.75 | 4.85 | 4.96 | 5.07 | 5.18 | 5.29 | 5.4 | 5.51 | 5.62 | 5.73 | 5.84 | 5.95 | 6.06 | 6.17 | 6.28 | 6.39 | 6.49 | 6.6 | 6.71 | 6.82 | 6.93 | 7.04 | 7.15 | 7.26 | 7.37 | 7.48 | 7.59 | 7.7 | 7.81 | 7.92 | 8.02 | 8.13 | 8.24 | 8.35 | 8.46
26.62 | 4.09 | 4.2 | 4.31 | 4.42 | 4.53 | 4.64 | 4.74 | 4.85 | 4.96 | 5.07 | 5.18 | 5.29 | 5.4 | 5.51 | 5.62 | 5.73 | 5.84 | 5.95 | 6.06 | 6.17 | 6.27 | 6.38 | 6.49 | 6.6 | 6.71 | 6.82 | 6.93 | 7.04 | 7.15 | 7.26 | 7.37 | 7.48 | 7.59 | 7.7 | 7.8 | 7.91 | 8.02 | 8.13 | 8.24 | 8.35 | 8.46
26.63 | 4.09 | 4.2 | 4.31 | 4.42 | 4.52 | 4.63 | 4.74 | 4.85 | 4.96 | 5.07 | 5.18 | 5.29 | 5.4 | 5.51 | 5.62 | 5.73 | 5.84 | 5.95 | 6.05 | 6.16 | 6.27 | 6.38 | 6.49 | 6.6 | 6.71 | 6.82 | 6.93 | 7.04 | 7.15 | 7.26 | 7.37 | 7.48 | 7.58 | 7.69 | 7.8 | 7.91 | 8.02 | 8.13 | 8.24 | 8.35 | 8.46
26.64 | 4.09 | 4.2 | 4.31 | 4.41 | 4.52 | 4.63 | 4.74 | 4.85 | 4.96 | 5.07 | 5.18 | 5.29 | 5.4 | 5.51 | 5.62 | 5.73 | 5.84 | 5.94 | 6.05 | 6.16 | 6.27 | 6.38 | 6.49 | 6.6 | 6.71 | 6.82 | 6.93 | 7.04 | 7.15 | 7.26 | 7.36 | 7.47 | 7.58 | 7.69 | 7.8 | 7.91 | 8.02 | 8.13 | 8.24 | 8.35 | 8.46
26.65 | 4.09 | 4.2 | 4.3 | 4.41 | 4.52 | 4.63 | 4.74 | 4.85 | 4.96 | 5.07 | 5.18 | 5.29 | 5.4 | 5.51 | 5.62 | 5.72 | 5.83 | 5.94 | 6.05 | 6.16 | 6.27 | 6.38 | 6.49 | 6.6 | 6.71 | 6.82 | 6.93 | 7.04 | 7.15 | 7.25 | 7.36 | 7.47 | 7.58 | 7.69 | 7.8 | 7.91 | 8.02 | 8.13 | 8.24 | 8.35 | 8.46
26.66 | 4.09 | 4.19 | 4.3 | 4.41 | 4.52 | 4.63 | 4.74 | 4.85 | 4.96 | 5.07 | 5.18 | 5.29 | 5.4 | 5.51 | 5.61 | 5.72 | 5.83 | 5.94 | 6.05 | 6.16 | 6.27 | 6.38 | 6.49 | 6.6 | 6.71 | 6.82 | 6.93 | 7.03 | 7.14 | 7.25 | 7.36 | 7.47 | 7.58 | 7.69 | 7.8 | 7.91 | 8.02 | 8.13 | 8.24 | 8.34 | 8.45
26.67 | 4.08 | 4.19 | 4.3 | 4.41 | 4.52 | 4.63 | 4.74 | 4.85 | 4.96 | 5.07 | 5.18 | 5.29 | 5.4 | 5.5 | 5.61 | 5.72 | 5.83 | 5.94 | 6.05 | 6.16 | 6.27 | 6.38 | 6.49 | 6.6 | 6.71 | 6.81 | 6.92 | 7.03 | 7.14 | 7.25 | 7.36 | 7.47 | 7.58 | 7.69 | 7.8 | 7.91 | 8.02 | 8.13 | 8.23 | 8.34 | 8.45
26.68 | 4.08 | 4.19 | 4.3 | 4.41 | 4.52 | 4.63 | 4.74 | 4.85 | 4.96 | 5.07 | 5.18 | 5.28 | 5.39 | 5.5 | 5.61 | 5.72 | 5.83 | 5.94 | 6.05 | 6.16 | 6.27 | 6.38 | 6.49 | 6.6 | 6.7 | 6.81 | 6.92 | 7.03 | 7.14 | 7.25 | 7.36 | 7.47 | 7.58 | 7.69 | 7.8 | 7.91 | 8.01 | 8.12 | 8.23 | 8.34 | 8.45
26.69 | 4.08 | 4.19 | 4.3 | 4.41 | 4.52 | 4.63 | 4.74 | 4.85 | 4.96 | 5.07 | 5.17 | 5.28 | 5.39 | 5.5 | 5.61 | 5.72 | 5.83 | 5.94 | 6.05 | 6.16 | 6.27 | 6.38 | 6.48 | 6.59 | 6.7 | 6.81 | 6.92 | 7.03 | 7.14 | 7.25 | 7.36 | 7.47 | 7.58 | 7.69 | 7.79 | 7.9 | 8.01 | 8.12 | 8.23 | 8.34 | 8.45
26.7 | 4.08 | 4.19 | 4.3 | 4.41 | 4.52 | 4.63 | 4.74 | 4.85 | 4.96 | 5.06 | 5.17 | 5.28 | 5.39 | 5.5 | 5.61 | 5.72 | 5.83 | 5.94 | 6.05 | 6.16 | 6.27 | 6.37 | 6.48 | 6.59 | 6.7 | 6.81 | 6.92 | 7.03 | 7.14 | 7.25 | 7.36 | 7.47 | 7.58 | 7.68 | 7.79 | 7.9 | 8.01 | 8.12 | 8.23 | 8.34 | 8.45
26.71 | 4.08 | 4.19 | 4.3 | 4.41 | 4.52 | 4.63 | 4.74 | 4.85 | 4.95 | 5.06 | 5.17 | 5.28 | 5.39 | 5.5 | 5.61 | 5.72 | 5.83 | 5.94 | 6.05 | 6.16 | 6.26 | 6.37 | 6.48 | 6.59 | 6.7 | 6.81 | 6.92 | 7.03 | 7.14 | 7.25 | 7.36 | 7.46 | 7.57 | 7.68 | 7.79 | 7.9 | 8.01 | 8.12 | 8.23 | 8.34 | 8.45
26.72 | 4.08 | 4.19 | 4.3 | 4.41 | 4.52 | 4.63 | 4.74 | 4.84 | 4.95 | 5.06 | 5.17 | 5.28 | 5.39 | 5.5 | 5.61 | 5.72 | 5.83 | 5.94 | 6.04 | 6.15 | 6.26 | 6.37 | 6.48 | 6.59 | 6.7 | 6.81 | 6.92 | 7.03 | 7.14 | 7.25 | 7.35 | 7.46 | 7.57 | 7.68 | 7.79 | 7.9 | 8.01 | 8.12 | 8.23 | 8.34 | 8.45
26.73 | 4.08 | 4.19 | 4.3 | 4.41 | 4.52 | 4.63 | 4.73 | 4.84 | 4.95 | 5.06 | 5.17 | 5.28 | 5.39 | 5.5 | 5.61 | 5.72 | 5.83 | 5.93 | 6.04 | 6.15 | 6.26 | 6.37 | 6.48 | 6.59 | 6.7 | 6.81 | 6.92 | 7.03 | 7.13 | 7.24 | 7.35 | 7.46 | 7.57 | 7.68 | 7.79 | 7.9 | 8.01 | 8.12 | 8.23 | 8.33 | 8.44
26.74 | 4.08 | 4.19 | 4.3 | 4.41 | 4.52 | 4.62 | 4.73 | 4.84 | 4.95 | 5.06 | 5.17 | 5.28 | 5.39 | 5.5 | 5.61 | 5.72 | 5.82 | 5.93 | 6.04 | 6.15 | 6.26 | 6.37 | 6.48 | 6.59 | 6.7 | 6.81 | 6.92 | 7.02 | 7.13 | 7.24 | 7.35 | 7.46 | 7.57 | 7.68 | 7.79 | 7.9 | 8.01 | 8.11 | 8.22 | 8.33 | 8.44
26.75 | 4.08 | 4.19 | 4.3 | 4.41 | 4.51 | 4.62 | 4.73 | 4.84 | 4.95 | 5.06 | 5.17 | 5.28 | 5.39 | 5.5 | 5.61 | 5.71 | 5.82 | 5.93 | 6.04 | 6.15 | 6.26 | 6.37 | 6.48 | 6.59 | 6.7 | 6.8 | 6.91 | 7.02 | 7.13 | 7.24 | 7.35 | 7.46 | 7.57 | 7.68 | 7.79 | 7.9 | 8 | 8.11 | 8.22 | 8.33 | 8.44
26.76 | 4.08 | 4.19 | 4.3 | 4.4 | 4.51 | 4.62 | 4.73 | 4.84 | 4.95 | 5.06 | 5.17 | 5.28 | 5.39 | 5.49 | 5.6 | 5.71 | 5.82 | 5.93 | 6.04 | 6.15 | 6.26 | 6.37 | 6.48 | 6.59 | 6.69 | 6.8 | 6.91 | 7.02 | 7.13 | 7.24 | 7.35 | 7.46 | 7.57 | 7.68 | 7.78 | 7.89 | 8 | 8.11 | 8.22 | 8.33 | 8.44
26.77 | 4.08 | 4.19 | 4.29 | 4.4 | 4.51 | 4.62 | 4.73 | 4.84 | 4.95 | 5.06 | 5.17 | 5.28 | 5.38 | 5.49 | 5.6 | 5.71 | 5.82 | 5.93 | 6.04 | 6.15 | 6.26 | 6.37 | 6.48 | 6.58 | 6.69 | 6.8 | 6.91 | 7.02 | 7.13 | 7.24 | 7.35 | 7.46 | 7.57 | 7.67 | 7.78 | 7.89 | 8 | 8.11 | 8.22 | 8.33 | 8.44
26.78 | 4.08 | 4.18 | 4.29 | 4.4 | 4.51 | 4.62 | 4.73 | 4.84 | 4.95 | 5.06 | 5.17 | 5.27 | 5.38 | 5.49 | 5.6 | 5.71 | 5.82 | 5.93 | 6.04 | 6.15 | 6.26 | 6.37 | 6.47 | 6.58 | 6.69 | 6.8 | 6.91 | 7.02 | 7.13 | 7.24 | 7.35 | 7.46 | 7.56 | 7.67 | 7.78 | 7.89 | 8 | 8.11 | 8.22 | 8.33 | 8.44
26.79 | 4.07 | 4.18 | 4.29 | 4.4 | 4.51 | 4.62 | 4.73 | 4.84 | 4.95 | 5.06 | 5.16 | 5.27 | 5.38 | 5.49 | 5.6 | 5.71 | 5.82 | 5.93 | 6.04 | 6.15 | 6.25 | 6.36 | 6.47 | 6.58 | 6.69 | 6.8 | 6.91 | 7.02 | 7.13 | 7.24 | 7.34 | 7.45 | 7.56 | 7.67 | 7.78 | 7.89 | 8 | 8.11 | 8.22 | 8.33 | 8.43
26.8 | 4.07 | 4.18 | 4.29 | 4.4 | 4.51 | 4.62 | 4.73 | 4.84 | 4.95 | 5.05 | 5.16 | 5.27 | 5.38 | 5.49 | 5.6 | 5.71 | 5.82 | 5.93 | 6.04 | 6.14 | 6.25 | 6.36 | 6.47 | 6.58 | 6.69 | 6.8 | 6.91 | 7.02 | 7.13 | 7.23 | 7.34 | 7.45 | 7.56 | 7.67 | 7.78 | 7.89 | 8 | 8.11 | 8.22 | 8.32 | 8.43
26.81 | 4.07 | 4.18 | 4.29 | 4.4 | 4.51 | 4.62 | 4.73 | 4.84 | 4.94 | 5.05 | 5.16 | 5.27 | 5.38 | 5.49 | 5.6 | 5.71 | 5.82 | 5.93 | 6.03 | 6.14 | 6.25 | 6.36 | 6.47 | 6.58 | 6.69 | 6.8 | 6.91 | 7.02 | 7.12 | 7.23 | 7.34 | 7.45 | 7.56 | 7.67 | 7.78 | 7.89 | 8 | 8.1 | 8.21 | 8.32 | 8.43
26.82 | 4.07 | 4.18 | 4.29 | 4.4 | 4.51 | 4.62 | 4.73 | 4.83 | 4.94 | 5.05 | 5.16 | 5.27 | 5.38 | 5.49 | 5.6 | 5.71 | 5.82 | 5.92 | 6.03 | 6.14 | 6.25 | 6.36 | 6.47 | 6.58 | 6.69 | 6.8 | 6.91 | 7.01 | 7.12 | 7.23 | 7.34 | 7.45 | 7.56 | 7.67 | 7.78 | 7.89 | 7.99 | 8.1 | 8.21 | 8.32 | 8.43
26.83 | 4.07 | 4.18 | 4.29 | 4.4 | 4.51 | 4.62 | 4.73 | 4.83 | 4.94 | 5.05 | 5.16 | 5.27 | 5.38 | 5.49 | 5.6 | 5.71 | 5.81 | 5.92 | 6.03 | 6.14 | 6.25 | 6.36 | 6.47 | 6.58 | 6.69 | 6.79 | 6.9 | 7.01 | 7.12 | 7.23 | 7.34 | 7.45 | 7.56 | 7.67 | 7.78 | 7.88 | 7.99 | 8.1 | 8.21 | 8.32 | 8.43
26.84 | 4.07 | 4.18 | 4.29 | 4.4 | 4.51 | 4.62 | 4.72 | 4.83 | 4.94 | 5.05 | 5.16 | 5.27 | 5.38 | 5.49 | 5.6 | 5.7 | 5.81 | 5.92 | 6.03 | 6.14 | 6.25 | 6.36 | 6.47 | 6.58 | 6.68 | 6.79 | 6.9 | 7.01 | 7.12 | 7.23 | 7.34 | 7.45 | 7.56 | 7.66 | 7.77 | 7.88 | 7.99 | 8.1 | 8.21 | 8.32 | 8.43
26.85 | 4.07 | 4.18 | 4.29 | 4.4 | 4.51 | 4.61 | 4.72 | 4.83 | 4.94 | 5.05 | 5.16 | 5.27 | 5.38 | 5.49 | 5.59 | 5.7 | 5.81 | 5.92 | 6.03 | 6.14 | 6.25 | 6.36 | 6.47 | 6.57 | 6.68 | 6.79 | 6.9 | 7.01 | 7.12 | 7.23 | 7.34 | 7.45 | 7.55 | 7.66 | 7.77 | 7.88 | 7.99 | 8.1 | 8.21 | 8.32 | 8.43
26.86 | 4.07 | 4.18 | 4.29 | 4.4 | 4.5 | 4.61 | 4.72 | 4.83 | 4.94 | 5.05 | 5.16 | 5.27 | 5.38 | 5.48 | 5.59 | 5.7 | 5.81 | 5.92 | 6.03 | 6.14 | 6.25 | 6.36 | 6.46 | 6.57 | 6.68 | 6.79 | 6.9 | 7.01 | 7.12 | 7.23 | 7.34 | 7.44 | 7.55 | 7.66 | 7.77 | 7.88 | 7.99 | 8.1 | 8.21 | 8.32 | 8.42
26.87 | 4.07 | 4.18 | 4.29 | 4.39 | 4.5 | 4.61 | 4.72 | 4.83 | 4.94 | 5.05 | 5.16 | 5.27 | 5.37 | 5.48 | 5.59 | 5.7 | 5.81 | 5.92 | 6.03 | 6.14 | 6.25 | 6.35 | 6.46 | 6.57 | 6.68 | 6.79 | 6.9 | 7.01 | 7.12 | 7.23 | 7.33 | 7.44 | 7.55 | 7.66 | 7.77 | 7.88 | 7.99 | 8.1 | 8.2 | 8.31 | 8.42
26.88 | 4.07 | 4.18 | 4.29 | 4.39 | 4.5 | 4.61 | 4.72 | 4.83 | 4.94 | 5.05 | 5.16 | 5.26 | 5.37 | 5.48 | 5.59 | 5.7 | 5.81 | 5.92 | 6.03 | 6.14 | 6.24 | 6.35 | 6.46 | 6.57 | 6.68 | 6.79 | 6.9 | 7.01 | 7.12 | 7.22 | 7.33 | 7.44 | 7.55 | 7.66 | 7.77 | 7.88 | 7.99 | 8.09 | 8.2 | 8.31 | 8.42
26.89 | 4.07 | 4.18 | 4.28 | 4.39 | 4.5 | 4.61 | 4.72 | 4.83 | 4.94 | 5.05 | 5.15 | 5.26 | 5.37 | 5.48 | 5.59 | 5.7 | 5.81 | 5.92 | 6.03 | 6.13 | 6.24 | 6.35 | 6.46 | 6.57 | 6.68 | 6.79 | 6.9 | 7.01 | 7.11 | 7.22 | 7.33 | 7.44 | 7.55 | 7.66 | 7.77 | 7.88 | 7.98 | 8.09 | 8.2 | 8.31 | 8.42
26.9 | 4.07 | 4.17 | 4.28 | 4.39 | 4.5 | 4.61 | 4.72 | 4.83 | 4.94 | 5.05 | 5.15 | 5.26 | 5.37 | 5.48 | 5.59 | 5.7 | 5.81 | 5.92 | 6.02 | 6.13 | 6.24 | 6.35 | 6.46 | 6.57 | 6.68 | 6.79 | 6.89 | 7 | 7.11 | 7.22 | 7.33 | 7.44 | 7.55 | 7.66 | 7.77 | 7.87 | 7.98 | 8.09 | 8.2 | 8.31 | 8.42
26.91 | 4.06 | 4.17 | 4.28 | 4.39 | 4.5 | 4.61 | 4.72 | 4.83 | 4.94 | 5.04 | 5.15 | 5.26 | 5.37 | 5.48 | 5.59 | 5.7 | 5.81 | 5.91 | 6.02 | 6.13 | 6.24 | 6.35 | 6.46 | 6.57 | 6.68 | 6.78 | 6.89 | 7 | 7.11 | 7.22 | 7.33 | 7.44 | 7.55 | 7.66 | 7.76 | 7.87 | 7.98 | 8.09 | 8.2 | 8.31 | 8.42
26.92 | 4.06 | 4.17 | 4.28 | 4.39 | 4.5 | 4.61 | 4.72 | 4.83 | 4.93 | 5.04 | 5.15 | 5.26 | 5.37 | 5.48 | 5.59 | 5.7 | 5.8 | 5.91 | 6.02 | 6.13 | 6.24 | 6.35 | 6.46 | 6.57 | 6.67 | 6.78 | 6.89 | 7 | 7.11 | 7.22 | 7.33 | 7.44 | 7.54 | 7.65 | 7.76 | 7.87 | 7.98 | 8.09 | 8.2 | 8.31 | 8.41
26.93 | 4.06 | 4.17 | 4.28 | 4.39 | 4.5 | 4.61 | 4.72 | 4.82 | 4.93 | 5.04 | 5.15 | 5.26 | 5.37 | 5.48 | 5.59 | 5.69 | 5.8 | 5.91 | 6.02 | 6.13 | 6.24 | 6.35 | 6.46 | 6.56 | 6.67 | 6.78 | 6.89 | 7 | 7.11 | 7.22 | 7.33 | 7.43 | 7.54 | 7.65 | 7.76 | 7.87 | 7.98 | 8.09 | 8.2 | 8.3 | 8.41
26.94 | 4.06 | 4.17 | 4.28 | 4.39 | 4.5 | 4.61 | 4.71 | 4.82 | 4.93 | 5.04 | 5.15 | 5.26 | 5.37 | 5.48 | 5.59 | 5.69 | 5.8 | 5.91 | 6.02 | 6.13 | 6.24 | 6.35 | 6.45 | 6.56 | 6.67 | 6.78 | 6.89 | 7 | 7.11 | 7.22 | 7.32 | 7.43 | 7.54 | 7.65 | 7.76 | 7.87 | 7.98 | 8.09 | 8.19 | 8.3 | 8.41
26.95 | 4.06 | 4.17 | 4.28 | 4.39 | 4.5 | 4.61 | 4.71 | 4.82 | 4.93 | 5.04 | 5.15 | 5.26 | 5.37 | 5.48 | 5.58 | 5.69 | 5.8 | 5.91 | 6.02 | 6.13 | 6.24 | 6.35 | 6.45 | 6.56 | 6.67 | 6.78 | 6.89 | 7 | 7.11 | 7.21 | 7.32 | 7.43 | 7.54 | 7.65 | 7.76 | 7.87 | 7.98 | 8.08 | 8.19 | 8.3 | 8.41
26.96 | 4.06 | 4.17 | 4.28 | 4.39 | 4.5 | 4.6 | 4.71 | 4.82 | 4.93 | 5.04 | 5.15 | 5.26 | 5.37 | 5.47 | 5.58 | 5.69 | 5.8 | 5.91 | 6.02 | 6.13 | 6.24 | 6.34 | 6.45 | 6.56 | 6.67 | 6.78 | 6.89 | 7 | 7.1 | 7.21 | 7.32 | 7.43 | 7.54 | 7.65 | 7.76 | 7.87 | 7.97 | 8.08 | 8.19 | 8.3 | 8.41
26.97 | 4.06 | 4.17 | 4.28 | 4.39 | 4.49 | 4.6 | 4.71 | 4.82 | 4.93 | 5.04 | 5.15 | 5.26 | 5.36 | 5.47 | 5.58 | 5.69 | 5.8 | 5.91 | 6.02 | 6.13 | 6.23 | 6.34 | 6.45 | 6.56 | 6.67 | 6.78 | 6.89 | 6.99 | 7.1 | 7.21 | 7.32 | 7.43 | 7.54 | 7.65 | 7.76 | 7.86 | 7.97 | 8.08 | 8.19 | 8.3 | 8.41
26.98 | 4.06 | 4.17 | 4.28 | 4.39 | 4.49 | 4.6 | 4.71 | 4.82 | 4.93 | 5.04 | 5.15 | 5.25 | 5.36 | 5.47 | 5.58 | 5.69 | 5.8 | 5.91 | 6.02 | 6.12 | 6.23 | 6.34 | 6.45 | 6.56 | 6.67 | 6.78 | 6.88 | 6.99 | 7.1 | 7.21 | 7.32 | 7.43 | 7.54 | 7.65 | 7.75 | 7.86 | 7.97 | 8.08 | 8.19 | 8.3 | 8.41
26.99 | 4.06 | 4.17 | 4.28 | 4.38 | 4.49 | 4.6 | 4.71 | 4.82 | 4.93 | 5.04 | 5.15 | 5.25 | 5.36 | 5.47 | 5.58 | 5.69 | 5.8 | 5.91 | 6.01 | 6.12 | 6.23 | 6.34 | 6.45 | 6.56 | 6.67 | 6.77 | 6.88 | 6.99 | 7.1 | 7.21 | 7.32 | 7.43 | 7.54 | 7.64 | 7.75 | 7.86 | 7.97 | 8.08 | 8.19 | 8.3 | 8.4
27 | 4.06 | 4.17 | 4.27 | 4.38 | 4.49 | 4.6 | 4.71 | 4.82 | 4.93 | 5.04 | 5.14 | 5.25 | 5.36 | 5.47 | 5.58 | 5.69 | 5.8 | 5.9 | 6.01 | 6.12 | 6.23 | 6.34 | 6.45 | 6.56 | 6.67 | 6.77 | 6.88 | 6.99 | 7.1 | 7.21 | 7.32 | 7.43 | 7.53 | 7.64 | 7.75 | 7.86 | 7.97 | 8.08 | 8.19 | 8.29 | 8.4
27.01 | 4.06 | 4.17 | 4.27 | 4.38 | 4.49 | 4.6 | 4.71 | 4.82 | 4.93 | 5.03 | 5.14 | 5.25 | 5.36 | 5.47 | 5.58 | 5.69 | 5.79 | 5.9 | 6.01 | 6.12 | 6.23 | 6.34 | 6.45 | 6.56 | 6.66 | 6.77 | 6.88 | 6.99 | 7.1 | 7.21 | 7.32 | 7.42 | 7.53 | 7.64 | 7.75 | 7.86 | 7.97 | 8.08 | 8.18 | 8.29 | 8.4
27.02 | 4.06 | 4.16 | 4.27 | 4.38 | 4.49 | 4.6 | 4.71 | 4.82 | 4.92 | 5.03 | 5.14 | 5.25 | 5.36 | 5.47 | 5.58 | 5.69 | 5.79 | 5.9 | 6.01 | 6.12 | 6.23 | 6.34 | 6.45 | 6.55 | 6.66 | 6.77 | 6.88 | 6.99 | 7.1 | 7.21 | 7.31 | 7.42 | 7.53 | 7.64 | 7.75 | 7.86 | 7.97 | 8.07 | 8.18 | 8.29 | 8.4
27.03 | 4.06 | 4.16 | 4.27 | 4.38 | 4.49 | 4.6 | 4.71 | 4.82 | 4.92 | 5.03 | 5.14 | 5.25 | 5.36 | 5.47 | 5.58 | 5.68 | 5.79 | 5.9 | 6.01 | 6.12 | 6.23 | 6.34 | 6.44 | 6.55 | 6.66 | 6.77 | 6.88 | 6.99 | 7.1 | 7.2 | 7.31 | 7.42 | 7.53 | 7.64 | 7.75 | 7.86 | 7.96 | 8.07 | 8.18 | 8.29 | 8.4
27.04 | 4.05 | 4.16 | 4.27 | 4.38 | 4.49 | 4.6 | 4.71 | 4.81 | 4.92 | 5.03 | 5.14 | 5.25 | 5.36 | 5.47 | 5.57 | 5.68 | 5.79 | 5.9 | 6.01 | 6.12 | 6.23 | 6.33 | 6.44 | 6.55 | 6.66 | 6.77 | 6.88 | 6.99 | 7.09 | 7.2 | 7.31 | 7.42 | 7.53 | 7.64 | 7.75 | 7.85 | 7.96 | 8.07 | 8.18 | 8.29 | 8.4
27.05 | 4.05 | 4.16 | 4.27 | 4.38 | 4.49 | 4.6 | 4.7 | 4.81 | 4.92 | 5.03 | 5.14 | 5.25 | 5.36 | 5.46 | 5.57 | 5.68 | 5.79 | 5.9 | 6.01 | 6.12 | 6.22 | 6.33 | 6.44 | 6.55 | 6.66 | 6.77 | 6.88 | 6.98 | 7.09 | 7.2 | 7.31 | 7.42 | 7.53 | 7.64 | 7.74 | 7.85 | 7.96 | 8.07 | 8.18 | 8.29 | 8.4
27.06 | 4.05 | 4.16 | 4.27 | 4.38 | 4.49 | 4.6 | 4.7 | 4.81 | 4.92 | 5.03 | 5.14 | 5.25 | 5.36 | 5.46 | 5.57 | 5.68 | 5.79 | 5.9 | 6.01 | 6.12 | 6.22 | 6.33 | 6.44 | 6.55 | 6.66 | 6.77 | 6.87 | 6.98 | 7.09 | 7.2 | 7.31 | 7.42 | 7.53 | 7.63 | 7.74 | 7.85 | 7.96 | 8.07 | 8.18 | 8.29 | 8.39
27.07 | 4.05 | 4.16 | 4.27 | 4.38 | 4.49 | 4.59 | 4.7 | 4.81 | 4.92 | 5.03 | 5.14 | 5.25 | 5.35 | 5.46 | 5.57 | 5.68 | 5.79 | 5.9 | 6.01 | 6.11 | 6.22 | 6.33 | 6.44 | 6.55 | 6.66 | 6.77 | 6.87 | 6.98 | 7.09 | 7.2 | 7.31 | 7.42 | 7.52 | 7.63 | 7.74 | 7.85 | 7.96 | 8.07 | 8.18 | 8.28 | 8.39
27.08 | 4.05 | 4.16 | 4.27 | 4.38 | 4.49 | 4.59 | 4.7 | 4.81 | 4.92 | 5.03 | 5.14 | 5.24 | 5.35 | 5.46 | 5.57 | 5.68 | 5.79 | 5.9 | 6 | 6.11 | 6.22 | 6.33 | 6.44 | 6.55 | 6.66 | 6.76 | 6.87 | 6.98 | 7.09 | 7.2 | 7.31 | 7.41 | 7.52 | 7.63 | 7.74 | 7.85 | 7.96 | 8.07 | 8.17 | 8.28 | 8.39
27.09 | 4.05 | 4.16 | 4.27 | 4.38 | 4.48 | 4.59 | 4.7 | 4.81 | 4.92 | 5.03 | 5.14 | 5.24 | 5.35 | 5.46 | 5.57 | 5.68 | 5.79 | 5.89 | 6 | 6.11 | 6.22 | 6.33 | 6.44 | 6.55 | 6.65 | 6.76 | 6.87 | 6.98 | 7.09 | 7.2 | 7.3 | 7.41 | 7.52 | 7.63 | 7.74 | 7.85 | 7.96 | 8.06 | 8.17 | 8.28 | 8.39
27.1 | 4.05 | 4.16 | 4.27 | 4.37 | 4.48 | 4.59 | 4.7 | 4.81 | 4.92 | 5.03 | 5.13 | 5.24 | 5.35 | 5.46 | 5.57 | 5.68 | 5.79 | 5.89 | 6 | 6.11 | 6.22 | 6.33 | 6.44 | 6.54 | 6.65 | 6.76 | 6.87 | 6.98 | 7.09 | 7.2 | 7.3 | 7.41 | 7.52 | 7.63 | 7.74 | 7.85 | 7.95 | 8.06 | 8.17 | 8.28 | 8.39
27.11 | 4.05 | 4.16 | 4.27 | 4.37 | 4.48 | 4.59 | 4.7 | 4.81 | 4.92 | 5.02 | 5.13 | 5.24 | 5.35 | 5.46 | 5.57 | 5.68 | 5.78 | 5.89 | 6 | 6.11 | 6.22 | 6.33 | 6.43 | 6.54 | 6.65 | 6.76 | 6.87 | 6.98 | 7.09 | 7.19 | 7.3 | 7.41 | 7.52 | 7.63 | 7.74 | 7.84 | 7.95 | 8.06 | 8.17 | 8.28 | 8.39
27.12 | 4.05 | 4.16 | 4.26 | 4.37 | 4.48 | 4.59 | 4.7 | 4.81 | 4.92 | 5.02 | 5.13 | 5.24 | 5.35 | 5.46 | 5.57 | 5.67 | 5.78 | 5.89 | 6 | 6.11 | 6.22 | 6.33 | 6.43 | 6.54 | 6.65 | 6.76 | 6.87 | 6.98 | 7.08 | 7.19 | 7.3 | 7.41 | 7.52 | 7.63 | 7.73 | 7.84 | 7.95 | 8.06 | 8.17 | 8.28 | 8.38
27.13 | 4.05 | 4.16 | 4.26 | 4.37 | 4.48 | 4.59 | 4.7 | 4.81 | 4.91 | 5.02 | 5.13 | 5.24 | 5.35 | 5.46 | 5.57 | 5.67 | 5.78 | 5.89 | 6 | 6.11 | 6.22 | 6.32 | 6.43 | 6.54 | 6.65 | 6.76 | 6.87 | 6.97 | 7.08 | 7.19 | 7.3 | 7.41 | 7.52 | 7.62 | 7.73 | 7.84 | 7.95 | 8.06 | 8.17 | 8.28 | 8.38
27.14 | 4.05 | 4.15 | 4.26 | 4.37 | 4.48 | 4.59 | 4.7 | 4.81 | 4.91 | 5.02 | 5.13 | 5.24 | 5.35 | 5.46 | 5.56 | 5.67 | 5.78 | 5.89 | 6 | 6.11 | 6.21 | 6.32 | 6.43 | 6.54 | 6.65 | 6.76 | 6.86 | 6.97 | 7.08 | 7.19 | 7.3 | 7.41 | 7.52 | 7.62 | 7.73 | 7.84 | 7.95 | 8.06 | 8.17 | 8.27 | 8.38
27.15 | 4.05 | 4.15 | 4.26 | 4.37 | 4.48 | 4.59 | 4.7 | 4.8 | 4.91 | 5.02 | 5.13 | 5.24 | 5.35 | 5.45 | 5.56 | 5.67 | 5.78 | 5.89 | 6 | 6.1 | 6.21 | 6.32 | 6.43 | 6.54 | 6.65 | 6.76 | 6.86 | 6.97 | 7.08 | 7.19 | 7.3 | 7.41 | 7.51 | 7.62 | 7.73 | 7.84 | 7.95 | 8.06 | 8.16 | 8.27 | 8.38
27.16 | 4.04 | 4.15 | 4.26 | 4.37 | 4.48 | 4.59 | 4.69 | 4.8 | 4.91 | 5.02 | 5.13 | 5.24 | 5.35 | 5.45 | 5.56 | 5.67 | 5.78 | 5.89 | 6 | 6.1 | 6.21 | 6.32 | 6.43 | 6.54 | 6.65 | 6.75 | 6.86 | 6.97 | 7.08 | 7.19 | 7.3 | 7.4 | 7.51 | 7.62 | 7.73 | 7.84 | 7.95 | 8.05 | 8.16 | 8.27 | 8.38
27.17 | 4.04 | 4.15 | 4.26 | 4.37 | 4.48 | 4.59 | 4.69 | 4.8 | 4.91 | 5.02 | 5.13 | 5.24 | 5.34 | 5.45 | 5.56 | 5.67 | 5.78 | 5.89 | 5.99 | 6.1 | 6.21 | 6.32 | 6.43 | 6.54 | 6.64 | 6.75 | 6.86 | 6.97 | 7.08 | 7.19 | 7.29 | 7.4 | 7.51 | 7.62 | 7.73 | 7.84 | 7.94 | 8.05 | 8.16 | 8.27 | 8.38
27.18 | 4.04 | 4.15 | 4.26 | 4.37 | 4.48 | 4.58 | 4.69 | 4.8 | 4.91 | 5.02 | 5.13 | 5.23 | 5.34 | 5.45 | 5.56 | 5.67 | 5.78 | 5.88 | 5.99 | 6.1 | 6.21 | 6.32 | 6.43 | 6.53 | 6.64 | 6.75 | 6.86 | 6.97 | 7.08 | 7.18 | 7.29 | 7.4 | 7.51 | 7.62 | 7.73 | 7.83 | 7.94 | 8.05 | 8.16 | 8.27 | 8.38
27.19 | 4.04 | 4.15 | 4.26 | 4.37 | 4.48 | 4.58 | 4.69 | 4.8 | 4.91 | 5.02 | 5.13 | 5.23 | 5.34 | 5.45 | 5.56 | 5.67 | 5.78 | 5.88 | 5.99 | 6.1 | 6.21 | 6.32 | 6.43 | 6.53 | 6.64 | 6.75 | 6.86 | 6.97 | 7.08 | 7.18 | 7.29 | 7.4 | 7.51 | 7.62 | 7.72 | 7.83 | 7.94 | 8.05 | 8.16 | 8.27 | 8.37
27.2 | 4.04 | 4.15 | 4.26 | 4.37 | 4.47 | 4.58 | 4.69 | 4.8 | 4.91 | 5.02 | 5.12 | 5.23 | 5.34 | 5.45 | 5.56 | 5.67 | 5.77 | 5.88 | 5.99 | 6.1 | 6.21 | 6.32 | 6.42 | 6.53 | 6.64 | 6.75 | 6.86 | 6.97 | 7.07 | 7.18 | 7.29 | 7.4 | 7.51 | 7.62 | 7.72 | 7.83 | 7.94 | 8.05 | 8.16 | 8.26 | 8.37
27.21 | 4.04 | 4.15 | 4.26 | 4.37 | 4.47 | 4.58 | 4.69 | 4.8 | 4.91 | 5.02 | 5.12 | 5.23 | 5.34 | 5.45 | 5.56 | 5.66 | 5.77 | 5.88 | 5.99 | 6.1 | 6.21 | 6.31 | 6.42 | 6.53 | 6.64 | 6.75 | 6.86 | 6.96 | 7.07 | 7.18 | 7.29 | 7.4 | 7.51 | 7.61 | 7.72 | 7.83 | 7.94 | 8.05 | 8.16 | 8.26 | 8.37
27.22 | 4.04 | 4.15 | 4.26 | 4.36 | 4.47 | 4.58 | 4.69 | 4.8 | 4.91 | 5.01 | 5.12 | 5.23 | 5.34 | 5.45 | 5.56 | 5.66 | 5.77 | 5.88 | 5.99 | 6.1 | 6.21 | 6.31 | 6.42 | 6.53 | 6.64 | 6.75 | 6.85 | 6.96 | 7.07 | 7.18 | 7.29 | 7.4 | 7.5 | 7.61 | 7.72 | 7.83 | 7.94 | 8.05 | 8.15 | 8.26 | 8.37
27.23 | 4.04 | 4.15 | 4.26 | 4.36 | 4.47 | 4.58 | 4.69 | 4.8 | 4.91 | 5.01 | 5.12 | 5.23 | 5.34 | 5.45 | 5.55 | 5.66 | 5.77 | 5.88 | 5.99 | 6.1 | 6.2 | 6.31 | 6.42 | 6.53 | 6.64 | 6.75 | 6.85 | 6.96 | 7.07 | 7.18 | 7.29 | 7.39 | 7.5 | 7.61 | 7.72 | 7.83 | 7.94 | 8.04 | 8.15 | 8.26 | 8.37
27.24 | 4.04 | 4.15 | 4.25 | 4.36 | 4.47 | 4.58 | 4.69 | 4.8 | 4.9 | 5.01 | 5.12 | 5.23 | 5.34 | 5.45 | 5.55 | 5.66 | 5.77 | 5.88 | 5.99 | 6.09 | 6.2 | 6.31 | 6.42 | 6.53 | 6.64 | 6.74 | 6.85 | 6.96 | 7.07 | 7.18 | 7.29 | 7.39 | 7.5 | 7.61 | 7.72 | 7.83 | 7.93 | 8.04 | 8.15 | 8.26 | 8.37
27.25 | 4.04 | 4.15 | 4.25 | 4.36 | 4.47 | 4.58 | 4.69 | 4.79 | 4.9 | 5.01 | 5.12 | 5.23 | 5.34 | 5.44 | 5.55 | 5.66 | 5.77 | 5.88 | 5.99 | 6.09 | 6.2 | 6.31 | 6.42 | 6.53 | 6.63 | 6.74 | 6.85 | 6.96 | 7.07 | 7.18 | 7.28 | 7.39 | 7.5 | 7.61 | 7.72 | 7.82 | 7.93 | 8.04 | 8.15 | 8.26 | 8.37
27.26 | 4.04 | 4.14 | 4.25 | 4.36 | 4.47 | 4.58 | 4.69 | 4.79 | 4.9 | 5.01 | 5.12 | 5.23 | 5.34 | 5.44 | 5.55 | 5.66 | 5.77 | 5.88 | 5.98 | 6.09 | 6.2 | 6.31 | 6.42 | 6.53 | 6.63 | 6.74 | 6.85 | 6.96 | 7.07 | 7.17 | 7.28 | 7.39 | 7.5 | 7.61 | 7.72 | 7.82 | 7.93 | 8.04 | 8.15 | 8.26 | 8.36
27.27 | 4.04 | 4.14 | 4.25 | 4.36 | 4.47 | 4.58 | 4.68 | 4.79 | 4.9 | 5.01 | 5.12 | 5.23 | 5.33 | 5.44 | 5.55 | 5.66 | 5.77 | 5.87 | 5.98 | 6.09 | 6.2 | 6.31 | 6.42 | 6.52 | 6.63 | 6.74 | 6.85 | 6.96 | 7.06 | 7.17 | 7.28 | 7.39 | 7.5 | 7.61 | 7.71 | 7.82 | 7.93 | 8.04 | 8.15 | 8.25 | 8.36
27.28 | 4.03 | 4.14 | 4.25 | 4.36 | 4.47 | 4.58 | 4.68 | 4.79 | 4.9 | 5.01 | 5.12 | 5.22 | 5.33 | 5.44 | 5.55 | 5.66 | 5.77 | 5.87 | 5.98 | 6.09 | 6.2 | 6.31 | 6.41 | 6.52 | 6.63 | 6.74 | 6.85 | 6.96 | 7.06 | 7.17 | 7.28 | 7.39 | 7.5 | 7.6 | 7.71 | 7.82 | 7.93 | 8.04 | 8.14 | 8.25 | 8.36
27.29 | 4.03 | 4.14 | 4.25 | 4.36 | 4.47 | 4.57 | 4.68 | 4.79 | 4.9 | 5.01 | 5.12 | 5.22 | 5.33 | 5.44 | 5.55 | 5.66 | 5.76 | 5.87 | 5.98 | 6.09 | 6.2 | 6.31 | 6.41 | 6.52 | 6.63 | 6.74 | 6.85 | 6.95 | 7.06 | 7.17 | 7.28 | 7.39 | 7.49 | 7.6 | 7.71 | 7.82 | 7.93 | 8.04 | 8.14 | 8.25 | 8.36
27.3 | 4.03 | 4.14 | 4.25 | 4.36 | 4.47 | 4.57 | 4.68 | 4.79 | 4.9 | 5.01 | 5.11 | 5.22 | 5.33 | 5.44 | 5.55 | 5.66 | 5.76 | 5.87 | 5.98 | 6.09 | 6.2 | 6.3 | 6.41 | 6.52 | 6.63 | 6.74 | 6.84 | 6.95 | 7.06 | 7.17 | 7.28 | 7.39 | 7.49 | 7.6 | 7.71 | 7.82 | 7.93 | 8.03 | 8.14 | 8.25 | 8.36
27.31 | 4.03 | 4.14 | 4.25 | 4.36 | 4.47 | 4.57 | 4.68 | 4.79 | 4.9 | 5.01 | 5.11 | 5.22 | 5.33 | 5.44 | 5.55 | 5.65 | 5.76 | 5.87 | 5.98 | 6.09 | 6.19 | 6.3 | 6.41 | 6.52 | 6.63 | 6.74 | 6.84 | 6.95 | 7.06 | 7.17 | 7.28 | 7.38 | 7.49 | 7.6 | 7.71 | 7.82 | 7.92 | 8.03 | 8.14 | 8.25 | 8.36
27.32 | 4.03 | 4.14 | 4.25 | 4.36 | 4.46 | 4.57 | 4.68 | 4.79 | 4.9 | 5 | 5.11 | 5.22 | 5.33 | 5.44 | 5.55 | 5.65 | 5.76 | 5.87 | 5.98 | 6.09 | 6.19 | 6.3 | 6.41 | 6.52 | 6.63 | 6.73 | 6.84 | 6.95 | 7.06 | 7.17 | 7.27 | 7.38 | 7.49 | 7.6 | 7.71 | 7.81 | 7.92 | 8.03 | 8.14 | 8.25 | 8.36
27.33 | 4.03 | 4.14 | 4.25 | 4.36 | 4.46 | 4.57 | 4.68 | 4.79 | 4.9 | 5 | 5.11 | 5.22 | 5.33 | 5.44 | 5.54 | 5.65 | 5.76 | 5.87 | 5.98 | 6.08 | 6.19 | 6.3 | 6.41 | 6.52 | 6.62 | 6.73 | 6.84 | 6.95 | 7.06 | 7.17 | 7.27 | 7.38 | 7.49 | 7.6 | 7.71 | 7.81 | 7.92 | 8.03 | 8.14 | 8.25 | 8.35
27.34 | 4.03 | 4.14 | 4.25 | 4.35 | 4.46 | 4.57 | 4.68 | 4.79 | 4.89 | 5 | 5.11 | 5.22 | 5.33 | 5.44 | 5.54 | 5.65 | 5.76 | 5.87 | 5.98 | 6.08 | 6.19 | 6.3 | 6.41 | 6.52 | 6.62 | 6.73 | 6.84 | 6.95 | 7.06 | 7.16 | 7.27 | 7.38 | 7.49 | 7.6 | 7.7 | 7.81 | 7.92 | 8.03 | 8.14 | 8.24 | 8.35
27.35 | 4.03 | 4.14 | 4.25 | 4.35 | 4.46 | 4.57 | 4.68 | 4.79 | 4.89 | 5 | 5.11 | 5.22 | 5.33 | 5.43 | 5.54 | 5.65 | 5.76 | 5.87 | 5.97 | 6.08 | 6.19 | 6.3 | 6.41 | 6.51 | 6.62 | 6.73 | 6.84 | 6.95 | 7.05 | 7.16 | 7.27 | 7.38 | 7.49 | 7.59 | 7.7 | 7.81 | 7.92 | 8.03 | 8.13 | 8.24 | 8.35
27.36 | 4.03 | 4.14 | 4.24 | 4.35 | 4.46 | 4.57 | 4.68 | 4.78 | 4.89 | 5 | 5.11 | 5.22 | 5.32 | 5.43 | 5.54 | 5.65 | 5.76 | 5.87 | 5.97 | 6.08 | 6.19 | 6.3 | 6.41 | 6.51 | 6.62 | 6.73 | 6.84 | 6.95 | 7.05 | 7.16 | 7.27 | 7.38 | 7.49 | 7.59 | 7.7 | 7.81 | 7.92 | 8.03 | 8.13 | 8.24 | 8.35
27.37 | 4.03 | 4.14 | 4.24 | 4.35 | 4.46 | 4.57 | 4.68 | 4.78 | 4.89 | 5 | 5.11 | 5.22 | 5.32 | 5.43 | 5.54 | 5.65 | 5.76 | 5.86 | 5.97 | 6.08 | 6.19 | 6.3 | 6.4 | 6.51 | 6.62 | 6.73 | 6.84 | 6.94 | 7.05 | 7.16 | 7.27 | 7.38 | 7.48 | 7.59 | 7.7 | 7.81 | 7.92 | 8.02 | 8.13 | 8.24 | 8.35
27.38 | 4.03 | 4.13 | 4.24 | 4.35 | 4.46 | 4.57 | 4.67 | 4.78 | 4.89 | 5 | 5.11 | 5.21 | 5.32 | 5.43 | 5.54 | 5.65 | 5.75 | 5.86 | 5.97 | 6.08 | 6.19 | 6.29 | 6.4 | 6.51 | 6.62 | 6.73 | 6.83 | 6.94 | 7.05 | 7.16 | 7.27 | 7.37 | 7.48 | 7.59 | 7.7 | 7.81 | 7.91 | 8.02 | 8.13 | 8.24 | 8.35
27.39 | 4.03 | 4.13 | 4.24 | 4.35 | 4.46 | 4.57 | 4.67 | 4.78 | 4.89 | 5 | 5.11 | 5.21 | 5.32 | 5.43 | 5.54 | 5.65 | 5.75 | 5.86 | 5.97 | 6.08 | 6.19 | 6.29 | 6.4 | 6.51 | 6.62 | 6.73 | 6.83 | 6.94 | 7.05 | 7.16 | 7.27 | 7.37 | 7.48 | 7.59 | 7.7 | 7.81 | 7.91 | 8.02 | 8.13 | 8.24 | 8.34
27.4 | 4.03 | 4.13 | 4.24 | 4.35 | 4.46 | 4.57 | 4.67 | 4.78 | 4.89 | 5 | 5.11 | 5.21 | 5.32 | 5.43 | 5.54 | 5.64 | 5.75 | 5.86 | 5.97 | 6.08 | 6.18 | 6.29 | 6.4 | 6.51 | 6.62 | 6.72 | 6.83 | 6.94 | 7.05 | 7.16 | 7.26 | 7.37 | 7.48 | 7.59 | 7.7 | 7.8 | 7.91 | 8.02 | 8.13 | 8.24 | 8.34
27.41 | 4.02 | 4.13 | 4.24 | 4.35 | 4.46 | 4.56 | 4.67 | 4.78 | 4.89 | 5 | 5.1 | 5.21 | 5.32 | 5.43 | 5.54 | 5.64 | 5.75 | 5.86 | 5.97 | 6.08 | 6.18 | 6.29 | 6.4 | 6.51 | 6.62 | 6.72 | 6.83 | 6.94 | 7.05 | 7.15 | 7.26 | 7.37 | 7.48 | 7.59 | 7.69 | 7.8 | 7.91 | 8.02 | 8.13 | 8.23 | 8.34
27.42 | 4.02 | 4.13 | 4.24 | 4.35 | 4.46 | 4.56 | 4.67 | 4.78 | 4.89 | 5 | 5.1 | 5.21 | 5.32 | 5.43 | 5.53 | 5.64 | 5.75 | 5.86 | 5.97 | 6.07 | 6.18 | 6.29 | 6.4 | 6.51 | 6.61 | 6.72 | 6.83 | 6.94 | 7.05 | 7.15 | 7.26 | 7.37 | 7.48 | 7.59 | 7.69 | 7.8 | 7.91 | 8.02 | 8.12 | 8.23 | 8.34
27.43 | 4.02 | 4.13 | 4.24 | 4.35 | 4.45 | 4.56 | 4.67 | 4.78 | 4.89 | 4.99 | 5.1 | 5.21 | 5.32 | 5.43 | 5.53 | 5.64 | 5.75 | 5.86 | 5.97 | 6.07 | 6.18 | 6.29 | 6.4 | 6.5 | 6.61 | 6.72 | 6.83 | 6.94 | 7.04 | 7.15 | 7.26 | 7.37 | 7.48 | 7.58 | 7.69 | 7.8 | 7.91 | 8.02 | 8.12 | 8.23 | 8.34
27.44 | 4.02 | 4.13 | 4.24 | 4.35 | 4.45 | 4.56 | 4.67 | 4.78 | 4.89 | 4.99 | 5.1 | 5.21 | 5.32 | 5.42 | 5.53 | 5.64 | 5.75 | 5.86 | 5.96 | 6.07 | 6.18 | 6.29 | 6.4 | 6.5 | 6.61 | 6.72 | 6.83 | 6.94 | 7.04 | 7.15 | 7.26 | 7.37 | 7.47 | 7.58 | 7.69 | 7.8 | 7.91 | 8.01 | 8.12 | 8.23 | 8.34
27.45 | 4.02 | 4.13 | 4.24 | 4.34 | 4.45 | 4.56 | 4.67 | 4.78 | 4.88 | 4.99 | 5.1 | 5.21 | 5.32 | 5.42 | 5.53 | 5.64 | 5.75 | 5.86 | 5.96 | 6.07 | 6.18 | 6.29 | 6.39 | 6.5 | 6.61 | 6.72 | 6.83 | 6.93 | 7.04 | 7.15 | 7.26 | 7.37 | 7.47 | 7.58 | 7.69 | 7.8 | 7.9 | 8.01 | 8.12 | 8.23 | 8.34
27.46 | 4.02 | 4.13 | 4.24 | 4.34 | 4.45 | 4.56 | 4.67 | 4.78 | 4.88 | 4.99 | 5.1 | 5.21 | 5.31 | 5.42 | 5.53 | 5.64 | 5.75 | 5.85 | 5.96 | 6.07 | 6.18 | 6.29 | 6.39 | 6.5 | 6.61 | 6.72 | 6.82 | 6.93 | 7.04 | 7.15 | 7.26 | 7.36 | 7.47 | 7.58 | 7.69 | 7.8 | 7.9 | 8.01 | 8.12 | 8.23 | 8.33
27.47 | 4.02 | 4.13 | 4.24 | 4.34 | 4.45 | 4.56 | 4.67 | 4.77 | 4.88 | 4.99 | 5.1 | 5.21 | 5.31 | 5.42 | 5.53 | 5.64 | 5.75 | 5.85 | 5.96 | 6.07 | 6.18 | 6.28 | 6.39 | 6.5 | 6.61 | 6.72 | 6.82 | 6.93 | 7.04 | 7.15 | 7.25 | 7.36 | 7.47 | 7.58 | 7.69 | 7.79 | 7.9 | 8.01 | 8.12 | 8.23 | 8.33
27.48 | 4.02 | 4.13 | 4.23 | 4.34 | 4.45 | 4.56 | 4.67 | 4.77 | 4.88 | 4.99 | 5.1 | 5.21 | 5.31 | 5.42 | 5.53 | 5.64 | 5.74 | 5.85 | 5.96 | 6.07 | 6.18 | 6.28 | 6.39 | 6.5 | 6.61 | 6.71 | 6.82 | 6.93 | 7.04 | 7.15 | 7.25 | 7.36 | 7.47 | 7.58 | 7.68 | 7.79 | 7.9 | 8.01 | 8.12 | 8.22 | 8.33
27.49 | 4.02 | 4.13 | 4.23 | 4.34 | 4.45 | 4.56 | 4.66 | 4.77 | 4.88 | 4.99 | 5.1 | 5.2 | 5.31 | 5.42 | 5.53 | 5.64 | 5.74 | 5.85 | 5.96 | 6.07 | 6.17 | 6.28 | 6.39 | 6.5 | 6.61 | 6.71 | 6.82 | 6.93 | 7.04 | 7.14 | 7.25 | 7.36 | 7.47 | 7.58 | 7.68 | 7.79 | 7.9 | 8.01 | 8.11 | 8.22 | 8.33
27.5 | 4.02 | 4.13 | 4.23 | 4.34 | 4.45 | 4.56 | 4.66 | 4.77 | 4.88 | 4.99 | 5.1 | 5.2 | 5.31 | 5.42 | 5.53 | 5.63 | 5.74 | 5.85 | 5.96 | 6.07 | 6.17 | 6.28 | 6.39 | 6.5 | 6.6 | 6.71 | 6.82 | 6.93 | 7.04 | 7.14 | 7.25 | 7.36 | 7.47 | 7.57 | 7.68 | 7.79 | 7.9 | 8.01 | 8.11 | 8.22 | 8.33
27.51 | 4.02 | 4.12 | 4.23 | 4.34 | 4.45 | 4.56 | 4.66 | 4.77 | 4.88 | 4.99 | 5.09 | 5.2 | 5.31 | 5.42 | 5.53 | 5.63 | 5.74 | 5.85 | 5.96 | 6.06 | 6.17 | 6.28 | 6.39 | 6.5 | 6.6 | 6.71 | 6.82 | 6.93 | 7.03 | 7.14 | 7.25 | 7.36 | 7.47 | 7.57 | 7.68 | 7.79 | 7.9 | 8 | 8.11 | 8.22 | 8.33
27.52 | 4.02 | 4.12 | 4.23 | 4.34 | 4.45 | 4.55 | 4.66 | 4.77 | 4.88 | 4.99 | 5.09 | 5.2 | 5.31 | 5.42 | 5.52 | 5.63 | 5.74 | 5.85 | 5.96 | 6.06 | 6.17 | 6.28 | 6.39 | 6.49 | 6.6 | 6.71 | 6.82 | 6.93 | 7.03 | 7.14 | 7.25 | 7.36 | 7.46 | 7.57 | 7.68 | 7.79 | 7.89 | 8 | 8.11 | 8.22 | 8.33
27.53 | 4.01 | 4.12 | 4.23 | 4.34 | 4.45 | 4.55 | 4.66 | 4.77 | 4.88 | 4.98 | 5.09 | 5.2 | 5.31 | 5.42 | 5.52 | 5.63 | 5.74 | 5.85 | 5.95 | 6.06 | 6.17 | 6.28 | 6.39 | 6.49 | 6.6 | 6.71 | 6.82 | 6.92 | 7.03 | 7.14 | 7.25 | 7.35 | 7.46 | 7.57 | 7.68 | 7.79 | 7.89 | 8 | 8.11 | 8.22 | 8.32
27.54 | 4.01 | 4.12 | 4.23 | 4.34 | 4.44 | 4.55 | 4.66 | 4.77 | 4.88 | 4.98 | 5.09 | 5.2 | 5.31 | 5.41 | 5.52 | 5.63 | 5.74 | 5.85 | 5.95 | 6.06 | 6.17 | 6.28 | 6.38 | 6.49 | 6.6 | 6.71 | 6.81 | 6.92 | 7.03 | 7.14 | 7.25 | 7.35 | 7.46 | 7.57 | 7.68 | 7.78 | 7.89 | 8 | 8.11 | 8.22 | 8.32
27.55 | 4.01 | 4.12 | 4.23 | 4.34 | 4.44 | 4.55 | 4.66 | 4.77 | 4.88 | 4.98 | 5.09 | 5.2 | 5.31 | 5.41 | 5.52 | 5.63 | 5.74 | 5.84 | 5.95 | 6.06 | 6.17 | 6.28 | 6.38 | 6.49 | 6.6 | 6.71 | 6.81 | 6.92 | 7.03 | 7.14 | 7.24 | 7.35 | 7.46 | 7.57 | 7.68 | 7.78 | 7.89 | 8 | 8.11 | 8.21 | 8.32
27.56 | 4.01 | 4.12 | 4.23 | 4.34 | 4.44 | 4.55 | 4.66 | 4.77 | 4.87 | 4.98 | 5.09 | 5.2 | 5.3 | 5.41 | 5.52 | 5.63 | 5.74 | 5.84 | 5.95 | 6.06 | 6.17 | 6.27 | 6.38 | 6.49 | 6.6 | 6.7 | 6.81 | 6.92 | 7.03 | 7.14 | 7.24 | 7.35 | 7.46 | 7.57 | 7.67 | 7.78 | 7.89 | 8 | 8.1 | 8.21 | 8.32
27.57 | 4.01 | 4.12 | 4.23 | 4.33 | 4.44 | 4.55 | 4.66 | 4.77 | 4.87 | 4.98 | 5.09 | 5.2 | 5.3 | 5.41 | 5.52 | 5.63 | 5.73 | 5.84 | 5.95 | 6.06 | 6.17 | 6.27 | 6.38 | 6.49 | 6.6 | 6.7 | 6.81 | 6.92 | 7.03 | 7.13 | 7.24 | 7.35 | 7.46 | 7.56 | 7.67 | 7.78 | 7.89 | 8 | 8.1 | 8.21 | 8.32
27.58 | 4.01 | 4.12 | 4.23 | 4.33 | 4.44 | 4.55 | 4.66 | 4.76 | 4.87 | 4.98 | 5.09 | 5.2 | 5.3 | 5.41 | 5.52 | 5.63 | 5.73 | 5.84 | 5.95 | 6.06 | 6.16 | 6.27 | 6.38 | 6.49 | 6.59 | 6.7 | 6.81 | 6.92 | 7.03 | 7.13 | 7.24 | 7.35 | 7.46 | 7.56 | 7.67 | 7.78 | 7.89 | 7.99 | 8.1 | 8.21 | 8.32
27.59 | 4.01 | 4.12 | 4.23 | 4.33 | 4.44 | 4.55 | 4.66 | 4.76 | 4.87 | 4.98 | 5.09 | 5.19 | 5.3 | 5.41 | 5.52 | 5.62 | 5.73 | 5.84 | 5.95 | 6.06 | 6.16 | 6.27 | 6.38 | 6.49 | 6.59 | 6.7 | 6.81 | 6.92 | 7.02 | 7.13 | 7.24 | 7.35 | 7.45 | 7.56 | 7.67 | 7.78 | 7.88 | 7.99 | 8.1 | 8.21 | 8.32
27.6 | 4.01 | 4.12 | 4.22 | 4.33 | 4.44 | 4.55 | 4.66 | 4.76 | 4.87 | 4.98 | 5.09 | 5.19 | 5.3 | 5.41 | 5.52 | 5.62 | 5.73 | 5.84 | 5.95 | 6.05 | 6.16 | 6.27 | 6.38 | 6.48 | 6.59 | 6.7 | 6.81 | 6.92 | 7.02 | 7.13 | 7.24 | 7.35 | 7.45 | 7.56 | 7.67 | 7.78 | 7.88 | 7.99 | 8.1 | 8.21 | 8.31
27.61 | 4.01 | 4.12 | 4.22 | 4.33 | 4.44 | 4.55 | 4.65 | 4.76 | 4.87 | 4.98 | 5.08 | 5.19 | 5.3 | 5.41 | 5.52 | 5.62 | 5.73 | 5.84 | 5.95 | 6.05 | 6.16 | 6.27 | 6.38 | 6.48 | 6.59 | 6.7 | 6.81 | 6.91 | 7.02 | 7.13 | 7.24 | 7.34 | 7.45 | 7.56 | 7.67 | 7.77 | 7.88 | 7.99 | 8.1 | 8.2 | 8.31
27.62 | 4.01 | 4.12 | 4.22 | 4.33 | 4.44 | 4.55 | 4.65 | 4.76 | 4.87 | 4.98 | 5.08 | 5.19 | 5.3 | 5.41 | 5.51 | 5.62 | 5.73 | 5.84 | 5.94 | 6.05 | 6.16 | 6.27 | 6.37 | 6.48 | 6.59 | 6.7 | 6.81 | 6.91 | 7.02 | 7.13 | 7.24 | 7.34 | 7.45 | 7.56 | 7.67 | 7.77 | 7.88 | 7.99 | 8.1 | 8.2 | 8.31
27.63 | 4.01 | 4.11 | 4.22 | 4.33 | 4.44 | 4.54 | 4.65 | 4.76 | 4.87 | 4.98 | 5.08 | 5.19 | 5.3 | 5.41 | 5.51 | 5.62 | 5.73 | 5.84 | 5.94 | 6.05 | 6.16 | 6.27 | 6.37 | 6.48 | 6.59 | 6.7 | 6.8 | 6.91 | 7.02 | 7.13 | 7.23 | 7.34 | 7.45 | 7.56 | 7.66 | 7.77 | 7.88 | 7.99 | 8.09 | 8.2 | 8.31
27.64 | 4.01 | 4.11 | 4.22 | 4.33 | 4.44 | 4.54 | 4.65 | 4.76 | 4.87 | 4.97 | 5.08 | 5.19 | 5.3 | 5.4 | 5.51 | 5.62 | 5.73 | 5.83 | 5.94 | 6.05 | 6.16 | 6.26 | 6.37 | 6.48 | 6.59 | 6.7 | 6.8 | 6.91 | 7.02 | 7.13 | 7.23 | 7.34 | 7.45 | 7.56 | 7.66 | 7.77 | 7.88 | 7.99 | 8.09 | 8.2 | 8.31
27.65 | 4.01 | 4.11 | 4.22 | 4.33 | 4.44 | 4.54 | 4.65 | 4.76 | 4.87 | 4.97 | 5.08 | 5.19 | 5.3 | 5.4 | 5.51 | 5.62 | 5.73 | 5.83 | 5.94 | 6.05 | 6.16 | 6.26 | 6.37 | 6.48 | 6.59 | 6.69 | 6.8 | 6.91 | 7.02 | 7.12 | 7.23 | 7.34 | 7.45 | 7.55 | 7.66 | 7.77 | 7.88 | 7.98 | 8.09 | 8.2 | 8.31
27.66 | 4 | 4.11 | 4.22 | 4.33 | 4.43 | 4.54 | 4.65 | 4.76 | 4.86 | 4.97 | 5.08 | 5.19 | 5.29 | 5.4 | 5.51 | 5.62 | 5.73 | 5.83 | 5.94 | 6.05 | 6.16 | 6.26 | 6.37 | 6.48 | 6.59 | 6.69 | 6.8 | 6.91 | 7.02 | 7.12 | 7.23 | 7.34 | 7.45 | 7.55 | 7.66 | 7.77 | 7.88 | 7.98 | 8.09 | 8.2 | 8.31
27.67 | 4 | 4.11 | 4.22 | 4.33 | 4.43 | 4.54 | 4.65 | 4.76 | 4.86 | 4.97 | 5.08 | 5.19 | 5.29 | 5.4 | 5.51 | 5.62 | 5.72 | 5.83 | 5.94 | 6.05 | 6.15 | 6.26 | 6.37 | 6.48 | 6.58 | 6.69 | 6.8 | 6.91 | 7.01 | 7.12 | 7.23 | 7.34 | 7.44 | 7.55 | 7.66 | 7.77 | 7.87 | 7.98 | 8.09 | 8.2 | 8.3
27.68 | 4 | 4.11 | 4.22 | 4.33 | 4.43 | 4.54 | 4.65 | 4.76 | 4.86 | 4.97 | 5.08 | 5.19 | 5.29 | 5.4 | 5.51 | 5.62 | 5.72 | 5.83 | 5.94 | 6.05 | 6.15 | 6.26 | 6.37 | 6.48 | 6.58 | 6.69 | 6.8 | 6.91 | 7.01 | 7.12 | 7.23 | 7.34 | 7.44 | 7.55 | 7.66 | 7.76 | 7.87 | 7.98 | 8.09 | 8.19 | 8.3
27.69 | 4 | 4.11 | 4.22 | 4.32 | 4.43 | 4.54 | 4.65 | 4.75 | 4.86 | 4.97 | 5.08 | 5.18 | 5.29 | 5.4 | 5.51 | 5.61 | 5.72 | 5.83 | 5.94 | 6.04 | 6.15 | 6.26 | 6.37 | 6.47 | 6.58 | 6.69 | 6.8 | 6.9 | 7.01 | 7.12 | 7.23 | 7.33 | 7.44 | 7.55 | 7.66 | 7.76 | 7.87 | 7.98 | 8.09 | 8.19 | 8.3
27.7 | 4 | 4.11 | 4.22 | 4.32 | 4.43 | 4.54 | 4.65 | 4.75 | 4.86 | 4.97 | 5.08 | 5.18 | 5.29 | 5.4 | 5.51 | 5.61 | 5.72 | 5.83 | 5.94 | 6.04 | 6.15 | 6.26 | 6.37 | 6.47 | 6.58 | 6.69 | 6.8 | 6.9 | 7.01 | 7.12 | 7.22 | 7.33 | 7.44 | 7.55 | 7.65 | 7.76 | 7.87 | 7.98 | 8.08 | 8.19 | 8.3
27.71 | 4 | 4.11 | 4.22 | 4.32 | 4.43 | 4.54 | 4.65 | 4.75 | 4.86 | 4.97 | 5.08 | 5.18 | 5.29 | 5.4 | 5.5 | 5.61 | 5.72 | 5.83 | 5.93 | 6.04 | 6.15 | 6.26 | 6.36 | 6.47 | 6.58 | 6.69 | 6.79 | 6.9 | 7.01 | 7.12 | 7.22 | 7.33 | 7.44 | 7.55 | 7.65 | 7.76 | 7.87 | 7.98 | 8.08 | 8.19 | 8.3
27.72 | 4 | 4.11 | 4.21 | 4.32 | 4.43 | 4.54 | 4.64 | 4.75 | 4.86 | 4.97 | 5.07 | 5.18 | 5.29 | 5.4 | 5.5 | 5.61 | 5.72 | 5.83 | 5.93 | 6.04 | 6.15 | 6.26 | 6.36 | 6.47 | 6.58 | 6.69 | 6.79 | 6.9 | 7.01 | 7.11 | 7.22 | 7.33 | 7.44 | 7.54 | 7.65 | 7.76 | 7.87 | 7.97 | 8.08 | 8.19 | 8.3
27.73 | 4 | 4.11 | 4.21 | 4.32 | 4.43 | 4.54 | 4.64 | 4.75 | 4.86 | 4.97 | 5.07 | 5.18 | 5.29 | 5.4 | 5.5 | 5.61 | 5.72 | 5.82 | 5.93 | 6.04 | 6.15 | 6.25 | 6.36 | 6.47 | 6.58 | 6.68 | 6.79 | 6.9 | 7.01 | 7.11 | 7.22 | 7.33 | 7.44 | 7.54 | 7.65 | 7.76 | 7.87 | 7.97 | 8.08 | 8.19 | 8.29
27.74 | 4 | 4.11 | 4.21 | 4.32 | 4.43 | 4.54 | 4.64 | 4.75 | 4.86 | 4.96 | 5.07 | 5.18 | 5.29 | 5.39 | 5.5 | 5.61 | 5.72 | 5.82 | 5.93 | 6.04 | 6.15 | 6.25 | 6.36 | 6.47 | 6.58 | 6.68 | 6.79 | 6.9 | 7.01 | 7.11 | 7.22 | 7.33 | 7.43 | 7.54 | 7.65 | 7.76 | 7.86 | 7.97 | 8.08 | 8.19 | 8.29
27.75 | 4 | 4.1 | 4.21 | 4.32 | 4.43 | 4.53 | 4.64 | 4.75 | 4.86 | 4.96 | 5.07 | 5.18 | 5.29 | 5.39 | 5.5 | 5.61 | 5.72 | 5.82 | 5.93 | 6.04 | 6.14 | 6.25 | 6.36 | 6.47 | 6.57 | 6.68 | 6.79 | 6.9 | 7 | 7.11 | 7.22 | 7.33 | 7.43 | 7.54 | 7.65 | 7.76 | 7.86 | 7.97 | 8.08 | 8.18 | 8.29
27.76 | 4 | 4.1 | 4.21 | 4.32 | 4.43 | 4.53 | 4.64 | 4.75 | 4.86 | 4.96 | 5.07 | 5.18 | 5.28 | 5.39 | 5.5 | 5.61 | 5.71 | 5.82 | 5.93 | 6.04 | 6.14 | 6.25 | 6.36 | 6.47 | 6.57 | 6.68 | 6.79 | 6.9 | 7 | 7.11 | 7.22 | 7.32 | 7.43 | 7.54 | 7.65 | 7.75 | 7.86 | 7.97 | 8.08 | 8.18 | 8.29
27.77 | 4 | 4.1 | 4.21 | 4.32 | 4.43 | 4.53 | 4.64 | 4.75 | 4.85 | 4.96 | 5.07 | 5.18 | 5.28 | 5.39 | 5.5 | 5.61 | 5.71 | 5.82 | 5.93 | 6.04 | 6.14 | 6.25 | 6.36 | 6.46 | 6.57 | 6.68 | 6.79 | 6.89 | 7 | 7.11 | 7.22 | 7.32 | 7.43 | 7.54 | 7.65 | 7.75 | 7.86 | 7.97 | 8.07 | 8.18 | 8.29
27.78 | 3.99 | 4.1 | 4.21 | 4.32 | 4.42 | 4.53 | 4.64 | 4.75 | 4.85 | 4.96 | 5.07 | 5.18 | 5.28 | 5.39 | 5.5 | 5.6 | 5.71 | 5.82 | 5.93 | 6.03 | 6.14 | 6.25 | 6.36 | 6.46 | 6.57 | 6.68 | 6.79 | 6.89 | 7 | 7.11 | 7.21 | 7.32 | 7.43 | 7.54 | 7.64 | 7.75 | 7.86 | 7.97 | 8.07 | 8.18 | 8.29
27.79 | 3.99 | 4.1 | 4.21 | 4.32 | 4.42 | 4.53 | 4.64 | 4.75 | 4.85 | 4.96 | 5.07 | 5.17 | 5.28 | 5.39 | 5.5 | 5.6 | 5.71 | 5.82 | 5.93 | 6.03 | 6.14 | 6.25 | 6.35 | 6.46 | 6.57 | 6.68 | 6.78 | 6.89 | 7 | 7.11 | 7.21 | 7.32 | 7.43 | 7.54 | 7.64 | 7.75 | 7.86 | 7.96 | 8.07 | 8.18 | 8.29
27.8 | 3.99 | 4.1 | 4.21 | 4.32 | 4.42 | 4.53 | 4.64 | 4.74 | 4.85 | 4.96 | 5.07 | 5.17 | 5.28 | 5.39 | 5.5 | 5.6 | 5.71 | 5.82 | 5.92 | 6.03 | 6.14 | 6.25 | 6.35 | 6.46 | 6.57 | 6.68 | 6.78 | 6.89 | 7 | 7.1 | 7.21 | 7.32 | 7.43 | 7.53 | 7.64 | 7.75 | 7.86 | 7.96 | 8.07 | 8.18 | 8.28
27.81 | 3.99 | 4.1 | 4.21 | 4.31 | 4.42 | 4.53 | 4.64 | 4.74 | 4.85 | 4.96 | 5.07 | 5.17 | 5.28 | 5.39 | 5.49 | 5.6 | 5.71 | 5.82 | 5.92 | 6.03 | 6.14 | 6.25 | 6.35 | 6.46 | 6.57 | 6.67 | 6.78 | 6.89 | 7 | 7.1 | 7.21 | 7.32 | 7.43 | 7.53 | 7.64 | 7.75 | 7.85 | 7.96 | 8.07 | 8.18 | 8.28
27.82 | 3.99 | 4.1 | 4.21 | 4.31 | 4.42 | 4.53 | 4.64 | 4.74 | 4.85 | 4.96 | 5.06 | 5.17 | 5.28 | 5.39 | 5.49 | 5.6 | 5.71 | 5.82 | 5.92 | 6.03 | 6.14 | 6.24 | 6.35 | 6.46 | 6.57 | 6.67 | 6.78 | 6.89 | 6.99 | 7.1 | 7.21 | 7.32 | 7.42 | 7.53 | 7.64 | 7.75 | 7.85 | 7.96 | 8.07 | 8.17 | 8.28
27.83 | 3.99 | 4.1 | 4.21 | 4.31 | 4.42 | 4.53 | 4.63 | 4.74 | 4.85 | 4.96 | 5.06 | 5.17 | 5.28 | 5.39 | 5.49 | 5.6 | 5.71 | 5.81 | 5.92 | 6.03 | 6.14 | 6.24 | 6.35 | 6.46 | 6.56 | 6.67 | 6.78 | 6.89 | 6.99 | 7.1 | 7.21 | 7.32 | 7.42 | 7.53 | 7.64 | 7.74 | 7.85 | 7.96 | 8.07 | 8.17 | 8.28
27.84 | 3.99 | 4.1 | 4.2 | 4.31 | 4.42 | 4.53 | 4.63 | 4.74 | 4.85 | 4.96 | 5.06 | 5.17 | 5.28 | 5.38 | 5.49 | 5.6 | 5.71 | 5.81 | 5.92 | 6.03 | 6.13 | 6.24 | 6.35 | 6.46 | 6.56 | 6.67 | 6.78 | 6.89 | 6.99 | 7.1 | 7.21 | 7.31 | 7.42 | 7.53 | 7.64 | 7.74 | 7.85 | 7.96 | 8.06 | 8.17 | 8.28
27.85 | 3.99 | 4.1 | 4.2 | 4.31 | 4.42 | 4.53 | 4.63 | 4.74 | 4.85 | 4.95 | 5.06 | 5.17 | 5.28 | 5.38 | 5.49 | 5.6 | 5.7 | 5.81 | 5.92 | 6.03 | 6.13 | 6.24 | 6.35 | 6.46 | 6.56 | 6.67 | 6.78 | 6.88 | 6.99 | 7.1 | 7.21 | 7.31 | 7.42 | 7.53 | 7.63 | 7.74 | 7.85 | 7.96 | 8.06 | 8.17 | 8.28
27.86 | 3.99 | 4.1 | 4.2 | 4.31 | 4.42 | 4.52 | 4.63 | 4.74 | 4.85 | 4.95 | 5.06 | 5.17 | 5.27 | 5.38 | 5.49 | 5.6 | 5.7 | 5.81 | 5.92 | 6.03 | 6.13 | 6.24 | 6.35 | 6.45 | 6.56 | 6.67 | 6.78 | 6.88 | 6.99 | 7.1 | 7.2 | 7.31 | 7.42 | 7.53 | 7.63 | 7.74 | 7.85 | 7.95 | 8.06 | 8.17 | 8.28
27.87 | 3.99 | 4.09 | 4.2 | 4.31 | 4.42 | 4.52 | 4.63 | 4.74 | 4.85 | 4.95 | 5.06 | 5.17 | 5.27 | 5.38 | 5.49 | 5.6 | 5.7 | 5.81 | 5.92 | 6.02 | 6.13 | 6.24 | 6.35 | 6.45 | 6.56 | 6.67 | 6.77 | 6.88 | 6.99 | 7.1 | 7.2 | 7.31 | 7.42 | 7.52 | 7.63 | 7.74 | 7.85 | 7.95 | 8.06 | 8.17 | 8.27
27.88 | 3.99 | 4.09 | 4.2 | 4.31 | 4.42 | 4.52 | 4.63 | 4.74 | 4.84 | 4.95 | 5.06 | 5.17 | 5.27 | 5.38 | 5.49 | 5.59 | 5.7 | 5.81 | 5.92 | 6.02 | 6.13 | 6.24 | 6.34 | 6.45 | 6.56 | 6.67 | 6.77 | 6.88 | 6.99 | 7.09 | 7.2 | 7.31 | 7.42 | 7.52 | 7.63 | 7.74 | 7.84 | 7.95 | 8.06 | 8.17 | 8.27
27.89 | 3.99 | 4.09 | 4.2 | 4.31 | 4.41 | 4.52 | 4.63 | 4.74 | 4.84 | 4.95 | 5.06 | 5.16 | 5.27 | 5.38 | 5.49 | 5.59 | 5.7 | 5.81 | 5.91 | 6.02 | 6.13 | 6.24 | 6.34 | 6.45 | 6.56 | 6.66 | 6.77 | 6.88 | 6.99 | 7.09 | 7.2 | 7.31 | 7.41 | 7.52 | 7.63 | 7.74 | 7.84 | 7.95 | 8.06 | 8.16 | 8.27
27.9 | 3.99 | 4.09 | 4.2 | 4.31 | 4.41 | 4.52 | 4.63 | 4.74 | 4.84 | 4.95 | 5.06 | 5.16 | 5.27 | 5.38 | 5.49 | 5.59 | 5.7 | 5.81 | 5.91 | 6.02 | 6.13 | 6.24 | 6.34 | 6.45 | 6.56 | 6.66 | 6.77 | 6.88 | 6.98 | 7.09 | 7.2 | 7.31 | 7.41 | 7.52 | 7.63 | 7.73 | 7.84 | 7.95 | 8.06 | 8.16 | 8.27
27.91 | 3.98 | 4.09 | 4.2 | 4.31 | 4.41 | 4.52 | 4.63 | 4.73 | 4.84 | 4.95 | 5.06 | 5.16 | 5.27 | 5.38 | 5.48 | 5.59 | 5.7 | 5.81 | 5.91 | 6.02 | 6.13 | 6.23 | 6.34 | 6.45 | 6.56 | 6.66 | 6.77 | 6.88 | 6.98 | 7.09 | 7.2 | 7.3 | 7.41 | 7.52 | 7.63 | 7.73 | 7.84 | 7.95 | 8.05 | 8.16 | 8.27
27.92 | 3.98 | 4.09 | 4.2 | 4.31 | 4.41 | 4.52 | 4.63 | 4.73 | 4.84 | 4.95 | 5.05 | 5.16 | 5.27 | 5.38 | 5.48 | 5.59 | 5.7 | 5.8 | 5.91 | 6.02 | 6.13 | 6.23 | 6.34 | 6.45 | 6.55 | 6.66 | 6.77 | 6.88 | 6.98 | 7.09 | 7.2 | 7.3 | 7.41 | 7.52 | 7.62 | 7.73 | 7.84 | 7.95 | 8.05 | 8.16 | 8.27
27.93 | 3.98 | 4.09 | 4.2 | 4.3 | 4.41 | 4.52 | 4.63 | 4.73 | 4.84 | 4.95 | 5.05 | 5.16 | 5.27 | 5.38 | 5.48 | 5.59 | 5.7 | 5.8 | 5.91 | 6.02 | 6.12 | 6.23 | 6.34 | 6.45 | 6.55 | 6.66 | 6.77 | 6.87 | 6.98 | 7.09 | 7.2 | 7.3 | 7.41 | 7.52 | 7.62 | 7.73 | 7.84 | 7.94 | 8.05 | 8.16 | 8.27
27.94 | 3.98 | 4.09 | 4.2 | 4.3 | 4.41 | 4.52 | 4.62 | 4.73 | 4.84 | 4.95 | 5.05 | 5.16 | 5.27 | 5.37 | 5.48 | 5.59 | 5.7 | 5.8 | 5.91 | 6.02 | 6.12 | 6.23 | 6.34 | 6.44 | 6.55 | 6.66 | 6.77 | 6.87 | 6.98 | 7.09 | 7.19 | 7.3 | 7.41 | 7.52 | 7.62 | 7.73 | 7.84 | 7.94 | 8.05 | 8.16 | 8.26
27.95 | 3.98 | 4.09 | 4.2 | 4.3 | 4.41 | 4.52 | 4.62 | 4.73 | 4.84 | 4.94 | 5.05 | 5.16 | 5.27 | 5.37 | 5.48 | 5.59 | 5.69 | 5.8 | 5.91 | 6.02 | 6.12 | 6.23 | 6.34 | 6.44 | 6.55 | 6.66 | 6.76 | 6.87 | 6.98 | 7.09 | 7.19 | 7.3 | 7.41 | 7.51 | 7.62 | 7.73 | 7.83 | 7.94 | 8.05 | 8.16 | 8.26
27.96 | 3.98 | 4.09 | 4.19 | 4.3 | 4.41 | 4.52 | 4.62 | 4.73 | 4.84 | 4.94 | 5.05 | 5.16 | 5.27 | 5.37 | 5.48 | 5.59 | 5.69 | 5.8 | 5.91 | 6.01 | 6.12 | 6.23 | 6.34 | 6.44 | 6.55 | 6.66 | 6.76 | 6.87 | 6.98 | 7.08 | 7.19 | 7.3 | 7.41 | 7.51 | 7.62 | 7.73 | 7.83 | 7.94 | 8.05 | 8.15 | 8.26
27.97 | 3.98 | 4.09 | 4.19 | 4.3 | 4.41 | 4.51 | 4.62 | 4.73 | 4.84 | 4.94 | 5.05 | 5.16 | 5.26 | 5.37 | 5.48 | 5.59 | 5.69 | 5.8 | 5.91 | 6.01 | 6.12 | 6.23 | 6.33 | 6.44 | 6.55 | 6.66 | 6.76 | 6.87 | 6.98 | 7.08 | 7.19 | 7.3 | 7.4 | 7.51 | 7.62 | 7.72 | 7.83 | 7.94 | 8.05 | 8.15 | 8.26
27.98 | 3.98 | 4.09 | 4.19 | 4.3 | 4.41 | 4.51 | 4.62 | 4.73 | 4.84 | 4.94 | 5.05 | 5.16 | 5.26 | 5.37 | 5.48 | 5.58 | 5.69 | 5.8 | 5.91 | 6.01 | 6.12 | 6.23 | 6.33 | 6.44 | 6.55 | 6.65 | 6.76 | 6.87 | 6.97 | 7.08 | 7.19 | 7.3 | 7.4 | 7.51 | 7.62 | 7.72 | 7.83 | 7.94 | 8.04 | 8.15 | 8.26
27.99 | 3.98 | 4.09 | 4.19 | 4.3 | 4.41 | 4.51 | 4.62 | 4.73 | 4.83 | 4.94 | 5.05 | 5.16 | 5.26 | 5.37 | 5.48 | 5.58 | 5.69 | 5.8 | 5.9 | 6.01 | 6.12 | 6.22 | 6.33 | 6.44 | 6.55 | 6.65 | 6.76 | 6.87 | 6.97 | 7.08 | 7.19 | 7.29 | 7.4 | 7.51 | 7.62 | 7.72 | 7.83 | 7.94 | 8.04 | 8.15 | 8.26
28 | 3.98 | 4.08 | 4.19 | 4.3 | 4.41 | 4.51 | 4.62 | 4.73 | 4.83 | 4.94 | 5.05 | 5.15 | 5.26 | 5.37 | 5.48 | 5.58 | 5.69 | 5.8 | 5.9 | 6.01 | 6.12 | 6.22 | 6.33 | 6.44 | 6.54 | 6.65 | 6.76 | 6.87 | 6.97 | 7.08 | 7.19 | 7.29 | 7.4 | 7.51 | 7.61 | 7.72 | 7.83 | 7.93 | 8.04 | 8.15 | 8.26
28.01 | 3.98 | 4.08 | 4.19 | 4.3 | 4.4 | 4.51 | 4.62 | 4.73 | 4.83 | 4.94 | 5.05 | 5.15 | 5.26 | 5.37 | 5.47 | 5.58 | 5.69 | 5.79 | 5.9 | 6.01 | 6.12 | 6.22 | 6.33 | 6.44 | 6.54 | 6.65 | 6.76 | 6.86 | 6.97 | 7.08 | 7.18 | 7.29 | 7.4 | 7.51 | 7.61 | 7.72 | 7.83 | 7.93 | 8.04 | 8.15 | 8.25
28.02 | 3.98 | 4.08 | 4.19 | 4.3 | 4.4 | 4.51 | 4.62 | 4.72 | 4.83 | 4.94 | 5.05 | 5.15 | 5.26 | 5.37 | 5.47 | 5.58 | 5.69 | 5.79 | 5.9 | 6.01 | 6.11 | 6.22 | 6.33 | 6.44 | 6.54 | 6.65 | 6.76 | 6.86 | 6.97 | 7.08 | 7.18 | 7.29 | 7.4 | 7.5 | 7.61 | 7.72 | 7.83 | 7.93 | 8.04 | 8.15 | 8.25
28.03 | 3.98 | 4.08 | 4.19 | 4.3 | 4.4 | 4.51 | 4.62 | 4.72 | 4.83 | 4.94 | 5.04 | 5.15 | 5.26 | 5.37 | 5.47 | 5.58 | 5.69 | 5.79 | 5.9 | 6.01 | 6.11 | 6.22 | 6.33 | 6.43 | 6.54 | 6.65 | 6.75 | 6.86 | 6.97 | 7.08 | 7.18 | 7.29 | 7.4 | 7.5 | 7.61 | 7.72 | 7.82 | 7.93 | 8.04 | 8.14 | 8.25
28.04 | 3.97 | 4.08 | 4.19 | 4.29 | 4.4 | 4.51 | 4.62 | 4.72 | 4.83 | 4.94 | 5.04 | 5.15 | 5.26 | 5.36 | 5.47 | 5.58 | 5.68 | 5.79 | 5.9 | 6.01 | 6.11 | 6.22 | 6.33 | 6.43 | 6.54 | 6.65 | 6.75 | 6.86 | 6.97 | 7.07 | 7.18 | 7.29 | 7.39 | 7.5 | 7.61 | 7.72 | 7.82 | 7.93 | 8.04 | 8.14 | 8.25
28.05 | 3.97 | 4.08 | 4.19 | 4.29 | 4.4 | 4.51 | 4.61 | 4.72 | 4.83 | 4.94 | 5.04 | 5.15 | 5.26 | 5.36 | 5.47 | 5.58 | 5.68 | 5.79 | 5.9 | 6 | 6.11 | 6.22 | 6.32 | 6.43 | 6.54 | 6.65 | 6.75 | 6.86 | 6.97 | 7.07 | 7.18 | 7.29 | 7.39 | 7.5 | 7.61 | 7.71 | 7.82 | 7.93 | 8.03 | 8.14 | 8.25
28.06 | 3.97 | 4.08 | 4.19 | 4.29 | 4.4 | 4.51 | 4.61 | 4.72 | 4.83 | 4.93 | 5.04 | 5.15 | 5.26 | 5.36 | 5.47 | 5.58 | 5.68 | 5.79 | 5.9 | 6 | 6.11 | 6.22 | 6.32 | 6.43 | 6.54 | 6.64 | 6.75 | 6.86 | 6.96 | 7.07 | 7.18 | 7.29 | 7.39 | 7.5 | 7.61 | 7.71 | 7.82 | 7.93 | 8.03 | 8.14 | 8.25
28.07 | 3.97 | 4.08 | 4.19 | 4.29 | 4.4 | 4.51 | 4.61 | 4.72 | 4.83 | 4.93 | 5.04 | 5.15 | 5.25 | 5.36 | 5.47 | 5.57 | 5.68 | 5.79 | 5.9 | 6 | 6.11 | 6.22 | 6.32 | 6.43 | 6.54 | 6.64 | 6.75 | 6.86 | 6.96 | 7.07 | 7.18 | 7.28 | 7.39 | 7.5 | 7.6 | 7.71 | 7.82 | 7.92 | 8.03 | 8.14 | 8.25
28.08 | 3.97 | 4.08 | 4.18 | 4.29 | 4.4 | 4.51 | 4.61 | 4.72 | 4.83 | 4.93 | 5.04 | 5.15 | 5.25 | 5.36 | 5.47 | 5.57 | 5.68 | 5.79 | 5.89 | 6 | 6.11 | 6.21 | 6.32 | 6.43 | 6.54 | 6.64 | 6.75 | 6.86 | 6.96 | 7.07 | 7.18 | 7.28 | 7.39 | 7.5 | 7.6 | 7.71 | 7.82 | 7.92 | 8.03 | 8.14 | 8.24
28.09 | 3.97 | 4.08 | 4.18 | 4.29 | 4.4 | 4.5 | 4.61 | 4.72 | 4.82 | 4.93 | 5.04 | 5.15 | 5.25 | 5.36 | 5.47 | 5.57 | 5.68 | 5.79 | 5.89 | 6 | 6.11 | 6.21 | 6.32 | 6.43 | 6.53 | 6.64 | 6.75 | 6.85 | 6.96 | 7.07 | 7.17 | 7.28 | 7.39 | 7.49 | 7.6 | 7.71 | 7.82 | 7.92 | 8.03 | 8.14 | 8.24
28.1 | 3.97 | 4.08 | 4.18 | 4.29 | 4.4 | 4.5 | 4.61 | 4.72 | 4.82 | 4.93 | 5.04 | 5.14 | 5.25 | 5.36 | 5.46 | 5.57 | 5.68 | 5.79 | 5.89 | 6 | 6.11 | 6.21 | 6.32 | 6.43 | 6.53 | 6.64 | 6.75 | 6.85 | 6.96 | 7.07 | 7.17 | 7.28 | 7.39 | 7.49 | 7.6 | 7.71 | 7.81 | 7.92 | 8.03 | 8.13 | 8.24
28.11 | 3.97 | 4.08 | 4.18 | 4.29 | 4.4 | 4.5 | 4.61 | 4.72 | 4.82 | 4.93 | 5.04 | 5.14 | 5.25 | 5.36 | 5.46 | 5.57 | 5.68 | 5.78 | 5.89 | 6 | 6.1 | 6.21 | 6.32 | 6.42 | 6.53 | 6.64 | 6.75 | 6.85 | 6.96 | 7.07 | 7.17 | 7.28 | 7.39 | 7.49 | 7.6 | 7.71 | 7.81 | 7.92 | 8.03 | 8.13 | 8.24
28.12 | 3.97 | 4.07 | 4.18 | 4.29 | 4.4 | 4.5 | 4.61 | 4.72 | 4.82 | 4.93 | 5.04 | 5.14 | 5.25 | 5.36 | 5.46 | 5.57 | 5.68 | 5.78 | 5.89 | 6 | 6.1 | 6.21 | 6.32 | 6.42 | 6.53 | 6.64 | 6.74 | 6.85 | 6.96 | 7.06 | 7.17 | 7.28 | 7.38 | 7.49 | 7.6 | 7.7 | 7.81 | 7.92 | 8.02 | 8.13 | 8.24
28.13 | 3.97 | 4.07 | 4.18 | 4.29 | 4.39 | 4.5 | 4.61 | 4.71 | 4.82 | 4.93 | 5.03 | 5.14 | 5.25 | 5.36 | 5.46 | 5.57 | 5.68 | 5.78 | 5.89 | 6 | 6.1 | 6.21 | 6.32 | 6.42 | 6.53 | 6.64 | 6.74 | 6.85 | 6.96 | 7.06 | 7.17 | 7.28 | 7.38 | 7.49 | 7.6 | 7.7 | 7.81 | 7.92 | 8.02 | 8.13 | 8.24
28.14 | 3.97 | 4.07 | 4.18 | 4.29 | 4.39 | 4.5 | 4.61 | 4.71 | 4.82 | 4.93 | 5.03 | 5.14 | 5.25 | 5.35 | 5.46 | 5.57 | 5.67 | 5.78 | 5.89 | 5.99 | 6.1 | 6.21 | 6.31 | 6.42 | 6.53 | 6.63 | 6.74 | 6.85 | 6.95 | 7.06 | 7.17 | 7.27 | 7.38 | 7.49 | 7.6 | 7.7 | 7.81 | 7.92 | 8.02 | 8.13 | 8.24
28.15 | 3.97 | 4.07 | 4.18 | 4.29 | 4.39 | 4.5 | 4.61 | 4.71 | 4.82 | 4.93 | 5.03 | 5.14 | 5.25 | 5.35 | 5.46 | 5.57 | 5.67 | 5.78 | 5.89 | 5.99 | 6.1 | 6.21 | 6.31 | 6.42 | 6.53 | 6.63 | 6.74 | 6.85 | 6.95 | 7.06 | 7.17 | 7.27 | 7.38 | 7.49 | 7.59 | 7.7 | 7.81 | 7.91 | 8.02 | 8.13 | 8.23
28.16 | 3.96 | 4.07 | 4.18 | 4.28 | 4.39 | 4.5 | 4.61 | 4.71 | 4.82 | 4.93 | 5.03 | 5.14 | 5.25 | 5.35 | 5.46 | 5.57 | 5.67 | 5.78 | 5.89 | 5.99 | 6.1 | 6.21 | 6.31 | 6.42 | 6.53 | 6.63 | 6.74 | 6.85 | 6.95 | 7.06 | 7.17 | 7.27 | 7.38 | 7.49 | 7.59 | 7.7 | 7.81 | 7.91 | 8.02 | 8.13 | 8.23
28.17 | 3.96 | 4.07 | 4.18 | 4.28 | 4.39 | 4.5 | 4.6 | 4.71 | 4.82 | 4.92 | 5.03 | 5.14 | 5.24 | 5.35 | 5.46 | 5.56 | 5.67 | 5.78 | 5.88 | 5.99 | 6.1 | 6.2 | 6.31 | 6.42 | 6.52 | 6.63 | 6.74 | 6.84 | 6.95 | 7.06 | 7.16 | 7.27 | 7.38 | 7.48 | 7.59 | 7.7 | 7.8 | 7.91 | 8.02 | 8.12 | 8.23
28.18 | 3.96 | 4.07 | 4.18 | 4.28 | 4.39 | 4.5 | 4.6 | 4.71 | 4.82 | 4.92 | 5.03 | 5.14 | 5.24 | 5.35 | 5.46 | 5.56 | 5.67 | 5.78 | 5.88 | 5.99 | 6.1 | 6.2 | 6.31 | 6.42 | 6.52 | 6.63 | 6.74 | 6.84 | 6.95 | 7.06 | 7.16 | 7.27 | 7.38 | 7.48 | 7.59 | 7.7 | 7.8 | 7.91 | 8.02 | 8.12 | 8.23
28.19 | 3.96 | 4.07 | 4.18 | 4.28 | 4.39 | 4.5 | 4.6 | 4.71 | 4.82 | 4.92 | 5.03 | 5.14 | 5.24 | 5.35 | 5.46 | 5.56 | 5.67 | 5.78 | 5.88 | 5.99 | 6.1 | 6.2 | 6.31 | 6.42 | 6.52 | 6.63 | 6.74 | 6.84 | 6.95 | 7.06 | 7.16 | 7.27 | 7.38 | 7.48 | 7.59 | 7.69 | 7.8 | 7.91 | 8.01 | 8.12 | 8.23
28.2 | 3.96 | 4.07 | 4.17 | 4.28 | 4.39 | 4.49 | 4.6 | 4.71 | 4.81 | 4.92 | 5.03 | 5.13 | 5.24 | 5.35 | 5.45 | 5.56 | 5.67 | 5.77 | 5.88 | 5.99 | 6.09 | 6.2 | 6.31 | 6.41 | 6.52 | 6.63 | 6.73 | 6.84 | 6.95 | 7.05 | 7.16 | 7.27 | 7.37 | 7.48 | 7.59 | 7.69 | 7.8 | 7.91 | 8.01 | 8.12 | 8.23
28.21 | 3.96 | 4.07 | 4.17 | 4.28 | 4.39 | 4.49 | 4.6 | 4.71 | 4.81 | 4.92 | 5.03 | 5.13 | 5.24 | 5.35 | 5.45 | 5.56 | 5.67 | 5.77 | 5.88 | 5.99 | 6.09 | 6.2 | 6.31 | 6.41 | 6.52 | 6.63 | 6.73 | 6.84 | 6.95 | 7.05 | 7.16 | 7.27 | 7.37 | 7.48 | 7.59 | 7.69 | 7.8 | 7.91 | 8.01 | 8.12 | 8.23
28.22 | 3.96 | 4.07 | 4.17 | 4.28 | 4.39 | 4.49 | 4.6 | 4.71 | 4.81 | 4.92 | 5.03 | 5.13 | 5.24 | 5.35 | 5.45 | 5.56 | 5.67 | 5.77 | 5.88 | 5.99 | 6.09 | 6.2 | 6.31 | 6.41 | 6.52 | 6.63 | 6.73 | 6.84 | 6.94 | 7.05 | 7.16 | 7.26 | 7.37 | 7.48 | 7.58 | 7.69 | 7.8 | 7.9 | 8.01 | 8.12 | 8.22
28.23 | 3.96 | 4.07 | 4.17 | 4.28 | 4.39 | 4.49 | 4.6 | 4.71 | 4.81 | 4.92 | 5.03 | 5.13 | 5.24 | 5.35 | 5.45 | 5.56 | 5.66 | 5.77 | 5.88 | 5.98 | 6.09 | 6.2 | 6.3 | 6.41 | 6.52 | 6.62 | 6.73 | 6.84 | 6.94 | 7.05 | 7.16 | 7.26 | 7.37 | 7.48 | 7.58 | 7.69 | 7.8 | 7.9 | 8.01 | 8.12 | 8.22
28.24 | 3.96 | 4.07 | 4.17 | 4.28 | 4.38 | 4.49 | 4.6 | 4.7 | 4.81 | 4.92 | 5.02 | 5.13 | 5.24 | 5.34 | 5.45 | 5.56 | 5.66 | 5.77 | 5.88 | 5.98 | 6.09 | 6.2 | 6.3 | 6.41 | 6.52 | 6.62 | 6.73 | 6.84 | 6.94 | 7.05 | 7.16 | 7.26 | 7.37 | 7.48 | 7.58 | 7.69 | 7.79 | 7.9 | 8.01 | 8.11 | 8.22
28.25 | 3.96 | 4.06 | 4.17 | 4.28 | 4.38 | 4.49 | 4.6 | 4.7 | 4.81 | 4.92 | 5.02 | 5.13 | 5.24 | 5.34 | 5.45 | 5.56 | 5.66 | 5.77 | 5.88 | 5.98 | 6.09 | 6.2 | 6.3 | 6.41 | 6.51 | 6.62 | 6.73 | 6.83 | 6.94 | 7.05 | 7.15 | 7.26 | 7.37 | 7.47 | 7.58 | 7.69 | 7.79 | 7.9 | 8.01 | 8.11 | 8.22
28.26 | 3.96 | 4.06 | 4.17 | 4.28 | 4.38 | 4.49 | 4.6 | 4.7 | 4.81 | 4.92 | 5.02 | 5.13 | 5.24 | 5.34 | 5.45 | 5.56 | 5.66 | 5.77 | 5.87 | 5.98 | 6.09 | 6.19 | 6.3 | 6.41 | 6.51 | 6.62 | 6.73 | 6.83 | 6.94 | 7.05 | 7.15 | 7.26 | 7.37 | 7.47 | 7.58 | 7.69 | 7.79 | 7.9 | 8 | 8.11 | 8.22
28.27 | 3.96 | 4.06 | 4.17 | 4.28 | 4.38 | 4.49 | 4.6 | 4.7 | 4.81 | 4.91 | 5.02 | 5.13 | 5.23 | 5.34 | 5.45 | 5.55 | 5.66 | 5.77 | 5.87 | 5.98 | 6.09 | 6.19 | 6.3 | 6.41 | 6.51 | 6.62 | 6.73 | 6.83 | 6.94 | 7.05 | 7.15 | 7.26 | 7.36 | 7.47 | 7.58 | 7.68 | 7.79 | 7.9 | 8 | 8.11 | 8.22
28.28 | 3.96 | 4.06 | 4.17 | 4.27 | 4.38 | 4.49 | 4.59 | 4.7 | 4.81 | 4.91 | 5.02 | 5.13 | 5.23 | 5.34 | 5.45 | 5.55 | 5.66 | 5.77 | 5.87 | 5.98 | 6.09 | 6.19 | 6.3 | 6.4 | 6.51 | 6.62 | 6.72 | 6.83 | 6.94 | 7.04 | 7.15 | 7.26 | 7.36 | 7.47 | 7.58 | 7.68 | 7.79 | 7.9 | 8 | 8.11 | 8.21
28.29 | 3.95 | 4.06 | 4.17 | 4.27 | 4.38 | 4.49 | 4.59 | 4.7 | 4.81 | 4.91 | 5.02 | 5.13 | 5.23 | 5.34 | 5.45 | 5.55 | 5.66 | 5.76 | 5.87 | 5.98 | 6.08 | 6.19 | 6.3 | 6.4 | 6.51 | 6.62 | 6.72 | 6.83 | 6.94 | 7.04 | 7.15 | 7.26 | 7.36 | 7.47 | 7.57 | 7.68 | 7.79 | 7.89 | 8 | 8.11 | 8.21
28.3 | 3.95 | 4.06 | 4.17 | 4.27 | 4.38 | 4.49 | 4.59 | 4.7 | 4.81 | 4.91 | 5.02 | 5.13 | 5.23 | 5.34 | 5.44 | 5.55 | 5.66 | 5.76 | 5.87 | 5.98 | 6.08 | 6.19 | 6.3 | 6.4 | 6.51 | 6.62 | 6.72 | 6.83 | 6.93 | 7.04 | 7.15 | 7.25 | 7.36 | 7.47 | 7.57 | 7.68 | 7.79 | 7.89 | 8 | 8.11 | 8.21
28.31 | 3.95 | 4.06 | 4.17 | 4.27 | 4.38 | 4.49 | 4.59 | 4.7 | 4.8 | 4.91 | 5.02 | 5.12 | 5.23 | 5.34 | 5.44 | 5.55 | 5.66 | 5.76 | 5.87 | 5.98 | 6.08 | 6.19 | 6.3 | 6.4 | 6.51 | 6.61 | 6.72 | 6.83 | 6.93 | 7.04 | 7.15 | 7.25 | 7.36 | 7.47 | 7.57 | 7.68 | 7.79 | 7.89 | 8 | 8.1 | 8.21
28.32 | 3.95 | 4.06 | 4.17 | 4.27 | 4.38 | 4.48 | 4.59 | 4.7 | 4.8 | 4.91 | 5.02 | 5.12 | 5.23 | 5.34 | 5.44 | 5.55 | 5.66 | 5.76 | 5.87 | 5.97 | 6.08 | 6.19 | 6.29 | 6.4 | 6.51 | 6.61 | 6.72 | 6.83 | 6.93 | 7.04 | 7.15 | 7.25 | 7.36 | 7.46 | 7.57 | 7.68 | 7.78 | 7.89 | 8 | 8.1 | 8.21
28.33 | 3.95 | 4.06 | 4.16 | 4.27 | 4.38 | 4.48 | 4.59 | 4.7 | 4.8 | 4.91 | 5.02 | 5.12 | 5.23 | 5.34 | 5.44 | 5.55 | 5.65 | 5.76 | 5.87 | 5.97 | 6.08 | 6.19 | 6.29 | 6.4 | 6.51 | 6.61 | 6.72 | 6.82 | 6.93 | 7.04 | 7.14 | 7.25 | 7.36 | 7.46 | 7.57 | 7.68 | 7.78 | 7.89 | 8 | 8.1 | 8.21
28.34 | 3.95 | 4.06 | 4.16 | 4.27 | 4.38 | 4.48 | 4.59 | 4.7 | 4.8 | 4.91 | 5.01 | 5.12 | 5.23 | 5.33 | 5.44 | 5.55 | 5.65 | 5.76 | 5.87 | 5.97 | 6.08 | 6.19 | 6.29 | 6.4 | 6.5 | 6.61 | 6.72 | 6.82 | 6.93 | 7.04 | 7.14 | 7.25 | 7.36 | 7.46 | 7.57 | 7.67 | 7.78 | 7.89 | 7.99 | 8.1 | 8.21
28.35 | 3.95 | 4.06 | 4.16 | 4.27 | 4.38 | 4.48 | 4.59 | 4.69 | 4.8 | 4.91 | 5.01 | 5.12 | 5.23 | 5.33 | 5.44 | 5.55 | 5.65 | 5.76 | 5.86 | 5.97 | 6.08 | 6.18 | 6.29 | 6.4 | 6.5 | 6.61 | 6.72 | 6.82 | 6.93 | 7.04 | 7.14 | 7.25 | 7.35 | 7.46 | 7.57 | 7.67 | 7.78 | 7.89 | 7.99 | 8.1 | 8.2
28.36 | 3.95 | 4.06 | 4.16 | 4.27 | 4.37 | 4.48 | 4.59 | 4.69 | 4.8 | 4.91 | 5.01 | 5.12 | 5.23 | 5.33 | 5.44 | 5.54 | 5.65 | 5.76 | 5.86 | 5.97 | 6.08 | 6.18 | 6.29 | 6.4 | 6.5 | 6.61 | 6.71 | 6.82 | 6.93 | 7.03 | 7.14 | 7.25 | 7.35 | 7.46 | 7.57 | 7.67 | 7.78 | 7.88 | 7.99 | 8.1 | 8.2
28.37 | 3.95 | 4.05 | 4.16 | 4.27 | 4.37 | 4.48 | 4.59 | 4.69 | 4.8 | 4.91 | 5.01 | 5.12 | 5.22 | 5.33 | 5.44 | 5.54 | 5.65 | 5.76 | 5.86 | 5.97 | 6.08 | 6.18 | 6.29 | 6.39 | 6.5 | 6.61 | 6.71 | 6.82 | 6.93 | 7.03 | 7.14 | 7.25 | 7.35 | 7.46 | 7.56 | 7.67 | 7.78 | 7.88 | 7.99 | 8.1 | 8.2
28.38 | 3.95 | 4.05 | 4.16 | 4.27 | 4.37 | 4.48 | 4.59 | 4.69 | 4.8 | 4.9 | 5.01 | 5.12 | 5.22 | 5.33 | 5.44 | 5.54 | 5.65 | 5.76 | 5.86 | 5.97 | 6.07 | 6.18 | 6.29 | 6.39 | 6.5 | 6.61 | 6.71 | 6.82 | 6.92 | 7.03 | 7.14 | 7.24 | 7.35 | 7.46 | 7.56 | 7.67 | 7.78 | 7.88 | 7.99 | 8.09 | 8.2
28.39 | 3.95 | 4.05 | 4.16 | 4.27 | 4.37 | 4.48 | 4.58 | 4.69 | 4.8 | 4.9 | 5.01 | 5.12 | 5.22 | 5.33 | 5.44 | 5.54 | 5.65 | 5.75 | 5.86 | 5.97 | 6.07 | 6.18 | 6.29 | 6.39 | 6.5 | 6.6 | 6.71 | 6.82 | 6.92 | 7.03 | 7.14 | 7.24 | 7.35 | 7.46 | 7.56 | 7.67 | 7.77 | 7.88 | 7.99 | 8.09 | 8.2
28.4 | 3.95 | 4.05 | 4.16 | 4.26 | 4.37 | 4.48 | 4.58 | 4.69 | 4.8 | 4.9 | 5.01 | 5.12 | 5.22 | 5.33 | 5.43 | 5.54 | 5.65 | 5.75 | 5.86 | 5.97 | 6.07 | 6.18 | 6.28 | 6.39 | 6.5 | 6.6 | 6.71 | 6.82 | 6.92 | 7.03 | 7.14 | 7.24 | 7.35 | 7.45 | 7.56 | 7.67 | 7.77 | 7.88 | 7.99 | 8.09 | 8.2
28.41 | 3.95 | 4.05 | 4.16 | 4.26 | 4.37 | 4.48 | 4.58 | 4.69 | 4.8 | 4.9 | 5.01 | 5.11 | 5.22 | 5.33 | 5.43 | 5.54 | 5.65 | 5.75 | 5.86 | 5.96 | 6.07 | 6.18 | 6.28 | 6.39 | 6.5 | 6.6 | 6.71 | 6.81 | 6.92 | 7.03 | 7.13 | 7.24 | 7.35 | 7.45 | 7.56 | 7.67 | 7.77 | 7.88 | 7.98 | 8.09 | 8.2
28.42 | 3.94 | 4.05 | 4.16 | 4.26 | 4.37 | 4.48 | 4.58 | 4.69 | 4.79 | 4.9 | 5.01 | 5.11 | 5.22 | 5.33 | 5.43 | 5.54 | 5.64 | 5.75 | 5.86 | 5.96 | 6.07 | 6.18 | 6.28 | 6.39 | 6.49 | 6.6 | 6.71 | 6.81 | 6.92 | 7.03 | 7.13 | 7.24 | 7.34 | 7.45 | 7.56 | 7.66 | 7.77 | 7.88 | 7.98 | 8.09 | 8.19
28.43 | 3.94 | 4.05 | 4.16 | 4.26 | 4.37 | 4.47 | 4.58 | 4.69 | 4.79 | 4.9 | 5.01 | 5.11 | 5.22 | 5.33 | 5.43 | 5.54 | 5.64 | 5.75 | 5.86 | 5.96 | 6.07 | 6.18 | 6.28 | 6.39 | 6.49 | 6.6 | 6.71 | 6.81 | 6.92 | 7.03 | 7.13 | 7.24 | 7.34 | 7.45 | 7.56 | 7.66 | 7.77 | 7.87 | 7.98 | 8.09 | 8.19
28.44 | 3.94 | 4.05 | 4.16 | 4.26 | 4.37 | 4.47 | 4.58 | 4.69 | 4.79 | 4.9 | 5.01 | 5.11 | 5.22 | 5.32 | 5.43 | 5.54 | 5.64 | 5.75 | 5.86 | 5.96 | 6.07 | 6.17 | 6.28 | 6.39 | 6.49 | 6.6 | 6.71 | 6.81 | 6.92 | 7.02 | 7.13 | 7.24 | 7.34 | 7.45 | 7.55 | 7.66 | 7.77 | 7.87 | 7.98 | 8.09 | 8.19
28.45 | 3.94 | 4.05 | 4.15 | 4.26 | 4.37 | 4.47 | 4.58 | 4.69 | 4.79 | 4.9 | 5 | 5.11 | 5.22 | 5.32 | 5.43 | 5.54 | 5.64 | 5.75 | 5.85 | 5.96 | 6.07 | 6.17 | 6.28 | 6.39 | 6.49 | 6.6 | 6.7 | 6.81 | 6.92 | 7.02 | 7.13 | 7.23 | 7.34 | 7.45 | 7.55 | 7.66 | 7.77 | 7.87 | 7.98 | 8.08 | 8.19
28.46 | 3.94 | 4.05 | 4.15 | 4.26 | 4.37 | 4.47 | 4.58 | 4.68 | 4.79 | 4.9 | 5 | 5.11 | 5.22 | 5.32 | 5.43 | 5.53 | 5.64 | 5.75 | 5.85 | 5.96 | 6.07 | 6.17 | 6.28 | 6.38 | 6.49 | 6.6 | 6.7 | 6.81 | 6.92 | 7.02 | 7.13 | 7.23 | 7.34 | 7.45 | 7.55 | 7.66 | 7.76 | 7.87 | 7.98 | 8.08 | 8.19
28.47 | 3.94 | 4.05 | 4.15 | 4.26 | 4.37 | 4.47 | 4.58 | 4.68 | 4.79 | 4.9 | 5 | 5.11 | 5.21 | 5.32 | 5.43 | 5.53 | 5.64 | 5.75 | 5.85 | 5.96 | 6.06 | 6.17 | 6.28 | 6.38 | 6.49 | 6.6 | 6.7 | 6.81 | 6.91 | 7.02 | 7.13 | 7.23 | 7.34 | 7.44 | 7.55 | 7.66 | 7.76 | 7.87 | 7.98 | 8.08 | 8.19
28.48 | 3.94 | 4.05 | 4.15 | 4.26 | 4.36 | 4.47 | 4.58 | 4.68 | 4.79 | 4.9 | 5 | 5.11 | 5.21 | 5.32 | 5.43 | 5.53 | 5.64 | 5.74 | 5.85 | 5.96 | 6.06 | 6.17 | 6.28 | 6.38 | 6.49 | 6.59 | 6.7 | 6.81 | 6.91 | 7.02 | 7.12 | 7.23 | 7.34 | 7.44 | 7.55 | 7.66 | 7.76 | 7.87 | 7.97 | 8.08 | 8.19
28.49 | 3.94 | 4.05 | 4.15 | 4.26 | 4.36 | 4.47 | 4.58 | 4.68 | 4.79 | 4.89 | 5 | 5.11 | 5.21 | 5.32 | 5.43 | 5.53 | 5.64 | 5.74 | 5.85 | 5.96 | 6.06 | 6.17 | 6.27 | 6.38 | 6.49 | 6.59 | 6.7 | 6.81 | 6.91 | 7.02 | 7.12 | 7.23 | 7.34 | 7.44 | 7.55 | 7.65 | 7.76 | 7.87 | 7.97 | 8.08 | 8.18
28.5 | 3.94 | 4.04 | 4.15 | 4.26 | 4.36 | 4.47 | 4.58 | 4.68 | 4.79 | 4.89 | 5 | 5.11 | 5.21 | 5.32 | 5.42 | 5.53 | 5.64 | 5.74 | 5.85 | 5.95 | 6.06 | 6.17 | 6.27 | 6.38 | 6.49 | 6.59 | 6.7 | 6.8 | 6.91 | 7.02 | 7.12 | 7.23 | 7.33 | 7.44 | 7.55 | 7.65 | 7.76 | 7.87 | 7.97 | 8.08 | 8.18
28.51 | 3.94 | 4.04 | 4.15 | 4.26 | 4.36 | 4.47 | 4.57 | 4.68 | 4.79 | 4.89 | 5 | 5.1 | 5.21 | 5.32 | 5.42 | 5.53 | 5.64 | 5.74 | 5.85 | 5.95 | 6.06 | 6.17 | 6.27 | 6.38 | 6.48 | 6.59 | 6.7 | 6.8 | 6.91 | 7.01 | 7.12 | 7.23 | 7.33 | 7.44 | 7.55 | 7.65 | 7.76 | 7.86 | 7.97 | 8.08 | 8.18
28.52 | 3.94 | 4.04 | 4.15 | 4.25 | 4.36 | 4.47 | 4.57 | 4.68 | 4.79 | 4.89 | 5 | 5.1 | 5.21 | 5.32 | 5.42 | 5.53 | 5.63 | 5.74 | 5.85 | 5.95 | 6.06 | 6.16 | 6.27 | 6.38 | 6.48 | 6.59 | 6.7 | 6.8 | 6.91 | 7.01 | 7.12 | 7.23 | 7.33 | 7.44 | 7.54 | 7.65 | 7.76 | 7.86 | 7.97 | 8.07 | 8.18
28.53 | 3.94 | 4.04 | 4.15 | 4.25 | 4.36 | 4.47 | 4.57 | 4.68 | 4.78 | 4.89 | 5 | 5.1 | 5.21 | 5.32 | 5.42 | 5.53 | 5.63 | 5.74 | 5.85 | 5.95 | 6.06 | 6.16 | 6.27 | 6.38 | 6.48 | 6.59 | 6.69 | 6.8 | 6.91 | 7.01 | 7.12 | 7.22 | 7.33 | 7.44 | 7.54 | 7.65 | 7.75 | 7.86 | 7.97 | 8.07 | 8.18
28.54 | 3.93 | 4.04 | 4.15 | 4.25 | 4.36 | 4.47 | 4.57 | 4.68 | 4.78 | 4.89 | 5 | 5.1 | 5.21 | 5.31 | 5.42 | 5.53 | 5.63 | 5.74 | 5.84 | 5.95 | 6.06 | 6.16 | 6.27 | 6.37 | 6.48 | 6.59 | 6.69 | 6.8 | 6.91 | 7.01 | 7.12 | 7.22 | 7.33 | 7.44 | 7.54 | 7.65 | 7.75 | 7.86 | 7.97 | 8.07 | 8.18
28.55 | 3.93 | 4.04 | 4.15 | 4.25 | 4.36 | 4.46 | 4.57 | 4.68 | 4.78 | 4.89 | 4.99 | 5.1 | 5.21 | 5.31 | 5.42 | 5.53 | 5.63 | 5.74 | 5.84 | 5.95 | 6.06 | 6.16 | 6.27 | 6.37 | 6.48 | 6.59 | 6.69 | 6.8 | 6.9 | 7.01 | 7.12 | 7.22 | 7.33 | 7.43 | 7.54 | 7.65 | 7.75 | 7.86 | 7.96 | 8.07 | 8.18
28.56 | 3.93 | 4.04 | 4.15 | 4.25 | 4.36 | 4.46 | 4.57 | 4.68 | 4.78 | 4.89 | 4.99 | 5.1 | 5.21 | 5.31 | 5.42 | 5.52 | 5.63 | 5.74 | 5.84 | 5.95 | 6.05 | 6.16 | 6.27 | 6.37 | 6.48 | 6.58 | 6.69 | 6.8 | 6.9 | 7.01 | 7.11 | 7.22 | 7.33 | 7.43 | 7.54 | 7.64 | 7.75 | 7.86 | 7.96 | 8.07 | 8.17
28.57 | 3.93 | 4.04 | 4.14 | 4.25 | 4.36 | 4.46 | 4.57 | 4.67 | 4.78 | 4.89 | 4.99 | 5.1 | 5.21 | 5.31 | 5.42 | 5.52 | 5.63 | 5.74 | 5.84 | 5.95 | 6.05 | 6.16 | 6.27 | 6.37 | 6.48 | 6.58 | 6.69 | 6.8 | 6.9 | 7.01 | 7.11 | 7.22 | 7.33 | 7.43 | 7.54 | 7.64 | 7.75 | 7.86 | 7.96 | 8.07 | 8.17
28.58 | 3.93 | 4.04 | 4.14 | 4.25 | 4.36 | 4.46 | 4.57 | 4.67 | 4.78 | 4.89 | 4.99 | 5.1 | 5.2 | 5.31 | 5.42 | 5.52 | 5.63 | 5.73 | 5.84 | 5.95 | 6.05 | 6.16 | 6.26 | 6.37 | 6.48 | 6.58 | 6.69 | 6.79 | 6.9 | 7.01 | 7.11 | 7.22 | 7.32 | 7.43 | 7.54 | 7.64 | 7.75 | 7.85 | 7.96 | 8.07 | 8.17
28.59 | 3.93 | 4.04 | 4.14 | 4.25 | 4.36 | 4.46 | 4.57 | 4.67 | 4.78 | 4.89 | 4.99 | 5.1 | 5.2 | 5.31 | 5.42 | 5.52 | 5.63 | 5.73 | 5.84 | 5.95 | 6.05 | 6.16 | 6.26 | 6.37 | 6.48 | 6.58 | 6.69 | 6.79 | 6.9 | 7 | 7.11 | 7.22 | 7.32 | 7.43 | 7.53 | 7.64 | 7.75 | 7.85 | 7.96 | 8.06 | 8.17
28.6 | 3.93 | 4.04 | 4.14 | 4.25 | 4.35 | 4.46 | 4.57 | 4.67 | 4.78 | 4.88 | 4.99 | 5.1 | 5.2 | 5.31 | 5.41 | 5.52 | 5.63 | 5.73 | 5.84 | 5.94 | 6.05 | 6.16 | 6.26 | 6.37 | 6.47 | 6.58 | 6.69 | 6.79 | 6.9 | 7 | 7.11 | 7.22 | 7.32 | 7.43 | 7.53 | 7.64 | 7.75 | 7.85 | 7.96 | 8.06 | 8.17
28.61 | 3.93 | 4.04 | 4.14 | 4.25 | 4.35 | 4.46 | 4.57 | 4.67 | 4.78 | 4.88 | 4.99 | 5.1 | 5.2 | 5.31 | 5.41 | 5.52 | 5.63 | 5.73 | 5.84 | 5.94 | 6.05 | 6.15 | 6.26 | 6.37 | 6.47 | 6.58 | 6.68 | 6.79 | 6.9 | 7 | 7.11 | 7.21 | 7.32 | 7.43 | 7.53 | 7.64 | 7.74 | 7.85 | 7.96 | 8.06 | 8.17
28.62 | 3.93 | 4.03 | 4.14 | 4.25 | 4.35 | 4.46 | 4.56 | 4.67 | 4.78 | 4.88 | 4.99 | 5.09 | 5.2 | 5.31 | 5.41 | 5.52 | 5.62 | 5.73 | 5.84 | 5.94 | 6.05 | 6.15 | 6.26 | 6.37 | 6.47 | 6.58 | 6.68 | 6.79 | 6.9 | 7 | 7.11 | 7.21 | 7.32 | 7.42 | 7.53 | 7.64 | 7.74 | 7.85 | 7.95 | 8.06 | 8.17
28.63 | 3.93 | 4.03 | 4.14 | 4.25 | 4.35 | 4.46 | 4.56 | 4.67 | 4.78 | 4.88 | 4.99 | 5.09 | 5.2 | 5.31 | 5.41 | 5.52 | 5.62 | 5.73 | 5.83 | 5.94 | 6.05 | 6.15 | 6.26 | 6.36 | 6.47 | 6.58 | 6.68 | 6.79 | 6.89 | 7 | 7.11 | 7.21 | 7.32 | 7.42 | 7.53 | 7.64 | 7.74 | 7.85 | 7.95 | 8.06 | 8.16
28.64 | 3.93 | 4.03 | 4.14 | 4.24 | 4.35 | 4.46 | 4.56 | 4.67 | 4.77 | 4.88 | 4.99 | 5.09 | 5.2 | 5.3 | 5.41 | 5.52 | 5.62 | 5.73 | 5.83 | 5.94 | 6.05 | 6.15 | 6.26 | 6.36 | 6.47 | 6.58 | 6.68 | 6.79 | 6.89 | 7 | 7.1 | 7.21 | 7.32 | 7.42 | 7.53 | 7.63 | 7.74 | 7.85 | 7.95 | 8.06 | 8.16
28.65 | 3.93 | 4.03 | 4.14 | 4.24 | 4.35 | 4.46 | 4.56 | 4.67 | 4.77 | 4.88 | 4.99 | 5.09 | 5.2 | 5.3 | 5.41 | 5.52 | 5.62 | 5.73 | 5.83 | 5.94 | 6.04 | 6.15 | 6.26 | 6.36 | 6.47 | 6.57 | 6.68 | 6.79 | 6.89 | 7 | 7.1 | 7.21 | 7.32 | 7.42 | 7.53 | 7.63 | 7.74 | 7.84 | 7.95 | 8.06 | 8.16
28.66 | 3.93 | 4.03 | 4.14 | 4.24 | 4.35 | 4.46 | 4.56 | 4.67 | 4.77 | 4.88 | 4.98 | 5.09 | 5.2 | 5.3 | 5.41 | 5.51 | 5.62 | 5.73 | 5.83 | 5.94 | 6.04 | 6.15 | 6.26 | 6.36 | 6.47 | 6.57 | 6.68 | 6.78 | 6.89 | 7 | 7.1 | 7.21 | 7.31 | 7.42 | 7.53 | 7.63 | 7.74 | 7.84 | 7.95 | 8.05 | 8.16
28.67 | 3.92 | 4.03 | 4.14 | 4.24 | 4.35 | 4.45 | 4.56 | 4.67 | 4.77 | 4.88 | 4.98 | 5.09 | 5.2 | 5.3 | 5.41 | 5.51 | 5.62 | 5.72 | 5.83 | 5.94 | 6.04 | 6.15 | 6.25 | 6.36 | 6.47 | 6.57 | 6.68 | 6.78 | 6.89 | 7 | 7.1 | 7.21 | 7.31 | 7.42 | 7.52 | 7.63 | 7.74 | 7.84 | 7.95 | 8.05 | 8.16
28.68 | 3.92 | 4.03 | 4.14 | 4.24 | 4.35 | 4.45 | 4.56 | 4.67 | 4.77 | 4.88 | 4.98 | 5.09 | 5.19 | 5.3 | 5.41 | 5.51 | 5.62 | 5.72 | 5.83 | 5.94 | 6.04 | 6.15 | 6.25 | 6.36 | 6.46 | 6.57 | 6.68 | 6.78 | 6.89 | 6.99 | 7.1 | 7.21 | 7.31 | 7.42 | 7.52 | 7.63 | 7.73 | 7.84 | 7.95 | 8.05 | 8.16
28.69 | 3.92 | 4.03 | 4.13 | 4.24 | 4.35 | 4.45 | 4.56 | 4.66 | 4.77 | 4.88 | 4.98 | 5.09 | 5.19 | 5.3 | 5.41 | 5.51 | 5.62 | 5.72 | 5.83 | 5.93 | 6.04 | 6.15 | 6.25 | 6.36 | 6.46 | 6.57 | 6.68 | 6.78 | 6.89 | 6.99 | 7.1 | 7.2 | 7.31 | 7.42 | 7.52 | 7.63 | 7.73 | 7.84 | 7.94 | 8.05 | 8.16
28.7 | 3.92 | 4.03 | 4.13 | 4.24 | 4.35 | 4.45 | 4.56 | 4.66 | 4.77 | 4.88 | 4.98 | 5.09 | 5.19 | 5.3 | 5.4 | 5.51 | 5.62 | 5.72 | 5.83 | 5.93 | 6.04 | 6.14 | 6.25 | 6.36 | 6.46 | 6.57 | 6.67 | 6.78 | 6.89 | 6.99 | 7.1 | 7.2 | 7.31 | 7.41 | 7.52 | 7.63 | 7.73 | 7.84 | 7.94 | 8.05 | 8.15
28.71 | 3.92 | 4.03 | 4.13 | 4.24 | 4.35 | 4.45 | 4.56 | 4.66 | 4.77 | 4.87 | 4.98 | 5.09 | 5.19 | 5.3 | 5.4 | 5.51 | 5.61 | 5.72 | 5.83 | 5.93 | 6.04 | 6.14 | 6.25 | 6.36 | 6.46 | 6.57 | 6.67 | 6.78 | 6.88 | 6.99 | 7.1 | 7.2 | 7.31 | 7.41 | 7.52 | 7.62 | 7.73 | 7.84 | 7.94 | 8.05 | 8.15
28.72 | 3.92 | 4.03 | 4.13 | 4.24 | 4.34 | 4.45 | 4.56 | 4.66 | 4.77 | 4.87 | 4.98 | 5.08 | 5.19 | 5.3 | 5.4 | 5.51 | 5.61 | 5.72 | 5.83 | 5.93 | 6.04 | 6.14 | 6.25 | 6.35 | 6.46 | 6.57 | 6.67 | 6.78 | 6.88 | 6.99 | 7.09 | 7.2 | 7.31 | 7.41 | 7.52 | 7.62 | 7.73 | 7.83 | 7.94 | 8.05 | 8.15
28.73 | 3.92 | 4.03 | 4.13 | 4.24 | 4.34 | 4.45 | 4.55 | 4.66 | 4.77 | 4.87 | 4.98 | 5.08 | 5.19 | 5.3 | 5.4 | 5.51 | 5.61 | 5.72 | 5.82 | 5.93 | 6.04 | 6.14 | 6.25 | 6.35 | 6.46 | 6.56 | 6.67 | 6.78 | 6.88 | 6.99 | 7.09 | 7.2 | 7.3 | 7.41 | 7.52 | 7.62 | 7.73 | 7.83 | 7.94 | 8.04 | 8.15
28.74 | 3.92 | 4.03 | 4.13 | 4.24 | 4.34 | 4.45 | 4.55 | 4.66 | 4.77 | 4.87 | 4.98 | 5.08 | 5.19 | 5.29 | 5.4 | 5.51 | 5.61 | 5.72 | 5.82 | 5.93 | 6.03 | 6.14 | 6.25 | 6.35 | 6.46 | 6.56 | 6.67 | 6.77 | 6.88 | 6.99 | 7.09 | 7.2 | 7.3 | 7.41 | 7.51 | 7.62 | 7.73 | 7.83 | 7.94 | 8.04 | 8.15
28.75 | 3.92 | 4.02 | 4.13 | 4.24 | 4.34 | 4.45 | 4.55 | 4.66 | 4.76 | 4.87 | 4.98 | 5.08 | 5.19 | 5.29 | 5.4 | 5.5 | 5.61 | 5.72 | 5.82 | 5.93 | 6.03 | 6.14 | 6.25 | 6.35 | 6.46 | 6.56 | 6.67 | 6.77 | 6.88 | 6.99 | 7.09 | 7.2 | 7.3 | 7.41 | 7.51 | 7.62 | 7.73 | 7.83 | 7.94 | 8.04 | 8.15
28.76 | 3.92 | 4.02 | 4.13 | 4.24 | 4.34 | 4.45 | 4.55 | 4.66 | 4.76 | 4.87 | 4.98 | 5.08 | 5.19 | 5.29 | 5.4 | 5.5 | 5.61 | 5.72 | 5.82 | 5.93 | 6.03 | 6.14 | 6.24 | 6.35 | 6.46 | 6.56 | 6.67 | 6.77 | 6.88 | 6.98 | 7.09 | 7.2 | 7.3 | 7.41 | 7.51 | 7.62 | 7.72 | 7.83 | 7.94 | 8.04 | 8.15
28.77 | 3.92 | 4.02 | 4.13 | 4.23 | 4.34 | 4.45 | 4.55 | 4.66 | 4.76 | 4.87 | 4.97 | 5.08 | 5.19 | 5.29 | 5.4 | 5.5 | 5.61 | 5.71 | 5.82 | 5.93 | 6.03 | 6.14 | 6.24 | 6.35 | 6.45 | 6.56 | 6.67 | 6.77 | 6.88 | 6.98 | 7.09 | 7.19 | 7.3 | 7.41 | 7.51 | 7.62 | 7.72 | 7.83 | 7.93 | 8.04 | 8.14
28.78 | 3.92 | 4.02 | 4.13 | 4.23 | 4.34 | 4.44 | 4.55 | 4.66 | 4.76 | 4.87 | 4.97 | 5.08 | 5.18 | 5.29 | 5.4 | 5.5 | 5.61 | 5.71 | 5.82 | 5.92 | 6.03 | 6.14 | 6.24 | 6.35 | 6.45 | 6.56 | 6.66 | 6.77 | 6.88 | 6.98 | 7.09 | 7.19 | 7.3 | 7.4 | 7.51 | 7.62 | 7.72 | 7.83 | 7.93 | 8.04 | 8.14
28.79 | 3.92 | 4.02 | 4.13 | 4.23 | 4.34 | 4.44 | 4.55 | 4.66 | 4.76 | 4.87 | 4.97 | 5.08 | 5.18 | 5.29 | 5.4 | 5.5 | 5.61 | 5.71 | 5.82 | 5.92 | 6.03 | 6.13 | 6.24 | 6.35 | 6.45 | 6.56 | 6.66 | 6.77 | 6.87 | 6.98 | 7.09 | 7.19 | 7.3 | 7.4 | 7.51 | 7.61 | 7.72 | 7.83 | 7.93 | 8.04 | 8.14
28.8 | 3.91 | 4.02 | 4.13 | 4.23 | 4.34 | 4.44 | 4.55 | 4.65 | 4.76 | 4.87 | 4.97 | 5.08 | 5.18 | 5.29 | 5.39 | 5.5 | 5.61 | 5.71 | 5.82 | 5.92 | 6.03 | 6.13 | 6.24 | 6.35 | 6.45 | 6.56 | 6.66 | 6.77 | 6.87 | 6.98 | 7.08 | 7.19 | 7.3 | 7.4 | 7.51 | 7.61 | 7.72 | 7.82 | 7.93 | 8.04 | 8.14
28.81 | 3.91 | 4.02 | 4.13 | 4.23 | 4.34 | 4.44 | 4.55 | 4.65 | 4.76 | 4.86 | 4.97 | 5.08 | 5.18 | 5.29 | 5.39 | 5.5 | 5.6 | 5.71 | 5.82 | 5.92 | 6.03 | 6.13 | 6.24 | 6.34 | 6.45 | 6.56 | 6.66 | 6.77 | 6.87 | 6.98 | 7.08 | 7.19 | 7.29 | 7.4 | 7.51 | 7.61 | 7.72 | 7.82 | 7.93 | 8.03 | 8.14
28.82 | 3.91 | 4.02 | 4.12 | 4.23 | 4.34 | 4.44 | 4.55 | 4.65 | 4.76 | 4.86 | 4.97 | 5.08 | 5.18 | 5.29 | 5.39 | 5.5 | 5.6 | 5.71 | 5.81 | 5.92 | 6.03 | 6.13 | 6.24 | 6.34 | 6.45 | 6.55 | 6.66 | 6.77 | 6.87 | 6.98 | 7.08 | 7.19 | 7.29 | 7.4 | 7.5 | 7.61 | 7.72 | 7.82 | 7.93 | 8.03 | 8.14
28.83 | 3.91 | 4.02 | 4.12 | 4.23 | 4.33 | 4.44 | 4.55 | 4.65 | 4.76 | 4.86 | 4.97 | 5.07 | 5.18 | 5.29 | 5.39 | 5.5 | 5.6 | 5.71 | 5.81 | 5.92 | 6.02 | 6.13 | 6.24 | 6.34 | 6.45 | 6.55 | 6.66 | 6.76 | 6.87 | 6.98 | 7.08 | 7.19 | 7.29 | 7.4 | 7.5 | 7.61 | 7.71 | 7.82 | 7.93 | 8.03 | 8.14
28.84 | 3.91 | 4.02 | 4.12 | 4.23 | 4.33 | 4.44 | 4.55 | 4.65 | 4.76 | 4.86 | 4.97 | 5.07 | 5.18 | 5.28 | 5.39 | 5.5 | 5.6 | 5.71 | 5.81 | 5.92 | 6.02 | 6.13 | 6.23 | 6.34 | 6.45 | 6.55 | 6.66 | 6.76 | 6.87 | 6.97 | 7.08 | 7.19 | 7.29 | 7.4 | 7.5 | 7.61 | 7.71 | 7.82 | 7.92 | 8.03 | 8.14
28.85 | 3.91 | 4.02 | 4.12 | 4.23 | 4.33 | 4.44 | 4.54 | 4.65 | 4.76 | 4.86 | 4.97 | 5.07 | 5.18 | 5.28 | 5.39 | 5.49 | 5.6 | 5.71 | 5.81 | 5.92 | 6.02 | 6.13 | 6.23 | 6.34 | 6.44 | 6.55 | 6.66 | 6.76 | 6.87 | 6.97 | 7.08 | 7.18 | 7.29 | 7.39 | 7.5 | 7.61 | 7.71 | 7.82 | 7.92 | 8.03 | 8.13
28.86 | 3.91 | 4.02 | 4.12 | 4.23 | 4.33 | 4.44 | 4.54 | 4.65 | 4.75 | 4.86 | 4.97 | 5.07 | 5.18 | 5.28 | 5.39 | 5.49 | 5.6 | 5.7 | 5.81 | 5.92 | 6.02 | 6.13 | 6.23 | 6.34 | 6.44 | 6.55 | 6.65 | 6.76 | 6.87 | 6.97 | 7.08 | 7.18 | 7.29 | 7.39 | 7.5 | 7.6 | 7.71 | 7.82 | 7.92 | 8.03 | 8.13
28.87 | 3.91 | 4.01 | 4.12 | 4.23 | 4.33 | 4.44 | 4.54 | 4.65 | 4.75 | 4.86 | 4.96 | 5.07 | 5.18 | 5.28 | 5.39 | 5.49 | 5.6 | 5.7 | 5.81 | 5.91 | 6.02 | 6.13 | 6.23 | 6.34 | 6.44 | 6.55 | 6.65 | 6.76 | 6.86 | 6.97 | 7.08 | 7.18 | 7.29 | 7.39 | 7.5 | 7.6 | 7.71 | 7.81 | 7.92 | 8.03 | 8.13
28.88 | 3.91 | 4.01 | 4.12 | 4.23 | 4.33 | 4.44 | 4.54 | 4.65 | 4.75 | 4.86 | 4.96 | 5.07 | 5.18 | 5.28 | 5.39 | 5.49 | 5.6 | 5.7 | 5.81 | 5.91 | 6.02 | 6.12 | 6.23 | 6.34 | 6.44 | 6.55 | 6.65 | 6.76 | 6.86 | 6.97 | 7.07 | 7.18 | 7.29 | 7.39 | 7.5 | 7.6 | 7.71 | 7.81 | 7.92 | 8.02 | 8.13
28.89 | 3.91 | 4.01 | 4.12 | 4.22 | 4.33 | 4.44 | 4.54 | 4.65 | 4.75 | 4.86 | 4.96 | 5.07 | 5.17 | 5.28 | 5.39 | 5.49 | 5.6 | 5.7 | 5.81 | 5.91 | 6.02 | 6.12 | 6.23 | 6.33 | 6.44 | 6.55 | 6.65 | 6.76 | 6.86 | 6.97 | 7.07 | 7.18 | 7.28 | 7.39 | 7.5 | 7.6 | 7.71 | 7.81 | 7.92 | 8.02 | 8.13
28.9 | 3.91 | 4.01 | 4.12 | 4.22 | 4.33 | 4.43 | 4.54 | 4.65 | 4.75 | 4.86 | 4.96 | 5.07 | 5.17 | 5.28 | 5.38 | 5.49 | 5.6 | 5.7 | 5.81 | 5.91 | 6.02 | 6.12 | 6.23 | 6.33 | 6.44 | 6.54 | 6.65 | 6.76 | 6.86 | 6.97 | 7.07 | 7.18 | 7.28 | 7.39 | 7.49 | 7.6 | 7.7 | 7.81 | 7.92 | 8.02 | 8.13
28.91 | 3.91 | 4.01 | 4.12 | 4.22 | 4.33 | 4.43 | 4.54 | 4.64 | 4.75 | 4.86 | 4.96 | 5.07 | 5.17 | 5.28 | 5.38 | 5.49 | 5.59 | 5.7 | 5.81 | 5.91 | 6.02 | 6.12 | 6.23 | 6.33 | 6.44 | 6.54 | 6.65 | 6.75 | 6.86 | 6.97 | 7.07 | 7.18 | 7.28 | 7.39 | 7.49 | 7.6 | 7.7 | 7.81 | 7.91 | 8.02 | 8.13
28.92 | 3.91 | 4.01 | 4.12 | 4.22 | 4.33 | 4.43 | 4.54 | 4.64 | 4.75 | 4.85 | 4.96 | 5.07 | 5.17 | 5.28 | 5.38 | 5.49 | 5.59 | 5.7 | 5.8 | 5.91 | 6.01 | 6.12 | 6.23 | 6.33 | 6.44 | 6.54 | 6.65 | 6.75 | 6.86 | 6.96 | 7.07 | 7.17 | 7.28 | 7.39 | 7.49 | 7.6 | 7.7 | 7.81 | 7.91 | 8.02 | 8.12
28.93 | 3.9 | 4.01 | 4.12 | 4.22 | 4.33 | 4.43 | 4.54 | 4.64 | 4.75 | 4.85 | 4.96 | 5.06 | 5.17 | 5.28 | 5.38 | 5.49 | 5.59 | 5.7 | 5.8 | 5.91 | 6.01 | 6.12 | 6.22 | 6.33 | 6.44 | 6.54 | 6.65 | 6.75 | 6.86 | 6.96 | 7.07 | 7.17 | 7.28 | 7.38 | 7.49 | 7.6 | 7.7 | 7.81 | 7.91 | 8.02 | 8.12
28.94 | 3.9 | 4.01 | 4.11 | 4.22 | 4.33 | 4.43 | 4.54 | 4.64 | 4.75 | 4.85 | 4.96 | 5.06 | 5.17 | 5.27 | 5.38 | 5.49 | 5.59 | 5.7 | 5.8 | 5.91 | 6.01 | 6.12 | 6.22 | 6.33 | 6.43 | 6.54 | 6.65 | 6.75 | 6.86 | 6.96 | 7.07 | 7.17 | 7.28 | 7.38 | 7.49 | 7.59 | 7.7 | 7.8 | 7.91 | 8.02 | 8.12
28.95 | 3.9 | 4.01 | 4.11 | 4.22 | 4.32 | 4.43 | 4.54 | 4.64 | 4.75 | 4.85 | 4.96 | 5.06 | 5.17 | 5.27 | 5.38 | 5.48 | 5.59 | 5.7 | 5.8 | 5.91 | 6.01 | 6.12 | 6.22 | 6.33 | 6.43 | 6.54 | 6.64 | 6.75 | 6.85 | 6.96 | 7.07 | 7.17 | 7.28 | 7.38 | 7.49 | 7.59 | 7.7 | 7.8 | 7.91 | 8.01 | 8.12
28.96 | 3.9 | 4.01 | 4.11 | 4.22 | 4.32 | 4.43 | 4.53 | 4.64 | 4.75 | 4.85 | 4.96 | 5.06 | 5.17 | 5.27 | 5.38 | 5.48 | 5.59 | 5.69 | 5.8 | 5.91 | 6.01 | 6.12 | 6.22 | 6.33 | 6.43 | 6.54 | 6.64 | 6.75 | 6.85 | 6.96 | 7.06 | 7.17 | 7.28 | 7.38 | 7.49 | 7.59 | 7.7 | 7.8 | 7.91 | 8.01 | 8.12
28.97 | 3.9 | 4.01 | 4.11 | 4.22 | 4.32 | 4.43 | 4.53 | 4.64 | 4.74 | 4.85 | 4.96 | 5.06 | 5.17 | 5.27 | 5.38 | 5.48 | 5.59 | 5.69 | 5.8 | 5.9 | 6.01 | 6.11 | 6.22 | 6.33 | 6.43 | 6.54 | 6.64 | 6.75 | 6.85 | 6.96 | 7.06 | 7.17 | 7.27 | 7.38 | 7.48 | 7.59 | 7.7 | 7.8 | 7.91 | 8.01 | 8.12
28.98 | 3.9 | 4.01 | 4.11 | 4.22 | 4.32 | 4.43 | 4.53 | 4.64 | 4.74 | 4.85 | 4.95 | 5.06 | 5.17 | 5.27 | 5.38 | 5.48 | 5.59 | 5.69 | 5.8 | 5.9 | 6.01 | 6.11 | 6.22 | 6.32 | 6.43 | 6.54 | 6.64 | 6.75 | 6.85 | 6.96 | 7.06 | 7.17 | 7.27 | 7.38 | 7.48 | 7.59 | 7.69 | 7.8 | 7.9 | 8.01 | 8.12
28.99 | 3.9 | 4.01 | 4.11 | 4.22 | 4.32 | 4.43 | 4.53 | 4.64 | 4.74 | 4.85 | 4.95 | 5.06 | 5.16 | 5.27 | 5.38 | 5.48 | 5.59 | 5.69 | 5.8 | 5.9 | 6.01 | 6.11 | 6.22 | 6.32 | 6.43 | 6.53 | 6.64 | 6.74 | 6.85 | 6.96 | 7.06 | 7.17 | 7.27 | 7.38 | 7.48 | 7.59 | 7.69 | 7.8 | 7.9 | 8.01 | 8.11
29 | 3.9 | 4 | 4.11 | 4.22 | 4.32 | 4.43 | 4.53 | 4.64 | 4.74 | 4.85 | 4.95 | 5.06 | 5.16 | 5.27 | 5.37 | 5.48 | 5.58 | 5.69 | 5.8 | 5.9 | 6.01 | 6.11 | 6.22 | 6.32 | 6.43 | 6.53 | 6.64 | 6.74 | 6.85 | 6.95 | 7.06 | 7.16 | 7.27 | 7.38 | 7.48 | 7.59 | 7.69 | 7.8 | 7.9 | 8.01 | 8.11
29.01 | 3.9 | 4 | 4.11 | 4.21 | 4.32 | 4.43 | 4.53 | 4.64 | 4.74 | 4.85 | 4.95 | 5.06 | 5.16 | 5.27 | 5.37 | 5.48 | 5.58 | 5.69 | 5.79 | 5.9 | 6.01 | 6.11 | 6.22 | 6.32 | 6.43 | 6.53 | 6.64 | 6.74 | 6.85 | 6.95 | 7.06 | 7.16 | 7.27 | 7.37 | 7.48 | 7.58 | 7.69 | 7.8 | 7.9 | 8.01 | 8.11
29.02 | 3.9 | 4 | 4.11 | 4.21 | 4.32 | 4.42 | 4.53 | 4.63 | 4.74 | 4.85 | 4.95 | 5.06 | 5.16 | 5.27 | 5.37 | 5.48 | 5.58 | 5.69 | 5.79 | 5.9 | 6 | 6.11 | 6.21 | 6.32 | 6.43 | 6.53 | 6.64 | 6.74 | 6.85 | 6.95 | 7.06 | 7.16 | 7.27 | 7.37 | 7.48 | 7.58 | 7.69 | 7.79 | 7.9 | 8 | 8.11
29.03 | 3.9 | 4 | 4.11 | 4.21 | 4.32 | 4.42 | 4.53 | 4.63 | 4.74 | 4.84 | 4.95 | 5.06 | 5.16 | 5.27 | 5.37 | 5.48 | 5.58 | 5.69 | 5.79 | 5.9 | 6 | 6.11 | 6.21 | 6.32 | 6.42 | 6.53 | 6.63 | 6.74 | 6.85 | 6.95 | 7.06 | 7.16 | 7.27 | 7.37 | 7.48 | 7.58 | 7.69 | 7.79 | 7.9 | 8 | 8.11
29.04 | 3.9 | 4 | 4.11 | 4.21 | 4.32 | 4.42 | 4.53 | 4.63 | 4.74 | 4.84 | 4.95 | 5.05 | 5.16 | 5.26 | 5.37 | 5.48 | 5.58 | 5.69 | 5.79 | 5.9 | 6 | 6.11 | 6.21 | 6.32 | 6.42 | 6.53 | 6.63 | 6.74 | 6.84 | 6.95 | 7.05 | 7.16 | 7.26 | 7.37 | 7.48 | 7.58 | 7.69 | 7.79 | 7.9 | 8 | 8.11
29.05 | 3.9 | 4 | 4.11 | 4.21 | 4.32 | 4.42 | 4.53 | 4.63 | 4.74 | 4.84 | 4.95 | 5.05 | 5.16 | 5.26 | 5.37 | 5.47 | 5.58 | 5.68 | 5.79 | 5.9 | 6 | 6.11 | 6.21 | 6.32 | 6.42 | 6.53 | 6.63 | 6.74 | 6.84 | 6.95 | 7.05 | 7.16 | 7.26 | 7.37 | 7.47 | 7.58 | 7.68 | 7.79 | 7.9 | 8 | 8.11
29.06 | 3.89 | 4 | 4.11 | 4.21 | 4.32 | 4.42 | 4.53 | 4.63 | 4.74 | 4.84 | 4.95 | 5.05 | 5.16 | 5.26 | 5.37 | 5.47 | 5.58 | 5.68 | 5.79 | 5.89 | 6 | 6.1 | 6.21 | 6.32 | 6.42 | 6.53 | 6.63 | 6.74 | 6.84 | 6.95 | 7.05 | 7.16 | 7.26 | 7.37 | 7.47 | 7.58 | 7.68 | 7.79 | 7.89 | 8 | 8.1
29.07 | 3.89 | 4 | 4.1 | 4.21 | 4.31 | 4.42 | 4.53 | 4.63 | 4.74 | 4.84 | 4.95 | 5.05 | 5.16 | 5.26 | 5.37 | 5.47 | 5.58 | 5.68 | 5.79 | 5.89 | 6 | 6.1 | 6.21 | 6.31 | 6.42 | 6.52 | 6.63 | 6.74 | 6.84 | 6.95 | 7.05 | 7.16 | 7.26 | 7.37 | 7.47 | 7.58 | 7.68 | 7.79 | 7.89 | 8 | 8.1
29.08 | 3.89 | 4 | 4.1 | 4.21 | 4.31 | 4.42 | 4.52 | 4.63 | 4.73 | 4.84 | 4.95 | 5.05 | 5.16 | 5.26 | 5.37 | 5.47 | 5.58 | 5.68 | 5.79 | 5.89 | 6 | 6.1 | 6.21 | 6.31 | 6.42 | 6.52 | 6.63 | 6.73 | 6.84 | 6.94 | 7.05 | 7.15 | 7.26 | 7.37 | 7.47 | 7.58 | 7.68 | 7.79 | 7.89 | 8 | 8.1
29.09 | 3.89 | 4 | 4.1 | 4.21 | 4.31 | 4.42 | 4.52 | 4.63 | 4.73 | 4.84 | 4.94 | 5.05 | 5.15 | 5.26 | 5.37 | 5.47 | 5.58 | 5.68 | 5.79 | 5.89 | 6 | 6.1 | 6.21 | 6.31 | 6.42 | 6.52 | 6.63 | 6.73 | 6.84 | 6.94 | 7.05 | 7.15 | 7.26 | 7.36 | 7.47 | 7.57 | 7.68 | 7.78 | 7.89 | 7.99 | 8.1
29.1 | 3.89 | 4 | 4.1 | 4.21 | 4.31 | 4.42 | 4.52 | 4.63 | 4.73 | 4.84 | 4.94 | 5.05 | 5.15 | 5.26 | 5.36 | 5.47 | 5.57 | 5.68 | 5.78 | 5.89 | 6 | 6.1 | 6.21 | 6.31 | 6.42 | 6.52 | 6.63 | 6.73 | 6.84 | 6.94 | 7.05 | 7.15 | 7.26 | 7.36 | 7.47 | 7.57 | 7.68 | 7.78 | 7.89 | 7.99 | 8.1
29.11 | 3.89 | 4 | 4.1 | 4.21 | 4.31 | 4.42 | 4.52 | 4.63 | 4.73 | 4.84 | 4.94 | 5.05 | 5.15 | 5.26 | 5.36 | 5.47 | 5.57 | 5.68 | 5.78 | 5.89 | 5.99 | 6.1 | 6.2 | 6.31 | 6.41 | 6.52 | 6.63 | 6.73 | 6.84 | 6.94 | 7.05 | 7.15 | 7.26 | 7.36 | 7.47 | 7.57 | 7.68 | 7.78 | 7.89 | 7.99 | 8.1
29.12 | 3.89 | 4 | 4.1 | 4.21 | 4.31 | 4.42 | 4.52 | 4.63 | 4.73 | 4.84 | 4.94 | 5.05 | 5.15 | 5.26 | 5.36 | 5.47 | 5.57 | 5.68 | 5.78 | 5.89 | 5.99 | 6.1 | 6.2 | 6.31 | 6.41 | 6.52 | 6.62 | 6.73 | 6.83 | 6.94 | 7.04 | 7.15 | 7.25 | 7.36 | 7.46 | 7.57 | 7.68 | 7.78 | 7.89 | 7.99 | 8.1
29.13 | 3.89 | 3.99 | 4.1 | 4.2 | 4.31 | 4.41 | 4.52 | 4.63 | 4.73 | 4.84 | 4.94 | 5.05 | 5.15 | 5.26 | 5.36 | 5.47 | 5.57 | 5.68 | 5.78 | 5.89 | 5.99 | 6.1 | 6.2 | 6.31 | 6.41 | 6.52 | 6.62 | 6.73 | 6.83 | 6.94 | 7.04 | 7.15 | 7.25 | 7.36 | 7.46 | 7.57 | 7.67 | 7.78 | 7.88 | 7.99 | 8.09
29.14 | 3.89 | 3.99 | 4.1 | 4.2 | 4.31 | 4.41 | 4.52 | 4.62 | 4.73 | 4.83 | 4.94 | 5.04 | 5.15 | 5.26 | 5.36 | 5.47 | 5.57 | 5.68 | 5.78 | 5.89 | 5.99 | 6.1 | 6.2 | 6.31 | 6.41 | 6.52 | 6.62 | 6.73 | 6.83 | 6.94 | 7.04 | 7.15 | 7.25 | 7.36 | 7.46 | 7.57 | 7.67 | 7.78 | 7.88 | 7.99 | 8.09
29.15 | 3.89 | 3.99 | 4.1 | 4.2 | 4.31 | 4.41 | 4.52 | 4.62 | 4.73 | 4.83 | 4.94 | 5.04 | 5.15 | 5.25 | 5.36 | 5.46 | 5.57 | 5.67 | 5.78 | 5.88 | 5.99 | 6.09 | 6.2 | 6.31 | 6.41 | 6.52 | 6.62 | 6.73 | 6.83 | 6.94 | 7.04 | 7.15 | 7.25 | 7.36 | 7.46 | 7.57 | 7.67 | 7.78 | 7.88 | 7.99 | 8.09
29.16 | 3.89 | 3.99 | 4.1 | 4.2 | 4.31 | 4.41 | 4.52 | 4.62 | 4.73 | 4.83 | 4.94 | 5.04 | 5.15 | 5.25 | 5.36 | 5.46 | 5.57 | 5.67 | 5.78 | 5.88 | 5.99 | 6.09 | 6.2 | 6.3 | 6.41 | 6.51 | 6.62 | 6.72 | 6.83 | 6.93 | 7.04 | 7.14 | 7.25 | 7.35 | 7.46 | 7.56 | 7.67 | 7.77 | 7.88 | 7.98 | 8.09
29.17 | 3.89 | 3.99 | 4.1 | 4.2 | 4.31 | 4.41 | 4.52 | 4.62 | 4.73 | 4.83 | 4.94 | 5.04 | 5.15 | 5.25 | 5.36 | 5.46 | 5.57 | 5.67 | 5.78 | 5.88 | 5.99 | 6.09 | 6.2 | 6.3 | 6.41 | 6.51 | 6.62 | 6.72 | 6.83 | 6.93 | 7.04 | 7.14 | 7.25 | 7.35 | 7.46 | 7.56 | 7.67 | 7.77 | 7.88 | 7.98 | 8.09
29.18 | 3.89 | 3.99 | 4.1 | 4.2 | 4.31 | 4.41 | 4.52 | 4.62 | 4.73 | 4.83 | 4.94 | 5.04 | 5.15 | 5.25 | 5.36 | 5.46 | 5.57 | 5.67 | 5.78 | 5.88 | 5.99 | 6.09 | 6.2 | 6.3 | 6.41 | 6.51 | 6.62 | 6.72 | 6.83 | 6.93 | 7.04 | 7.14 | 7.25 | 7.35 | 7.46 | 7.56 | 7.67 | 7.77 | 7.88 | 7.98 | 8.09
29.19 | 3.88 | 3.99 | 4.09 | 4.2 | 4.3 | 4.41 | 4.51 | 4.62 | 4.73 | 4.83 | 4.94 | 5.04 | 5.15 | 5.25 | 5.36 | 5.46 | 5.57 | 5.67 | 5.78 | 5.88 | 5.99 | 6.09 | 6.2 | 6.3 | 6.41 | 6.51 | 6.62 | 6.72 | 6.83 | 6.93 | 7.04 | 7.14 | 7.25 | 7.35 | 7.46 | 7.56 | 7.67 | 7.77 | 7.88 | 7.98 | 8.09
29.2 | 3.88 | 3.99 | 4.09 | 4.2 | 4.3 | 4.41 | 4.51 | 4.62 | 4.72 | 4.83 | 4.93 | 5.04 | 5.14 | 5.25 | 5.35 | 5.46 | 5.56 | 5.67 | 5.77 | 5.88 | 5.98 | 6.09 | 6.19 | 6.3 | 6.4 | 6.51 | 6.61 | 6.72 | 6.82 | 6.93 | 7.03 | 7.14 | 7.24 | 7.35 | 7.45 | 7.56 | 7.66 | 7.77 | 7.87 | 7.98 | 8.08
29.21 | 3.88 | 3.99 | 4.09 | 4.2 | 4.3 | 4.41 | 4.51 | 4.62 | 4.72 | 4.83 | 4.93 | 5.04 | 5.14 | 5.25 | 5.35 | 5.46 | 5.56 | 5.67 | 5.77 | 5.88 | 5.98 | 6.09 | 6.19 | 6.3 | 6.4 | 6.51 | 6.61 | 6.72 | 6.82 | 6.93 | 7.03 | 7.14 | 7.24 | 7.35 | 7.45 | 7.56 | 7.66 | 7.77 | 7.87 | 7.98 | 8.08
29.22 | 3.88 | 3.99 | 4.09 | 4.2 | 4.3 | 4.41 | 4.51 | 4.62 | 4.72 | 4.83 | 4.93 | 5.04 | 5.14 | 5.25 | 5.35 | 5.46 | 5.56 | 5.67 | 5.77 | 5.88 | 5.98 | 6.09 | 6.19 | 6.3 | 6.4 | 6.51 | 6.61 | 6.72 | 6.82 | 6.93 | 7.03 | 7.14 | 7.24 | 7.35 | 7.45 | 7.56 | 7.66 | 7.77 | 7.87 | 7.98 | 8.08
29.23 | 3.88 | 3.99 | 4.09 | 4.2 | 4.3 | 4.41 | 4.51 | 4.62 | 4.72 | 4.83 | 4.93 | 5.04 | 5.14 | 5.25 | 5.35 | 5.46 | 5.56 | 5.67 | 5.77 | 5.88 | 5.98 | 6.09 | 6.19 | 6.3 | 6.4 | 6.51 | 6.61 | 6.72 | 6.82 | 6.93 | 7.03 | 7.14 | 7.24 | 7.35 | 7.45 | 7.56 | 7.66 | 7.77 | 7.87 | 7.98 | 8.08
29.24 | 3.88 | 3.99 | 4.09 | 4.2 | 4.3 | 4.41 | 4.51 | 4.62 | 4.72 | 4.83 | 4.93 | 5.04 | 5.14 | 5.25 | 5.35 | 5.46 | 5.56 | 5.67 | 5.77 | 5.88 | 5.98 | 6.09 | 6.19 | 6.29 | 6.4 | 6.5 | 6.61 | 6.71 | 6.82 | 6.92 | 7.03 | 7.13 | 7.24 | 7.34 | 7.45 | 7.55 | 7.66 | 7.76 | 7.87 | 7.97 | 8.08
29.25 | 3.88 | 3.98 | 4.09 | 4.19 | 4.3 | 4.4 | 4.51 | 4.61 | 4.72 | 4.82 | 4.93 | 5.03 | 5.14 | 5.24 | 5.35 | 5.45 | 5.56 | 5.66 | 5.77 | 5.87 | 5.98 | 6.08 | 6.19 | 6.29 | 6.4 | 6.5 | 6.61 | 6.71 | 6.82 | 6.92 | 7.03 | 7.13 | 7.24 | 7.34 | 7.45 | 7.55 | 7.66 | 7.76 | 7.87 | 7.97 | 8.08
29.26 | 3.88 | 3.98 | 4.09 | 4.19 | 4.3 | 4.4 | 4.51 | 4.61 | 4.72 | 4.82 | 4.93 | 5.03 | 5.14 | 5.24 | 5.35 | 5.45 | 5.56 | 5.66 | 5.77 | 5.87 | 5.98 | 6.08 | 6.19 | 6.29 | 6.4 | 6.5 | 6.61 | 6.71 | 6.82 | 6.92 | 7.03 | 7.13 | 7.24 | 7.34 | 7.45 | 7.55 | 7.66 | 7.76 | 7.87 | 7.97 | 8.08
29.27 | 3.88 | 3.98 | 4.09 | 4.19 | 4.3 | 4.4 | 4.51 | 4.61 | 4.72 | 4.82 | 4.93 | 5.03 | 5.14 | 5.24 | 5.35 | 5.45 | 5.56 | 5.66 | 5.77 | 5.87 | 5.98 | 6.08 | 6.19 | 6.29 | 6.4 | 6.5 | 6.61 | 6.71 | 6.82 | 6.92 | 7.03 | 7.13 | 7.24 | 7.34 | 7.45 | 7.55 | 7.66 | 7.76 | 7.86 | 7.97 | 8.07
29.28 | 3.88 | 3.98 | 4.09 | 4.19 | 4.3 | 4.4 | 4.51 | 4.61 | 4.72 | 4.82 | 4.93 | 5.03 | 5.14 | 5.24 | 5.35 | 5.45 | 5.56 | 5.66 | 5.77 | 5.87 | 5.98 | 6.08 | 6.19 | 6.29 | 6.4 | 6.5 | 6.61 | 6.71 | 6.81 | 6.92 | 7.02 | 7.13 | 7.23 | 7.34 | 7.44 | 7.55 | 7.65 | 7.76 | 7.86 | 7.97 | 8.07
29.29 | 3.88 | 3.98 | 4.09 | 4.19 | 4.3 | 4.4 | 4.51 | 4.61 | 4.72 | 4.82 | 4.93 | 5.03 | 5.14 | 5.24 | 5.35 | 5.45 | 5.56 | 5.66 | 5.76 | 5.87 | 5.97 | 6.08 | 6.18 | 6.29 | 6.39 | 6.5 | 6.6 | 6.71 | 6.81 | 6.92 | 7.02 | 7.13 | 7.23 | 7.34 | 7.44 | 7.55 | 7.65 | 7.76 | 7.86 | 7.97 | 8.07
29.3 | 3.88 | 3.98 | 4.09 | 4.19 | 4.3 | 4.4 | 4.51 | 4.61 | 4.72 | 4.82 | 4.92 | 5.03 | 5.13 | 5.24 | 5.34 | 5.45 | 5.55 | 5.66 | 5.76 | 5.87 | 5.97 | 6.08 | 6.18 | 6.29 | 6.39 | 6.5 | 6.6 | 6.71 | 6.81 | 6.92 | 7.02 | 7.13 | 7.23 | 7.34 | 7.44 | 7.55 | 7.65 | 7.76 | 7.86 | 7.97 | 8.07
29.31 | 3.88 | 3.98 | 4.09 | 4.19 | 4.29 | 4.4 | 4.5 | 4.61 | 4.71 | 4.82 | 4.92 | 5.03 | 5.13 | 5.24 | 5.34 | 5.45 | 5.55 | 5.66 | 5.76 | 5.87 | 5.97 | 6.08 | 6.18 | 6.29 | 6.39 | 6.5 | 6.6 | 6.71 | 6.81 | 6.92 | 7.02 | 7.13 | 7.23 | 7.34 | 7.44 | 7.54 | 7.65 | 7.75 | 7.86 | 7.96 | 8.07
29.32 | 3.87 | 3.98 | 4.08 | 4.19 | 4.29 | 4.4 | 4.5 | 4.61 | 4.71 | 4.82 | 4.92 | 5.03 | 5.13 | 5.24 | 5.34 | 5.45 | 5.55 | 5.66 | 5.76 | 5.87 | 5.97 | 6.08 | 6.18 | 6.29 | 6.39 | 6.5 | 6.6 | 6.71 | 6.81 | 6.91 | 7.02 | 7.12 | 7.23 | 7.33 | 7.44 | 7.54 | 7.65 | 7.75 | 7.86 | 7.96 | 8.07
29.33 | 3.87 | 3.98 | 4.08 | 4.19 | 4.29 | 4.4 | 4.5 | 4.61 | 4.71 | 4.82 | 4.92 | 5.03 | 5.13 | 5.24 | 5.34 | 5.45 | 5.55 | 5.66 | 5.76 | 5.87 | 5.97 | 6.08 | 6.18 | 6.28 | 6.39 | 6.49 | 6.6 | 6.7 | 6.81 | 6.91 | 7.02 | 7.12 | 7.23 | 7.33 | 7.44 | 7.54 | 7.65 | 7.75 | 7.86 | 7.96 | 8.07
29.34 | 3.87 | 3.98 | 4.08 | 4.19 | 4.29 | 4.4 | 4.5 | 4.61 | 4.71 | 4.82 | 4.92 | 5.03 | 5.13 | 5.24 | 5.34 | 5.45 | 5.55 | 5.65 | 5.76 | 5.86 | 5.97 | 6.07 | 6.18 | 6.28 | 6.39 | 6.49 | 6.6 | 6.7 | 6.81 | 6.91 | 7.02 | 7.12 | 7.23 | 7.33 | 7.44 | 7.54 | 7.65 | 7.75 | 7.86 | 7.96 | 8.06
29.35 | 3.87 | 3.98 | 4.08 | 4.19 | 4.29 | 4.4 | 4.5 | 4.61 | 4.71 | 4.82 | 4.92 | 5.03 | 5.13 | 5.23 | 5.34 | 5.44 | 5.55 | 5.65 | 5.76 | 5.86 | 5.97 | 6.07 | 6.18 | 6.28 | 6.39 | 6.49 | 6.6 | 6.7 | 6.81 | 6.91 | 7.02 | 7.12 | 7.23 | 7.33 | 7.43 | 7.54 | 7.64 | 7.75 | 7.85 | 7.96 | 8.06
29.36 | 3.87 | 3.98 | 4.08 | 4.19 | 4.29 | 4.4 | 4.5 | 4.61 | 4.71 | 4.81 | 4.92 | 5.02 | 5.13 | 5.23 | 5.34 | 5.44 | 5.55 | 5.65 | 5.76 | 5.86 | 5.97 | 6.07 | 6.18 | 6.28 | 6.39 | 6.49 | 6.6 | 6.7 | 6.81 | 6.91 | 7.01 | 7.12 | 7.22 | 7.33 | 7.43 | 7.54 | 7.64 | 7.75 | 7.85 | 7.96 | 8.06
29.37 | 3.87 | 3.98 | 4.08 | 4.19 | 4.29 | 4.39 | 4.5 | 4.6 | 4.71 | 4.81 | 4.92 | 5.02 | 5.13 | 5.23 | 5.34 | 5.44 | 5.55 | 5.65 | 5.76 | 5.86 | 5.97 | 6.07 | 6.18 | 6.28 | 6.39 | 6.49 | 6.59 | 6.7 | 6.8 | 6.91 | 7.01 | 7.12 | 7.22 | 7.33 | 7.43 | 7.54 | 7.64 | 7.75 | 7.85 | 7.96 | 8.06
29.38 | 3.87 | 3.97 | 4.08 | 4.18 | 4.29 | 4.39 | 4.5 | 4.6 | 4.71 | 4.81 | 4.92 | 5.02 | 5.13 | 5.23 | 5.34 | 5.44 | 5.55 | 5.65 | 5.76 | 5.86 | 5.96 | 6.07 | 6.17 | 6.28 | 6.38 | 6.49 | 6.59 | 6.7 | 6.8 | 6.91 | 7.01 | 7.12 | 7.22 | 7.33 | 7.43 | 7.54 | 7.64 | 7.75 | 7.85 | 7.95 | 8.06
29.39 | 3.87 | 3.97 | 4.08 | 4.18 | 4.29 | 4.39 | 4.5 | 4.6 | 4.71 | 4.81 | 4.92 | 5.02 | 5.13 | 5.23 | 5.34 | 5.44 | 5.55 | 5.65 | 5.75 | 5.86 | 5.96 | 6.07 | 6.17 | 6.28 | 6.38 | 6.49 | 6.59 | 6.7 | 6.8 | 6.91 | 7.01 | 7.12 | 7.22 | 7.33 | 7.43 | 7.53 | 7.64 | 7.74 | 7.85 | 7.95 | 8.06
29.4 | 3.87 | 3.97 | 4.08 | 4.18 | 4.29 | 4.39 | 4.5 | 4.6 | 4.71 | 4.81 | 4.92 | 5.02 | 5.13 | 5.23 | 5.33 | 5.44 | 5.54 | 5.65 | 5.75 | 5.86 | 5.96 | 6.07 | 6.17 | 6.28 | 6.38 | 6.49 | 6.59 | 6.7 | 6.8 | 6.91 | 7.01 | 7.11 | 7.22 | 7.32 | 7.43 | 7.53 | 7.64 | 7.74 | 7.85 | 7.95 | 8.06
29.41 | 3.87 | 3.97 | 4.08 | 4.18 | 4.29 | 4.39 | 4.5 | 4.6 | 4.71 | 4.81 | 4.91 | 5.02 | 5.12 | 5.23 | 5.33 | 5.44 | 5.54 | 5.65 | 5.75 | 5.86 | 5.96 | 6.07 | 6.17 | 6.28 | 6.38 | 6.49 | 6.59 | 6.69 | 6.8 | 6.9 | 7.01 | 7.11 | 7.22 | 7.32 | 7.43 | 7.53 | 7.64 | 7.74 | 7.85 | 7.95 | 8.06
29.42 | 3.87 | 3.97 | 4.08 | 4.18 | 4.29 | 4.39 | 4.5 | 4.6 | 4.7 | 4.81 | 4.91 | 5.02 | 5.12 | 5.23 | 5.33 | 5.44 | 5.54 | 5.65 | 5.75 | 5.86 | 5.96 | 6.07 | 6.17 | 6.27 | 6.38 | 6.48 | 6.59 | 6.69 | 6.8 | 6.9 | 7.01 | 7.11 | 7.22 | 7.32 | 7.43 | 7.53 | 7.64 | 7.74 | 7.84 | 7.95 | 8.05
29.43 | 3.87 | 3.97 | 4.08 | 4.18 | 4.28 | 4.39 | 4.49 | 4.6 | 4.7 | 4.81 | 4.91 | 5.02 | 5.12 | 5.23 | 5.33 | 5.44 | 5.54 | 5.65 | 5.75 | 5.85 | 5.96 | 6.06 | 6.17 | 6.27 | 6.38 | 6.48 | 6.59 | 6.69 | 6.8 | 6.9 | 7.01 | 7.11 | 7.22 | 7.32 | 7.42 | 7.53 | 7.63 | 7.74 | 7.84 | 7.95 | 8.05
29.44 | 3.87 | 3.97 | 4.07 | 4.18 | 4.28 | 4.39 | 4.49 | 4.6 | 4.7 | 4.81 | 4.91 | 5.02 | 5.12 | 5.23 | 5.33 | 5.44 | 5.54 | 5.64 | 5.75 | 5.85 | 5.96 | 6.06 | 6.17 | 6.27 | 6.38 | 6.48 | 6.59 | 6.69 | 6.8 | 6.9 | 7 | 7.11 | 7.21 | 7.32 | 7.42 | 7.53 | 7.63 | 7.74 | 7.84 | 7.95 | 8.05
29.45 | 3.86 | 3.97 | 4.07 | 4.18 | 4.28 | 4.39 | 4.49 | 4.6 | 4.7 | 4.81 | 4.91 | 5.02 | 5.12 | 5.23 | 5.33 | 5.43 | 5.54 | 5.64 | 5.75 | 5.85 | 5.96 | 6.06 | 6.17 | 6.27 | 6.38 | 6.48 | 6.59 | 6.69 | 6.79 | 6.9 | 7 | 7.11 | 7.21 | 7.32 | 7.42 | 7.53 | 7.63 | 7.74 | 7.84 | 7.94 | 8.05
29.46 | 3.86 | 3.97 | 4.07 | 4.18 | 4.28 | 4.39 | 4.49 | 4.6 | 4.7 | 4.81 | 4.91 | 5.01 | 5.12 | 5.22 | 5.33 | 5.43 | 5.54 | 5.64 | 5.75 | 5.85 | 5.96 | 6.06 | 6.17 | 6.27 | 6.37 | 6.48 | 6.58 | 6.69 | 6.79 | 6.9 | 7 | 7.11 | 7.21 | 7.32 | 7.42 | 7.53 | 7.63 | 7.73 | 7.84 | 7.94 | 8.05
29.47 | 3.86 | 3.97 | 4.07 | 4.18 | 4.28 | 4.39 | 4.49 | 4.6 | 4.7 | 4.8 | 4.91 | 5.01 | 5.12 | 5.22 | 5.33 | 5.43 | 5.54 | 5.64 | 5.75 | 5.85 | 5.96 | 6.06 | 6.16 | 6.27 | 6.37 | 6.48 | 6.58 | 6.69 | 6.79 | 6.9 | 7 | 7.11 | 7.21 | 7.31 | 7.42 | 7.52 | 7.63 | 7.73 | 7.84 | 7.94 | 8.05
29.48 | 3.86 | 3.97 | 4.07 | 4.18 | 4.28 | 4.39 | 4.49 | 4.59 | 4.7 | 4.8 | 4.91 | 5.01 | 5.12 | 5.22 | 5.33 | 5.43 | 5.54 | 5.64 | 5.75 | 5.85 | 5.95 | 6.06 | 6.16 | 6.27 | 6.37 | 6.48 | 6.58 | 6.69 | 6.79 | 6.9 | 7 | 7.1 | 7.21 | 7.31 | 7.42 | 7.52 | 7.63 | 7.73 | 7.84 | 7.94 | 8.05
29.49 | 3.86 | 3.97 | 4.07 | 4.18 | 4.28 | 4.38 | 4.49 | 4.59 | 4.7 | 4.8 | 4.91 | 5.01 | 5.12 | 5.22 | 5.33 | 5.43 | 5.53 | 5.64 | 5.74 | 5.85 | 5.95 | 6.06 | 6.16 | 6.27 | 6.37 | 6.48 | 6.58 | 6.69 | 6.79 | 6.89 | 7 | 7.1 | 7.21 | 7.31 | 7.42 | 7.52 | 7.63 | 7.73 | 7.83 | 7.94 | 8.04
29.5 | 3.86 | 3.97 | 4.07 | 4.17 | 4.28 | 4.38 | 4.49 | 4.59 | 4.7 | 4.8 | 4.91 | 5.01 | 5.12 | 5.22 | 5.32 | 5.43 | 5.53 | 5.64 | 5.74 | 5.85 | 5.95 | 6.06 | 6.16 | 6.27 | 6.37 | 6.47 | 6.58 | 6.68 | 6.79 | 6.89 | 7 | 7.1 | 7.21 | 7.31 | 7.42 | 7.52 | 7.62 | 7.73 | 7.83 | 7.94 | 8.04
29.51 | 3.86 | 3.96 | 4.07 | 4.17 | 4.28 | 4.38 | 4.49 | 4.59 | 4.7 | 4.8 | 4.91 | 5.01 | 5.11 | 5.22 | 5.32 | 5.43 | 5.53 | 5.64 | 5.74 | 5.85 | 5.95 | 6.06 | 6.16 | 6.26 | 6.37 | 6.47 | 6.58 | 6.68 | 6.79 | 6.89 | 7 | 7.1 | 7.21 | 7.31 | 7.41 | 7.52 | 7.62 | 7.73 | 7.83 | 7.94 | 8.04
29.52 | 3.86 | 3.96 | 4.07 | 4.17 | 4.28 | 4.38 | 4.49 | 4.59 | 4.7 | 4.8 | 4.9 | 5.01 | 5.11 | 5.22 | 5.32 | 5.43 | 5.53 | 5.64 | 5.74 | 5.85 | 5.95 | 6.05 | 6.16 | 6.26 | 6.37 | 6.47 | 6.58 | 6.68 | 6.79 | 6.89 | 6.99 | 7.1 | 7.2 | 7.31 | 7.41 | 7.52 | 7.62 | 7.73 | 7.83 | 7.94 | 8.04
29.53 | 3.86 | 3.96 | 4.07 | 4.17 | 4.28 | 4.38 | 4.49 | 4.59 | 4.69 | 4.8 | 4.9 | 5.01 | 5.11 | 5.22 | 5.32 | 5.43 | 5.53 | 5.64 | 5.74 | 5.84 | 5.95 | 6.05 | 6.16 | 6.26 | 6.37 | 6.47 | 6.58 | 6.68 | 6.78 | 6.89 | 6.99 | 7.1 | 7.2 | 7.31 | 7.41 | 7.52 | 7.62 | 7.73 | 7.83 | 7.93 | 8.04
29.54 | 3.86 | 3.96 | 4.07 | 4.17 | 4.28 | 4.38 | 4.48 | 4.59 | 4.69 | 4.8 | 4.9 | 5.01 | 5.11 | 5.22 | 5.32 | 5.43 | 5.53 | 5.63 | 5.74 | 5.84 | 5.95 | 6.05 | 6.16 | 6.26 | 6.37 | 6.47 | 6.57 | 6.68 | 6.78 | 6.89 | 6.99 | 7.1 | 7.2 | 7.31 | 7.41 | 7.51 | 7.62 | 7.72 | 7.83 | 7.93 | 8.04
29.55 | 3.86 | 3.96 | 4.07 | 4.17 | 4.28 | 4.38 | 4.48 | 4.59 | 4.69 | 4.8 | 4.9 | 5.01 | 5.11 | 5.22 | 5.32 | 5.42 | 5.53 | 5.63 | 5.74 | 5.84 | 5.95 | 6.05 | 6.16 | 6.26 | 6.36 | 6.47 | 6.57 | 6.68 | 6.78 | 6.89 | 6.99 | 7.1 | 7.2 | 7.3 | 7.41 | 7.51 | 7.62 | 7.72 | 7.83 | 7.93 | 8.04
29.56 | 3.86 | 3.96 | 4.07 | 4.17 | 4.27 | 4.38 | 4.48 | 4.59 | 4.69 | 4.8 | 4.9 | 5.01 | 5.11 | 5.21 | 5.32 | 5.42 | 5.53 | 5.63 | 5.74 | 5.84 | 5.95 | 6.05 | 6.15 | 6.26 | 6.36 | 6.47 | 6.57 | 6.68 | 6.78 | 6.89 | 6.99 | 7.09 | 7.2 | 7.3 | 7.41 | 7.51 | 7.62 | 7.72 | 7.83 | 7.93 | 8.03
29.57 | 3.86 | 3.96 | 4.06 | 4.17 | 4.27 | 4.38 | 4.48 | 4.59 | 4.69 | 4.8 | 4.9 | 5 | 5.11 | 5.21 | 5.32 | 5.42 | 5.53 | 5.63 | 5.74 | 5.84 | 5.94 | 6.05 | 6.15 | 6.26 | 6.36 | 6.47 | 6.57 | 6.68 | 6.78 | 6.88 | 6.99 | 7.09 | 7.2 | 7.3 | 7.41 | 7.51 | 7.62 | 7.72 | 7.82 | 7.93 | 8.03
29.58 | 3.85 | 3.96 | 4.06 | 4.17 | 4.27 | 4.38 | 4.48 | 4.59 | 4.69 | 4.79 | 4.9 | 5 | 5.11 | 5.21 | 5.32 | 5.42 | 5.53 | 5.63 | 5.73 | 5.84 | 5.94 | 6.05 | 6.15 | 6.26 | 6.36 | 6.47 | 6.57 | 6.67 | 6.78 | 6.88 | 6.99 | 7.09 | 7.2 | 7.3 | 7.41 | 7.51 | 7.61 | 7.72 | 7.82 | 7.93 | 8.03
29.59 | 3.85 | 3.96 | 4.06 | 4.17 | 4.27 | 4.38 | 4.48 | 4.59 | 4.69 | 4.79 | 4.9 | 5 | 5.11 | 5.21 | 5.32 | 5.42 | 5.52 | 5.63 | 5.73 | 5.84 | 5.94 | 6.05 | 6.15 | 6.26 | 6.36 | 6.46 | 6.57 | 6.67 | 6.78 | 6.88 | 6.99 | 7.09 | 7.2 | 7.3 | 7.4 | 7.51 | 7.61 | 7.72 | 7.82 | 7.93 | 8.03
29.6 | 3.85 | 3.96 | 4.06 | 4.17 | 4.27 | 4.38 | 4.48 | 4.58 | 4.69 | 4.79 | 4.9 | 5 | 5.11 | 5.21 | 5.31 | 5.42 | 5.52 | 5.63 | 5.73 | 5.84 | 5.94 | 6.05 | 6.15 | 6.25 | 6.36 | 6.46 | 6.57 | 6.67 | 6.78 | 6.88 | 6.99 | 7.09 | 7.19 | 7.3 | 7.4 | 7.51 | 7.61 | 7.72 | 7.82 | 7.92 | 8.03
29.61 | 3.85 | 3.96 | 4.06 | 4.17 | 4.27 | 4.37 | 4.48 | 4.58 | 4.69 | 4.79 | 4.9 | 5 | 5.11 | 5.21 | 5.31 | 5.42 | 5.52 | 5.63 | 5.73 | 5.84 | 5.94 | 6.04 | 6.15 | 6.25 | 6.36 | 6.46 | 6.57 | 6.67 | 6.78 | 6.88 | 6.98 | 7.09 | 7.19 | 7.3 | 7.4 | 7.51 | 7.61 | 7.71 | 7.82 | 7.92 | 8.03
29.62 | 3.85 | 3.96 | 4.06 | 4.16 | 4.27 | 4.37 | 4.48 | 4.58 | 4.69 | 4.79 | 4.9 | 5 | 5.1 | 5.21 | 5.31 | 5.42 | 5.52 | 5.63 | 5.73 | 5.83 | 5.94 | 6.04 | 6.15 | 6.25 | 6.36 | 6.46 | 6.57 | 6.67 | 6.77 | 6.88 | 6.98 | 7.09 | 7.19 | 7.3 | 7.4 | 7.5 | 7.61 | 7.71 | 7.82 | 7.92 | 8.03
29.63 | 3.85 | 3.96 | 4.06 | 4.16 | 4.27 | 4.37 | 4.48 | 4.58 | 4.69 | 4.79 | 4.89 | 5 | 5.1 | 5.21 | 5.31 | 5.42 | 5.52 | 5.63 | 5.73 | 5.83 | 5.94 | 6.04 | 6.15 | 6.25 | 6.36 | 6.46 | 6.56 | 6.67 | 6.77 | 6.88 | 6.98 | 7.09 | 7.19 | 7.29 | 7.4 | 7.5 | 7.61 | 7.71 | 7.82 | 7.92 | 8.02
29.64 | 3.85 | 3.95 | 4.06 | 4.16 | 4.27 | 4.37 | 4.48 | 4.58 | 4.69 | 4.79 | 4.89 | 5 | 5.1 | 5.21 | 5.31 | 5.42 | 5.52 | 5.62 | 5.73 | 5.83 | 5.94 | 6.04 | 6.15 | 6.25 | 6.35 | 6.46 | 6.56 | 6.67 | 6.77 | 6.88 | 6.98 | 7.08 | 7.19 | 7.29 | 7.4 | 7.5 | 7.61 | 7.71 | 7.81 | 7.92 | 8.02
29.65 | 3.85 | 3.95 | 4.06 | 4.16 | 4.27 | 4.37 | 4.48 | 4.58 | 4.68 | 4.79 | 4.89 | 5 | 5.1 | 5.21 | 5.31 | 5.41 | 5.52 | 5.62 | 5.73 | 5.83 | 5.94 | 6.04 | 6.14 | 6.25 | 6.35 | 6.46 | 6.56 | 6.67 | 6.77 | 6.87 | 6.98 | 7.08 | 7.19 | 7.29 | 7.4 | 7.5 | 7.6 | 7.71 | 7.81 | 7.92 | 8.02
29.66 | 3.85 | 3.95 | 4.06 | 4.16 | 4.27 | 4.37 | 4.47 | 4.58 | 4.68 | 4.79 | 4.89 | 5 | 5.1 | 5.2 | 5.31 | 5.41 | 5.52 | 5.62 | 5.73 | 5.83 | 5.93 | 6.04 | 6.14 | 6.25 | 6.35 | 6.46 | 6.56 | 6.66 | 6.77 | 6.87 | 6.98 | 7.08 | 7.19 | 7.29 | 7.39 | 7.5 | 7.6 | 7.71 | 7.81 | 7.92 | 8.02
29.67 | 3.85 | 3.95 | 4.06 | 4.16 | 4.27 | 4.37 | 4.47 | 4.58 | 4.68 | 4.79 | 4.89 | 5 | 5.1 | 5.2 | 5.31 | 5.41 | 5.52 | 5.62 | 5.73 | 5.83 | 5.93 | 6.04 | 6.14 | 6.25 | 6.35 | 6.46 | 6.56 | 6.66 | 6.77 | 6.87 | 6.98 | 7.08 | 7.19 | 7.29 | 7.39 | 7.5 | 7.6 | 7.71 | 7.81 | 7.91 | 8.02
29.68 | 3.85 | 3.95 | 4.06 | 4.16 | 4.26 | 4.37 | 4.47 | 4.58 | 4.68 | 4.79 | 4.89 | 4.99 | 5.1 | 5.2 | 5.31 | 5.41 | 5.52 | 5.62 | 5.72 | 5.83 | 5.93 | 6.04 | 6.14 | 6.25 | 6.35 | 6.45 | 6.56 | 6.66 | 6.77 | 6.87 | 6.98 | 7.08 | 7.18 | 7.29 | 7.39 | 7.5 | 7.6 | 7.71 | 7.81 | 7.91 | 8.02
29.69 | 3.85 | 3.95 | 4.05 | 4.16 | 4.26 | 4.37 | 4.47 | 4.58 | 4.68 | 4.78 | 4.89 | 4.99 | 5.1 | 5.2 | 5.31 | 5.41 | 5.51 | 5.62 | 5.72 | 5.83 | 5.93 | 6.04 | 6.14 | 6.24 | 6.35 | 6.45 | 6.56 | 6.66 | 6.77 | 6.87 | 6.97 | 7.08 | 7.18 | 7.29 | 7.39 | 7.5 | 7.6 | 7.7 | 7.81 | 7.91 | 8.02
29.7 | 3.85 | 3.95 | 4.05 | 4.16 | 4.26 | 4.37 | 4.47 | 4.58 | 4.68 | 4.78 | 4.89 | 4.99 | 5.1 | 5.2 | 5.31 | 5.41 | 5.51 | 5.62 | 5.72 | 5.83 | 5.93 | 6.03 | 6.14 | 6.24 | 6.35 | 6.45 | 6.56 | 6.66 | 6.76 | 6.87 | 6.97 | 7.08 | 7.18 | 7.29 | 7.39 | 7.49 | 7.6 | 7.7 | 7.81 | 7.91 | 8.01
29.71 | 3.84 | 3.95 | 4.05 | 4.16 | 4.26 | 4.37 | 4.47 | 4.57 | 4.68 | 4.78 | 4.89 | 4.99 | 5.1 | 5.2 | 5.3 | 5.41 | 5.51 | 5.62 | 5.72 | 5.83 | 5.93 | 6.03 | 6.14 | 6.24 | 6.35 | 6.45 | 6.55 | 6.66 | 6.76 | 6.87 | 6.97 | 7.08 | 7.18 | 7.28 | 7.39 | 7.49 | 7.6 | 7.7 | 7.81 | 7.91 | 8.01
29.72 | 3.84 | 3.95 | 4.05 | 4.16 | 4.26 | 4.37 | 4.47 | 4.57 | 4.68 | 4.78 | 4.89 | 4.99 | 5.09 | 5.2 | 5.3 | 5.41 | 5.51 | 5.62 | 5.72 | 5.82 | 5.93 | 6.03 | 6.14 | 6.24 | 6.35 | 6.45 | 6.55 | 6.66 | 6.76 | 6.87 | 6.97 | 7.07 | 7.18 | 7.28 | 7.39 | 7.49 | 7.6 | 7.7 | 7.8 | 7.91 | 8.01
29.73 | 3.84 | 3.95 | 4.05 | 4.16 | 4.26 | 4.36 | 4.47 | 4.57 | 4.68 | 4.78 | 4.89 | 4.99 | 5.09 | 5.2 | 5.3 | 5.41 | 5.51 | 5.61 | 5.72 | 5.82 | 5.93 | 6.03 | 6.14 | 6.24 | 6.34 | 6.45 | 6.55 | 6.66 | 6.76 | 6.87 | 6.97 | 7.07 | 7.18 | 7.28 | 7.39 | 7.49 | 7.59 | 7.7 | 7.8 | 7.91 | 8.01
29.74 | 3.84 | 3.95 | 4.05 | 4.16 | 4.26 | 4.36 | 4.47 | 4.57 | 4.68 | 4.78 | 4.88 | 4.99 | 5.09 | 5.2 | 5.3 | 5.41 | 5.51 | 5.61 | 5.72 | 5.82 | 5.93 | 6.03 | 6.13 | 6.24 | 6.34 | 6.45 | 6.55 | 6.66 | 6.76 | 6.86 | 6.97 | 7.07 | 7.18 | 7.28 | 7.38 | 7.49 | 7.59 | 7.7 | 7.8 | 7.91 | 8.01
29.75 | 3.84 | 3.95 | 4.05 | 4.15 | 4.26 | 4.36 | 4.47 | 4.57 | 4.68 | 4.78 | 4.88 | 4.99 | 5.09 | 5.2 | 5.3 | 5.4 | 5.51 | 5.61 | 5.72 | 5.82 | 5.93 | 6.03 | 6.13 | 6.24 | 6.34 | 6.45 | 6.55 | 6.65 | 6.76 | 6.86 | 6.97 | 7.07 | 7.18 | 7.28 | 7.38 | 7.49 | 7.59 | 7.7 | 7.8 | 7.9 | 8.01
29.76 | 3.84 | 3.95 | 4.05 | 4.15 | 4.26 | 4.36 | 4.47 | 4.57 | 4.67 | 4.78 | 4.88 | 4.99 | 5.09 | 5.2 | 5.3 | 5.4 | 5.51 | 5.61 | 5.72 | 5.82 | 5.92 | 6.03 | 6.13 | 6.24 | 6.34 | 6.44 | 6.55 | 6.65 | 6.76 | 6.86 | 6.97 | 7.07 | 7.17 | 7.28 | 7.38 | 7.49 | 7.59 | 7.69 | 7.8 | 7.9 | 8.01
29.77 | 3.84 | 3.94 | 4.05 | 4.15 | 4.26 | 4.36 | 4.47 | 4.57 | 4.67 | 4.78 | 4.88 | 4.99 | 5.09 | 5.19 | 5.3 | 5.4 | 5.51 | 5.61 | 5.71 | 5.82 | 5.92 | 6.03 | 6.13 | 6.24 | 6.34 | 6.44 | 6.55 | 6.65 | 6.76 | 6.86 | 6.96 | 7.07 | 7.17 | 7.28 | 7.38 | 7.48 | 7.59 | 7.69 | 7.8 | 7.9 | 8.01
29.78 | 3.84 | 3.94 | 4.05 | 4.15 | 4.26 | 4.36 | 4.46 | 4.57 | 4.67 | 4.78 | 4.88 | 4.99 | 5.09 | 5.19 | 5.3 | 5.4 | 5.51 | 5.61 | 5.71 | 5.82 | 5.92 | 6.03 | 6.13 | 6.23 | 6.34 | 6.44 | 6.55 | 6.65 | 6.75 | 6.86 | 6.96 | 7.07 | 7.17 | 7.28 | 7.38 | 7.48 | 7.59 | 7.69 | 7.8 | 7.9 | 8
29.79 | 3.84 | 3.94 | 4.05 | 4.15 | 4.26 | 4.36 | 4.46 | 4.57 | 4.67 | 4.78 | 4.88 | 4.98 | 5.09 | 5.19 | 5.3 | 5.4 | 5.5 | 5.61 | 5.71 | 5.82 | 5.92 | 6.03 | 6.13 | 6.23 | 6.34 | 6.44 | 6.55 | 6.65 | 6.75 | 6.86 | 6.96 | 7.07 | 7.17 | 7.27 | 7.38 | 7.48 | 7.59 | 7.69 | 7.79 | 7.9 | 8
29.8 | 3.84 | 3.94 | 4.05 | 4.15 | 4.25 | 4.36 | 4.46 | 4.57 | 4.67 | 4.77 | 4.88 | 4.98 | 5.09 | 5.19 | 5.3 | 5.4 | 5.5 | 5.61 | 5.71 | 5.82 | 5.92 | 6.02 | 6.13 | 6.23 | 6.34 | 6.44 | 6.54 | 6.65 | 6.75 | 6.86 | 6.96 | 7.06 | 7.17 | 7.27 | 7.38 | 7.48 | 7.58 | 7.69 | 7.79 | 7.9 | 8
29.81 | 3.84 | 3.94 | 4.05 | 4.15 | 4.25 | 4.36 | 4.46 | 4.57 | 4.67 | 4.77 | 4.88 | 4.98 | 5.09 | 5.19 | 5.29 | 5.4 | 5.5 | 5.61 | 5.71 | 5.81 | 5.92 | 6.02 | 6.13 | 6.23 | 6.34 | 6.44 | 6.54 | 6.65 | 6.75 | 6.86 | 6.96 | 7.06 | 7.17 | 7.27 | 7.38 | 7.48 | 7.58 | 7.69 | 7.79 | 7.9 | 8
29.82 | 3.84 | 3.94 | 4.04 | 4.15 | 4.25 | 4.36 | 4.46 | 4.57 | 4.67 | 4.77 | 4.88 | 4.98 | 5.09 | 5.19 | 5.29 | 5.4 | 5.5 | 5.61 | 5.71 | 5.81 | 5.92 | 6.02 | 6.13 | 6.23 | 6.33 | 6.44 | 6.54 | 6.65 | 6.75 | 6.85 | 6.96 | 7.06 | 7.17 | 7.27 | 7.37 | 7.48 | 7.58 | 7.69 | 7.79 | 7.89 | 8
29.83 | 3.84 | 3.94 | 4.04 | 4.15 | 4.25 | 4.36 | 4.46 | 4.56 | 4.67 | 4.77 | 4.88 | 4.98 | 5.08 | 5.19 | 5.29 | 5.4 | 5.5 | 5.6 | 5.71 | 5.81 | 5.92 | 6.02 | 6.12 | 6.23 | 6.33 | 6.44 | 6.54 | 6.64 | 6.75 | 6.85 | 6.96 | 7.06 | 7.17 | 7.27 | 7.37 | 7.48 | 7.58 | 7.69 | 7.79 | 7.89 | 8
29.84 | 3.84 | 3.94 | 4.04 | 4.15 | 4.25 | 4.36 | 4.46 | 4.56 | 4.67 | 4.77 | 4.88 | 4.98 | 5.08 | 5.19 | 5.29 | 5.4 | 5.5 | 5.6 | 5.71 | 5.81 | 5.92 | 6.02 | 6.12 | 6.23 | 6.33 | 6.44 | 6.54 | 6.64 | 6.75 | 6.85 | 6.96 | 7.06 | 7.16 | 7.27 | 7.37 | 7.48 | 7.58 | 7.68 | 7.79 | 7.89 | 8
29.85 | 3.83 | 3.94 | 4.04 | 4.15 | 4.25 | 4.35 | 4.46 | 4.56 | 4.67 | 4.77 | 4.87 | 4.98 | 5.08 | 5.19 | 5.29 | 5.39 | 5.5 | 5.6 | 5.71 | 5.81 | 5.91 | 6.02 | 6.12 | 6.23 | 6.33 | 6.43 | 6.54 | 6.64 | 6.75 | 6.85 | 6.95 | 7.06 | 7.16 | 7.27 | 7.37 | 7.47 | 7.58 | 7.68 | 7.79 | 7.89 | 7.99
29.86 | 3.83 | 3.94 | 4.04 | 4.15 | 4.25 | 4.35 | 4.46 | 4.56 | 4.67 | 4.77 | 4.87 | 4.98 | 5.08 | 5.19 | 5.29 | 5.39 | 5.5 | 5.6 | 5.71 | 5.81 | 5.91 | 6.02 | 6.12 | 6.23 | 6.33 | 6.43 | 6.54 | 6.64 | 6.75 | 6.85 | 6.95 | 7.06 | 7.16 | 7.27 | 7.37 | 7.47 | 7.58 | 7.68 | 7.79 | 7.89 | 7.99
29.87 | 3.83 | 3.94 | 4.04 | 4.14 | 4.25 | 4.35 | 4.46 | 4.56 | 4.66 | 4.77 | 4.87 | 4.98 | 5.08 | 5.18 | 5.29 | 5.39 | 5.5 | 5.6 | 5.7 | 5.81 | 5.91 | 6.02 | 6.12 | 6.22 | 6.33 | 6.43 | 6.54 | 6.64 | 6.74 | 6.85 | 6.95 | 7.06 | 7.16 | 7.26 | 7.37 | 7.47 | 7.58 | 7.68 | 7.78 | 7.89 | 7.99
29.88 | 3.83 | 3.94 | 4.04 | 4.14 | 4.25 | 4.35 | 4.46 | 4.56 | 4.66 | 4.77 | 4.87 | 4.98 | 5.08 | 5.18 | 5.29 | 5.39 | 5.5 | 5.6 | 5.7 | 5.81 | 5.91 | 6.02 | 6.12 | 6.22 | 6.33 | 6.43 | 6.54 | 6.64 | 6.74 | 6.85 | 6.95 | 7.05 | 7.16 | 7.26 | 7.37 | 7.47 | 7.57 | 7.68 | 7.78 | 7.89 | 7.99
29.89 | 3.83 | 3.94 | 4.04 | 4.14 | 4.25 | 4.35 | 4.46 | 4.56 | 4.66 | 4.77 | 4.87 | 4.97 | 5.08 | 5.18 | 5.29 | 5.39 | 5.49 | 5.6 | 5.7 | 5.81 | 5.91 | 6.01 | 6.12 | 6.22 | 6.33 | 6.43 | 6.53 | 6.64 | 6.74 | 6.85 | 6.95 | 7.05 | 7.16 | 7.26 | 7.37 | 7.47 | 7.57 | 7.68 | 7.78 | 7.88 | 7.99
29.9 | 3.83 | 3.93 | 4.04 | 4.14 | 4.25 | 4.35 | 4.45 | 4.56 | 4.66 | 4.77 | 4.87 | 4.97 | 5.08 | 5.18 | 5.29 | 5.39 | 5.49 | 5.6 | 5.7 | 5.81 | 5.91 | 6.01 | 6.12 | 6.22 | 6.32 | 6.43 | 6.53 | 6.64 | 6.74 | 6.84 | 6.95 | 7.05 | 7.16 | 7.26 | 7.36 | 7.47 | 7.57 | 7.68 | 7.78 | 7.88 | 7.99
29.91 | 3.83 | 3.93 | 4.04 | 4.14 | 4.25 | 4.35 | 4.45 | 4.56 | 4.66 | 4.77 | 4.87 | 4.97 | 5.08 | 5.18 | 5.28 | 5.39 | 5.49 | 5.6 | 5.7 | 5.8 | 5.91 | 6.01 | 6.12 | 6.22 | 6.32 | 6.43 | 6.53 | 6.64 | 6.74 | 6.84 | 6.95 | 7.05 | 7.16 | 7.26 | 7.36 | 7.47 | 7.57 | 7.67 | 7.78 | 7.88 | 7.99
29.92 | 3.83 | 3.93 | 4.04 | 4.14 | 4.24 | 4.35 | 4.45 | 4.56 | 4.66 | 4.76 | 4.87 | 4.97 | 5.08 | 5.18 | 5.28 | 5.39 | 5.49 | 5.6 | 5.7 | 5.8 | 5.91 | 6.01 | 6.11 | 6.22 | 6.32 | 6.43 | 6.53 | 6.63 | 6.74 | 6.84 | 6.95 | 7.05 | 7.15 | 7.26 | 7.36 | 7.47 | 7.57 | 7.67 | 7.78 | 7.88 | 7.98
29.93 | 3.83 | 3.93 | 4.04 | 4.14 | 4.24 | 4.35 | 4.45 | 4.56 | 4.66 | 4.76 | 4.87 | 4.97 | 5.08 | 5.18 | 5.28 | 5.39 | 5.49 | 5.59 | 5.7 | 5.8 | 5.91 | 6.01 | 6.11 | 6.22 | 6.32 | 6.43 | 6.53 | 6.63 | 6.74 | 6.84 | 6.94 | 7.05 | 7.15 | 7.26 | 7.36 | 7.46 | 7.57 | 7.67 | 7.78 | 7.88 | 7.98
29.94 | 3.83 | 3.93 | 4.04 | 4.14 | 4.24 | 4.35 | 4.45 | 4.55 | 4.66 | 4.76 | 4.87 | 4.97 | 5.07 | 5.18 | 5.28 | 5.39 | 5.49 | 5.59 | 5.7 | 5.8 | 5.91 | 6.01 | 6.11 | 6.22 | 6.32 | 6.42 | 6.53 | 6.63 | 6.74 | 6.84 | 6.94 | 7.05 | 7.15 | 7.26 | 7.36 | 7.46 | 7.57 | 7.67 | 7.77 | 7.88 | 7.98
29.95 | 3.83 | 3.93 | 4.03 | 4.14 | 4.24 | 4.35 | 4.45 | 4.55 | 4.66 | 4.76 | 4.87 | 4.97 | 5.07 | 5.18 | 5.28 | 5.38 | 5.49 | 5.59 | 5.7 | 5.8 | 5.9 | 6.01 | 6.11 | 6.22 | 6.32 | 6.42 | 6.53 | 6.63 | 6.73 | 6.84 | 6.94 | 7.05 | 7.15 | 7.25 | 7.36 | 7.46 | 7.57 | 7.67 | 7.77 | 7.88 | 7.98
29.96 | 3.83 | 3.93 | 4.03 | 4.14 | 4.24 | 4.35 | 4.45 | 4.55 | 4.66 | 4.76 | 4.86 | 4.97 | 5.07 | 5.18 | 5.28 | 5.38 | 5.49 | 5.59 | 5.7 | 5.8 | 5.9 | 6.01 | 6.11 | 6.21 | 6.32 | 6.42 | 6.53 | 6.63 | 6.73 | 6.84 | 6.94 | 7.04 | 7.15 | 7.25 | 7.36 | 7.46 | 7.56 | 7.67 | 7.77 | 7.88 | 7.98
29.97 | 3.83 | 3.93 | 4.03 | 4.14 | 4.24 | 4.34 | 4.45 | 4.55 | 4.66 | 4.76 | 4.86 | 4.97 | 5.07 | 5.18 | 5.28 | 5.38 | 5.49 | 5.59 | 5.69 | 5.8 | 5.9 | 6.01 | 6.11 | 6.21 | 6.32 | 6.42 | 6.52 | 6.63 | 6.73 | 6.84 | 6.94 | 7.04 | 7.15 | 7.25 | 7.36 | 7.46 | 7.56 | 7.67 | 7.77 | 7.87 | 7.98
29.98 | 3.82 | 3.93 | 4.03 | 4.14 | 4.24 | 4.34 | 4.45 | 4.55 | 4.66 | 4.76 | 4.86 | 4.97 | 5.07 | 5.17 | 5.28 | 5.38 | 5.49 | 5.59 | 5.69 | 5.8 | 5.9 | 6 | 6.11 | 6.21 | 6.32 | 6.42 | 6.52 | 6.63 | 6.73 | 6.83 | 6.94 | 7.04 | 7.15 | 7.25 | 7.35 | 7.46 | 7.56 | 7.67 | 7.77 | 7.87 | 7.98
29.99 | 3.82 | 3.93 | 4.03 | 4.14 | 4.24 | 4.34 | 4.45 | 4.55 | 4.65 | 4.76 | 4.86 | 4.97 | 5.07 | 5.17 | 5.28 | 5.38 | 5.48 | 5.59 | 5.69 | 5.8 | 5.9 | 6 | 6.11 | 6.21 | 6.31 | 6.42 | 6.52 | 6.63 | 6.73 | 6.83 | 6.94 | 7.04 | 7.15 | 7.25 | 7.35 | 7.46 | 7.56 | 7.66 | 7.77 | 7.87 | 7.98
30 | 3.82 | 3.93 | 4.03 | 4.13 | 4.24 | 4.34 | 4.45 | 4.55 | 4.65 | 4.76 | 4.86 | 4.96 | 5.07 | 5.17 | 5.28 | 5.38 | 5.48 | 5.59 | 5.69 | 5.79 | 5.9 | 6 | 6.11 | 6.21 | 6.31 | 6.42 | 6.52 | 6.63 | 6.73 | 6.83 | 6.94 | 7.04 | 7.14 | 7.25 | 7.35 | 7.46 | 7.56 | 7.66 | 7.77 | 7.87 | 7.97
30.01 | 3.82 | 3.93 | 4.03 | 4.13 | 4.24 | 4.34 | 4.44 | 4.55 | 4.65 | 4.76 | 4.86 | 4.96 | 5.07 | 5.17 | 5.28 | 5.38 | 5.48 | 5.59 | 5.69 | 5.79 | 5.9 | 6 | 6.11 | 6.21 | 6.31 | 6.42 | 6.52 | 6.62 | 6.73 | 6.83 | 6.94 | 7.04 | 7.14 | 7.25 | 7.35 | 7.45 | 7.56 | 7.66 | 7.76 | 7.87 | 7.97
30.02 | 3.82 | 3.93 | 4.03 | 4.13 | 4.24 | 4.34 | 4.44 | 4.55 | 4.65 | 4.76 | 4.86 | 4.96 | 5.07 | 5.17 | 5.27 | 5.38 | 5.48 | 5.59 | 5.69 | 5.79 | 5.9 | 6 | 6.1 | 6.21 | 6.31 | 6.42 | 6.52 | 6.62 | 6.73 | 6.83 | 6.93 | 7.04 | 7.14 | 7.25 | 7.35 | 7.45 | 7.56 | 7.66 | 7.76 | 7.87 | 7.97
30.03 | 3.82 | 3.92 | 4.03 | 4.13 | 4.24 | 4.34 | 4.44 | 4.55 | 4.65 | 4.75 | 4.86 | 4.96 | 5.07 | 5.17 | 5.27 | 5.38 | 5.48 | 5.58 | 5.69 | 5.79 | 5.9 | 6 | 6.1 | 6.21 | 6.31 | 6.41 | 6.52 | 6.62 | 6.73 | 6.83 | 6.93 | 7.04 | 7.14 | 7.24 | 7.35 | 7.45 | 7.55 | 7.66 | 7.76 | 7.87 | 7.97
30.04 | 3.82 | 3.92 | 4.03 | 4.13 | 4.23 | 4.34 | 4.44 | 4.55 | 4.65 | 4.75 | 4.86 | 4.96 | 5.06 | 5.17 | 5.27 | 5.38 | 5.48 | 5.58 | 5.69 | 5.79 | 5.89 | 6 | 6.1 | 6.21 | 6.31 | 6.41 | 6.52 | 6.62 | 6.72 | 6.83 | 6.93 | 7.04 | 7.14 | 7.24 | 7.35 | 7.45 | 7.55 | 7.66 | 7.76 | 7.86 | 7.97
30.05 | 3.82 | 3.92 | 4.03 | 4.13 | 4.23 | 4.34 | 4.44 | 4.55 | 4.65 | 4.75 | 4.86 | 4.96 | 5.06 | 5.17 | 5.27 | 5.37 | 5.48 | 5.58 | 5.69 | 5.79 | 5.89 | 6 | 6.1 | 6.2 | 6.31 | 6.41 | 6.52 | 6.62 | 6.72 | 6.83 | 6.93 | 7.03 | 7.14 | 7.24 | 7.34 | 7.45 | 7.55 | 7.66 | 7.76 | 7.86 | 7.97
30.06 | 3.82 | 3.92 | 4.03 | 4.13 | 4.23 | 4.34 | 4.44 | 4.54 | 4.65 | 4.75 | 4.86 | 4.96 | 5.06 | 5.17 | 5.27 | 5.37 | 5.48 | 5.58 | 5.68 | 5.79 | 5.89 | 6 | 6.1 | 6.2 | 6.31 | 6.41 | 6.51 | 6.62 | 6.72 | 6.83 | 6.93 | 7.03 | 7.14 | 7.24 | 7.34 | 7.45 | 7.55 | 7.65 | 7.76 | 7.86 | 7.97
30.07 | 3.82 | 3.92 | 4.03 | 4.13 | 4.23 | 4.34 | 4.44 | 4.54 | 4.65 | 4.75 | 4.85 | 4.96 | 5.06 | 5.17 | 5.27 | 5.37 | 5.48 | 5.58 | 5.68 | 5.79 | 5.89 | 5.99 | 6.1 | 6.2 | 6.31 | 6.41 | 6.51 | 6.62 | 6.72 | 6.82 | 6.93 | 7.03 | 7.14 | 7.24 | 7.34 | 7.45 | 7.55 | 7.65 | 7.76 | 7.86 | 7.96
30.08 | 3.82 | 3.92 | 4.02 | 4.13 | 4.23 | 4.34 | 4.44 | 4.54 | 4.65 | 4.75 | 4.85 | 4.96 | 5.06 | 5.16 | 5.27 | 5.37 | 5.48 | 5.58 | 5.68 | 5.79 | 5.89 | 5.99 | 6.1 | 6.2 | 6.3 | 6.41 | 6.51 | 6.62 | 6.72 | 6.82 | 6.93 | 7.03 | 7.13 | 7.24 | 7.34 | 7.44 | 7.55 | 7.65 | 7.76 | 7.86 | 7.96
30.09 | 3.82 | 3.92 | 4.02 | 4.13 | 4.23 | 4.33 | 4.44 | 4.54 | 4.65 | 4.75 | 4.85 | 4.96 | 5.06 | 5.16 | 5.27 | 5.37 | 5.47 | 5.58 | 5.68 | 5.79 | 5.89 | 5.99 | 6.1 | 6.2 | 6.3 | 6.41 | 6.51 | 6.61 | 6.72 | 6.82 | 6.93 | 7.03 | 7.13 | 7.24 | 7.34 | 7.44 | 7.55 | 7.65 | 7.75 | 7.86 | 7.96
30.1 | 3.82 | 3.92 | 4.02 | 4.13 | 4.23 | 4.33 | 4.44 | 4.54 | 4.64 | 4.75 | 4.85 | 4.96 | 5.06 | 5.16 | 5.27 | 5.37 | 5.47 | 5.58 | 5.68 | 5.78 | 5.89 | 5.99 | 6.1 | 6.2 | 6.3 | 6.41 | 6.51 | 6.61 | 6.72 | 6.82 | 6.92 | 7.03 | 7.13 | 7.23 | 7.34 | 7.44 | 7.55 | 7.65 | 7.75 | 7.86 | 7.96
30.11 | 3.81 | 3.92 | 4.02 | 4.13 | 4.23 | 4.33 | 4.44 | 4.54 | 4.64 | 4.75 | 4.85 | 4.95 | 5.06 | 5.16 | 5.27 | 5.37 | 5.47 | 5.58 | 5.68 | 5.78 | 5.89 | 5.99 | 6.09 | 6.2 | 6.3 | 6.41 | 6.51 | 6.61 | 6.72 | 6.82 | 6.92 | 7.03 | 7.13 | 7.23 | 7.34 | 7.44 | 7.54 | 7.65 | 7.75 | 7.86 | 7.96
30.12 | 3.81 | 3.92 | 4.02 | 4.12 | 4.23 | 4.33 | 4.44 | 4.54 | 4.64 | 4.75 | 4.85 | 4.95 | 5.06 | 5.16 | 5.26 | 5.37 | 5.47 | 5.58 | 5.68 | 5.78 | 5.89 | 5.99 | 6.09 | 6.2 | 6.3 | 6.4 | 6.51 | 6.61 | 6.71 | 6.82 | 6.92 | 7.03 | 7.13 | 7.23 | 7.34 | 7.44 | 7.54 | 7.65 | 7.75 | 7.85 | 7.96
30.13 | 3.81 | 3.92 | 4.02 | 4.12 | 4.23 | 4.33 | 4.43 | 4.54 | 4.64 | 4.75 | 4.85 | 4.95 | 5.06 | 5.16 | 5.26 | 5.37 | 5.47 | 5.57 | 5.68 | 5.78 | 5.88 | 5.99 | 6.09 | 6.2 | 6.3 | 6.4 | 6.51 | 6.61 | 6.71 | 6.82 | 6.92 | 7.02 | 7.13 | 7.23 | 7.33 | 7.44 | 7.54 | 7.65 | 7.75 | 7.85 | 7.96
30.14 | 3.81 | 3.92 | 4.02 | 4.12 | 4.23 | 4.33 | 4.43 | 4.54 | 4.64 | 4.74 | 4.85 | 4.95 | 5.06 | 5.16 | 5.26 | 5.37 | 5.47 | 5.57 | 5.68 | 5.78 | 5.88 | 5.99 | 6.09 | 6.19 | 6.3 | 6.4 | 6.51 | 6.61 | 6.71 | 6.82 | 6.92 | 7.02 | 7.13 | 7.23 | 7.33 | 7.44 | 7.54 | 7.64 | 7.75 | 7.85 | 7.95
30.15 | 3.81 | 3.92 | 4.02 | 4.12 | 4.23 | 4.33 | 4.43 | 4.54 | 4.64 | 4.74 | 4.85 | 4.95 | 5.05 | 5.16 | 5.26 | 5.37 | 5.47 | 5.57 | 5.68 | 5.78 | 5.88 | 5.99 | 6.09 | 6.19 | 6.3 | 6.4 | 6.5 | 6.61 | 6.71 | 6.81 | 6.92 | 7.02 | 7.13 | 7.23 | 7.33 | 7.44 | 7.54 | 7.64 | 7.75 | 7.85 | 7.95
30.16 | 3.81 | 3.91 | 4.02 | 4.12 | 4.23 | 4.33 | 4.43 | 4.54 | 4.64 | 4.74 | 4.85 | 4.95 | 5.05 | 5.16 | 5.26 | 5.36 | 5.47 | 5.57 | 5.67 | 5.78 | 5.88 | 5.99 | 6.09 | 6.19 | 6.3 | 6.4 | 6.5 | 6.61 | 6.71 | 6.81 | 6.92 | 7.02 | 7.12 | 7.23 | 7.33 | 7.43 | 7.54 | 7.64 | 7.74 | 7.85 | 7.95
30.17 | 3.81 | 3.91 | 4.02 | 4.12 | 4.22 | 4.33 | 4.43 | 4.53 | 4.64 | 4.74 | 4.85 | 4.95 | 5.05 | 5.16 | 5.26 | 5.36 | 5.47 | 5.57 | 5.67 | 5.78 | 5.88 | 5.98 | 6.09 | 6.19 | 6.29 | 6.4 | 6.5 | 6.61 | 6.71 | 6.81 | 6.92 | 7.02 | 7.12 | 7.23 | 7.33 | 7.43 | 7.54 | 7.64 | 7.74 | 7.85 | 7.95
30.18 | 3.81 | 3.91 | 4.02 | 4.12 | 4.22 | 4.33 | 4.43 | 4.53 | 4.64 | 4.74 | 4.84 | 4.95 | 5.05 | 5.16 | 5.26 | 5.36 | 5.47 | 5.57 | 5.67 | 5.78 | 5.88 | 5.98 | 6.09 | 6.19 | 6.29 | 6.4 | 6.5 | 6.6 | 6.71 | 6.81 | 6.91 | 7.02 | 7.12 | 7.22 | 7.33 | 7.43 | 7.54 | 7.64 | 7.74 | 7.85 | 7.95
30.19 | 3.81 | 3.91 | 4.02 | 4.12 | 4.22 | 4.33 | 4.43 | 4.53 | 4.64 | 4.74 | 4.84 | 4.95 | 5.05 | 5.15 | 5.26 | 5.36 | 5.46 | 5.57 | 5.67 | 5.78 | 5.88 | 5.98 | 6.09 | 6.19 | 6.29 | 6.4 | 6.5 | 6.6 | 6.71 | 6.81 | 6.91 | 7.02 | 7.12 | 7.22 | 7.33 | 7.43 | 7.53 | 7.64 | 7.74 | 7.84 | 7.95
30.2 | 3.81 | 3.91 | 4.01 | 4.12 | 4.22 | 4.33 | 4.43 | 4.53 | 4.64 | 4.74 | 4.84 | 4.95 | 5.05 | 5.15 | 5.26 | 5.36 | 5.46 | 5.57 | 5.67 | 5.77 | 5.88 | 5.98 | 6.08 | 6.19 | 6.29 | 6.39 | 6.5 | 6.6 | 6.71 | 6.81 | 6.91 | 7.02 | 7.12 | 7.22 | 7.33 | 7.43 | 7.53 | 7.64 | 7.74 | 7.84 | 7.95
30.21 | 3.81 | 3.91 | 4.01 | 4.12 | 4.22 | 4.32 | 4.43 | 4.53 | 4.63 | 4.74 | 4.84 | 4.95 | 5.05 | 5.15 | 5.26 | 5.36 | 5.46 | 5.57 | 5.67 | 5.77 | 5.88 | 5.98 | 6.08 | 6.19 | 6.29 | 6.39 | 6.5 | 6.6 | 6.7 | 6.81 | 6.91 | 7.01 | 7.12 | 7.22 | 7.32 | 7.43 | 7.53 | 7.63 | 7.74 | 7.84 | 7.95
30.22 | 3.81 | 3.91 | 4.01 | 4.12 | 4.22 | 4.32 | 4.43 | 4.53 | 4.63 | 4.74 | 4.84 | 4.94 | 5.05 | 5.15 | 5.25 | 5.36 | 5.46 | 5.57 | 5.67 | 5.77 | 5.88 | 5.98 | 6.08 | 6.19 | 6.29 | 6.39 | 6.5 | 6.6 | 6.7 | 6.81 | 6.91 | 7.01 | 7.12 | 7.22 | 7.32 | 7.43 | 7.53 | 7.63 | 7.74 | 7.84 | 7.94
30.23 | 3.81 | 3.91 | 4.01 | 4.12 | 4.22 | 4.32 | 4.43 | 4.53 | 4.63 | 4.74 | 4.84 | 4.94 | 5.05 | 5.15 | 5.25 | 5.36 | 5.46 | 5.56 | 5.67 | 5.77 | 5.87 | 5.98 | 6.08 | 6.18 | 6.29 | 6.39 | 6.49 | 6.6 | 6.7 | 6.81 | 6.91 | 7.01 | 7.12 | 7.22 | 7.32 | 7.43 | 7.53 | 7.63 | 7.74 | 7.84 | 7.94
30.24 | 3.8 | 3.91 | 4.01 | 4.12 | 4.22 | 4.32 | 4.43 | 4.53 | 4.63 | 4.74 | 4.84 | 4.94 | 5.05 | 5.15 | 5.25 | 5.36 | 5.46 | 5.56 | 5.67 | 5.77 | 5.87 | 5.98 | 6.08 | 6.18 | 6.29 | 6.39 | 6.49 | 6.6 | 6.7 | 6.8 | 6.91 | 7.01 | 7.11 | 7.22 | 7.32 | 7.42 | 7.53 | 7.63 | 7.73 | 7.84 | 7.94
30.25 | 3.8 | 3.91 | 4.01 | 4.11 | 4.22 | 4.32 | 4.42 | 4.53 | 4.63 | 4.73 | 4.84 | 4.94 | 5.05 | 5.15 | 5.25 | 5.36 | 5.46 | 5.56 | 5.67 | 5.77 | 5.87 | 5.98 | 6.08 | 6.18 | 6.29 | 6.39 | 6.49 | 6.6 | 6.7 | 6.8 | 6.91 | 7.01 | 7.11 | 7.22 | 7.32 | 7.42 | 7.53 | 7.63 | 7.73 | 7.84 | 7.94
30.26 | 3.8 | 3.91 | 4.01 | 4.11 | 4.22 | 4.32 | 4.42 | 4.53 | 4.63 | 4.73 | 4.84 | 4.94 | 5.04 | 5.15 | 5.25 | 5.35 | 5.46 | 5.56 | 5.66 | 5.77 | 5.87 | 5.97 | 6.08 | 6.18 | 6.28 | 6.39 | 6.49 | 6.59 | 6.7 | 6.8 | 6.9 | 7.01 | 7.11 | 7.21 | 7.32 | 7.42 | 7.52 | 7.63 | 7.73 | 7.84 | 7.94
30.27 | 3.8 | 3.91 | 4.01 | 4.11 | 4.22 | 4.32 | 4.42 | 4.53 | 4.63 | 4.73 | 4.84 | 4.94 | 5.04 | 5.15 | 5.25 | 5.35 | 5.46 | 5.56 | 5.66 | 5.77 | 5.87 | 5.97 | 6.08 | 6.18 | 6.28 | 6.39 | 6.49 | 6.59 | 6.7 | 6.8 | 6.9 | 7.01 | 7.11 | 7.21 | 7.32 | 7.42 | 7.52 | 7.63 | 7.73 | 7.83 | 7.94
30.28 | 3.8 | 3.91 | 4.01 | 4.11 | 4.22 | 4.32 | 4.42 | 4.53 | 4.63 | 4.73 | 4.84 | 4.94 | 5.04 | 5.15 | 5.25 | 5.35 | 5.46 | 5.56 | 5.66 | 5.77 | 5.87 | 5.97 | 6.08 | 6.18 | 6.28 | 6.39 | 6.49 | 6.59 | 6.7 | 6.8 | 6.9 | 7.01 | 7.11 | 7.21 | 7.32 | 7.42 | 7.52 | 7.63 | 7.73 | 7.83 | 7.94
30.29 | 3.8 | 3.9 | 4.01 | 4.11 | 4.21 | 4.32 | 4.42 | 4.52 | 4.63 | 4.73 | 4.83 | 4.94 | 5.04 | 5.14 | 5.25 | 5.35 | 5.45 | 5.56 | 5.66 | 5.76 | 5.87 | 5.97 | 6.07 | 6.18 | 6.28 | 6.38 | 6.49 | 6.59 | 6.69 | 6.8 | 6.9 | 7 | 7.11 | 7.21 | 7.31 | 7.42 | 7.52 | 7.62 | 7.73 | 7.83 | 7.93
30.3 | 3.8 | 3.9 | 4.01 | 4.11 | 4.21 | 4.32 | 4.42 | 4.52 | 4.63 | 4.73 | 4.83 | 4.94 | 5.04 | 5.14 | 5.25 | 5.35 | 5.45 | 5.56 | 5.66 | 5.76 | 5.87 | 5.97 | 6.07 | 6.18 | 6.28 | 6.38 | 6.49 | 6.59 | 6.69 | 6.8 | 6.9 | 7 | 7.11 | 7.21 | 7.31 | 7.42 | 7.52 | 7.62 | 7.73 | 7.83 | 7.93
30.31 | 3.8 | 3.9 | 4.01 | 4.11 | 4.21 | 4.32 | 4.42 | 4.52 | 4.63 | 4.73 | 4.83 | 4.94 | 5.04 | 5.14 | 5.25 | 5.35 | 5.45 | 5.56 | 5.66 | 5.76 | 5.87 | 5.97 | 6.07 | 6.18 | 6.28 | 6.38 | 6.49 | 6.59 | 6.69 | 6.8 | 6.9 | 7 | 7.11 | 7.21 | 7.31 | 7.42 | 7.52 | 7.62 | 7.73 | 7.83 | 7.93
30.32 | 3.8 | 3.9 | 4.01 | 4.11 | 4.21 | 4.32 | 4.42 | 4.52 | 4.63 | 4.73 | 4.83 | 4.94 | 5.04 | 5.14 | 5.25 | 5.35 | 5.45 | 5.56 | 5.66 | 5.76 | 5.86 | 5.97 | 6.07 | 6.17 | 6.28 | 6.38 | 6.48 | 6.59 | 6.69 | 6.79 | 6.9 | 7 | 7.1 | 7.21 | 7.31 | 7.41 | 7.52 | 7.62 | 7.72 | 7.83 | 7.93
30.33 | 3.8 | 3.9 | 4 | 4.11 | 4.21 | 4.31 | 4.42 | 4.52 | 4.62 | 4.73 | 4.83 | 4.93 | 5.04 | 5.14 | 5.24 | 5.35 | 5.45 | 5.55 | 5.66 | 5.76 | 5.86 | 5.97 | 6.07 | 6.17 | 6.28 | 6.38 | 6.48 | 6.59 | 6.69 | 6.79 | 6.9 | 7 | 7.1 | 7.21 | 7.31 | 7.41 | 7.52 | 7.62 | 7.72 | 7.83 | 7.93
30.34 | 3.8 | 3.9 | 4 | 4.11 | 4.21 | 4.31 | 4.42 | 4.52 | 4.62 | 4.73 | 4.83 | 4.93 | 5.04 | 5.14 | 5.24 | 5.35 | 5.45 | 5.55 | 5.66 | 5.76 | 5.86 | 5.97 | 6.07 | 6.17 | 6.28 | 6.38 | 6.48 | 6.59 | 6.69 | 6.79 | 6.9 | 7 | 7.1 | 7.2 | 7.31 | 7.41 | 7.51 | 7.62 | 7.72 | 7.82 | 7.93
30.35 | 3.8 | 3.9 | 4 | 4.11 | 4.21 | 4.31 | 4.42 | 4.52 | 4.62 | 4.73 | 4.83 | 4.93 | 5.04 | 5.14 | 5.24 | 5.35 | 5.45 | 5.55 | 5.66 | 5.76 | 5.86 | 5.96 | 6.07 | 6.17 | 6.27 | 6.38 | 6.48 | 6.58 | 6.69 | 6.79 | 6.89 | 7 | 7.1 | 7.2 | 7.31 | 7.41 | 7.51 | 7.62 | 7.72 | 7.82 | 7.93
30.36 | 3.8 | 3.9 | 4 | 4.11 | 4.21 | 4.31 | 4.42 | 4.52 | 4.62 | 4.73 | 4.83 | 4.93 | 5.03 | 5.14 | 5.24 | 5.34 | 5.45 | 5.55 | 5.65 | 5.76 | 5.86 | 5.96 | 6.07 | 6.17 | 6.27 | 6.38 | 6.48 | 6.58 | 6.69 | 6.79 | 6.89 | 7 | 7.1 | 7.2 | 7.31 | 7.41 | 7.51 | 7.62 | 7.72 | 7.82 | 7.92
30.37 | 3.8 | 3.9 | 4 | 4.1 | 4.21 | 4.31 | 4.41 | 4.52 | 4.62 | 4.72 | 4.83 | 4.93 | 5.03 | 5.14 | 5.24 | 5.34 | 5.45 | 5.55 | 5.65 | 5.76 | 5.86 | 5.96 | 6.07 | 6.17 | 6.27 | 6.38 | 6.48 | 6.58 | 6.69 | 6.79 | 6.89 | 6.99 | 7.1 | 7.2 | 7.3 | 7.41 | 7.51 | 7.61 | 7.72 | 7.82 | 7.92
30.38 | 3.79 | 3.9 | 4 | 4.1 | 4.21 | 4.31 | 4.41 | 4.52 | 4.62 | 4.72 | 4.83 | 4.93 | 5.03 | 5.14 | 5.24 | 5.34 | 5.45 | 5.55 | 5.65 | 5.76 | 5.86 | 5.96 | 6.06 | 6.17 | 6.27 | 6.37 | 6.48 | 6.58 | 6.68 | 6.79 | 6.89 | 6.99 | 7.1 | 7.2 | 7.3 | 7.41 | 7.51 | 7.61 | 7.72 | 7.82 | 7.92
30.39 | 3.79 | 3.9 | 4 | 4.1 | 4.21 | 4.31 | 4.41 | 4.52 | 4.62 | 4.72 | 4.83 | 4.93 | 5.03 | 5.14 | 5.24 | 5.34 | 5.44 | 5.55 | 5.65 | 5.75 | 5.86 | 5.96 | 6.06 | 6.17 | 6.27 | 6.37 | 6.48 | 6.58 | 6.68 | 6.79 | 6.89 | 6.99 | 7.1 | 7.2 | 7.3 | 7.41 | 7.51 | 7.61 | 7.71 | 7.82 | 7.92
30.4 | 3.79 | 3.9 | 4 | 4.1 | 4.21 | 4.31 | 4.41 | 4.52 | 4.62 | 4.72 | 4.82 | 4.93 | 5.03 | 5.13 | 5.24 | 5.34 | 5.44 | 5.55 | 5.65 | 5.75 | 5.86 | 5.96 | 6.06 | 6.17 | 6.27 | 6.37 | 6.48 | 6.58 | 6.68 | 6.78 | 6.89 | 6.99 | 7.09 | 7.2 | 7.3 | 7.4 | 7.51 | 7.61 | 7.71 | 7.82 | 7.92
30.41 | 3.79 | 3.9 | 4 | 4.1 | 4.2 | 4.31 | 4.41 | 4.51 | 4.62 | 4.72 | 4.82 | 4.93 | 5.03 | 5.13 | 5.24 | 5.34 | 5.44 | 5.55 | 5.65 | 5.75 | 5.86 | 5.96 | 6.06 | 6.16 | 6.27 | 6.37 | 6.47 | 6.58 | 6.68 | 6.78 | 6.89 | 6.99 | 7.09 | 7.2 | 7.3 | 7.4 | 7.51 | 7.61 | 7.71 | 7.81 | 7.92
30.42 | 3.79 | 3.89 | 4 | 4.1 | 4.2 | 4.31 | 4.41 | 4.51 | 4.62 | 4.72 | 4.82 | 4.93 | 5.03 | 5.13 | 5.24 | 5.34 | 5.44 | 5.54 | 5.65 | 5.75 | 5.85 | 5.96 | 6.06 | 6.16 | 6.27 | 6.37 | 6.47 | 6.58 | 6.68 | 6.78 | 6.89 | 6.99 | 7.09 | 7.19 | 7.3 | 7.4 | 7.5 | 7.61 | 7.71 | 7.81 | 7.92
30.43 | 3.79 | 3.89 | 4 | 4.1 | 4.2 | 4.31 | 4.41 | 4.51 | 4.62 | 4.72 | 4.82 | 4.93 | 5.03 | 5.13 | 5.23 | 5.34 | 5.44 | 5.54 | 5.65 | 5.75 | 5.85 | 5.96 | 6.06 | 6.16 | 6.27 | 6.37 | 6.47 | 6.58 | 6.68 | 6.78 | 6.88 | 6.99 | 7.09 | 7.19 | 7.3 | 7.4 | 7.5 | 7.61 | 7.71 | 7.81 | 7.92
30.44 | 3.79 | 3.89 | 4 | 4.1 | 4.2 | 4.31 | 4.41 | 4.51 | 4.61 | 4.72 | 4.82 | 4.92 | 5.03 | 5.13 | 5.23 | 5.34 | 5.44 | 5.54 | 5.65 | 5.75 | 5.85 | 5.96 | 6.06 | 6.16 | 6.26 | 6.37 | 6.47 | 6.57 | 6.68 | 6.78 | 6.88 | 6.99 | 7.09 | 7.19 | 7.3 | 7.4 | 7.5 | 7.6 | 7.71 | 7.81 | 7.91
30.45 | 3.79 | 3.89 | 4 | 4.1 | 4.2 | 4.3 | 4.41 | 4.51 | 4.61 | 4.72 | 4.82 | 4.92 | 5.03 | 5.13 | 5.23 | 5.34 | 5.44 | 5.54 | 5.65 | 5.75 | 5.85 | 5.95 | 6.06 | 6.16 | 6.26 | 6.37 | 6.47 | 6.57 | 6.68 | 6.78 | 6.88 | 6.99 | 7.09 | 7.19 | 7.29 | 7.4 | 7.5 | 7.6 | 7.71 | 7.81 | 7.91
30.46 | 3.79 | 3.89 | 3.99 | 4.1 | 4.2 | 4.3 | 4.41 | 4.51 | 4.61 | 4.72 | 4.82 | 4.92 | 5.03 | 5.13 | 5.23 | 5.33 | 5.44 | 5.54 | 5.64 | 5.75 | 5.85 | 5.95 | 6.06 | 6.16 | 6.26 | 6.37 | 6.47 | 6.57 | 6.67 | 6.78 | 6.88 | 6.98 | 7.09 | 7.19 | 7.29 | 7.4 | 7.5 | 7.6 | 7.71 | 7.81 | 7.91
30.47 | 3.79 | 3.89 | 3.99 | 4.1 | 4.2 | 4.3 | 4.41 | 4.51 | 4.61 | 4.72 | 4.82 | 4.92 | 5.02 | 5.13 | 5.23 | 5.33 | 5.44 | 5.54 | 5.64 | 5.75 | 5.85 | 5.95 | 6.06 | 6.16 | 6.26 | 6.36 | 6.47 | 6.57 | 6.67 | 6.78 | 6.88 | 6.98 | 7.09 | 7.19 | 7.29 | 7.39 | 7.5 | 7.6 | 7.7 | 7.81 | 7.91
30.48 | 3.79 | 3.89 | 3.99 | 4.1 | 4.2 | 4.3 | 4.41 | 4.51 | 4.61 | 4.71 | 4.82 | 4.92 | 5.02 | 5.13 | 5.23 | 5.33 | 5.44 | 5.54 | 5.64 | 5.75 | 5.85 | 5.95 | 6.05 | 6.16 | 6.26 | 6.36 | 6.47 | 6.57 | 6.67 | 6.78 | 6.88 | 6.98 | 7.08 | 7.19 | 7.29 | 7.39 | 7.5 | 7.6 | 7.7 | 7.81 | 7.91
30.49 | 3.79 | 3.89 | 3.99 | 4.1 | 4.2 | 4.3 | 4.4 | 4.51 | 4.61 | 4.71 | 4.82 | 4.92 | 5.02 | 5.13 | 5.23 | 5.33 | 5.43 | 5.54 | 5.64 | 5.74 | 5.85 | 5.95 | 6.05 | 6.16 | 6.26 | 6.36 | 6.47 | 6.57 | 6.67 | 6.77 | 6.88 | 6.98 | 7.08 | 7.19 | 7.29 | 7.39 | 7.5 | 7.6 | 7.7 | 7.8 | 7.91
30.5 | 3.79 | 3.89 | 3.99 | 4.09 | 4.2 | 4.3 | 4.4 | 4.51 | 4.61 | 4.71 | 4.82 | 4.92 | 5.02 | 5.12 | 5.23 | 5.33 | 5.43 | 5.54 | 5.64 | 5.74 | 5.85 | 5.95 | 6.05 | 6.16 | 6.26 | 6.36 | 6.46 | 6.57 | 6.67 | 6.77 | 6.88 | 6.98 | 7.08 | 7.19 | 7.29 | 7.39 | 7.49 | 7.6 | 7.7 | 7.8 | 7.91
30.51 | 3.78 | 3.89 | 3.99 | 4.09 | 4.2 | 4.3 | 4.4 | 4.51 | 4.61 | 4.71 | 4.81 | 4.92 | 5.02 | 5.12 | 5.23 | 5.33 | 5.43 | 5.54 | 5.64 | 5.74 | 5.84 | 5.95 | 6.05 | 6.15 | 6.26 | 6.36 | 6.46 | 6.57 | 6.67 | 6.77 | 6.87 | 6.98 | 7.08 | 7.18 | 7.29 | 7.39 | 7.49 | 7.6 | 7.7 | 7.8 | 7.9
30.52 | 3.78 | 3.89 | 3.99 | 4.09 | 4.2 | 4.3 | 4.4 | 4.5 | 4.61 | 4.71 | 4.81 | 4.92 | 5.02 | 5.12 | 5.23 | 5.33 | 5.43 | 5.53 | 5.64 | 5.74 | 5.84 | 5.95 | 6.05 | 6.15 | 6.26 | 6.36 | 6.46 | 6.56 | 6.67 | 6.77 | 6.87 | 6.98 | 7.08 | 7.18 | 7.29 | 7.39 | 7.49 | 7.59 | 7.7 | 7.8 | 7.9
30.53 | 3.78 | 3.89 | 3.99 | 4.09 | 4.2 | 4.3 | 4.4 | 4.5 | 4.61 | 4.71 | 4.81 | 4.92 | 5.02 | 5.12 | 5.23 | 5.33 | 5.43 | 5.53 | 5.64 | 5.74 | 5.84 | 5.95 | 6.05 | 6.15 | 6.25 | 6.36 | 6.46 | 6.56 | 6.67 | 6.77 | 6.87 | 6.98 | 7.08 | 7.18 | 7.28 | 7.39 | 7.49 | 7.59 | 7.7 | 7.8 | 7.9
30.54 | 3.78 | 3.89 | 3.99 | 4.09 | 4.19 | 4.3 | 4.4 | 4.5 | 4.61 | 4.71 | 4.81 | 4.92 | 5.02 | 5.12 | 5.22 | 5.33 | 5.43 | 5.53 | 5.64 | 5.74 | 5.84 | 5.94 | 6.05 | 6.15 | 6.25 | 6.36 | 6.46 | 6.56 | 6.67 | 6.77 | 6.87 | 6.97 | 7.08 | 7.18 | 7.28 | 7.39 | 7.49 | 7.59 | 7.69 | 7.8 | 7.9
30.55 | 3.78 | 3.88 | 3.99 | 4.09 | 4.19 | 4.3 | 4.4 | 4.5 | 4.61 | 4.71 | 4.81 | 4.91 | 5.02 | 5.12 | 5.22 | 5.33 | 5.43 | 5.53 | 5.63 | 5.74 | 5.84 | 5.94 | 6.05 | 6.15 | 6.25 | 6.36 | 6.46 | 6.56 | 6.66 | 6.77 | 6.87 | 6.97 | 7.08 | 7.18 | 7.28 | 7.38 | 7.49 | 7.59 | 7.69 | 7.8 | 7.9
30.56 | 3.78 | 3.88 | 3.99 | 4.09 | 4.19 | 4.3 | 4.4 | 4.5 | 4.6 | 4.71 | 4.81 | 4.91 | 5.02 | 5.12 | 5.22 | 5.33 | 5.43 | 5.53 | 5.63 | 5.74 | 5.84 | 5.94 | 6.05 | 6.15 | 6.25 | 6.35 | 6.46 | 6.56 | 6.66 | 6.77 | 6.87 | 6.97 | 7.07 | 7.18 | 7.28 | 7.38 | 7.49 | 7.59 | 7.69 | 7.79 | 7.9
30.57 | 3.78 | 3.88 | 3.99 | 4.09 | 4.19 | 4.29 | 4.4 | 4.5 | 4.6 | 4.71 | 4.81 | 4.91 | 5.02 | 5.12 | 5.22 | 5.32 | 5.43 | 5.53 | 5.63 | 5.74 | 5.84 | 5.94 | 6.04 | 6.15 | 6.25 | 6.35 | 6.46 | 6.56 | 6.66 | 6.76 | 6.87 | 6.97 | 7.07 | 7.18 | 7.28 | 7.38 | 7.48 | 7.59 | 7.69 | 7.79 | 7.9
30.58 | 3.78 | 3.88 | 3.99 | 4.09 | 4.19 | 4.29 | 4.4 | 4.5 | 4.6 | 4.71 | 4.81 | 4.91 | 5.01 | 5.12 | 5.22 | 5.32 | 5.43 | 5.53 | 5.63 | 5.73 | 5.84 | 5.94 | 6.04 | 6.15 | 6.25 | 6.35 | 6.45 | 6.56 | 6.66 | 6.76 | 6.87 | 6.97 | 7.07 | 7.18 | 7.28 | 7.38 | 7.48 | 7.59 | 7.69 | 7.79 | 7.9
30.59 | 3.78 | 3.88 | 3.98 | 4.09 | 4.19 | 4.29 | 4.4 | 4.5 | 4.6 | 4.7 | 4.81 | 4.91 | 5.01 | 5.12 | 5.22 | 5.32 | 5.43 | 5.53 | 5.63 | 5.73 | 5.84 | 5.94 | 6.04 | 6.15 | 6.25 | 6.35 | 6.45 | 6.56 | 6.66 | 6.76 | 6.87 | 6.97 | 7.07 | 7.17 | 7.28 | 7.38 | 7.48 | 7.59 | 7.69 | 7.79 | 7.89
30.6 | 3.78 | 3.88 | 3.98 | 4.09 | 4.19 | 4.29 | 4.4 | 4.5 | 4.6 | 4.7 | 4.81 | 4.91 | 5.01 | 5.12 | 5.22 | 5.32 | 5.42 | 5.53 | 5.63 | 5.73 | 5.84 | 5.94 | 6.04 | 6.14 | 6.25 | 6.35 | 6.45 | 6.56 | 6.66 | 6.76 | 6.86 | 6.97 | 7.07 | 7.17 | 7.28 | 7.38 | 7.48 | 7.58 | 7.69 | 7.79 | 7.89
30.61 | 3.78 | 3.88 | 3.98 | 4.09 | 4.19 | 4.29 | 4.39 | 4.5 | 4.6 | 4.7 | 4.81 | 4.91 | 5.01 | 5.11 | 5.22 | 5.32 | 5.42 | 5.53 | 5.63 | 5.73 | 5.83 | 5.94 | 6.04 | 6.14 | 6.25 | 6.35 | 6.45 | 6.55 | 6.66 | 6.76 | 6.86 | 6.97 | 7.07 | 7.17 | 7.27 | 7.38 | 7.48 | 7.58 | 7.69 | 7.79 | 7.89
30.62 | 3.78 | 3.88 | 3.98 | 4.09 | 4.19 | 4.29 | 4.39 | 4.5 | 4.6 | 4.7 | 4.81 | 4.91 | 5.01 | 5.11 | 5.22 | 5.32 | 5.42 | 5.52 | 5.63 | 5.73 | 5.83 | 5.94 | 6.04 | 6.14 | 6.24 | 6.35 | 6.45 | 6.55 | 6.66 | 6.76 | 6.86 | 6.96 | 7.07 | 7.17 | 7.27 | 7.38 | 7.48 | 7.58 | 7.68 | 7.79 | 7.89
30.63 | 3.78 | 3.88 | 3.98 | 4.08 | 4.19 | 4.29 | 4.39 | 4.5 | 4.6 | 4.7 | 4.8 | 4.91 | 5.01 | 5.11 | 5.22 | 5.32 | 5.42 | 5.52 | 5.63 | 5.73 | 5.83 | 5.94 | 6.04 | 6.14 | 6.24 | 6.35 | 6.45 | 6.55 | 6.65 | 6.76 | 6.86 | 6.96 | 7.07 | 7.17 | 7.27 | 7.37 | 7.48 | 7.58 | 7.68 | 7.79 | 7.89
30.64 | 3.77 | 3.88 | 3.98 | 4.08 | 4.19 | 4.29 | 4.39 | 4.49 | 4.6 | 4.7 | 4.8 | 4.91 | 5.01 | 5.11 | 5.21 | 5.32 | 5.42 | 5.52 | 5.63 | 5.73 | 5.83 | 5.93 | 6.04 | 6.14 | 6.24 | 6.35 | 6.45 | 6.55 | 6.65 | 6.76 | 6.86 | 6.96 | 7.06 | 7.17 | 7.27 | 7.37 | 7.48 | 7.58 | 7.68 | 7.78 | 7.89
30.65 | 3.77 | 3.88 | 3.98 | 4.08 | 4.19 | 4.29 | 4.39 | 4.49 | 4.6 | 4.7 | 4.8 | 4.91 | 5.01 | 5.11 | 5.21 | 5.32 | 5.42 | 5.52 | 5.62 | 5.73 | 5.83 | 5.93 | 6.04 | 6.14 | 6.24 | 6.34 | 6.45 | 6.55 | 6.65 | 6.76 | 6.86 | 6.96 | 7.06 | 7.17 | 7.27 | 7.37 | 7.47 | 7.58 | 7.68 | 7.78 | 7.89
30.66 | 3.77 | 3.88 | 3.98 | 4.08 | 4.18 | 4.29 | 4.39 | 4.49 | 4.6 | 4.7 | 4.8 | 4.9 | 5.01 | 5.11 | 5.21 | 5.32 | 5.42 | 5.52 | 5.62 | 5.73 | 5.83 | 5.93 | 6.03 | 6.14 | 6.24 | 6.34 | 6.45 | 6.55 | 6.65 | 6.75 | 6.86 | 6.96 | 7.06 | 7.17 | 7.27 | 7.37 | 7.47 | 7.58 | 7.68 | 7.78 | 7.88
30.67 | 3.77 | 3.88 | 3.98 | 4.08 | 4.18 | 4.29 | 4.39 | 4.49 | 4.6 | 4.7 | 4.8 | 4.9 | 5.01 | 5.11 | 5.21 | 5.31 | 5.42 | 5.52 | 5.62 | 5.73 | 5.83 | 5.93 | 6.03 | 6.14 | 6.24 | 6.34 | 6.44 | 6.55 | 6.65 | 6.75 | 6.86 | 6.96 | 7.06 | 7.16 | 7.27 | 7.37 | 7.47 | 7.57 | 7.68 | 7.78 | 7.88
30.68 | 3.77 | 3.87 | 3.98 | 4.08 | 4.18 | 4.29 | 4.39 | 4.49 | 4.59 | 4.7 | 4.8 | 4.9 | 5.01 | 5.11 | 5.21 | 5.31 | 5.42 | 5.52 | 5.62 | 5.72 | 5.83 | 5.93 | 6.03 | 6.14 | 6.24 | 6.34 | 6.44 | 6.55 | 6.65 | 6.75 | 6.85 | 6.96 | 7.06 | 7.16 | 7.27 | 7.37 | 7.47 | 7.57 | 7.68 | 7.78 | 7.88
30.69 | 3.77 | 3.87 | 3.98 | 4.08 | 4.18 | 4.29 | 4.39 | 4.49 | 4.59 | 4.7 | 4.8 | 4.9 | 5 | 5.11 | 5.21 | 5.31 | 5.42 | 5.52 | 5.62 | 5.72 | 5.83 | 5.93 | 6.03 | 6.13 | 6.24 | 6.34 | 6.44 | 6.55 | 6.65 | 6.75 | 6.85 | 6.96 | 7.06 | 7.16 | 7.26 | 7.37 | 7.47 | 7.57 | 7.67 | 7.78 | 7.88
30.7 | 3.77 | 3.87 | 3.98 | 4.08 | 4.18 | 4.28 | 4.39 | 4.49 | 4.59 | 4.7 | 4.8 | 4.9 | 5 | 5.11 | 5.21 | 5.31 | 5.41 | 5.52 | 5.62 | 5.72 | 5.83 | 5.93 | 6.03 | 6.13 | 6.24 | 6.34 | 6.44 | 6.54 | 6.65 | 6.75 | 6.85 | 6.95 | 7.06 | 7.16 | 7.26 | 7.37 | 7.47 | 7.57 | 7.67 | 7.78 | 7.88
30.71 | 3.77 | 3.87 | 3.98 | 4.08 | 4.18 | 4.28 | 4.39 | 4.49 | 4.59 | 4.69 | 4.8 | 4.9 | 5 | 5.11 | 5.21 | 5.31 | 5.41 | 5.52 | 5.62 | 5.72 | 5.82 | 5.93 | 6.03 | 6.13 | 6.23 | 6.34 | 6.44 | 6.54 | 6.65 | 6.75 | 6.85 | 6.95 | 7.06 | 7.16 | 7.26 | 7.36 | 7.47 | 7.57 | 7.67 | 7.78 | 7.88
30.72 | 3.77 | 3.87 | 3.97 | 4.08 | 4.18 | 4.28 | 4.39 | 4.49 | 4.59 | 4.69 | 4.8 | 4.9 | 5 | 5.1 | 5.21 | 5.31 | 5.41 | 5.51 | 5.62 | 5.72 | 5.82 | 5.93 | 6.03 | 6.13 | 6.23 | 6.34 | 6.44 | 6.54 | 6.64 | 6.75 | 6.85 | 6.95 | 7.06 | 7.16 | 7.26 | 7.36 | 7.47 | 7.57 | 7.67 | 7.77 | 7.88
30.73 | 3.77 | 3.87 | 3.97 | 4.08 | 4.18 | 4.28 | 4.38 | 4.49 | 4.59 | 4.69 | 4.8 | 4.9 | 5 | 5.1 | 5.21 | 5.31 | 5.41 | 5.51 | 5.62 | 5.72 | 5.82 | 5.92 | 6.03 | 6.13 | 6.23 | 6.34 | 6.44 | 6.54 | 6.64 | 6.75 | 6.85 | 6.95 | 7.05 | 7.16 | 7.26 | 7.36 | 7.46 | 7.57 | 7.67 | 7.77 | 7.88
30.74 | 3.77 | 3.87 | 3.97 | 4.08 | 4.18 | 4.28 | 4.38 | 4.49 | 4.59 | 4.69 | 4.79 | 4.9 | 5 | 5.1 | 5.21 | 5.31 | 5.41 | 5.51 | 5.62 | 5.72 | 5.82 | 5.92 | 6.03 | 6.13 | 6.23 | 6.33 | 6.44 | 6.54 | 6.64 | 6.74 | 6.85 | 6.95 | 7.05 | 7.16 | 7.26 | 7.36 | 7.46 | 7.57 | 7.67 | 7.77 | 7.87
30.75 | 3.77 | 3.87 | 3.97 | 4.07 | 4.18 | 4.28 | 4.38 | 4.49 | 4.59 | 4.69 | 4.79 | 4.9 | 5 | 5.1 | 5.2 | 5.31 | 5.41 | 5.51 | 5.61 | 5.72 | 5.82 | 5.92 | 6.03 | 6.13 | 6.23 | 6.33 | 6.44 | 6.54 | 6.64 | 6.74 | 6.85 | 6.95 | 7.05 | 7.15 | 7.26 | 7.36 | 7.46 | 7.56 | 7.67 | 7.77 | 7.87
30.76 | 3.77 | 3.87 | 3.97 | 4.07 | 4.18 | 4.28 | 4.38 | 4.48 | 4.59 | 4.69 | 4.79 | 4.9 | 5 | 5.1 | 5.2 | 5.31 | 5.41 | 5.51 | 5.61 | 5.72 | 5.82 | 5.92 | 6.02 | 6.13 | 6.23 | 6.33 | 6.43 | 6.54 | 6.64 | 6.74 | 6.85 | 6.95 | 7.05 | 7.15 | 7.26 | 7.36 | 7.46 | 7.56 | 7.67 | 7.77 | 7.87
30.77 | 3.77 | 3.87 | 3.97 | 4.07 | 4.18 | 4.28 | 4.38 | 4.48 | 4.59 | 4.69 | 4.79 | 4.89 | 5 | 5.1 | 5.2 | 5.3 | 5.41 | 5.51 | 5.61 | 5.72 | 5.82 | 5.92 | 6.02 | 6.13 | 6.23 | 6.33 | 6.43 | 6.54 | 6.64 | 6.74 | 6.84 | 6.95 | 7.05 | 7.15 | 7.25 | 7.36 | 7.46 | 7.56 | 7.66 | 7.77 | 7.87
30.78 | 3.76 | 3.87 | 3.97 | 4.07 | 4.18 | 4.28 | 4.38 | 4.48 | 4.59 | 4.69 | 4.79 | 4.89 | 5 | 5.1 | 5.2 | 5.3 | 5.41 | 5.51 | 5.61 | 5.71 | 5.82 | 5.92 | 6.02 | 6.12 | 6.23 | 6.33 | 6.43 | 6.53 | 6.64 | 6.74 | 6.84 | 6.95 | 7.05 | 7.15 | 7.25 | 7.36 | 7.46 | 7.56 | 7.66 | 7.77 | 7.87
30.79 | 3.76 | 3.87 | 3.97 | 4.07 | 4.17 | 4.28 | 4.38 | 4.48 | 4.58 | 4.69 | 4.79 | 4.89 | 5 | 5.1 | 5.2 | 5.3 | 5.41 | 5.51 | 5.61 | 5.71 | 5.82 | 5.92 | 6.02 | 6.12 | 6.23 | 6.33 | 6.43 | 6.53 | 6.64 | 6.74 | 6.84 | 6.94 | 7.05 | 7.15 | 7.25 | 7.35 | 7.46 | 7.56 | 7.66 | 7.76 | 7.87
30.8 | 3.76 | 3.87 | 3.97 | 4.07 | 4.17 | 4.28 | 4.38 | 4.48 | 4.58 | 4.69 | 4.79 | 4.89 | 4.99 | 5.1 | 5.2 | 5.3 | 5.4 | 5.51 | 5.61 | 5.71 | 5.81 | 5.92 | 6.02 | 6.12 | 6.22 | 6.33 | 6.43 | 6.53 | 6.64 | 6.74 | 6.84 | 6.94 | 7.05 | 7.15 | 7.25 | 7.35 | 7.46 | 7.56 | 7.66 | 7.76 | 7.87
30.81 | 3.76 | 3.86 | 3.97 | 4.07 | 4.17 | 4.28 | 4.38 | 4.48 | 4.58 | 4.69 | 4.79 | 4.89 | 4.99 | 5.1 | 5.2 | 5.3 | 5.4 | 5.51 | 5.61 | 5.71 | 5.81 | 5.92 | 6.02 | 6.12 | 6.22 | 6.33 | 6.43 | 6.53 | 6.63 | 6.74 | 6.84 | 6.94 | 7.04 | 7.15 | 7.25 | 7.35 | 7.45 | 7.56 | 7.66 | 7.76 | 7.86
30.82 | 3.76 | 3.86 | 3.97 | 4.07 | 4.17 | 4.27 | 4.38 | 4.48 | 4.58 | 4.68 | 4.79 | 4.89 | 4.99 | 5.09 | 5.2 | 5.3 | 5.4 | 5.51 | 5.61 | 5.71 | 5.81 | 5.92 | 6.02 | 6.12 | 6.22 | 6.33 | 6.43 | 6.53 | 6.63 | 6.74 | 6.84 | 6.94 | 7.04 | 7.15 | 7.25 | 7.35 | 7.45 | 7.56 | 7.66 | 7.76 | 7.86
30.83 | 3.76 | 3.86 | 3.97 | 4.07 | 4.17 | 4.27 | 4.38 | 4.48 | 4.58 | 4.68 | 4.79 | 4.89 | 4.99 | 5.09 | 5.2 | 5.3 | 5.4 | 5.5 | 5.61 | 5.71 | 5.81 | 5.91 | 6.02 | 6.12 | 6.22 | 6.32 | 6.43 | 6.53 | 6.63 | 6.73 | 6.84 | 6.94 | 7.04 | 7.14 | 7.25 | 7.35 | 7.45 | 7.55 | 7.66 | 7.76 | 7.86
30.84 | 3.76 | 3.86 | 3.97 | 4.07 | 4.17 | 4.27 | 4.38 | 4.48 | 4.58 | 4.68 | 4.79 | 4.89 | 4.99 | 5.09 | 5.2 | 5.3 | 5.4 | 5.5 | 5.61 | 5.71 | 5.81 | 5.91 | 6.02 | 6.12 | 6.22 | 6.32 | 6.43 | 6.53 | 6.63 | 6.73 | 6.84 | 6.94 | 7.04 | 7.14 | 7.25 | 7.35 | 7.45 | 7.55 | 7.66 | 7.76 | 7.86
30.85 | 3.76 | 3.86 | 3.96 | 4.07 | 4.17 | 4.27 | 4.37 | 4.48 | 4.58 | 4.68 | 4.78 | 4.89 | 4.99 | 5.09 | 5.19 | 5.3 | 5.4 | 5.5 | 5.6 | 5.71 | 5.81 | 5.91 | 6.01 | 6.12 | 6.22 | 6.32 | 6.42 | 6.53 | 6.63 | 6.73 | 6.83 | 6.94 | 7.04 | 7.14 | 7.24 | 7.35 | 7.45 | 7.55 | 7.65 | 7.76 | 7.86
30.86 | 3.76 | 3.86 | 3.96 | 4.07 | 4.17 | 4.27 | 4.37 | 4.48 | 4.58 | 4.68 | 4.78 | 4.89 | 4.99 | 5.09 | 5.19 | 5.3 | 5.4 | 5.5 | 5.6 | 5.71 | 5.81 | 5.91 | 6.01 | 6.12 | 6.22 | 6.32 | 6.42 | 6.53 | 6.63 | 6.73 | 6.83 | 6.94 | 7.04 | 7.14 | 7.24 | 7.35 | 7.45 | 7.55 | 7.65 | 7.76 | 7.86
30.87 | 3.76 | 3.86 | 3.96 | 4.07 | 4.17 | 4.27 | 4.37 | 4.48 | 4.58 | 4.68 | 4.78 | 4.89 | 4.99 | 5.09 | 5.19 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.81 | 5.91 | 6.01 | 6.11 | 6.22 | 6.32 | 6.42 | 6.52 | 6.63 | 6.73 | 6.83 | 6.93 | 7.04 | 7.14 | 7.24 | 7.34 | 7.45 | 7.55 | 7.65 | 7.75 | 7.86
30.88 | 3.76 | 3.86 | 3.96 | 4.06 | 4.17 | 4.27 | 4.37 | 4.47 | 4.58 | 4.68 | 4.78 | 4.88 | 4.99 | 5.09 | 5.19 | 5.29 | 5.4 | 5.5 | 5.6 | 5.7 | 5.81 | 5.91 | 6.01 | 6.11 | 6.22 | 6.32 | 6.42 | 6.52 | 6.63 | 6.73 | 6.83 | 6.93 | 7.04 | 7.14 | 7.24 | 7.34 | 7.45 | 7.55 | 7.65 | 7.75 | 7.86
30.89 | 3.76 | 3.86 | 3.96 | 4.06 | 4.17 | 4.27 | 4.37 | 4.47 | 4.58 | 4.68 | 4.78 | 4.88 | 4.99 | 5.09 | 5.19 | 5.29 | 5.4 | 5.5 | 5.6 | 5.7 | 5.81 | 5.91 | 6.01 | 6.11 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.73 | 6.83 | 6.93 | 7.03 | 7.14 | 7.24 | 7.34 | 7.44 | 7.55 | 7.65 | 7.75 | 7.85
30.9 | 3.76 | 3.86 | 3.96 | 4.06 | 4.17 | 4.27 | 4.37 | 4.47 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.09 | 5.19 | 5.29 | 5.39 | 5.5 | 5.6 | 5.7 | 5.8 | 5.91 | 6.01 | 6.11 | 6.21 | 6.32 | 6.42 | 6.52 | 6.62 | 6.73 | 6.83 | 6.93 | 7.03 | 7.14 | 7.24 | 7.34 | 7.44 | 7.55 | 7.65 | 7.75 | 7.85
30.91 | 3.75 | 3.86 | 3.96 | 4.06 | 4.16 | 4.27 | 4.37 | 4.47 | 4.57 | 4.68 | 4.78 | 4.88 | 4.98 | 5.09 | 5.19 | 5.29 | 5.39 | 5.5 | 5.6 | 5.7 | 5.8 | 5.91 | 6.01 | 6.11 | 6.21 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.83 | 6.93 | 7.03 | 7.13 | 7.24 | 7.34 | 7.44 | 7.54 | 7.65 | 7.75 | 7.85
30.92 | 3.75 | 3.86 | 3.96 | 4.06 | 4.16 | 4.27 | 4.37 | 4.47 | 4.57 | 4.68 | 4.78 | 4.88 | 4.98 | 5.09 | 5.19 | 5.29 | 5.39 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 6.01 | 6.11 | 6.21 | 6.31 | 6.42 | 6.52 | 6.62 | 6.72 | 6.83 | 6.93 | 7.03 | 7.13 | 7.24 | 7.34 | 7.44 | 7.54 | 7.64 | 7.75 | 7.85
30.93 | 3.75 | 3.86 | 3.96 | 4.06 | 4.16 | 4.27 | 4.37 | 4.47 | 4.57 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.19 | 5.29 | 5.39 | 5.49 | 5.6 | 5.7 | 5.8 | 5.9 | 6.01 | 6.11 | 6.21 | 6.31 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.93 | 7.03 | 7.13 | 7.23 | 7.34 | 7.44 | 7.54 | 7.64 | 7.75 | 7.85
30.94 | 3.75 | 3.86 | 3.96 | 4.06 | 4.16 | 4.26 | 4.37 | 4.47 | 4.57 | 4.67 | 4.78 | 4.88 | 4.98 | 5.08 | 5.19 | 5.29 | 5.39 | 5.49 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.11 | 6.21 | 6.31 | 6.41 | 6.52 | 6.62 | 6.72 | 6.82 | 6.93 | 7.03 | 7.13 | 7.23 | 7.34 | 7.44 | 7.54 | 7.64 | 7.74 | 7.85
30.95 | 3.75 | 3.85 | 3.96 | 4.06 | 4.16 | 4.26 | 4.37 | 4.47 | 4.57 | 4.67 | 4.78 | 4.88 | 4.98 | 5.08 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.7 | 5.8 | 5.9 | 6 | 6.11 | 6.21 | 6.31 | 6.41 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.03 | 7.13 | 7.23 | 7.33 | 7.44 | 7.54 | 7.64 | 7.74 | 7.85
30.96 | 3.75 | 3.85 | 3.96 | 4.06 | 4.16 | 4.26 | 4.37 | 4.47 | 4.57 | 4.67 | 4.77 | 4.88 | 4.98 | 5.08 | 5.18 | 5.29 | 5.39 | 5.49 | 5.59 | 5.7 | 5.8 | 5.9 | 6 | 6.11 | 6.21 | 6.31 | 6.41 | 6.51 | 6.62 | 6.72 | 6.82 | 6.92 | 7.03 | 7.13 | 7.23 | 7.33 | 7.44 | 7.54 | 7.64 | 7.74 | 7.84
30.97 | 3.75 | 3.85 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.47 | 4.57 | 4.67 | 4.77 | 4.88 | 4.98 | 5.08 | 5.18 | 5.29 | 5.39 | 5.49 | 5.59 | 5.69 | 5.8 | 5.9 | 6 | 6.1 | 6.21 | 6.31 | 6.41 | 6.51 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.13 | 7.23 | 7.33 | 7.43 | 7.54 | 7.64 | 7.74 | 7.84
30.98 | 3.75 | 3.85 | 3.95 | 4.06 | 4.16 | 4.26 | 4.36 | 4.47 | 4.57 | 4.67 | 4.77 | 4.88 | 4.98 | 5.08 | 5.18 | 5.28 | 5.39 | 5.49 | 5.59 | 5.69 | 5.8 | 5.9 | 6 | 6.1 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.72 | 6.82 | 6.92 | 7.02 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53 | 7.64 | 7.74 | 7.84
30.99 | 3.75 | 3.85 | 3.95 | 4.06 | 4.16 | 4.26 | 4.36 | 4.47 | 4.57 | 4.67 | 4.77 | 4.87 | 4.98 | 5.08 | 5.18 | 5.28 | 5.39 | 5.49 | 5.59 | 5.69 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.31 | 6.41 | 6.51 | 6.61 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.23 | 7.33 | 7.43 | 7.53 | 7.64 | 7.74 | 7.84
31 | 3.75 | 3.85 | 3.95 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.57 | 4.67 | 4.77 | 4.87 | 4.98 | 5.08 | 5.18 | 5.28 | 5.38 | 5.49 | 5.59 | 5.69 | 5.79 | 5.9 | 6 | 6.1 | 6.2 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.82 | 6.92 | 7.02 | 7.12 | 7.23 | 7.33 | 7.43 | 7.53 | 7.63 | 7.74 | 7.84
31.01 | 3.75 | 3.85 | 3.95 | 4.05 | 4.16 | 4.26 | 4.36 | 4.46 | 4.57 | 4.67 | 4.77 | 4.87 | 4.97 | 5.08 | 5.18 | 5.28 | 5.38 | 5.49 | 5.59 | 5.69 | 5.79 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.41 | 6.51 | 6.61 | 6.71 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.33 | 7.43 | 7.53 | 7.63 | 7.74 | 7.84
31.02 | 3.75 | 3.85 | 3.95 | 4.05 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.67 | 4.77 | 4.87 | 4.97 | 5.08 | 5.18 | 5.28 | 5.38 | 5.49 | 5.59 | 5.69 | 5.79 | 5.89 | 6 | 6.1 | 6.2 | 6.3 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.92 | 7.02 | 7.12 | 7.22 | 7.33 | 7.43 | 7.53 | 7.63 | 7.73 | 7.84
31.03 | 3.75 | 3.85 | 3.95 | 4.05 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.67 | 4.77 | 4.87 | 4.97 | 5.08 | 5.18 | 5.28 | 5.38 | 5.48 | 5.59 | 5.69 | 5.79 | 5.89 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.51 | 6.61 | 6.71 | 6.81 | 6.92 | 7.02 | 7.12 | 7.22 | 7.32 | 7.43 | 7.53 | 7.63 | 7.73 | 7.84
31.04 | 3.75 | 3.85 | 3.95 | 4.05 | 4.15 | 4.26 | 4.36 | 4.46 | 4.56 | 4.67 | 4.77 | 4.87 | 4.97 | 5.07 | 5.18 | 5.28 | 5.38 | 5.48 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.1 | 6.2 | 6.3 | 6.4 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7.02 | 7.12 | 7.22 | 7.32 | 7.43 | 7.53 | 7.63 | 7.73 | 7.83
31.05 | 3.74 | 3.85 | 3.95 | 4.05 | 4.15 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.77 | 4.87 | 4.97 | 5.07 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.69 | 5.79 | 5.89 | 5.99 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.61 | 6.71 | 6.81 | 6.91 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.53 | 7.63 | 7.73 | 7.83
31.06 | 3.74 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.36 | 4.46 | 4.56 | 4.66 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.28 | 5.38 | 5.48 | 5.58 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.2 | 6.3 | 6.4 | 6.5 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52 | 7.63 | 7.73 | 7.83
31.07 | 3.74 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.36 | 4.46 | 4.56 | 4.66 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.79 | 5.89 | 5.99 | 6.09 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.22 | 7.32 | 7.42 | 7.52 | 7.63 | 7.73 | 7.83
31.08 | 3.74 | 3.84 | 3.95 | 4.05 | 4.15 | 4.25 | 4.36 | 4.46 | 4.56 | 4.66 | 4.76 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.38 | 5.48 | 5.58 | 5.68 | 5.79 | 5.89 | 5.99 | 6.09 | 6.19 | 6.3 | 6.4 | 6.5 | 6.6 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.22 | 7.32 | 7.42 | 7.52 | 7.62 | 7.73 | 7.83
31.09 | 3.74 | 3.84 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.46 | 4.56 | 4.66 | 4.76 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.89 | 5.99 | 6.09 | 6.19 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.81 | 6.91 | 7.01 | 7.11 | 7.21 | 7.32 | 7.42 | 7.52 | 7.62 | 7.73 | 7.83
31.1 | 3.74 | 3.84 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.97 | 5.07 | 5.17 | 5.27 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.89 | 5.99 | 6.09 | 6.19 | 6.29 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.91 | 7.01 | 7.11 | 7.21 | 7.32 | 7.42 | 7.52 | 7.62 | 7.72 | 7.83
31.11 | 3.74 | 3.84 | 3.94 | 4.05 | 4.15 | 4.25 | 4.35 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.97 | 5.07 | 5.17 | 5.27 | 5.37 | 5.48 | 5.58 | 5.68 | 5.78 | 5.88 | 5.99 | 6.09 | 6.19 | 6.29 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.42 | 7.52 | 7.62 | 7.72 | 7.82
31.12 | 3.74 | 3.84 | 3.94 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.07 | 5.17 | 5.27 | 5.37 | 5.48 | 5.58 | 5.68 | 5.78 | 5.88 | 5.99 | 6.09 | 6.19 | 6.29 | 6.39 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.52 | 7.62 | 7.72 | 7.82
31.13 | 3.74 | 3.84 | 3.94 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.58 | 5.68 | 5.78 | 5.88 | 5.98 | 6.09 | 6.19 | 6.29 | 6.39 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.52 | 7.62 | 7.72 | 7.82
31.14 | 3.74 | 3.84 | 3.94 | 4.04 | 4.15 | 4.25 | 4.35 | 4.45 | 4.55 | 4.66 | 4.76 | 4.86 | 4.96 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.58 | 5.68 | 5.78 | 5.88 | 5.98 | 6.09 | 6.19 | 6.29 | 6.39 | 6.49 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51 | 7.62 | 7.72 | 7.82
31.15 | 3.74 | 3.84 | 3.94 | 4.04 | 4.15 | 4.25 | 4.35 | 4.45 | 4.55 | 4.66 | 4.76 | 4.86 | 4.96 | 5.06 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.7 | 6.8 | 6.9 | 7 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51 | 7.62 | 7.72 | 7.82
31.16 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.25 | 4.35 | 4.45 | 4.55 | 4.66 | 4.76 | 4.86 | 4.96 | 5.06 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.7 | 6.8 | 6.9 | 7 | 7.1 | 7.21 | 7.31 | 7.41 | 7.51 | 7.61 | 7.72 | 7.82
31.17 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.76 | 4.86 | 4.96 | 5.06 | 5.16 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.29 | 6.39 | 6.49 | 6.59 | 6.69 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.31 | 7.41 | 7.51 | 7.61 | 7.71 | 7.82
31.18 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.76 | 4.86 | 4.96 | 5.06 | 5.16 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.29 | 6.39 | 6.49 | 6.59 | 6.69 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.31 | 7.41 | 7.51 | 7.61 | 7.71 | 7.82
31.19 | 3.73 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.41 | 7.51 | 7.61 | 7.71 | 7.81
31.2 | 3.73 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.51 | 7.61 | 7.71 | 7.81
31.21 | 3.73 | 3.83 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.96 | 5.06 | 5.16 | 5.26 | 5.36 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.98 | 6.08 | 6.18 | 6.28 | 6.38 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.51 | 7.61 | 7.71 | 7.81
31.22 | 3.73 | 3.83 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.96 | 5.06 | 5.16 | 5.26 | 5.36 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.98 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5 | 7.61 | 7.71 | 7.81
31.23 | 3.73 | 3.83 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.06 | 5.16 | 5.26 | 5.36 | 5.46 | 5.57 | 5.67 | 5.77 | 5.87 | 5.97 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5 | 7.6 | 7.71 | 7.81
31.24 | 3.73 | 3.83 | 3.93 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.06 | 5.16 | 5.26 | 5.36 | 5.46 | 5.57 | 5.67 | 5.77 | 5.87 | 5.97 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.2 | 7.3 | 7.4 | 7.5 | 7.6 | 7.71 | 7.81
31.25 | 3.73 | 3.83 | 3.93 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.67 | 5.77 | 5.87 | 5.97 | 6.07 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.3 | 7.4 | 7.5 | 7.6 | 7.7 | 7.81
31.26 | 3.73 | 3.83 | 3.93 | 4.03 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.67 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.28 | 6.38 | 6.48 | 6.58 | 6.68 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.3 | 7.4 | 7.5 | 7.6 | 7.7 | 7.8
31.27 | 3.73 | 3.83 | 3.93 | 4.03 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.26 | 5.36 | 5.46 | 5.56 | 5.66 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.28 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.89 | 6.99 | 7.09 | 7.19 | 7.29 | 7.4 | 7.5 | 7.6 | 7.7 | 7.8
31.28 | 3.73 | 3.83 | 3.93 | 4.03 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.26 | 5.36 | 5.46 | 5.56 | 5.66 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.27 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.89 | 6.99 | 7.09 | 7.19 | 7.29 | 7.39 | 7.5 | 7.6 | 7.7 | 7.8
31.29 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.87 | 5.97 | 6.07 | 6.17 | 6.27 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.88 | 6.99 | 7.09 | 7.19 | 7.29 | 7.39 | 7.5 | 7.6 | 7.7 | 7.8
31.3 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.48 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49 | 7.6 | 7.7 | 7.8
31.31 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49 | 7.59 | 7.7 | 7.8
31.32 | 3.72 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.08 | 7.19 | 7.29 | 7.39 | 7.49 | 7.59 | 7.69 | 7.8
31.33 | 3.72 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.68 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.29 | 7.39 | 7.49 | 7.59 | 7.69 | 7.8
31.34 | 3.72 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.29 | 7.39 | 7.49 | 7.59 | 7.69 | 7.79
31.35 | 3.72 | 3.82 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.28 | 7.39 | 7.49 | 7.59 | 7.69 | 7.79
31.36 | 3.72 | 3.82 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.16 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.88 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.49 | 7.59 | 7.69 | 7.79
31.37 | 3.72 | 3.82 | 3.92 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.76 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.49 | 7.59 | 7.69 | 7.79
31.38 | 3.72 | 3.82 | 3.92 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.48 | 7.59 | 7.69 | 7.79
31.39 | 3.72 | 3.82 | 3.92 | 4.02 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.08 | 7.18 | 7.28 | 7.38 | 7.48 | 7.58 | 7.69 | 7.79
31.4 | 3.72 | 3.82 | 3.92 | 4.02 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.18 | 7.28 | 7.38 | 7.48 | 7.58 | 7.68 | 7.79
31.41 | 3.72 | 3.82 | 3.92 | 4.02 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.18 | 7.28 | 7.38 | 7.48 | 7.58 | 7.68 | 7.78
31.42 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.95 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.28 | 7.38 | 7.48 | 7.58 | 7.68 | 7.78
31.43 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.55 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.38 | 7.48 | 7.58 | 7.68 | 7.78
31.44 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.55 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.76 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.37 | 7.48 | 7.58 | 7.68 | 7.78
31.45 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.15 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.76 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.37 | 7.47 | 7.58 | 7.68 | 7.78
31.46 | 3.71 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.36 | 6.46 | 6.56 | 6.66 | 6.76 | 6.86 | 6.97 | 7.07 | 7.17 | 7.27 | 7.37 | 7.47 | 7.58 | 7.68 | 7.78
31.47 | 3.71 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.75 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.36 | 6.46 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.07 | 7.17 | 7.27 | 7.37 | 7.47 | 7.57 | 7.68 | 7.78
31.48 | 3.71 | 3.81 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.46 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.17 | 7.27 | 7.37 | 7.47 | 7.57 | 7.67 | 7.78
31.49 | 3.71 | 3.81 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.45 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.17 | 7.27 | 7.37 | 7.47 | 7.57 | 7.67 | 7.77
31.5 | 3.71 | 3.81 | 3.91 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.16 | 7.27 | 7.37 | 7.47 | 7.57 | 7.67 | 7.77
31.51 | 3.71 | 3.81 | 3.91 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.05 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.16 | 7.26 | 7.37 | 7.47 | 7.57 | 7.67 | 7.77
31.52 | 3.71 | 3.81 | 3.91 | 4.01 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.05 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.65 | 6.76 | 6.86 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.47 | 7.57 | 7.67 | 7.77
31.53 | 3.71 | 3.81 | 3.91 | 4.01 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.86 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46 | 7.57 | 7.67 | 7.77
31.54 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.25 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46 | 7.56 | 7.67 | 7.77
31.55 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46 | 7.56 | 7.67 | 7.77
31.56 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.95 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46 | 7.56 | 7.66 | 7.77
31.57 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.95 | 7.05 | 7.16 | 7.26 | 7.36 | 7.46 | 7.56 | 7.66 | 7.76
31.58 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.55 | 6.65 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.26 | 7.36 | 7.46 | 7.56 | 7.66 | 7.76
31.59 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.65 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.36 | 7.46 | 7.56 | 7.66 | 7.76
31.6 | 3.7 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.46 | 7.56 | 7.66 | 7.76
31.61 | 3.7 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.93 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.45 | 7.56 | 7.66 | 7.76
31.62 | 3.7 | 3.8 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.45 | 7.55 | 7.66 | 7.76
31.63 | 3.7 | 3.8 | 3.9 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.45 | 7.55 | 7.65 | 7.76
31.64 | 3.7 | 3.8 | 3.9 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.53 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.05 | 7.15 | 7.25 | 7.35 | 7.45 | 7.55 | 7.65 | 7.75
31.65 | 3.7 | 3.8 | 3.9 | 4 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.23 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.15 | 7.25 | 7.35 | 7.45 | 7.55 | 7.65 | 7.75
31.66 | 3.7 | 3.8 | 3.9 | 4 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.25 | 7.35 | 7.45 | 7.55 | 7.65 | 7.75
31.67 | 3.7 | 3.8 | 3.9 | 4 | 4.1 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.35 | 7.45 | 7.55 | 7.65 | 7.75
31.68 | 3.7 | 3.8 | 3.9 | 4 | 4.1 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.83 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.45 | 7.55 | 7.65 | 7.75
31.69 | 3.7 | 3.8 | 3.9 | 4 | 4.1 | 4.2 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.82 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44 | 7.55 | 7.65 | 7.75
31.7 | 3.7 | 3.8 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44 | 7.54 | 7.65 | 7.75
31.71 | 3.7 | 3.8 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44 | 7.54 | 7.64 | 7.75
31.72 | 3.7 | 3.8 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.4 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44 | 7.54 | 7.64 | 7.74
31.73 | 3.69 | 3.8 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.42 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.14 | 7.24 | 7.34 | 7.44 | 7.54 | 7.64 | 7.74
31.74 | 3.69 | 3.8 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.24 | 7.34 | 7.44 | 7.54 | 7.64 | 7.74
31.75 | 3.69 | 3.79 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.6 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.34 | 7.44 | 7.54 | 7.64 | 7.74
31.76 | 3.69 | 3.79 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.6 | 4.7 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.44 | 7.54 | 7.64 | 7.74
31.77 | 3.69 | 3.79 | 3.89 | 4 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.61 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43 | 7.54 | 7.64 | 7.74
31.78 | 3.69 | 3.79 | 3.89 | 3.99 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.61 | 5.71 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53 | 7.64 | 7.74
31.79 | 3.69 | 3.79 | 3.89 | 3.99 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.9 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53 | 7.63 | 7.74
31.8 | 3.69 | 3.79 | 3.89 | 3.99 | 4.09 | 4.2 | 4.3 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.9 | 5 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53 | 7.63 | 7.73
31.81 | 3.69 | 3.79 | 3.89 | 3.99 | 4.09 | 4.19 | 4.3 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.21 | 5.31 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6.01 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53 | 7.63 | 7.73
31.82 | 3.69 | 3.79 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.31 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6.01 | 6.11 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53 | 7.63 | 7.73
31.83 | 3.69 | 3.79 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.33 | 7.43 | 7.53 | 7.63 | 7.73
31.84 | 3.69 | 3.79 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.39 | 4.5 | 4.6 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.31 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.32 | 7.43 | 7.53 | 7.63 | 7.73
31.85 | 3.69 | 3.79 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.39 | 4.49 | 4.6 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.4 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.53 | 7.63 | 7.73
31.86 | 3.69 | 3.79 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.39 | 4.49 | 4.59 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.4 | 5.5 | 5.61 | 5.71 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.51 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52 | 7.63 | 7.73
31.87 | 3.68 | 3.79 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.39 | 4.49 | 4.59 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.4 | 5.5 | 5.6 | 5.71 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52 | 7.62 | 7.72
31.88 | 3.68 | 3.79 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.39 | 4.49 | 4.59 | 4.69 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.92 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52 | 7.62 | 7.72
31.89 | 3.68 | 3.78 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.39 | 4.49 | 4.59 | 4.69 | 4.79 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.8 | 5.91 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52 | 7.62 | 7.72
31.9 | 3.68 | 3.78 | 3.88 | 3.99 | 4.09 | 4.19 | 4.29 | 4.39 | 4.49 | 4.59 | 4.69 | 4.79 | 4.89 | 5 | 5.1 | 5.2 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52 | 7.62 | 7.72
31.91 | 3.68 | 3.78 | 3.88 | 3.98 | 4.09 | 4.19 | 4.29 | 4.39 | 4.49 | 4.59 | 4.69 | 4.79 | 4.89 | 4.99 | 5.1 | 5.2 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.22 | 7.32 | 7.42 | 7.52 | 7.62 | 7.72
31.92 | 3.68 | 3.78 | 3.88 | 3.98 | 4.08 | 4.19 | 4.29 | 4.39 | 4.49 | 4.59 | 4.69 | 4.79 | 4.89 | 4.99 | 5.09 | 5.2 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.42 | 7.52 | 7.62 | 7.72
31.93 | 3.68 | 3.78 | 3.88 | 3.98 | 4.08 | 4.19 | 4.29 | 4.39 | 4.49 | 4.59 | 4.69 | 4.79 | 4.89 | 4.99 | 5.09 | 5.19 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.52 | 7.62 | 7.72
31.94 | 3.68 | 3.78 | 3.88 | 3.98 | 4.08 | 4.18 | 4.29 | 4.39 | 4.49 | 4.59 | 4.69 | 4.79 | 4.89 | 4.99 | 5.09 | 5.19 | 5.29 | 5.4 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51 | 7.62 | 7.72
31.95 | 3.68 | 3.78 | 3.88 | 3.98 | 4.08 | 4.18 | 4.28 | 4.39 | 4.49 | 4.59 | 4.69 | 4.79 | 4.89 | 4.99 | 5.09 | 5.19 | 5.29 | 5.39 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51 | 7.61 | 7.71
31.96 | 3.68 | 3.78 | 3.88 | 3.98 | 4.08 | 4.18 | 4.28 | 4.38 | 4.49 | 4.59 | 4.69 | 4.79 | 4.89 | 4.99 | 5.09 | 5.19 | 5.29 | 5.39 | 5.49 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.81 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51 | 7.61 | 7.71
31.97 | 3.68 | 3.78 | 3.88 | 3.98 | 4.08 | 4.18 | 4.28 | 4.38 | 4.48 | 4.59 | 4.69 | 4.79 | 4.89 | 4.99 | 5.09 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51 | 7.61 | 7.71
31.98 | 3.68 | 3.78 | 3.88 | 3.98 | 4.08 | 4.18 | 4.28 | 4.38 | 4.48 | 4.58 | 4.69 | 4.79 | 4.89 | 4.99 | 5.09 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.69 | 5.79 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51 | 7.61 | 7.71
31.99 | 3.68 | 3.78 | 3.88 | 3.98 | 4.08 | 4.18 | 4.28 | 4.38 | 4.48 | 4.58 | 4.68 | 4.79 | 4.89 | 4.99 | 5.09 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.69 | 5.79 | 5.89 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.1 | 7.21 | 7.31 | 7.41 | 7.51 | 7.61 | 7.71
32 | 3.68 | 3.78 | 3.88 | 3.98 | 4.08 | 4.18 | 4.28 | 4.38 | 4.48 | 4.58 | 4.68 | 4.78 | 4.89 | 4.99 | 5.09 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.41 | 7.51 | 7.61 | 7.71
32.01 | 3.67 | 3.78 | 3.88 | 3.98 | 4.08 | 4.18 | 4.28 | 4.38 | 4.48 | 4.58 | 4.68 | 4.78 | 4.88 | 4.99 | 5.09 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.51 | 7.61 | 7.71
32.02 | 3.67 | 3.77 | 3.88 | 3.98 | 4.08 | 4.18 | 4.28 | 4.38 | 4.48 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.09 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.19 | 6.29 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5 | 7.6 | 7.71
32.03 | 3.67 | 3.77 | 3.87 | 3.98 | 4.08 | 4.18 | 4.28 | 4.38 | 4.48 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.19 | 6.29 | 6.39 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5 | 7.6 | 7.7
32.04 | 3.67 | 3.77 | 3.87 | 3.97 | 4.08 | 4.18 | 4.28 | 4.38 | 4.48 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.18 | 5.28 | 5.39 | 5.49 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.7 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5 | 7.6 | 7.7
32.05 | 3.67 | 3.77 | 3.87 | 3.97 | 4.07 | 4.18 | 4.28 | 4.38 | 4.48 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.18 | 5.28 | 5.38 | 5.49 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.69 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5 | 7.6 | 7.7
32.06 | 3.67 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.28 | 4.38 | 4.48 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.18 | 5.28 | 5.38 | 5.48 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5 | 7.6 | 7.7
32.07 | 3.67 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.27 | 4.38 | 4.48 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.2 | 7.3 | 7.4 | 7.5 | 7.6 | 7.7
32.08 | 3.67 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.27 | 4.37 | 4.48 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.89 | 5.99 | 6.09 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.3 | 7.4 | 7.5 | 7.6 | 7.7
32.09 | 3.67 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.27 | 4.37 | 4.47 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.88 | 5.99 | 6.09 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.29 | 7.39 | 7.5 | 7.6 | 7.7
32.1 | 3.67 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.27 | 4.37 | 4.47 | 4.57 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49 | 7.59 | 7.7
32.11 | 3.67 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.27 | 4.37 | 4.47 | 4.57 | 4.67 | 4.77 | 4.88 | 4.98 | 5.08 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49 | 7.59 | 7.69
32.12 | 3.67 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.27 | 4.37 | 4.47 | 4.57 | 4.67 | 4.77 | 4.87 | 4.98 | 5.08 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.38 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49 | 7.59 | 7.69
32.13 | 3.67 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.27 | 4.37 | 4.47 | 4.57 | 4.67 | 4.77 | 4.87 | 4.97 | 5.08 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49 | 7.59 | 7.69
32.14 | 3.67 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.27 | 4.37 | 4.47 | 4.57 | 4.67 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.89 | 6.99 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49 | 7.59 | 7.69
32.15 | 3.66 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.27 | 4.37 | 4.47 | 4.57 | 4.67 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49 | 7.59 | 7.69
32.16 | 3.66 | 3.76 | 3.87 | 3.97 | 4.07 | 4.17 | 4.27 | 4.37 | 4.47 | 4.57 | 4.67 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.58 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.29 | 7.39 | 7.49 | 7.59 | 7.69
32.17 | 3.66 | 3.76 | 3.86 | 3.96 | 4.07 | 4.17 | 4.27 | 4.37 | 4.47 | 4.57 | 4.67 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.49 | 7.59 | 7.69
32.18 | 3.66 | 3.76 | 3.86 | 3.96 | 4.06 | 4.17 | 4.27 | 4.37 | 4.47 | 4.57 | 4.67 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.48 | 7.58 | 7.68
32.19 | 3.66 | 3.76 | 3.86 | 3.96 | 4.06 | 4.16 | 4.27 | 4.37 | 4.47 | 4.57 | 4.67 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.97 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.48 | 7.58 | 7.68
32.2 | 3.66 | 3.76 | 3.86 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.47 | 4.57 | 4.67 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.28 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.48 | 7.58 | 7.68
32.21 | 3.66 | 3.76 | 3.86 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.57 | 4.67 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.48 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.48 | 7.58 | 7.68
32.22 | 3.66 | 3.76 | 3.86 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.48 | 7.58 | 7.68
32.23 | 3.66 | 3.76 | 3.86 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.97 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.48 | 7.58 | 7.68
32.24 | 3.66 | 3.76 | 3.86 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.28 | 7.38 | 7.48 | 7.58 | 7.68
32.25 | 3.66 | 3.76 | 3.86 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.06 | 5.16 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.37 | 7.48 | 7.58 | 7.68
32.26 | 3.66 | 3.76 | 3.86 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.36 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.37 | 7.47 | 7.57 | 7.67
32.27 | 3.66 | 3.76 | 3.86 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.66 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.37 | 7.47 | 7.57 | 7.67
32.28 | 3.66 | 3.76 | 3.86 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.37 | 7.47 | 7.57 | 7.67
32.29 | 3.65 | 3.75 | 3.86 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.37 | 7.47 | 7.57 | 7.67
32.3 | 3.65 | 3.75 | 3.85 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.37 | 7.47 | 7.57 | 7.67
32.31 | 3.65 | 3.75 | 3.85 | 3.95 | 4.05 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.37 | 7.47 | 7.57 | 7.67
32.32 | 3.65 | 3.75 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.36 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.17 | 7.27 | 7.37 | 7.47 | 7.57 | 7.67
32.33 | 3.65 | 3.75 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.47 | 7.57 | 7.67
32.34 | 3.65 | 3.75 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46 | 7.56 | 7.66
32.35 | 3.65 | 3.75 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.06 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46 | 7.56 | 7.66
32.36 | 3.65 | 3.75 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46 | 7.56 | 7.66
32.37 | 3.65 | 3.75 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.76 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46 | 7.56 | 7.66
32.38 | 3.65 | 3.75 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.95 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46 | 7.56 | 7.66
32.39 | 3.65 | 3.75 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.46 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46 | 7.56 | 7.66
32.4 | 3.65 | 3.75 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46 | 7.56 | 7.66
32.41 | 3.65 | 3.75 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.46 | 7.56 | 7.66
32.42 | 3.65 | 3.75 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.45 | 7.55 | 7.65
32.43 | 3.64 | 3.74 | 3.84 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.45 | 7.55 | 7.65
32.44 | 3.64 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.45 | 7.55 | 7.65
32.45 | 3.64 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.45 | 7.55 | 7.65
32.46 | 3.64 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.45 | 7.55 | 7.65
32.47 | 3.64 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.45 | 7.55 | 7.65
32.48 | 3.64 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.45 | 7.55 | 7.65
32.49 | 3.64 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.35 | 7.45 | 7.55 | 7.65
32.5 | 3.64 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44 | 7.54 | 7.64
32.51 | 3.64 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44 | 7.54 | 7.64
32.52 | 3.64 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44 | 7.54 | 7.64
32.53 | 3.64 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44 | 7.54 | 7.64
32.54 | 3.64 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44 | 7.54 | 7.64
32.55 | 3.64 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44 | 7.54 | 7.64
32.56 | 3.64 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44 | 7.54 | 7.64
32.57 | 3.63 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44 | 7.54 | 7.64
32.58 | 3.63 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53 | 7.63
32.59 | 3.63 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53 | 7.63
32.6 | 3.63 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53 | 7.63
32.61 | 3.63 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53 | 7.63
32.62 | 3.63 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53 | 7.63
32.63 | 3.63 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53 | 7.63
32.64 | 3.63 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53 | 7.63
32.65 | 3.63 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53 | 7.62
32.66 | 3.63 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52 | 7.62
32.67 | 3.63 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52 | 7.62
32.68 | 3.63 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52 | 7.62
32.69 | 3.63 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52 | 7.62
32.7 | 3.63 | 3.73 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52 | 7.62
32.71 | 3.62 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52 | 7.62
32.72 | 3.62 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52 | 7.62
32.73 | 3.62 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52 | 7.61
32.74 | 3.62 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51 | 7.61
32.75 | 3.62 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51 | 7.61
32.76 | 3.62 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51 | 7.61
32.77 | 3.62 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51 | 7.61
32.78 | 3.62 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51 | 7.61
32.79 | 3.62 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51 | 7.61
32.8 | 3.62 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51 | 7.61
32.81 | 3.62 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51 | 7.6
32.82 | 3.62 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.2 | 7.3 | 7.4 | 7.5 | 7.6
32.83 | 3.62 | 3.72 | 3.82 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5 | 7.6
32.84 | 3.62 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.51 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5 | 7.6
32.85 | 3.61 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5 | 7.6
32.86 | 3.61 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5 | 7.6
32.87 | 3.61 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.61 | 5.71 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5 | 7.6
32.88 | 3.61 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5 | 7.6
32.89 | 3.61 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5 | 7.59
32.9 | 3.61 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.1 | 5.2 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.29 | 7.39 | 7.49 | 7.59
32.91 | 3.61 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49 | 7.59
32.92 | 3.61 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49 | 7.59
32.93 | 3.61 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.5 | 4.6 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49 | 7.59
32.94 | 3.61 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.3 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49 | 7.59
32.95 | 3.61 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.2 | 4.3 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49 | 7.59
32.96 | 3.61 | 3.71 | 3.81 | 3.91 | 4 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 5.99 | 6.09 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49 | 7.59
32.97 | 3.61 | 3.71 | 3.81 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.79 | 5.89 | 5.99 | 6.09 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49 | 7.58
32.98 | 3.61 | 3.7 | 3.8 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.4 | 5.5 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.29 | 7.38 | 7.48 | 7.58
32.99 | 3.6 | 3.7 | 3.8 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.39 | 5.49 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.18 | 7.28 | 7.38 | 7.48 | 7.58
33 | 3.6 | 3.7 | 3.8 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.48 | 7.58
33.01 | 3.6 | 3.7 | 3.8 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.9 | 5 | 5.09 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.69 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.48 | 7.58
33.02 | 3.6 | 3.7 | 3.8 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.89 | 4.99 | 5.09 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.19 | 6.29 | 6.39 | 6.49 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.48 | 7.58
33.03 | 3.6 | 3.7 | 3.8 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.6 | 4.7 | 4.79 | 4.89 | 4.99 | 5.09 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.19 | 6.29 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.48 | 7.58
33.04 | 3.6 | 3.7 | 3.8 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.59 | 4.69 | 4.79 | 4.89 | 4.99 | 5.09 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.48 | 7.58
33.05 | 3.6 | 3.7 | 3.8 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.4 | 4.49 | 4.59 | 4.69 | 4.79 | 4.89 | 4.99 | 5.09 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.48 | 7.57
33.06 | 3.6 | 3.7 | 3.8 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.39 | 4.49 | 4.59 | 4.69 | 4.79 | 4.89 | 4.99 | 5.09 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.69 | 5.79 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.28 | 7.37 | 7.47 | 7.57
33.07 | 3.6 | 3.7 | 3.8 | 3.9 | 4 | 4.1 | 4.2 | 4.29 | 4.39 | 4.49 | 4.59 | 4.69 | 4.79 | 4.89 | 4.99 | 5.09 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.69 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.08 | 7.17 | 7.27 | 7.37 | 7.47 | 7.57
33.08 | 3.6 | 3.7 | 3.8 | 3.9 | 4 | 4.09 | 4.19 | 4.29 | 4.39 | 4.49 | 4.59 | 4.69 | 4.79 | 4.89 | 4.99 | 5.09 | 5.19 | 5.29 | 5.39 | 5.49 | 5.58 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.07 | 7.17 | 7.27 | 7.37 | 7.47 | 7.57
33.09 | 3.6 | 3.7 | 3.8 | 3.9 | 3.99 | 4.09 | 4.19 | 4.29 | 4.39 | 4.49 | 4.59 | 4.69 | 4.79 | 4.89 | 4.99 | 5.09 | 5.19 | 5.29 | 5.39 | 5.48 | 5.58 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.37 | 7.47 | 7.57
33.1 | 3.6 | 3.7 | 3.8 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.39 | 4.49 | 4.59 | 4.69 | 4.79 | 4.89 | 4.99 | 5.09 | 5.19 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.37 | 7.47 | 7.57
33.11 | 3.6 | 3.7 | 3.79 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.39 | 4.49 | 4.59 | 4.69 | 4.79 | 4.89 | 4.99 | 5.09 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.37 | 7.47 | 7.57
33.12 | 3.6 | 3.69 | 3.79 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.39 | 4.49 | 4.59 | 4.69 | 4.79 | 4.89 | 4.99 | 5.08 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.37 | 7.47 | 7.57
33.13 | 3.59 | 3.69 | 3.79 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.39 | 4.49 | 4.59 | 4.69 | 4.79 | 4.89 | 4.98 | 5.08 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.37 | 7.47 | 7.56
33.14 | 3.59 | 3.69 | 3.79 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.39 | 4.49 | 4.59 | 4.69 | 4.79 | 4.88 | 4.98 | 5.08 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.36 | 7.46 | 7.56
33.15 | 3.59 | 3.69 | 3.79 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.39 | 4.49 | 4.59 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.88 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.26 | 7.36 | 7.46 | 7.56
33.16 | 3.59 | 3.69 | 3.79 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.39 | 4.49 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.87 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46 | 7.56
33.17 | 3.59 | 3.69 | 3.79 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.39 | 4.48 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46 | 7.56
33.18 | 3.59 | 3.69 | 3.79 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.38 | 4.48 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.67 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.86 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46 | 7.56
33.19 | 3.59 | 3.69 | 3.79 | 3.89 | 3.99 | 4.09 | 4.19 | 4.28 | 4.38 | 4.48 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.18 | 5.28 | 5.38 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46 | 7.56
33.2 | 3.59 | 3.69 | 3.79 | 3.89 | 3.99 | 4.09 | 4.18 | 4.28 | 4.38 | 4.48 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.18 | 5.28 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46 | 7.56
33.21 | 3.59 | 3.69 | 3.79 | 3.89 | 3.99 | 4.08 | 4.18 | 4.28 | 4.38 | 4.48 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.18 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.46 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46 | 7.55
33.22 | 3.59 | 3.69 | 3.79 | 3.89 | 3.98 | 4.08 | 4.18 | 4.28 | 4.38 | 4.48 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.45 | 7.55
33.23 | 3.59 | 3.69 | 3.79 | 3.89 | 3.98 | 4.08 | 4.18 | 4.28 | 4.38 | 4.48 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.97 | 6.07 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.16 | 7.25 | 7.35 | 7.45 | 7.55
33.24 | 3.59 | 3.69 | 3.79 | 3.88 | 3.98 | 4.08 | 4.18 | 4.28 | 4.38 | 4.48 | 4.58 | 4.68 | 4.78 | 4.88 | 4.97 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.97 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.15 | 7.25 | 7.35 | 7.45 | 7.55
33.25 | 3.59 | 3.69 | 3.78 | 3.88 | 3.98 | 4.08 | 4.18 | 4.28 | 4.38 | 4.48 | 4.58 | 4.68 | 4.78 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.05 | 7.15 | 7.25 | 7.35 | 7.45 | 7.55
33.26 | 3.59 | 3.68 | 3.78 | 3.88 | 3.98 | 4.08 | 4.18 | 4.28 | 4.38 | 4.48 | 4.58 | 4.68 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.76 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.45 | 7.55
33.27 | 3.58 | 3.68 | 3.78 | 3.88 | 3.98 | 4.08 | 4.18 | 4.28 | 4.38 | 4.48 | 4.58 | 4.67 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.76 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.45 | 7.55
33.28 | 3.58 | 3.68 | 3.78 | 3.88 | 3.98 | 4.08 | 4.18 | 4.28 | 4.38 | 4.48 | 4.57 | 4.67 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.65 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.45 | 7.55
33.29 | 3.58 | 3.68 | 3.78 | 3.88 | 3.98 | 4.08 | 4.18 | 4.28 | 4.38 | 4.47 | 4.57 | 4.67 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.55 | 6.65 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.45 | 7.54
33.3 | 3.58 | 3.68 | 3.78 | 3.88 | 3.98 | 4.08 | 4.18 | 4.28 | 4.38 | 4.47 | 4.57 | 4.67 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.37 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.44 | 7.54
33.31 | 3.58 | 3.68 | 3.78 | 3.88 | 3.98 | 4.08 | 4.18 | 4.28 | 4.37 | 4.47 | 4.57 | 4.67 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.34 | 7.44 | 7.54
33.32 | 3.58 | 3.68 | 3.78 | 3.88 | 3.98 | 4.08 | 4.18 | 4.27 | 4.37 | 4.47 | 4.57 | 4.67 | 4.77 | 4.87 | 4.97 | 5.07 | 5.17 | 5.26 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.16 | 6.25 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.95 | 7.05 | 7.14 | 7.24 | 7.34 | 7.44 | 7.54
33.33 | 3.58 | 3.68 | 3.78 | 3.88 | 3.98 | 4.08 | 4.17 | 4.27 | 4.37 | 4.47 | 4.57 | 4.67 | 4.77 | 4.87 | 4.97 | 5.07 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.95 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44 | 7.54
33.34 | 3.58 | 3.68 | 3.78 | 3.88 | 3.98 | 4.07 | 4.17 | 4.27 | 4.37 | 4.47 | 4.57 | 4.67 | 4.77 | 4.87 | 4.97 | 5.06 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.05 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44 | 7.54
33.35 | 3.58 | 3.68 | 3.78 | 3.88 | 3.98 | 4.07 | 4.17 | 4.27 | 4.37 | 4.47 | 4.57 | 4.67 | 4.77 | 4.87 | 4.96 | 5.06 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.86 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44 | 7.54
33.36 | 3.58 | 3.68 | 3.78 | 3.88 | 3.97 | 4.07 | 4.17 | 4.27 | 4.37 | 4.47 | 4.57 | 4.67 | 4.77 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.65 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44 | 7.54
33.37 | 3.58 | 3.68 | 3.78 | 3.87 | 3.97 | 4.07 | 4.17 | 4.27 | 4.37 | 4.47 | 4.57 | 4.67 | 4.77 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.66 | 5.75 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44 | 7.53
33.38 | 3.58 | 3.68 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.27 | 4.37 | 4.47 | 4.57 | 4.67 | 4.76 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.45 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.43 | 7.53
33.39 | 3.58 | 3.68 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.27 | 4.37 | 4.47 | 4.57 | 4.66 | 4.76 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.36 | 5.46 | 5.55 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.33 | 7.43 | 7.53
33.4 | 3.58 | 3.67 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.27 | 4.37 | 4.47 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.36 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.23 | 7.33 | 7.43 | 7.53
33.41 | 3.57 | 3.67 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.27 | 4.37 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.15 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53
33.42 | 3.57 | 3.67 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.27 | 4.37 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.06 | 5.16 | 5.26 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.95 | 6.05 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53
33.43 | 3.57 | 3.67 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.27 | 4.36 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.06 | 5.16 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.95 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53
33.44 | 3.57 | 3.67 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.06 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.74 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43 | 7.53
33.45 | 3.57 | 3.67 | 3.77 | 3.87 | 3.97 | 4.07 | 4.17 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.96 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.85 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.64 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43 | 7.52
33.46 | 3.57 | 3.67 | 3.77 | 3.87 | 3.97 | 4.07 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.75 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.54 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.42 | 7.52
33.47 | 3.57 | 3.67 | 3.77 | 3.87 | 3.97 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.76 | 4.86 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.65 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.32 | 7.42 | 7.52
33.48 | 3.57 | 3.67 | 3.77 | 3.87 | 3.97 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.76 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.55 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.34 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.22 | 7.32 | 7.42 | 7.52
33.49 | 3.57 | 3.67 | 3.77 | 3.87 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52
33.5 | 3.57 | 3.67 | 3.77 | 3.87 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.66 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.45 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.24 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52
33.51 | 3.57 | 3.67 | 3.77 | 3.86 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.56 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.35 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.14 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52
33.52 | 3.57 | 3.67 | 3.76 | 3.86 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.25 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.04 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.83 | 6.92 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.52
33.53 | 3.57 | 3.67 | 3.76 | 3.86 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.46 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.94 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.73 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.32 | 7.42 | 7.51
33.54 | 3.57 | 3.66 | 3.76 | 3.86 | 3.96 | 4.06 | 4.16 | 4.26 | 4.36 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.15 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.63 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.32 | 7.41 | 7.51
33.55 | 3.57 | 3.66 | 3.76 | 3.86 | 3.96 | 4.06 | 4.16 | 4.26 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.05 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.84 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.53 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.31 | 7.41 | 7.51
33.56 | 3.56 | 3.66 | 3.76 | 3.86 | 3.96 | 4.06 | 4.16 | 4.26 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.74 | 5.83 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.43 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.22 | 7.31 | 7.41 | 7.51
33.57 | 3.56 | 3.66 | 3.76 | 3.86 | 3.96 | 4.06 | 4.16 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.95 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.64 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.12 | 7.21 | 7.31 | 7.41 | 7.51
33.58 | 3.56 | 3.66 | 3.76 | 3.86 | 3.96 | 4.06 | 4.16 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.85 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.23 | 6.33 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.02 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51
33.59 | 3.56 | 3.66 | 3.76 | 3.86 | 3.96 | 4.06 | 4.15 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.54 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.23 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.92 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51
33.6 | 3.56 | 3.66 | 3.76 | 3.86 | 3.96 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.75 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.44 | 5.53 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.13 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.82 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.51
33.61 | 3.56 | 3.66 | 3.76 | 3.86 | 3.96 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.55 | 4.65 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.81 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.5
33.62 | 3.56 | 3.66 | 3.76 | 3.86 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.55 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.34 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.93 | 6.03 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.72 | 6.81 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.41 | 7.5
33.63 | 3.56 | 3.66 | 3.76 | 3.86 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.55 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.24 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.93 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.62 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.21 | 7.31 | 7.4 | 7.5
33.64 | 3.56 | 3.66 | 3.76 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.83 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.52 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.21 | 7.3 | 7.4 | 7.5
33.65 | 3.56 | 3.66 | 3.76 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.45 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.14 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.11 | 7.2 | 7.3 | 7.4 | 7.5
33.66 | 3.56 | 3.66 | 3.75 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.04 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.73 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.42 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5
33.67 | 3.56 | 3.66 | 3.75 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.35 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.63 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.32 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7.01 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5
33.68 | 3.56 | 3.65 | 3.75 | 3.85 | 3.95 | 4.05 | 4.15 | 4.25 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.94 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.22 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.91 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5
33.69 | 3.56 | 3.65 | 3.75 | 3.85 | 3.95 | 4.05 | 4.15 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.53 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.81 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.5
33.7 | 3.55 | 3.65 | 3.75 | 3.85 | 3.95 | 4.05 | 4.15 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.84 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.12 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.4 | 7.49
33.71 | 3.55 | 3.65 | 3.75 | 3.85 | 3.95 | 4.05 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.74 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.43 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.02 | 6.11 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.71 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.3 | 7.39 | 7.49
33.72 | 3.55 | 3.65 | 3.75 | 3.85 | 3.95 | 4.05 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.33 | 5.42 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.51 | 6.61 | 6.7 | 6.8 | 6.9 | 7 | 7.1 | 7.2 | 7.29 | 7.39 | 7.49
33.73 | 3.55 | 3.65 | 3.75 | 3.85 | 3.95 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.64 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.82 | 5.92 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.51 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.1 | 7.19 | 7.29 | 7.39 | 7.49
33.74 | 3.55 | 3.65 | 3.75 | 3.85 | 3.95 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.23 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.82 | 5.91 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.1 | 7.19 | 7.29 | 7.39 | 7.49
33.75 | 3.55 | 3.65 | 3.75 | 3.85 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.54 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.31 | 6.41 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49
33.76 | 3.55 | 3.65 | 3.75 | 3.85 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.13 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.72 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.31 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 6.99 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49
33.77 | 3.55 | 3.65 | 3.75 | 3.85 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.44 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.03 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.71 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 6.99 | 7.09 | 7.19 | 7.29 | 7.39 | 7.49
33.78 | 3.55 | 3.65 | 3.75 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.62 | 5.71 | 5.81 | 5.91 | 6.01 | 6.11 | 6.21 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.89 | 6.99 | 7.09 | 7.19 | 7.29 | 7.39 | 7.48
33.79 | 3.55 | 3.65 | 3.75 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.34 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.93 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.52 | 5.61 | 5.71 | 5.81 | 5.91 | 6.01 | 6.11 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.29 | 7.38 | 7.48
33.8 | 3.55 | 3.65 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6.01 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.28 | 7.38 | 7.48
33.81 | 3.55 | 3.65 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.24 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.83 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.42 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6.01 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.19 | 7.28 | 7.38 | 7.48
33.82 | 3.55 | 3.64 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.09 | 7.18 | 7.28 | 7.38 | 7.48
33.83 | 3.55 | 3.64 | 3.74 | 3.84 | 3.94 | 4.04 | 4.14 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.73 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.32 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.91 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.08 | 7.18 | 7.28 | 7.38 | 7.48
33.84 | 3.54 | 3.64 | 3.74 | 3.84 | 3.94 | 4.04 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.31 | 5.41 | 5.51 | 5.61 | 5.71 | 5.81 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 6.99 | 7.08 | 7.18 | 7.28 | 7.38 | 7.48
33.85 | 3.54 | 3.64 | 3.74 | 3.84 | 3.94 | 4.04 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.63 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.22 | 5.31 | 5.41 | 5.51 | 5.61 | 5.71 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.89 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.48
33.86 | 3.54 | 3.64 | 3.74 | 3.84 | 3.94 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.21 | 5.31 | 5.41 | 5.51 | 5.61 | 5.71 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.88 | 6.98 | 7.08 | 7.18 | 7.28 | 7.38 | 7.47
33.87 | 3.54 | 3.64 | 3.74 | 3.84 | 3.94 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.53 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.12 | 5.21 | 5.31 | 5.41 | 5.51 | 5.61 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.29 | 6.39 | 6.49 | 6.59 | 6.69 | 6.79 | 6.88 | 6.98 | 7.08 | 7.18 | 7.28 | 7.37 | 7.47
33.88 | 3.54 | 3.64 | 3.74 | 3.84 | 3.94 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.61 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.69 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.28 | 7.37 | 7.47
33.89 | 3.54 | 3.64 | 3.74 | 3.84 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.43 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.02 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.68 | 6.78 | 6.88 | 6.98 | 7.08 | 7.18 | 7.27 | 7.37 | 7.47
33.9 | 3.54 | 3.64 | 3.74 | 3.84 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.51 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.09 | 6.19 | 6.29 | 6.39 | 6.49 | 6.59 | 6.68 | 6.78 | 6.88 | 6.98 | 7.08 | 7.17 | 7.27 | 7.37 | 7.47
33.91 | 3.54 | 3.64 | 3.74 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.92 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.09 | 6.19 | 6.29 | 6.39 | 6.49 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.08 | 7.17 | 7.27 | 7.37 | 7.47
33.92 | 3.54 | 3.64 | 3.74 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.33 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.41 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 5.99 | 6.09 | 6.19 | 6.29 | 6.39 | 6.49 | 6.58 | 6.68 | 6.78 | 6.88 | 6.98 | 7.07 | 7.17 | 7.27 | 7.37 | 7.47
33.93 | 3.54 | 3.64 | 3.74 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.82 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.4 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 5.99 | 6.09 | 6.19 | 6.29 | 6.39 | 6.48 | 6.58 | 6.68 | 6.78 | 6.88 | 6.97 | 7.07 | 7.17 | 7.27 | 7.37 | 7.47
33.94 | 3.54 | 3.64 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.23 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.31 | 5.4 | 5.5 | 5.6 | 5.7 | 5.8 | 5.89 | 5.99 | 6.09 | 6.19 | 6.29 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.88 | 6.97 | 7.07 | 7.17 | 7.27 | 7.37 | 7.46
33.95 | 3.54 | 3.64 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.72 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.79 | 5.89 | 5.99 | 6.09 | 6.19 | 6.29 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.36 | 7.46
33.96 | 3.54 | 3.63 | 3.73 | 3.83 | 3.93 | 4.03 | 4.13 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.21 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.79 | 5.89 | 5.99 | 6.09 | 6.19 | 6.28 | 6.38 | 6.48 | 6.58 | 6.68 | 6.78 | 6.87 | 6.97 | 7.07 | 7.17 | 7.27 | 7.36 | 7.46
33.97 | 3.54 | 3.63 | 3.73 | 3.83 | 3.93 | 4.03 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.62 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.2 | 5.3 | 5.4 | 5.5 | 5.6 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.19 | 6.28 | 6.38 | 6.48 | 6.58 | 6.68 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.26 | 7.36 | 7.46
33.98 | 3.54 | 3.63 | 3.73 | 3.83 | 3.93 | 4.03 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.11 | 5.2 | 5.3 | 5.4 | 5.5 | 5.6 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.17 | 7.26 | 7.36 | 7.46
33.99 | 3.53 | 3.63 | 3.73 | 3.83 | 3.93 | 4.03 | 4.12 | 4.22 | 4.32 | 4.42 | 4.52 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.1 | 5.2 | 5.3 | 5.4 | 5.5 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.09 | 6.18 | 6.28 | 6.38 | 6.48 | 6.58 | 6.67 | 6.77 | 6.87 | 6.97 | 7.07 | 7.16 | 7.26 | 7.36 | 7.46
34 | 3.53 | 3.63 | 3.73 | 3.83 | 3.93 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5.01 | 5.1 | 5.2 | 5.3 | 5.4 | 5.5 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46
34.01 | 3.53 | 3.63 | 3.73 | 3.83 | 3.93 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5 | 5.1 | 5.2 | 5.3 | 5.4 | 5.49 | 5.59 | 5.69 | 5.79 | 5.89 | 5.99 | 6.08 | 6.18 | 6.28 | 6.38 | 6.48 | 6.57 | 6.67 | 6.77 | 6.87 | 6.97 | 7.06 | 7.16 | 7.26 | 7.36 | 7.46
34.02 | 3.53 | 3.63 | 3.73 | 3.83 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.42 | 4.51 | 4.61 | 4.71 | 4.81 | 4.91 | 5 | 5.1 | 5.2 | 5.3 | 5.4 | 5.49 | 5.59 | 5.69 | 5.79 | 5.89 | 5.98 | 6.08 | 6.18 | 6.28 | 6.38 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.45
34.03 | 3.53 | 3.63 | 3.73 | 3.83 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.39 | 5.49 | 5.59 | 5.69 | 5.79 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.38 | 6.47 | 6.57 | 6.67 | 6.77 | 6.87 | 6.96 | 7.06 | 7.16 | 7.26 | 7.36 | 7.45
34.04 | 3.53 | 3.63 | 3.73 | 3.83 | 3.92 | 4.02 | 4.12 | 4.22 | 4.32 | 4.41 | 4.51 | 4.61 | 4.71 | 4.81 | 4.9 | 5 | 5.1 | 5.2 | 5.3 | 5.39 | 5.49 | 5.59 | 5.69 | 5.79 | 5.88 | 5.98 | 6.08 | 6.18 | 6.28 | 6.37 | 6.47 | 6.57 | 6.67 | 6.77 | 6.86 | 6.96 | 7.06 | 7.16 | 7.26 | 7.35 | 7.45
34.05 | 3.53 | 3.63 | 3.73 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.29 | 5.39 | 5.49 | 5.59 | 5.69 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.76 | 6.86 | 6.96 | 7.06 | 7.16 | 7.25 | 7.35 | 7.45
34.06 | 3.53 | 3.63 | 3.73 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.29 | 5.39 | 5.49 | 5.59 | 5.69 | 5.78 | 5.88 | 5.98 | 6.08 | 6.18 | 6.27 | 6.37 | 6.47 | 6.57 | 6.67 | 6.76 | 6.86 | 6.96 | 7.06 | 7.16 | 7.25 | 7.35 | 7.45
34.07 | 3.53 | 3.63 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.22 | 4.31 | 4.41 | 4.51 | 4.61 | 4.71 | 4.8 | 4.9 | 5 | 5.1 | 5.2 | 5.29 | 5.39 | 5.49 | 5.59 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.15 | 7.25 | 7.35 | 7.45
34.08 | 3.53 | 3.63 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.68 | 5.78 | 5.88 | 5.98 | 6.08 | 6.17 | 6.27 | 6.37 | 6.47 | 6.57 | 6.66 | 6.76 | 6.86 | 6.96 | 7.06 | 7.15 | 7.25 | 7.35 | 7.45
34.09 | 3.53 | 3.63 | 3.72 | 3.82 | 3.92 | 4.02 | 4.12 | 4.21 | 4.31 | 4.41 | 4.51 | 4.61 | 4.7 | 4.8 | 4.9 | 5 | 5.1 | 5.19 | 5.29 | 5.39 | 5.49 | 5.59 | 5.68 | 5.78 | 5.88 | 5.98 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.05 | 7.15 | 7.25 | 7.35 | 7.45
34.1 | 3.53 | 3.62 | 3.72 | 3.82 | 3.92 | 4.02 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.6 | 4.7 | 4.8 | 4.9 | 5 | 5.09 | 5.19 | 5.29 | 5.39 | 5.49 | 5.58 | 5.68 | 5.78 | 5.88 | 5.98 | 6.07 | 6.17 | 6.27 | 6.37 | 6.47 | 6.56 | 6.66 | 6.76 | 6.86 | 6.96 | 7.05 | 7.15 | 7.25 | 7.35 | 7.44
34.11 | 3.53 | 3.62 | 3.72 | 3.82 | 3.92 | 4.02 | 4.11 | 4.21 | 4.31 | 4.41 | 4.51 | 4.6 | 4.7 | 4.8 | 4.9 | 5 | 5.09 | 5.19 | 5.29 | 5.39 | 5.49 | 5.58 | 5.68 | 5.78 | 5.88 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.46 | 6.56 | 6.66 | 6.76 | 6.86 | 6.95 | 7.05 | 7.15 | 7.25 | 7.35 | 7.44
34.12 | 3.53 | 3.62 | 3.72 | 3.82 | 3.92 | 4.02 | 4.11 | 4.21 | 4.31 | 4.41 | 4.5 | 4.6 | 4.7 | 4.8 | 4.9 | 4.99 | 5.09 | 5.19 | 5.29 | 5.39 | 5.48 | 5.58 | 5.68 | 5.78 | 5.88 | 5.97 | 6.07 | 6.17 | 6.27 | 6.37 | 6.46 | 6.56 | 6.66 | 6.76 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.34 | 7.44
34.13 | 3.52 | 3.62 | 3.72 | 3.82 | 3.92 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.5 | 4.6 | 4.7 | 4.8 | 4.9 | 4.99 | 5.09 | 5.19 | 5.29 | 5.39 | 5.48 | 5.58 | 5.68 | 5.78 | 5.87 | 5.97 | 6.07 | 6.17 | 6.27 | 6.36 | 6.46 | 6.56 | 6.66 | 6.76 | 6.85 | 6.95 | 7.05 | 7.15 | 7.25 | 7.34 | 7.44
34.14 | 3.52 | 3.62 | 3.72 | 3.82 | 3.92 | 4.01 | 4.11 | 4.21 | 4.31 | 4.41 | 4.5 | 4.6 | 4.7 | 4.8 | 4.89 | 4.99 | 5.09 | 5.19 | 5.29 | 5.38 | 5.48 | 5.58 | 5.68 | 5.78 | 5.87 | 5.97 | 6.07 | 6.17 | 6.27 | 6.36 | 6.46 | 6.56 | 6.66 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.24 | 7.34 | 7.44
34.15 | 3.52 | 3.62 | 3.72 | 3.82 | 3.92 | 4.01 | 4.11 | 4.21 | 4.31 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.89 | 4.99 | 5.09 | 5.19 | 5.29 | 5.38 | 5.48 | 5.58 | 5.68 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.26 | 6.36 | 6.46 | 6.56 | 6.66 | 6.75 | 6.85 | 6.95 | 7.05 | 7.15 | 7.24 | 7.34 | 7.44
34.16 | 3.52 | 3.62 | 3.72 | 3.82 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.4 | 4.5 | 4.6 | 4.7 | 4.8 | 4.89 | 4.99 | 5.09 | 5.19 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.77 | 5.87 | 5.97 | 6.07 | 6.17 | 6.26 | 6.36 | 6.46 | 6.56 | 6.65 | 6.75 | 6.85 | 6.95 | 7.05 | 7.14 | 7.24 | 7.34 | 7.44
34.17 | 3.52 | 3.62 | 3.72 | 3.82 | 3.91 | 4.01 | 4.11 | 4.21 | 4.31 | 4.4 | 4.5 | 4.6 | 4.7 | 4.79 | 4.89 | 4.99 | 5.09 | 5.19 | 5.28 | 5.38 | 5.48 | 5.58 | 5.68 | 5.77 | 5.87 | 5.97 | 6.07 | 6.16 | 6.26 | 6.36 | 6.46 | 6.56 | 6.65 | 6.75 | 6.85 | 6.95 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44
34.18 | 3.52 | 3.62 | 3.72 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.3 | 4.4 | 4.5 | 4.6 | 4.7 | 4.79 | 4.89 | 4.99 | 5.09 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.67 | 5.77 | 5.87 | 5.97 | 6.07 | 6.16 | 6.26 | 6.36 | 6.46 | 6.55 | 6.65 | 6.75 | 6.85 | 6.95 | 7.04 | 7.14 | 7.24 | 7.34 | 7.44
34.19 | 3.52 | 3.62 | 3.72 | 3.81 | 3.91 | 4.01 | 4.11 | 4.21 | 4.3 | 4.4 | 4.5 | 4.6 | 4.69 | 4.79 | 4.89 | 4.99 | 5.09 | 5.18 | 5.28 | 5.38 | 5.48 | 5.58 | 5.67 | 5.77 | 5.87 | 5.97 | 6.06 | 6.16 | 6.26 | 6.36 | 6.46 | 6.55 | 6.65 | 6.75 | 6.85 | 6.94 | 7.04 | 7.14 | 7.24 | 7.34 | 7.43
34.2 | 3.52 | 3.62 | 3.72 | 3.81 | 3.91 | 4.01 | 4.11 | 4.2 | 4.3 | 4.4 | 4.5 | 4.6 | 4.69 | 4.79 | 4.89 | 4.99 | 5.09 | 5.18 | 5.28 | 5.38 | 5.48 | 5.57 | 5.67 | 5.77 | 5.87 | 5.97 | 6.06 | 6.16 | 6.26 | 6.36 | 6.45 | 6.55 | 6.65 | 6.75 | 6.85 | 6.94 | 7.04 | 7.14 | 7.24 | 7.33 | 7.43
34.21 | 3.52 | 3.62 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.2 | 4.3 | 4.4 | 4.5 | 4.6 | 4.69 | 4.79 | 4.89 | 4.99 | 5.08 | 5.18 | 5.28 | 5.38 | 5.48 | 5.57 | 5.67 | 5.77 | 5.87 | 5.96 | 6.06 | 6.16 | 6.26 | 6.36 | 6.45 | 6.55 | 6.65 | 6.75 | 6.84 | 6.94 | 7.04 | 7.14 | 7.24 | 7.33 | 7.43
34.22 | 3.52 | 3.62 | 3.71 | 3.81 | 3.91 | 4.01 | 4.11 | 4.2 | 4.3 | 4.4 | 4.5 | 4.59 | 4.69 | 4.79 | 4.89 | 4.99 | 5.08 | 5.18 | 5.28 | 5.38 | 5.47 | 5.57 | 5.67 | 5.77 | 5.87 | 5.96 | 6.06 | 6.16 | 6.26 | 6.35 | 6.45 | 6.55 | 6.65 | 6.75 | 6.84 | 6.94 | 7.04 | 7.14 | 7.23 | 7.33 | 7.43
34.23 | 3.52 | 3.62 | 3.71 | 3.81 | 3.91 | 4.01 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.59 | 4.69 | 4.79 | 4.89 | 4.98 | 5.08 | 5.18 | 5.28 | 5.38 | 5.47 | 5.57 | 5.67 | 5.77 | 5.86 | 5.96 | 6.06 | 6.16 | 6.26 | 6.35 | 6.45 | 6.55 | 6.65 | 6.74 | 6.84 | 6.94 | 7.04 | 7.14 | 7.23 | 7.33 | 7.43
34.24 | 3.52 | 3.61 | 3.71 | 3.81 | 3.91 | 4.01 | 4.1 | 4.2 | 4.3 | 4.4 | 4.5 | 4.59 | 4.69 | 4.79 | 4.89 | 4.98 | 5.08 | 5.18 | 5.28 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.86 | 5.96 | 6.06 | 6.16 | 6.25 | 6.35 | 6.45 | 6.55 | 6.65 | 6.74 | 6.84 | 6.94 | 7.04 | 7.13 | 7.23 | 7.33 | 7.43
34.25 | 3.52 | 3.61 | 3.71 | 3.81 | 3.91 | 4.01 | 4.1 | 4.2 | 4.3 | 4.4 | 4.49 | 4.59 | 4.69 | 4.79 | 4.89 | 4.98 | 5.08 | 5.18 | 5.28 | 5.37 | 5.47 | 5.57 | 5.67 | 5.77 | 5.86 | 5.96 | 6.06 | 6.16 | 6.25 | 6.35 | 6.45 | 6.55 | 6.64 | 6.74 | 6.84 | 6.94 | 7.04 | 7.13 | 7.23 | 7.33 | 7.43
34.26 | 3.52 | 3.61 | 3.71 | 3.81 | 3.91 | 4 | 4.1 | 4.2 | 4.3 | 4.4 | 4.49 | 4.59 | 4.69 | 4.79 | 4.88 | 4.98 | 5.08 | 5.18 | 5.28 | 5.37 | 5.47 | 5.57 | 5.67 | 5.76 | 5.86 | 5.96 | 6.06 | 6.15 | 6.25 | 6.35 | 6.45 | 6.55 | 6.64 | 6.74 | 6.84 | 6.94 | 7.03 | 7.13 | 7.23 | 7.33 | 7.43
34.27 | 3.52 | 3.61 | 3.71 | 3.81 | 3.91 | 4 | 4.1 | 4.2 | 4.3 | 4.39 | 4.49 | 4.59 | 4.69 | 4.79 | 4.88 | 4.98 | 5.08 | 5.18 | 5.27 | 5.37 | 5.47 | 5.57 | 5.67 | 5.76 | 5.86 | 5.96 | 6.06 | 6.15 | 6.25 | 6.35 | 6.45 | 6.54 | 6.64 | 6.74 | 6.84 | 6.94 | 7.03 | 7.13 | 7.23 | 7.33 | 7.42
34.28 | 3.51 | 3.61 | 3.71 | 3.81 | 3.91 | 4 | 4.1 | 4.2 | 4.3 | 4.39 | 4.49 | 4.59 | 4.69 | 4.78 | 4.88 | 4.98 | 5.08 | 5.18 | 5.27 | 5.37 | 5.47 | 5.57 | 5.66 | 5.76 | 5.86 | 5.96 | 6.06 | 6.15 | 6.25 | 6.35 | 6.45 | 6.54 | 6.64 | 6.74 | 6.84 | 6.93 | 7.03 | 7.13 | 7.23 | 7.33 | 7.42
34.29 | 3.51 | 3.61 | 3.71 | 3.81 | 3.9 | 4 | 4.1 | 4.2 | 4.3 | 4.39 | 4.49 | 4.59 | 4.69 | 4.78 | 4.88 | 4.98 | 5.08 | 5.17 | 5.27 | 5.37 | 5.47 | 5.57 | 5.66 | 5.76 | 5.86 | 5.96 | 6.05 | 6.15 | 6.25 | 6.35 | 6.44 | 6.54 | 6.64 | 6.74 | 6.84 | 6.93 | 7.03 | 7.13 | 7.23 | 7.32 | 7.42
34.3 | 3.51 | 3.61 | 3.71 | 3.81 | 3.9 | 4 | 4.1 | 4.2 | 4.29 | 4.39 | 4.49 | 4.59 | 4.69 | 4.78 | 4.88 | 4.98 | 5.08 | 5.17 | 5.27 | 5.37 | 5.47 | 5.56 | 5.66 | 5.76 | 5.86 | 5.96 | 6.05 | 6.15 | 6.25 | 6.35 | 6.44 | 6.54 | 6.64 | 6.74 | 6.83 | 6.93 | 7.03 | 7.13 | 7.23 | 7.32 | 7.42
34.31 | 3.51 | 3.61 | 3.71 | 3.81 | 3.9 | 4 | 4.1 | 4.2 | 4.29 | 4.39 | 4.49 | 4.59 | 4.68 | 4.78 | 4.88 | 4.98 | 5.08 | 5.17 | 5.27 | 5.37 | 5.47 | 5.56 | 5.66 | 5.76 | 5.86 | 5.95 | 6.05 | 6.15 | 6.25 | 6.35 | 6.44 | 6.54 | 6.64 | 6.74 | 6.83 | 6.93 | 7.03 | 7.13 | 7.22 | 7.32 | 7.42
34.32 | 3.51 | 3.61 | 3.71 | 3.8 | 3.9 | 4 | 4.1 | 4.2 | 4.29 | 4.39 | 4.49 | 4.59 | 4.68 | 4.78 | 4.88 | 4.98 | 5.07 | 5.17 | 5.27 | 5.37 | 5.47 | 5.56 | 5.66 | 5.76 | 5.86 | 5.95 | 6.05 | 6.15 | 6.25 | 6.34 | 6.44 | 6.54 | 6.64 | 6.73 | 6.83 | 6.93 | 7.03 | 7.13 | 7.22 | 7.32 | 7.42
34.33 | 3.51 | 3.61 | 3.71 | 3.8 | 3.9 | 4 | 4.1 | 4.19 | 4.29 | 4.39 | 4.49 | 4.59 | 4.68 | 4.78 | 4.88 | 4.98 | 5.07 | 5.17 | 5.27 | 5.37 | 5.46 | 5.56 | 5.66 | 5.76 | 5.85 | 5.95 | 6.05 | 6.15 | 6.25 | 6.34 | 6.44 | 6.54 | 6.64 | 6.73 | 6.83 | 6.93 | 7.03 | 7.12 | 7.22 | 7.32 | 7.42
34.34 | 3.51 | 3.61 | 3.71 | 3.8 | 3.9 | 4 | 4.1 | 4.19 | 4.29 | 4.39 | 4.49 | 4.58 | 4.68 | 4.78 | 4.88 | 4.98 | 5.07 | 5.17 | 5.27 | 5.37 | 5.46 | 5.56 | 5.66 | 5.76 | 5.85 | 5.95 | 6.05 | 6.15 | 6.24 | 6.34 | 6.44 | 6.54 | 6.63 | 6.73 | 6.83 | 6.93 | 7.03 | 7.12 | 7.22 | 7.32 | 7.42
34.35 | 3.51 | 3.61 | 3.7 | 3.8 | 3.9 | 4 | 4.1 | 4.19 | 4.29 | 4.39 | 4.49 | 4.58 | 4.68 | 4.78 | 4.88 | 4.97 | 5.07 | 5.17 | 5.27 | 5.36 | 5.46 | 5.56 | 5.66 | 5.76 | 5.85 | 5.95 | 6.05 | 6.15 | 6.24 | 6.34 | 6.44 | 6.54 | 6.63 | 6.73 | 6.83 | 6.93 | 7.02 | 7.12 | 7.22 | 7.32 | 7.41
34.36 | 3.51 | 3.61 | 3.7 | 3.8 | 3.9 | 4 | 4.09 | 4.19 | 4.29 | 4.39 | 4.49 | 4.58 | 4.68 | 4.78 | 4.88 | 4.97 | 5.07 | 5.17 | 5.27 | 5.36 | 5.46 | 5.56 | 5.66 | 5.75 | 5.85 | 5.95 | 6.05 | 6.14 | 6.24 | 6.34 | 6.44 | 6.53 | 6.63 | 6.73 | 6.83 | 6.93 | 7.02 | 7.12 | 7.22 | 7.32 | 7.41
34.37 | 3.51 | 3.61 | 3.7 | 3.8 | 3.9 | 4 | 4.09 | 4.19 | 4.29 | 4.39 | 4.48 | 4.58 | 4.68 | 4.78 | 4.87 | 4.97 | 5.07 | 5.17 | 5.27 | 5.36 | 5.46 | 5.56 | 5.66 | 5.75 | 5.85 | 5.95 | 6.05 | 6.14 | 6.24 | 6.34 | 6.44 | 6.53 | 6.63 | 6.73 | 6.83 | 6.92 | 7.02 | 7.12 | 7.22 | 7.31 | 7.41
34.38 | 3.51 | 3.61 | 3.7 | 3.8 | 3.9 | 4 | 4.09 | 4.19 | 4.29 | 4.39 | 4.48 | 4.58 | 4.68 | 4.78 | 4.87 | 4.97 | 5.07 | 5.17 | 5.26 | 5.36 | 5.46 | 5.56 | 5.65 | 5.75 | 5.85 | 5.95 | 6.04 | 6.14 | 6.24 | 6.34 | 6.44 | 6.53 | 6.63 | 6.73 | 6.83 | 6.92 | 7.02 | 7.12 | 7.22 | 7.31 | 7.41
34.39 | 3.51 | 3.6 | 3.7 | 3.8 | 3.9 | 3.99 | 4.09 | 4.19 | 4.29 | 4.39 | 4.48 | 4.58 | 4.68 | 4.78 | 4.87 | 4.97 | 5.07 | 5.17 | 5.26 | 5.36 | 5.46 | 5.56 | 5.65 | 5.75 | 5.85 | 5.95 | 6.04 | 6.14 | 6.24 | 6.34 | 6.43 | 6.53 | 6.63 | 6.73 | 6.82 | 6.92 | 7.02 | 7.12 | 7.21 | 7.31 | 7.41
34.4 | 3.51 | 3.6 | 3.7 | 3.8 | 3.9 | 3.99 | 4.09 | 4.19 | 4.29 | 4.38 | 4.48 | 4.58 | 4.68 | 4.77 | 4.87 | 4.97 | 5.07 | 5.16 | 5.26 | 5.36 | 5.46 | 5.56 | 5.65 | 5.75 | 5.85 | 5.95 | 6.04 | 6.14 | 6.24 | 6.34 | 6.43 | 6.53 | 6.63 | 6.73 | 6.82 | 6.92 | 7.02 | 7.12 | 7.21 | 7.31 | 7.41
34.41 | 3.51 | 3.6 | 3.7 | 3.8 | 3.9 | 3.99 | 4.09 | 4.19 | 4.29 | 4.38 | 4.48 | 4.58 | 4.68 | 4.77 | 4.87 | 4.97 | 5.07 | 5.16 | 5.26 | 5.36 | 5.46 | 5.55 | 5.65 | 5.75 | 5.85 | 5.94 | 6.04 | 6.14 | 6.24 | 6.33 | 6.43 | 6.53 | 6.63 | 6.72 | 6.82 | 6.92 | 7.02 | 7.11 | 7.21 | 7.31 | 7.41
34.42 | 3.5 | 3.6 | 3.7 | 3.8 | 3.89 | 3.99 | 4.09 | 4.19 | 4.29 | 4.38 | 4.48 | 4.58 | 4.68 | 4.77 | 4.87 | 4.97 | 5.07 | 5.16 | 5.26 | 5.36 | 5.46 | 5.55 | 5.65 | 5.75 | 5.85 | 5.94 | 6.04 | 6.14 | 6.24 | 6.33 | 6.43 | 6.53 | 6.63 | 6.72 | 6.82 | 6.92 | 7.02 | 7.11 | 7.21 | 7.31 | 7.41
34.43 | 3.5 | 3.6 | 3.7 | 3.8 | 3.89 | 3.99 | 4.09 | 4.19 | 4.28 | 4.38 | 4.48 | 4.58 | 4.67 | 4.77 | 4.87 | 4.97 | 5.06 | 5.16 | 5.26 | 5.36 | 5.45 | 5.55 | 5.65 | 5.75 | 5.84 | 5.94 | 6.04 | 6.14 | 6.23 | 6.33 | 6.43 | 6.53 | 6.62 | 6.72 | 6.82 | 6.92 | 7.01 | 7.11 | 7.21 | 7.31 | 7.4
34.44 | 3.5 | 3.6 | 3.7 | 3.8 | 3.89 | 3.99 | 4.09 | 4.19 | 4.28 | 4.38 | 4.48 | 4.58 | 4.67 | 4.77 | 4.87 | 4.97 | 5.06 | 5.16 | 5.26 | 5.36 | 5.45 | 5.55 | 5.65 | 5.75 | 5.84 | 5.94 | 6.04 | 6.14 | 6.23 | 6.33 | 6.43 | 6.53 | 6.62 | 6.72 | 6.82 | 6.92 | 7.01 | 7.11 | 7.21 | 7.31 | 7.4
34.45 | 3.5 | 3.6 | 3.7 | 3.8 | 3.89 | 3.99 | 4.09 | 4.19 | 4.28 | 4.38 | 4.48 | 4.58 | 4.67 | 4.77 | 4.87 | 4.97 | 5.06 | 5.16 | 5.26 | 5.36 | 5.45 | 5.55 | 5.65 | 5.75 | 5.84 | 5.94 | 6.04 | 6.14 | 6.23 | 6.33 | 6.43 | 6.53 | 6.62 | 6.72 | 6.82 | 6.91 | 7.01 | 7.11 | 7.21 | 7.3 | 7.4
34.46 | 3.5 | 3.6 | 3.7 | 3.79 | 3.89 | 3.99 | 4.09 | 4.18 | 4.28 | 4.38 | 4.48 | 4.57 | 4.67 | 4.77 | 4.87 | 4.96 | 5.06 | 5.16 | 5.26 | 5.35 | 5.45 | 5.55 | 5.65 | 5.74 | 5.84 | 5.94 | 6.04 | 6.13 | 6.23 | 6.33 | 6.43 | 6.52 | 6.62 | 6.72 | 6.82 | 6.91 | 7.01 | 7.11 | 7.21 | 7.3 | 7.4
34.47 | 3.5 | 3.6 | 3.7 | 3.79 | 3.89 | 3.99 | 4.09 | 4.18 | 4.28 | 4.38 | 4.48 | 4.57 | 4.67 | 4.77 | 4.87 | 4.96 | 5.06 | 5.16 | 5.26 | 5.35 | 5.45 | 5.55 | 5.65 | 5.74 | 5.84 | 5.94 | 6.04 | 6.13 | 6.23 | 6.33 | 6.43 | 6.52 | 6.62 | 6.72 | 6.82 | 6.91 | 7.01 | 7.11 | 7.21 | 7.3 | 7.4
34.48 | 3.5 | 3.6 | 3.7 | 3.79 | 3.89 | 3.99 | 4.09 | 4.18 | 4.28 | 4.38 | 4.48 | 4.57 | 4.67 | 4.77 | 4.87 | 4.96 | 5.06 | 5.16 | 5.26 | 5.35 | 5.45 | 5.55 | 5.64 | 5.74 | 5.84 | 5.94 | 6.03 | 6.13 | 6.23 | 6.33 | 6.42 | 6.52 | 6.62 | 6.72 | 6.81 | 6.91 | 7.01 | 7.11 | 7.2 | 7.3 | 7.4
34.49 | 3.5 | 3.6 | 3.69 | 3.79 | 3.89 | 3.99 | 4.08 | 4.18 | 4.28 | 4.38 | 4.47 | 4.57 | 4.67 | 4.77 | 4.86 | 4.96 | 5.06 | 5.16 | 5.25 | 5.35 | 5.45 | 5.55 | 5.64 | 5.74 | 5.84 | 5.94 | 6.03 | 6.13 | 6.23 | 6.33 | 6.42 | 6.52 | 6.62 | 6.72 | 6.81 | 6.91 | 7.01 | 7.11 | 7.2 | 7.3 | 7.4
34.5 | 3.5 | 3.6 | 3.69 | 3.79 | 3.89 | 3.99 | 4.08 | 4.18 | 4.28 | 4.38 | 4.47 | 4.57 | 4.67 | 4.77 | 4.86 | 4.96 | 5.06 | 5.16 | 5.25 | 5.35 | 5.45 | 5.55 | 5.64 | 5.74 | 5.84 | 5.94 | 6.03 | 6.13 | 6.23 | 6.32 | 6.42 | 6.52 | 6.62 | 6.71 | 6.81 | 6.91 | 7.01 | 7.1 | 7.2 | 7.3 | 7.4
34.51 | 3.5 | 3.6 | 3.69 | 3.79 | 3.89 | 3.99 | 4.08 | 4.18 | 4.28 | 4.38 | 4.47 | 4.57 | 4.67 | 4.77 | 4.86 | 4.96 | 5.06 | 5.15 | 5.25 | 5.35 | 5.45 | 5.54 | 5.64 | 5.74 | 5.84 | 5.93 | 6.03 | 6.13 | 6.23 | 6.32 | 6.42 | 6.52 | 6.62 | 6.71 | 6.81 | 6.91 | 7.01 | 7.1 | 7.2 | 7.3 | 7.4
34.52 | 3.5 | 3.6 | 3.69 | 3.79 | 3.89 | 3.99 | 4.08 | 4.18 | 4.28 | 4.37 | 4.47 | 4.57 | 4.67 | 4.76 | 4.86 | 4.96 | 5.06 | 5.15 | 5.25 | 5.35 | 5.45 | 5.54 | 5.64 | 5.74 | 5.84 | 5.93 | 6.03 | 6.13 | 6.23 | 6.32 | 6.42 | 6.52 | 6.61 | 6.71 | 6.81 | 6.91 | 7 | 7.1 | 7.2 | 7.3 | 7.39
34.53 | 3.5 | 3.59 | 3.69 | 3.79 | 3.89 | 3.98 | 4.08 | 4.18 | 4.28 | 4.37 | 4.47 | 4.57 | 4.67 | 4.76 | 4.86 | 4.96 | 5.06 | 5.15 | 5.25 | 5.35 | 5.45 | 5.54 | 5.64 | 5.74 | 5.83 | 5.93 | 6.03 | 6.13 | 6.22 | 6.32 | 6.42 | 6.52 | 6.61 | 6.71 | 6.81 | 6.91 | 7 | 7.1 | 7.2 | 7.3 | 7.39
34.54 | 3.5 | 3.59 | 3.69 | 3.79 | 3.89 | 3.98 | 4.08 | 4.18 | 4.28 | 4.37 | 4.47 | 4.57 | 4.67 | 4.76 | 4.86 | 4.96 | 5.05 | 5.15 | 5.25 | 5.35 | 5.44 | 5.54 | 5.64 | 5.74 | 5.83 | 5.93 | 6.03 | 6.13 | 6.22 | 6.32 | 6.42 | 6.52 | 6.61 | 6.71 | 6.81 | 6.9 | 7 | 7.1 | 7.2 | 7.29 | 7.39
34.55 | 3.5 | 3.59 | 3.69 | 3.79 | 3.89 | 3.98 | 4.08 | 4.18 | 4.27 | 4.37 | 4.47 | 4.57 | 4.66 | 4.76 | 4.86 | 4.96 | 5.05 | 5.15 | 5.25 | 5.35 | 5.44 | 5.54 | 5.64 | 5.74 | 5.83 | 5.93 | 6.03 | 6.12 | 6.22 | 6.32 | 6.42 | 6.51 | 6.61 | 6.71 | 6.81 | 6.9 | 7 | 7.1 | 7.2 | 7.29 | 7.39
34.56 | 3.5 | 3.59 | 3.69 | 3.79 | 3.88 | 3.98 | 4.08 | 4.18 | 4.27 | 4.37 | 4.47 | 4.57 | 4.66 | 4.76 | 4.86 | 4.96 | 5.05 | 5.15 | 5.25 | 5.35 | 5.44 | 5.54 | 5.64 | 5.73 | 5.83 | 5.93 | 6.03 | 6.12 | 6.22 | 6.32 | 6.42 | 6.51 | 6.61 | 6.71 | 6.81 | 6.9 | 7 | 7.1 | 7.19 | 7.29 | 7.39
34.57 | 3.49 | 3.59 | 3.69 | 3.79 | 3.88 | 3.98 | 4.08 | 4.18 | 4.27 | 4.37 | 4.47 | 4.57 | 4.66 | 4.76 | 4.86 | 4.95 | 5.05 | 5.15 | 5.25 | 5.34 | 5.44 | 5.54 | 5.64 | 5.73 | 5.83 | 5.93 | 6.03 | 6.12 | 6.22 | 6.32 | 6.41 | 6.51 | 6.61 | 6.71 | 6.8 | 6.9 | 7 | 7.1 | 7.19 | 7.29 | 7.39
34.58 | 3.49 | 3.59 | 3.69 | 3.79 | 3.88 | 3.98 | 4.08 | 4.18 | 4.27 | 4.37 | 4.47 | 4.56 | 4.66 | 4.76 | 4.86 | 4.95 | 5.05 | 5.15 | 5.25 | 5.34 | 5.44 | 5.54 | 5.64 | 5.73 | 5.83 | 5.93 | 6.02 | 6.12 | 6.22 | 6.32 | 6.41 | 6.51 | 6.61 | 6.71 | 6.8 | 6.9 | 7 | 7.09 | 7.19 | 7.29 | 7.39
34.59 | 3.49 | 3.59 | 3.69 | 3.79 | 3.88 | 3.98 | 4.08 | 4.17 | 4.27 | 4.37 | 4.47 | 4.56 | 4.66 | 4.76 | 4.86 | 4.95 | 5.05 | 5.15 | 5.24 | 5.34 | 5.44 | 5.54 | 5.63 | 5.73 | 5.83 | 5.93 | 6.02 | 6.12 | 6.22 | 6.32 | 6.41 | 6.51 | 6.61 | 6.7 | 6.8 | 6.9 | 7 | 7.09 | 7.19 | 7.29 | 7.39
34.6 | 3.49 | 3.59 | 3.69 | 3.78 | 3.88 | 3.98 | 4.08 | 4.17 | 4.27 | 4.37 | 4.47 | 4.56 | 4.66 | 4.76 | 4.85 | 4.95 | 5.05 | 5.15 | 5.24 | 5.34 | 5.44 | 5.54 | 5.63 | 5.73 | 5.83 | 5.93 | 6.02 | 6.12 | 6.22 | 6.31 | 6.41 | 6.51 | 6.61 | 6.7 | 6.8 | 6.9 | 7 | 7.09 | 7.19 | 7.29 | 7.38
34.61 | 3.49 | 3.59 | 3.69 | 3.78 | 3.88 | 3.98 | 4.08 | 4.17 | 4.27 | 4.37 | 4.46 | 4.56 | 4.66 | 4.76 | 4.85 | 4.95 | 5.05 | 5.15 | 5.24 | 5.34 | 5.44 | 5.53 | 5.63 | 5.73 | 5.83 | 5.92 | 6.02 | 6.12 | 6.22 | 6.31 | 6.41 | 6.51 | 6.6 | 6.7 | 6.8 | 6.9 | 6.99 | 7.09 | 7.19 | 7.29 | 7.38
34.62 | 3.49 | 3.59 | 3.69 | 3.78 | 3.88 | 3.98 | 4.07 | 4.17 | 4.27 | 4.37 | 4.46 | 4.56 | 4.66 | 4.76 | 4.85 | 4.95 | 5.05 | 5.14 | 5.24 | 5.34 | 5.44 | 5.53 | 5.63 | 5.73 | 5.83 | 5.92 | 6.02 | 6.12 | 6.21 | 6.31 | 6.41 | 6.51 | 6.6 | 6.7 | 6.8 | 6.9 | 6.99 | 7.09 | 7.19 | 7.28 | 7.38
34.63 | 3.49 | 3.59 | 3.68 | 3.78 | 3.88 | 3.98 | 4.07 | 4.17 | 4.27 | 4.37 | 4.46 | 4.56 | 4.66 | 4.75 | 4.85 | 4.95 | 5.05 | 5.14 | 5.24 | 5.34 | 5.44 | 5.53 | 5.63 | 5.73 | 5.82 | 5.92 | 6.02 | 6.12 | 6.21 | 6.31 | 6.41 | 6.51 | 6.6 | 6.7 | 6.8 | 6.89 | 6.99 | 7.09 | 7.19 | 7.28 | 7.38
34.64 | 3.49 | 3.59 | 3.68 | 3.78 | 3.88 | 3.98 | 4.07 | 4.17 | 4.27 | 4.37 | 4.46 | 4.56 | 4.66 | 4.75 | 4.85 | 4.95 | 5.05 | 5.14 | 5.24 | 5.34 | 5.43 | 5.53 | 5.63 | 5.73 | 5.82 | 5.92 | 6.02 | 6.12 | 6.21 | 6.31 | 6.41 | 6.5 | 6.6 | 6.7 | 6.8 | 6.89 | 6.99 | 7.09 | 7.18 | 7.28 | 7.38
34.65 | 3.49 | 3.59 | 3.68 | 3.78 | 3.88 | 3.98 | 4.07 | 4.17 | 4.27 | 4.36 | 4.46 | 4.56 | 4.66 | 4.75 | 4.85 | 4.95 | 5.04 | 5.14 | 5.24 | 5.34 | 5.43 | 5.53 | 5.63 | 5.73 | 5.82 | 5.92 | 6.02 | 6.11 | 6.21 | 6.31 | 6.41 | 6.5 | 6.6 | 6.7 | 6.79 | 6.89 | 6.99 | 7.09 | 7.18 | 7.28 | 7.38
34.66 | 3.49 | 3.59 | 3.68 | 3.78 | 3.88 | 3.97 | 4.07 | 4.17 | 4.27 | 4.36 | 4.46 | 4.56 | 4.66 | 4.75 | 4.85 | 4.95 | 5.04 | 5.14 | 5.24 | 5.34 | 5.43 | 5.53 | 5.63 | 5.72 | 5.82 | 5.92 | 6.02 | 6.11 | 6.21 | 6.31 | 6.41 | 6.5 | 6.6 | 6.7 | 6.79 | 6.89 | 6.99 | 7.09 | 7.18 | 7.28 | 7.38
34.67 | 3.49 | 3.58 | 3.68 | 3.78 | 3.88 | 3.97 | 4.07 | 4.17 | 4.27 | 4.36 | 4.46 | 4.56 | 4.65 | 4.75 | 4.85 | 4.95 | 5.04 | 5.14 | 5.24 | 5.33 | 5.43 | 5.53 | 5.63 | 5.72 | 5.82 | 5.92 | 6.02 | 6.11 | 6.21 | 6.31 | 6.4 | 6.5 | 6.6 | 6.7 | 6.79 | 6.89 | 6.99 | 7.08 | 7.18 | 7.28 | 7.38
34.68 | 3.49 | 3.58 | 3.68 | 3.78 | 3.88 | 3.97 | 4.07 | 4.17 | 4.26 | 4.36 | 4.46 | 4.56 | 4.65 | 4.75 | 4.85 | 4.95 | 5.04 | 5.14 | 5.24 | 5.33 | 5.43 | 5.53 | 5.63 | 5.72 | 5.82 | 5.92 | 6.01 | 6.11 | 6.21 | 6.31 | 6.4 | 6.5 | 6.6 | 6.69 | 6.79 | 6.89 | 6.99 | 7.08 | 7.18 | 7.28 | 7.37
34.69 | 3.49 | 3.58 | 3.68 | 3.78 | 3.88 | 3.97 | 4.07 | 4.17 | 4.26 | 4.36 | 4.46 | 4.56 | 4.65 | 4.75 | 4.85 | 4.94 | 5.04 | 5.14 | 5.24 | 5.33 | 5.43 | 5.53 | 5.62 | 5.72 | 5.82 | 5.92 | 6.01 | 6.11 | 6.21 | 6.3 | 6.4 | 6.5 | 6.6 | 6.69 | 6.79 | 6.89 | 6.98 | 7.08 | 7.18 | 7.28 | 7.37
34.7 | 3.49 | 3.58 | 3.68 | 3.78 | 3.87 | 3.97 | 4.07 | 4.17 | 4.26 | 4.36 | 4.46 | 4.55 | 4.65 | 4.75 | 4.85 | 4.94 | 5.04 | 5.14 | 5.23 | 5.33 | 5.43 | 5.53 | 5.62 | 5.72 | 5.82 | 5.91 | 6.01 | 6.11 | 6.21 | 6.3 | 6.4 | 6.5 | 6.6 | 6.69 | 6.79 | 6.89 | 6.98 | 7.08 | 7.18 | 7.28 | 7.37
34.71 | 3.48 | 3.58 | 3.68 | 3.78 | 3.87 | 3.97 | 4.07 | 4.17 | 4.26 | 4.36 | 4.46 | 4.55 | 4.65 | 4.75 | 4.85 | 4.94 | 5.04 | 5.14 | 5.23 | 5.33 | 5.43 | 5.53 | 5.62 | 5.72 | 5.82 | 5.91 | 6.01 | 6.11 | 6.21 | 6.3 | 6.4 | 6.5 | 6.59 | 6.69 | 6.79 | 6.89 | 6.98 | 7.08 | 7.18 | 7.27 | 7.37
34.72 | 3.48 | 3.58 | 3.68 | 3.78 | 3.87 | 3.97 | 4.07 | 4.16 | 4.26 | 4.36 | 4.46 | 4.55 | 4.65 | 4.75 | 4.84 | 4.94 | 5.04 | 5.14 | 5.23 | 5.33 | 5.43 | 5.52 | 5.62 | 5.72 | 5.82 | 5.91 | 6.01 | 6.11 | 6.2 | 6.3 | 6.4 | 6.5 | 6.59 | 6.69 | 6.79 | 6.88 | 6.98 | 7.08 | 7.18 | 7.27 | 7.37
34.73 | 3.48 | 3.58 | 3.68 | 3.77 | 3.87 | 3.97 | 4.07 | 4.16 | 4.26 | 4.36 | 4.45 | 4.55 | 4.65 | 4.75 | 4.84 | 4.94 | 5.04 | 5.13 | 5.23 | 5.33 | 5.43 | 5.52 | 5.62 | 5.72 | 5.81 | 5.91 | 6.01 | 6.11 | 6.2 | 6.3 | 6.4 | 6.49 | 6.59 | 6.69 | 6.79 | 6.88 | 6.98 | 7.08 | 7.17 | 7.27 | 7.37
34.74 | 3.48 | 3.58 | 3.68 | 3.77 | 3.87 | 3.97 | 4.07 | 4.16 | 4.26 | 4.36 | 4.45 | 4.55 | 4.65 | 4.75 | 4.84 | 4.94 | 5.04 | 5.13 | 5.23 | 5.33 | 5.43 | 5.52 | 5.62 | 5.72 | 5.81 | 5.91 | 6.01 | 6.11 | 6.2 | 6.3 | 6.4 | 6.49 | 6.59 | 6.69 | 6.78 | 6.88 | 6.98 | 7.08 | 7.17 | 7.27 | 7.37
34.75 | 3.48 | 3.58 | 3.68 | 3.77 | 3.87 | 3.97 | 4.06 | 4.16 | 4.26 | 4.36 | 4.45 | 4.55 | 4.65 | 4.74 | 4.84 | 4.94 | 5.04 | 5.13 | 5.23 | 5.33 | 5.42 | 5.52 | 5.62 | 5.72 | 5.81 | 5.91 | 6.01 | 6.1 | 6.2 | 6.3 | 6.4 | 6.49 | 6.59 | 6.69 | 6.78 | 6.88 | 6.98 | 7.07 | 7.17 | 7.27 | 7.37
34.76 | 3.48 | 3.58 | 3.68 | 3.77 | 3.87 | 3.97 | 4.06 | 4.16 | 4.26 | 4.36 | 4.45 | 4.55 | 4.65 | 4.74 | 4.84 | 4.94 | 5.04 | 5.13 | 5.23 | 5.33 | 5.42 | 5.52 | 5.62 | 5.71 | 5.81 | 5.91 | 6.01 | 6.1 | 6.2 | 6.3 | 6.39 | 6.49 | 6.59 | 6.69 | 6.78 | 6.88 | 6.98 | 7.07 | 7.17 | 7.27 | 7.37
34.77 | 3.48 | 3.58 | 3.67 | 3.77 | 3.87 | 3.97 | 4.06 | 4.16 | 4.26 | 4.35 | 4.45 | 4.55 | 4.65 | 4.74 | 4.84 | 4.94 | 5.03 | 5.13 | 5.23 | 5.33 | 5.42 | 5.52 | 5.62 | 5.71 | 5.81 | 5.91 | 6 | 6.1 | 6.2 | 6.3 | 6.39 | 6.49 | 6.59 | 6.68 | 6.78 | 6.88 | 6.98 | 7.07 | 7.17 | 7.27 | 7.36
34.78 | 3.48 | 3.58 | 3.67 | 3.77 | 3.87 | 3.97 | 4.06 | 4.16 | 4.26 | 4.35 | 4.45 | 4.55 | 4.65 | 4.74 | 4.84 | 4.94 | 5.03 | 5.13 | 5.23 | 5.32 | 5.42 | 5.52 | 5.62 | 5.71 | 5.81 | 5.91 | 6 | 6.1 | 6.2 | 6.3 | 6.39 | 6.49 | 6.59 | 6.68 | 6.78 | 6.88 | 6.97 | 7.07 | 7.17 | 7.27 | 7.36
34.79 | 3.48 | 3.58 | 3.67 | 3.77 | 3.87 | 3.96 | 4.06 | 4.16 | 4.26 | 4.35 | 4.45 | 4.55 | 4.64 | 4.74 | 4.84 | 4.94 | 5.03 | 5.13 | 5.23 | 5.32 | 5.42 | 5.52 | 5.61 | 5.71 | 5.81 | 5.91 | 6 | 6.1 | 6.2 | 6.29 | 6.39 | 6.49 | 6.59 | 6.68 | 6.78 | 6.88 | 6.97 | 7.07 | 7.17 | 7.26 | 7.36
34.8 | 3.48 | 3.58 | 3.67 | 3.77 | 3.87 | 3.96 | 4.06 | 4.16 | 4.26 | 4.35 | 4.45 | 4.55 | 4.64 | 4.74 | 4.84 | 4.93 | 5.03 | 5.13 | 5.23 | 5.32 | 5.42 | 5.52 | 5.61 | 5.71 | 5.81 | 5.9 | 6 | 6.1 | 6.2 | 6.29 | 6.39 | 6.49 | 6.58 | 6.68 | 6.78 | 6.88 | 6.97 | 7.07 | 7.17 | 7.26 | 7.36
34.81 | 3.48 | 3.58 | 3.67 | 3.77 | 3.87 | 3.96 | 4.06 | 4.16 | 4.25 | 4.35 | 4.45 | 4.55 | 4.64 | 4.74 | 4.84 | 4.93 | 5.03 | 5.13 | 5.22 | 5.32 | 5.42 | 5.52 | 5.61 | 5.71 | 5.81 | 5.9 | 6 | 6.1 | 6.19 | 6.29 | 6.39 | 6.49 | 6.58 | 6.68 | 6.78 | 6.87 | 6.97 | 7.07 | 7.17 | 7.26 | 7.36
34.82 | 3.48 | 3.57 | 3.67 | 3.77 | 3.87 | 3.96 | 4.06 | 4.16 | 4.25 | 4.35 | 4.45 | 4.54 | 4.64 | 4.74 | 4.84 | 4.93 | 5.03 | 5.13 | 5.22 | 5.32 | 5.42 | 5.51 | 5.61 | 5.71 | 5.81 | 5.9 | 6 | 6.1 | 6.19 | 6.29 | 6.39 | 6.48 | 6.58 | 6.68 | 6.78 | 6.87 | 6.97 | 7.07 | 7.16 | 7.26 | 7.36
34.83 | 3.48 | 3.57 | 3.67 | 3.77 | 3.86 | 3.96 | 4.06 | 4.16 | 4.25 | 4.35 | 4.45 | 4.54 | 4.64 | 4.74 | 4.83 | 4.93 | 5.03 | 5.13 | 5.22 | 5.32 | 5.42 | 5.51 | 5.61 | 5.71 | 5.8 | 5.9 | 6 | 6.1 | 6.19 | 6.29 | 6.39 | 6.48 | 6.58 | 6.68 | 6.77 | 6.87 | 6.97 | 7.07 | 7.16 | 7.26 | 7.36
34.84 | 3.48 | 3.57 | 3.67 | 3.77 | 3.86 | 3.96 | 4.06 | 4.16 | 4.25 | 4.35 | 4.45 | 4.54 | 4.64 | 4.74 | 4.83 | 4.93 | 5.03 | 5.13 | 5.22 | 5.32 | 5.42 | 5.51 | 5.61 | 5.71 | 5.8 | 5.9 | 6 | 6.09 | 6.19 | 6.29 | 6.39 | 6.48 | 6.58 | 6.68 | 6.77 | 6.87 | 6.97 | 7.06 | 7.16 | 7.26 | 7.36
34.85 | 3.48 | 3.57 | 3.67 | 3.77 | 3.86 | 3.96 | 4.06 | 4.15 | 4.25 | 4.35 | 4.45 | 4.54 | 4.64 | 4.74 | 4.83 | 4.93 | 5.03 | 5.12 | 5.22 | 5.32 | 5.42 | 5.51 | 5.61 | 5.71 | 5.8 | 5.9 | 6 | 6.09 | 6.19 | 6.29 | 6.38 | 6.48 | 6.58 | 6.68 | 6.77 | 6.87 | 6.97 | 7.06 | 7.16 | 7.26 | 7.35
34.86 | 3.47 | 3.57 | 3.67 | 3.77 | 3.86 | 3.96 | 4.06 | 4.15 | 4.25 | 4.35 | 4.44 | 4.54 | 4.64 | 4.74 | 4.83 | 4.93 | 5.03 | 5.12 | 5.22 | 5.32 | 5.41 | 5.51 | 5.61 | 5.7 | 5.8 | 5.9 | 6 | 6.09 | 6.19 | 6.29 | 6.38 | 6.48 | 6.58 | 6.67 | 6.77 | 6.87 | 6.97 | 7.06 | 7.16 | 7.26 | 7.35
34.87 | 3.47 | 3.57 | 3.67 | 3.76 | 3.86 | 3.96 | 4.06 | 4.15 | 4.25 | 4.35 | 4.44 | 4.54 | 4.64 | 4.73 | 4.83 | 4.93 | 5.03 | 5.12 | 5.22 | 5.32 | 5.41 | 5.51 | 5.61 | 5.7 | 5.8 | 5.9 | 5.99 | 6.09 | 6.19 | 6.29 | 6.38 | 6.48 | 6.58 | 6.67 | 6.77 | 6.87 | 6.96 | 7.06 | 7.16 | 7.25 | 7.35
34.88 | 3.47 | 3.57 | 3.67 | 3.76 | 3.86 | 3.96 | 4.05 | 4.15 | 4.25 | 4.35 | 4.44 | 4.54 | 4.64 | 4.73 | 4.83 | 4.93 | 5.02 | 5.12 | 5.22 | 5.32 | 5.41 | 5.51 | 5.61 | 5.7 | 5.8 | 5.9 | 5.99 | 6.09 | 6.19 | 6.28 | 6.38 | 6.48 | 6.58 | 6.67 | 6.77 | 6.87 | 6.96 | 7.06 | 7.16 | 7.25 | 7.35
34.89 | 3.47 | 3.57 | 3.67 | 3.76 | 3.86 | 3.96 | 4.05 | 4.15 | 4.25 | 4.35 | 4.44 | 4.54 | 4.64 | 4.73 | 4.83 | 4.93 | 5.02 | 5.12 | 5.22 | 5.31 | 5.41 | 5.51 | 5.61 | 5.7 | 5.8 | 5.9 | 5.99 | 6.09 | 6.19 | 6.28 | 6.38 | 6.48 | 6.57 | 6.67 | 6.77 | 6.86 | 6.96 | 7.06 | 7.16 | 7.25 | 7.35
34.9 | 3.47 | 3.57 | 3.67 | 3.76 | 3.86 | 3.96 | 4.05 | 4.15 | 4.25 | 4.34 | 4.44 | 4.54 | 4.64 | 4.73 | 4.83 | 4.93 | 5.02 | 5.12 | 5.22 | 5.31 | 5.41 | 5.51 | 5.6 | 5.7 | 5.8 | 5.89 | 5.99 | 6.09 | 6.19 | 6.28 | 6.38 | 6.48 | 6.57 | 6.67 | 6.77 | 6.86 | 6.96 | 7.06 | 7.15 | 7.25 | 7.35
34.91 | 3.47 | 3.57 | 3.67 | 3.76 | 3.86 | 3.96 | 4.05 | 4.15 | 4.25 | 4.34 | 4.44 | 4.54 | 4.63 | 4.73 | 4.83 | 4.92 | 5.02 | 5.12 | 5.22 | 5.31 | 5.41 | 5.51 | 5.6 | 5.7 | 5.8 | 5.89 | 5.99 | 6.09 | 6.18 | 6.28 | 6.38 | 6.48 | 6.57 | 6.67 | 6.77 | 6.86 | 6.96 | 7.06 | 7.15 | 7.25 | 7.35
34.92 | 3.47 | 3.57 | 3.66 | 3.76 | 3.86 | 3.96 | 4.05 | 4.15 | 4.25 | 4.34 | 4.44 | 4.54 | 4.63 | 4.73 | 4.83 | 4.92 | 5.02 | 5.12 | 5.21 | 5.31 | 5.41 | 5.51 | 5.6 | 5.7 | 5.8 | 5.89 | 5.99 | 6.09 | 6.18 | 6.28 | 6.38 | 6.47 | 6.57 | 6.67 | 6.76 | 6.86 | 6.96 | 7.06 | 7.15 | 7.25 | 7.35
34.93 | 3.47 | 3.57 | 3.66 | 3.76 | 3.86 | 3.95 | 4.05 | 4.15 | 4.24 | 4.34 | 4.44 | 4.54 | 4.63 | 4.73 | 4.83 | 4.92 | 5.02 | 5.12 | 5.21 | 5.31 | 5.41 | 5.5 | 5.6 | 5.7 | 5.8 | 5.89 | 5.99 | 6.09 | 6.18 | 6.28 | 6.38 | 6.47 | 6.57 | 6.67 | 6.76 | 6.86 | 6.96 | 7.05 | 7.15 | 7.25 | 7.34
34.94 | 3.47 | 3.57 | 3.66 | 3.76 | 3.86 | 3.95 | 4.05 | 4.15 | 4.24 | 4.34 | 4.44 | 4.53 | 4.63 | 4.73 | 4.83 | 4.92 | 5.02 | 5.12 | 5.21 | 5.31 | 5.41 | 5.5 | 5.6 | 5.7 | 5.79 | 5.89 | 5.99 | 6.08 | 6.18 | 6.28 | 6.38 | 6.47 | 6.57 | 6.67 | 6.76 | 6.86 | 6.96 | 7.05 | 7.15 | 7.25 | 7.34
34.95 | 3.47 | 3.57 | 3.66 | 3.76 | 3.86 | 3.95 | 4.05 | 4.15 | 4.24 | 4.34 | 4.44 | 4.53 | 4.63 | 4.73 | 4.82 | 4.92 | 5.02 | 5.12 | 5.21 | 5.31 | 5.41 | 5.5 | 5.6 | 5.7 | 5.79 | 5.89 | 5.99 | 6.08 | 6.18 | 6.28 | 6.37 | 6.47 | 6.57 | 6.66 | 6.76 | 6.86 | 6.96 | 7.05 | 7.15 | 7.25 | 7.34
34.96 | 3.47 | 3.56 | 3.66 | 3.76 | 3.86 | 3.95 | 4.05 | 4.15 | 4.24 | 4.34 | 4.44 | 4.53 | 4.63 | 4.73 | 4.82 | 4.92 | 5.02 | 5.11 | 5.21 | 5.31 | 5.4 | 5.5 | 5.6 | 5.7 | 5.79 | 5.89 | 5.99 | 6.08 | 6.18 | 6.28 | 6.37 | 6.47 | 6.57 | 6.66 | 6.76 | 6.86 | 6.95 | 7.05 | 7.15 | 7.24 | 7.34
34.97 | 3.47 | 3.56 | 3.66 | 3.76 | 3.85 | 3.95 | 4.05 | 4.14 | 4.24 | 4.34 | 4.44 | 4.53 | 4.63 | 4.73 | 4.82 | 4.92 | 5.02 | 5.11 | 5.21 | 5.31 | 5.4 | 5.5 | 5.6 | 5.69 | 5.79 | 5.89 | 5.98 | 6.08 | 6.18 | 6.28 | 6.37 | 6.47 | 6.57 | 6.66 | 6.76 | 6.86 | 6.95 | 7.05 | 7.15 | 7.24 | 7.34
34.98 | 3.47 | 3.56 | 3.66 | 3.76 | 3.85 | 3.95 | 4.05 | 4.14 | 4.24 | 4.34 | 4.43 | 4.53 | 4.63 | 4.73 | 4.82 | 4.92 | 5.02 | 5.11 | 5.21 | 5.31 | 5.4 | 5.5 | 5.6 | 5.69 | 5.79 | 5.89 | 5.98 | 6.08 | 6.18 | 6.27 | 6.37 | 6.47 | 6.56 | 6.66 | 6.76 | 6.85 | 6.95 | 7.05 | 7.15 | 7.24 | 7.34
34.99 | 3.47 | 3.56 | 3.66 | 3.76 | 3.85 | 3.95 | 4.05 | 4.14 | 4.24 | 4.34 | 4.43 | 4.53 | 4.63 | 4.72 | 4.82 | 4.92 | 5.01 | 5.11 | 5.21 | 5.31 | 5.4 | 5.5 | 5.6 | 5.69 | 5.79 | 5.89 | 5.98 | 6.08 | 6.18 | 6.27 | 6.37 | 6.47 | 6.56 | 6.66 | 6.76 | 6.85 | 6.95 | 7.05 | 7.14 | 7.24 | 7.34
35 | 3.47 | 3.56 | 3.66 | 3.76 | 3.85 | 3.95 | 4.05 | 4.14 | 4.24 | 4.34 | 4.43 | 4.53 | 4.63 | 4.72 | 4.82 | 4.92 | 5.01 | 5.11 | 5.21 | 5.3 | 5.4 | 5.5 | 5.59 | 5.69 | 5.79 | 5.88 | 5.98 | 6.08 | 6.18 | 6.27 | 6.37 | 6.47 | 6.56 | 6.66 | 6.76 | 6.85 | 6.95 | 7.05 | 7.14 | 7.24 | 7.34
35.01 | 3.46 | 3.56 | 3.66 | 3.75 | 3.85 | 3.95 | 4.05 | 4.14 | 4.24 | 4.34 | 4.43 | 4.53 | 4.63 | 4.72 | 4.82 | 4.92 | 5.01 | 5.11 | 5.21 | 5.3 | 5.4 | 5.5 | 5.59 | 5.69 | 5.79 | 5.88 | 5.98 | 6.08 | 6.17 | 6.27 | 6.37 | 6.46 | 6.56 | 6.66 | 6.75 | 6.85 | 6.95 | 7.04 | 7.14 | 7.24 | 7.34
35.02 | 3.46 | 3.56 | 3.66 | 3.75 | 3.85 | 3.95 | 4.04 | 4.14 | 4.24 | 4.33 | 4.43 | 4.53 | 4.63 | 4.72 | 4.82 | 4.92 | 5.01 | 5.11 | 5.21 | 5.3 | 5.4 | 5.5 | 5.59 | 5.69 | 5.79 | 5.88 | 5.98 | 6.08 | 6.17 | 6.27 | 6.37 | 6.46 | 6.56 | 6.66 | 6.75 | 6.85 | 6.95 | 7.04 | 7.14 | 7.24 | 7.33
35.03 | 3.46 | 3.56 | 3.66 | 3.75 | 3.85 | 3.95 | 4.04 | 4.14 | 4.24 | 4.33 | 4.43 | 4.53 | 4.62 | 4.72 | 4.82 | 4.91 | 5.01 | 5.11 | 5.2 | 5.3 | 5.4 | 5.49 | 5.59 | 5.69 | 5.79 | 5.88 | 5.98 | 6.08 | 6.17 | 6.27 | 6.37 | 6.46 | 6.56 | 6.66 | 6.75 | 6.85 | 6.95 | 7.04 | 7.14 | 7.24 | 7.33
35.04 | 3.46 | 3.56 | 3.66 | 3.75 | 3.85 | 3.95 | 4.04 | 4.14 | 4.24 | 4.33 | 4.43 | 4.53 | 4.62 | 4.72 | 4.82 | 4.91 | 5.01 | 5.11 | 5.2 | 5.3 | 5.4 | 5.49 | 5.59 | 5.69 | 5.78 | 5.88 | 5.98 | 6.07 | 6.17 | 6.27 | 6.36 | 6.46 | 6.56 | 6.65 | 6.75 | 6.85 | 6.94 | 7.04 | 7.14 | 7.23 | 7.33
35.05 | 3.46 | 3.56 | 3.66 | 3.75 | 3.85 | 3.95 | 4.04 | 4.14 | 4.24 | 4.33 | 4.43 | 4.53 | 4.62 | 4.72 | 4.82 | 4.91 | 5.01 | 5.11 | 5.2 | 5.3 | 5.4 | 5.49 | 5.59 | 5.69 | 5.78 | 5.88 | 5.98 | 6.07 | 6.17 | 6.27 | 6.36 | 6.46 | 6.56 | 6.65 | 6.75 | 6.85 | 6.94 | 7.04 | 7.14 | 7.23 | 7.33
35.06 | 3.46 | 3.56 | 3.65 | 3.75 | 3.85 | 3.94 | 4.04 | 4.14 | 4.23 | 4.33 | 4.43 | 4.52 | 4.62 | 4.72 | 4.82 | 4.91 | 5.01 | 5.11 | 5.2 | 5.3 | 5.4 | 5.49 | 5.59 | 5.69 | 5.78 | 5.88 | 5.98 | 6.07 | 6.17 | 6.27 | 6.36 | 6.46 | 6.56 | 6.65 | 6.75 | 6.85 | 6.94 | 7.04 | 7.14 | 7.23 | 7.33
35.07 | 3.46 | 3.56 | 3.65 | 3.75 | 3.85 | 3.94 | 4.04 | 4.14 | 4.23 | 4.33 | 4.43 | 4.52 | 4.62 | 4.72 | 4.81 | 4.91 | 5.01 | 5.1 | 5.2 | 5.3 | 5.39 | 5.49 | 5.59 | 5.68 | 5.78 | 5.88 | 5.97 | 6.07 | 6.17 | 6.26 | 6.36 | 6.46 | 6.55 | 6.65 | 6.75 | 6.84 | 6.94 | 7.04 | 7.13 | 7.23 | 7.33
35.08 | 3.46 | 3.56 | 3.65 | 3.75 | 3.85 | 3.94 | 4.04 | 4.14 | 4.23 | 4.33 | 4.43 | 4.52 | 4.62 | 4.72 | 4.81 | 4.91 | 5.01 | 5.1 | 5.2 | 5.3 | 5.39 | 5.49 | 5.59 | 5.68 | 5.78 | 5.88 | 5.97 | 6.07 | 6.17 | 6.26 | 6.36 | 6.46 | 6.55 | 6.65 | 6.75 | 6.84 | 6.94 | 7.04 | 7.13 | 7.23 | 7.33
35.09 | 3.46 | 3.56 | 3.65 | 3.75 | 3.85 | 3.94 | 4.04 | 4.14 | 4.23 | 4.33 | 4.43 | 4.52 | 4.62 | 4.72 | 4.81 | 4.91 | 5.01 | 5.1 | 5.2 | 5.3 | 5.39 | 5.49 | 5.59 | 5.68 | 5.78 | 5.88 | 5.97 | 6.07 | 6.17 | 6.26 | 6.36 | 6.46 | 6.55 | 6.65 | 6.75 | 6.84 | 6.94 | 7.04 | 7.13 | 7.23 | 7.33
35.1 | 3.46 | 3.55 | 3.65 | 3.75 | 3.84 | 3.94 | 4.04 | 4.13 | 4.23 | 4.33 | 4.42 | 4.52 | 4.62 | 4.72 | 4.81 | 4.91 | 5 | 5.1 | 5.2 | 5.29 | 5.39 | 5.49 | 5.58 | 5.68 | 5.78 | 5.87 | 5.97 | 6.07 | 6.16 | 6.26 | 6.36 | 6.45 | 6.55 | 6.65 | 6.74 | 6.84 | 6.94 | 7.03 | 7.13 | 7.23 | 7.32
35.11 | 3.46 | 3.55 | 3.65 | 3.75 | 3.84 | 3.94 | 4.04 | 4.13 | 4.23 | 4.33 | 4.42 | 4.52 | 4.62 | 4.71 | 4.81 | 4.91 | 5 | 5.1 | 5.2 | 5.29 | 5.39 | 5.49 | 5.58 | 5.68 | 5.78 | 5.87 | 5.97 | 6.07 | 6.16 | 6.26 | 6.36 | 6.45 | 6.55 | 6.65 | 6.74 | 6.84 | 6.94 | 7.03 | 7.13 | 7.23 | 7.32
35.12 | 3.46 | 3.55 | 3.65 | 3.75 | 3.84 | 3.94 | 4.04 | 4.13 | 4.23 | 4.33 | 4.42 | 4.52 | 4.62 | 4.71 | 4.81 | 4.91 | 5 | 5.1 | 5.2 | 5.29 | 5.39 | 5.49 | 5.58 | 5.68 | 5.78 | 5.87 | 5.97 | 6.07 | 6.16 | 6.26 | 6.36 | 6.45 | 6.55 | 6.65 | 6.74 | 6.84 | 6.94 | 7.03 | 7.13 | 7.23 | 7.32
35.13 | 3.46 | 3.55 | 3.65 | 3.75 | 3.84 | 3.94 | 4.04 | 4.13 | 4.23 | 4.33 | 4.42 | 4.52 | 4.62 | 4.71 | 4.81 | 4.91 | 5 | 5.1 | 5.2 | 5.29 | 5.39 | 5.49 | 5.58 | 5.68 | 5.78 | 5.87 | 5.97 | 6.07 | 6.16 | 6.26 | 6.35 | 6.45 | 6.55 | 6.64 | 6.74 | 6.84 | 6.93 | 7.03 | 7.13 | 7.22 | 7.32
35.14 | 3.46 | 3.55 | 3.65 | 3.75 | 3.84 | 3.94 | 4.04 | 4.13 | 4.23 | 4.33 | 4.42 | 4.52 | 4.62 | 4.71 | 4.81 | 4.9 | 5 | 5.1 | 5.19 | 5.29 | 5.39 | 5.48 | 5.58 | 5.68 | 5.77 | 5.87 | 5.97 | 6.06 | 6.16 | 6.26 | 6.35 | 6.45 | 6.55 | 6.64 | 6.74 | 6.84 | 6.93 | 7.03 | 7.13 | 7.22 | 7.32
35.15 | 3.45 | 3.55 | 3.65 | 3.74 | 3.84 | 3.94 | 4.03 | 4.13 | 4.23 | 4.32 | 4.42 | 4.52 | 4.61 | 4.71 | 4.81 | 4.9 | 5 | 5.1 | 5.19 | 5.29 | 5.39 | 5.48 | 5.58 | 5.68 | 5.77 | 5.87 | 5.97 | 6.06 | 6.16 | 6.26 | 6.35 | 6.45 | 6.55 | 6.64 | 6.74 | 6.84 | 6.93 | 7.03 | 7.13 | 7.22 | 7.32
35.16 | 3.45 | 3.55 | 3.65 | 3.74 | 3.84 | 3.94 | 4.03 | 4.13 | 4.23 | 4.32 | 4.42 | 4.52 | 4.61 | 4.71 | 4.81 | 4.9 | 5 | 5.1 | 5.19 | 5.29 | 5.39 | 5.48 | 5.58 | 5.68 | 5.77 | 5.87 | 5.97 | 6.06 | 6.16 | 6.26 | 6.35 | 6.45 | 6.54 | 6.64 | 6.74 | 6.83 | 6.93 | 7.03 | 7.12 | 7.22 | 7.32
35.17 | 3.45 | 3.55 | 3.65 | 3.74 | 3.84 | 3.94 | 4.03 | 4.13 | 4.23 | 4.32 | 4.42 | 4.52 | 4.61 | 4.71 | 4.81 | 4.9 | 5 | 5.1 | 5.19 | 5.29 | 5.39 | 5.48 | 5.58 | 5.67 | 5.77 | 5.87 | 5.96 | 6.06 | 6.16 | 6.25 | 6.35 | 6.45 | 6.54 | 6.64 | 6.74 | 6.83 | 6.93 | 7.03 | 7.12 | 7.22 | 7.32
35.18 | 3.45 | 3.55 | 3.65 | 3.74 | 3.84 | 3.94 | 4.03 | 4.13 | 4.23 | 4.32 | 4.42 | 4.52 | 4.61 | 4.71 | 4.8 | 4.9 | 5 | 5.09 | 5.19 | 5.29 | 5.38 | 5.48 | 5.58 | 5.67 | 5.77 | 5.87 | 5.96 | 6.06 | 6.16 | 6.25 | 6.35 | 6.45 | 6.54 | 6.64 | 6.74 | 6.83 | 6.93 | 7.03 | 7.12 | 7.22 | 7.31
35.19 | 3.45 | 3.55 | 3.65 | 3.74 | 3.84 | 3.93 | 4.03 | 4.13 | 4.22 | 4.32 | 4.42 | 4.51 | 4.61 | 4.71 | 4.8 | 4.9 | 5 | 5.09 | 5.19 | 5.29 | 5.38 | 5.48 | 5.58 | 5.67 | 5.77 | 5.87 | 5.96 | 6.06 | 6.16 | 6.25 | 6.35 | 6.45 | 6.54 | 6.64 | 6.73 | 6.83 | 6.93 | 7.02 | 7.12 | 7.22 | 7.31
35.2 | 3.45 | 3.55 | 3.64 | 3.74 | 3.84 | 3.93 | 4.03 | 4.13 | 4.22 | 4.32 | 4.42 | 4.51 | 4.61 | 4.71 | 4.8 | 4.9 | 5 | 5.09 | 5.19 | 5.29 | 5.38 | 5.48 | 5.58 | 5.67 | 5.77 | 5.86 | 5.96 | 6.06 | 6.15 | 6.25 | 6.35 | 6.44 | 6.54 | 6.64 | 6.73 | 6.83 | 6.93 | 7.02 | 7.12 | 7.22 | 7.31
35.21 | 3.45 | 3.55 | 3.64 | 3.74 | 3.84 | 3.93 | 4.03 | 4.13 | 4.22 | 4.32 | 4.42 | 4.51 | 4.61 | 4.71 | 4.8 | 4.9 | 5 | 5.09 | 5.19 | 5.28 | 5.38 | 5.48 | 5.57 | 5.67 | 5.77 | 5.86 | 5.96 | 6.06 | 6.15 | 6.25 | 6.35 | 6.44 | 6.54 | 6.64 | 6.73 | 6.83 | 6.93 | 7.02 | 7.12 | 7.21 | 7.31
35.22 | 3.45 | 3.55 | 3.64 | 3.74 | 3.84 | 3.93 | 4.03 | 4.13 | 4.22 | 4.32 | 4.42 | 4.51 | 4.61 | 4.7 | 4.8 | 4.9 | 4.99 | 5.09 | 5.19 | 5.28 | 5.38 | 5.48 | 5.57 | 5.67 | 5.77 | 5.86 | 5.96 | 6.06 | 6.15 | 6.25 | 6.35 | 6.44 | 6.54 | 6.63 | 6.73 | 6.83 | 6.92 | 7.02 | 7.12 | 7.21 | 7.31
35.23 | 3.45 | 3.55 | 3.64 | 3.74 | 3.84 | 3.93 | 4.03 | 4.12 | 4.22 | 4.32 | 4.41 | 4.51 | 4.61 | 4.7 | 4.8 | 4.9 | 4.99 | 5.09 | 5.19 | 5.28 | 5.38 | 5.48 | 5.57 | 5.67 | 5.77 | 5.86 | 5.96 | 6.05 | 6.15 | 6.25 | 6.34 | 6.44 | 6.54 | 6.63 | 6.73 | 6.83 | 6.92 | 7.02 | 7.12 | 7.21 | 7.31
35.24 | 3.45 | 3.55 | 3.64 | 3.74 | 3.83 | 3.93 | 4.03 | 4.12 | 4.22 | 4.32 | 4.41 | 4.51 | 4.61 | 4.7 | 4.8 | 4.9 | 4.99 | 5.09 | 5.19 | 5.28 | 5.38 | 5.48 | 5.57 | 5.67 | 5.76 | 5.86 | 5.96 | 6.05 | 6.15 | 6.25 | 6.34 | 6.44 | 6.54 | 6.63 | 6.73 | 6.83 | 6.92 | 7.02 | 7.11 | 7.21 | 7.31
35.25 | 3.45 | 3.54 | 3.64 | 3.74 | 3.83 | 3.93 | 4.03 | 4.12 | 4.22 | 4.32 | 4.41 | 4.51 | 4.61 | 4.7 | 4.8 | 4.9 | 4.99 | 5.09 | 5.18 | 5.28 | 5.38 | 5.47 | 5.57 | 5.67 | 5.76 | 5.86 | 5.96 | 6.05 | 6.15 | 6.25 | 6.34 | 6.44 | 6.54 | 6.63 | 6.73 | 6.82 | 6.92 | 7.02 | 7.11 | 7.21 | 7.31
35.26 | 3.45 | 3.54 | 3.64 | 3.74 | 3.83 | 3.93 | 4.03 | 4.12 | 4.22 | 4.32 | 4.41 | 4.51 | 4.61 | 4.7 | 4.8 | 4.89 | 4.99 | 5.09 | 5.18 | 5.28 | 5.38 | 5.47 | 5.57 | 5.67 | 5.76 | 5.86 | 5.96 | 6.05 | 6.15 | 6.24 | 6.34 | 6.44 | 6.53 | 6.63 | 6.73 | 6.82 | 6.92 | 7.02 | 7.11 | 7.21 | 7.31
35.27 | 3.45 | 3.54 | 3.64 | 3.74 | 3.83 | 3.93 | 4.03 | 4.12 | 4.22 | 4.31 | 4.41 | 4.51 | 4.6 | 4.7 | 4.8 | 4.89 | 4.99 | 5.09 | 5.18 | 5.28 | 5.38 | 5.47 | 5.57 | 5.67 | 5.76 | 5.86 | 5.95 | 6.05 | 6.15 | 6.24 | 6.34 | 6.44 | 6.53 | 6.63 | 6.73 | 6.82 | 6.92 | 7.02 | 7.11 | 7.21 | 7.3
35.28 | 3.45 | 3.54 | 3.64 | 3.74 | 3.83 | 3.93 | 4.02 | 4.12 | 4.22 | 4.31 | 4.41 | 4.51 | 4.6 | 4.7 | 4.8 | 4.89 | 4.99 | 5.09 | 5.18 | 5.28 | 5.37 | 5.47 | 5.57 | 5.66 | 5.76 | 5.86 | 5.95 | 6.05 | 6.15 | 6.24 | 6.34 | 6.44 | 6.53 | 6.63 | 6.72 | 6.82 | 6.92 | 7.01 | 7.11 | 7.21 | 7.3
35.29 | 3.45 | 3.54 | 3.64 | 3.73 | 3.83 | 3.93 | 4.02 | 4.12 | 4.22 | 4.31 | 4.41 | 4.51 | 4.6 | 4.7 | 4.8 | 4.89 | 4.99 | 5.08 | 5.18 | 5.28 | 5.37 | 5.47 | 5.57 | 5.66 | 5.76 | 5.86 | 5.95 | 6.05 | 6.15 | 6.24 | 6.34 | 6.43 | 6.53 | 6.63 | 6.72 | 6.82 | 6.92 | 7.01 | 7.11 | 7.21 | 7.3
35.3 | 3.44 | 3.54 | 3.64 | 3.73 | 3.83 | 3.93 | 4.02 | 4.12 | 4.22 | 4.31 | 4.41 | 4.51 | 4.6 | 4.7 | 4.79 | 4.89 | 4.99 | 5.08 | 5.18 | 5.28 | 5.37 | 5.47 | 5.57 | 5.66 | 5.76 | 5.85 | 5.95 | 6.05 | 6.14 | 6.24 | 6.34 | 6.43 | 6.53 | 6.63 | 6.72 | 6.82 | 6.92 | 7.01 | 7.11 | 7.2 | 7.3
35.31 | 3.44 | 3.54 | 3.64 | 3.73 | 3.83 | 3.93 | 4.02 | 4.12 | 4.22 | 4.31 | 4.41 | 4.5 | 4.6 | 4.7 | 4.79 | 4.89 | 4.99 | 5.08 | 5.18 | 5.28 | 5.37 | 5.47 | 5.56 | 5.66 | 5.76 | 5.85 | 5.95 | 6.05 | 6.14 | 6.24 | 6.34 | 6.43 | 6.53 | 6.63 | 6.72 | 6.82 | 6.91 | 7.01 | 7.11 | 7.2 | 7.3
35.32 | 3.44 | 3.54 | 3.64 | 3.73 | 3.83 | 3.93 | 4.02 | 4.12 | 4.21 | 4.31 | 4.41 | 4.5 | 4.6 | 4.7 | 4.79 | 4.89 | 4.99 | 5.08 | 5.18 | 5.27 | 5.37 | 5.47 | 5.56 | 5.66 | 5.76 | 5.85 | 5.95 | 6.05 | 6.14 | 6.24 | 6.33 | 6.43 | 6.53 | 6.62 | 6.72 | 6.82 | 6.91 | 7.01 | 7.11 | 7.2 | 7.3
35.33 | 3.44 | 3.54 | 3.64 | 3.73 | 3.83 | 3.92 | 4.02 | 4.12 | 4.21 | 4.31 | 4.41 | 4.5 | 4.6 | 4.7 | 4.79 | 4.89 | 4.98 | 5.08 | 5.18 | 5.27 | 5.37 | 5.47 | 5.56 | 5.66 | 5.76 | 5.85 | 5.95 | 6.04 | 6.14 | 6.24 | 6.33 | 6.43 | 6.53 | 6.62 | 6.72 | 6.82 | 6.91 | 7.01 | 7.1 | 7.2 | 7.3
35.34 | 3.44 | 3.54 | 3.63 | 3.73 | 3.83 | 3.92 | 4.02 | 4.12 | 4.21 | 4.31 | 4.41 | 4.5 | 4.6 | 4.69 | 4.79 | 4.89 | 4.98 | 5.08 | 5.18 | 5.27 | 5.37 | 5.47 | 5.56 | 5.66 | 5.75 | 5.85 | 5.95 | 6.04 | 6.14 | 6.24 | 6.33 | 6.43 | 6.53 | 6.62 | 6.72 | 6.81 | 6.91 | 7.01 | 7.1 | 7.2 | 7.3
35.35 | 3.44 | 3.54 | 3.63 | 3.73 | 3.83 | 3.92 | 4.02 | 4.12 | 4.21 | 4.31 | 4.4 | 4.5 | 4.6 | 4.69 | 4.79 | 4.89 | 4.98 | 5.08 | 5.18 | 5.27 | 5.37 | 5.46 | 5.56 | 5.66 | 5.75 | 5.85 | 5.95 | 6.04 | 6.14 | 6.24 | 6.33 | 6.43 | 6.52 | 6.62 | 6.72 | 6.81 | 6.91 | 7.01 | 7.1 | 7.2 | 7.29
35.36 | 3.44 | 3.54 | 3.63 | 3.73 | 3.83 | 3.92 | 4.02 | 4.12 | 4.21 | 4.31 | 4.4 | 4.5 | 4.6 | 4.69 | 4.79 | 4.89 | 4.98 | 5.08 | 5.17 | 5.27 | 5.37 | 5.46 | 5.56 | 5.66 | 5.75 | 5.85 | 5.95 | 6.04 | 6.14 | 6.23 | 6.33 | 6.43 | 6.52 | 6.62 | 6.72 | 6.81 | 6.91 | 7 | 7.1 | 7.2 | 7.29
35.37 | 3.44 | 3.54 | 3.63 | 3.73 | 3.83 | 3.92 | 4.02 | 4.11 | 4.21 | 4.31 | 4.4 | 4.5 | 4.6 | 4.69 | 4.79 | 4.88 | 4.98 | 5.08 | 5.17 | 5.27 | 5.37 | 5.46 | 5.56 | 5.66 | 5.75 | 5.85 | 5.94 | 6.04 | 6.14 | 6.23 | 6.33 | 6.43 | 6.52 | 6.62 | 6.71 | 6.81 | 6.91 | 7 | 7.1 | 7.2 | 7.29
35.38 | 3.44 | 3.54 | 3.63 | 3.73 | 3.82 | 3.92 | 4.02 | 4.11 | 4.21 | 4.31 | 4.4 | 4.5 | 4.6 | 4.69 | 4.79 | 4.88 | 4.98 | 5.08 | 5.17 | 5.27 | 5.37 | 5.46 | 5.56 | 5.65 | 5.75 | 5.85 | 5.94 | 6.04 | 6.14 | 6.23 | 6.33 | 6.42 | 6.52 | 6.62 | 6.71 | 6.81 | 6.91 | 7 | 7.1 | 7.2 | 7.29
35.39 | 3.44 | 3.53 | 3.63 | 3.73 | 3.82 | 3.92 | 4.02 | 4.11 | 4.21 | 4.31 | 4.4 | 4.5 | 4.59 | 4.69 | 4.79 | 4.88 | 4.98 | 5.08 | 5.17 | 5.27 | 5.36 | 5.46 | 5.56 | 5.65 | 5.75 | 5.85 | 5.94 | 6.04 | 6.13 | 6.23 | 6.33 | 6.42 | 6.52 | 6.62 | 6.71 | 6.81 | 6.91 | 7 | 7.1 | 7.19 | 7.29
35.4 | 3.44 | 3.53 | 3.63 | 3.73 | 3.82 | 3.92 | 4.02 | 4.11 | 4.21 | 4.3 | 4.4 | 4.5 | 4.59 | 4.69 | 4.79 | 4.88 | 4.98 | 5.07 | 5.17 | 5.27 | 5.36 | 5.46 | 5.56 | 5.65 | 5.75 | 5.85 | 5.94 | 6.04 | 6.13 | 6.23 | 6.33 | 6.42 | 6.52 | 6.62 | 6.71 | 6.81 | 6.9 | 7 | 7.1 | 7.19 | 7.29
35.41 | 3.44 | 3.53 | 3.63 | 3.73 | 3.82 | 3.92 | 4.01 | 4.11 | 4.21 | 4.3 | 4.4 | 4.5 | 4.59 | 4.69 | 4.79 | 4.88 | 4.98 | 5.07 | 5.17 | 5.27 | 5.36 | 5.46 | 5.56 | 5.65 | 5.75 | 5.84 | 5.94 | 6.04 | 6.13 | 6.23 | 6.33 | 6.42 | 6.52 | 6.61 | 6.71 | 6.81 | 6.9 | 7 | 7.1 | 7.19 | 7.29
35.42 | 3.44 | 3.53 | 3.63 | 3.73 | 3.82 | 3.92 | 4.01 | 4.11 | 4.21 | 4.3 | 4.4 | 4.5 | 4.59 | 4.69 | 4.78 | 4.88 | 4.98 | 5.07 | 5.17 | 5.27 | 5.36 | 5.46 | 5.55 | 5.65 | 5.75 | 5.84 | 5.94 | 6.04 | 6.13 | 6.23 | 6.32 | 6.42 | 6.52 | 6.61 | 6.71 | 6.81 | 6.9 | 7 | 7.09 | 7.19 | 7.29
35.43 | 3.44 | 3.53 | 3.63 | 3.72 | 3.82 | 3.92 | 4.01 | 4.11 | 4.21 | 4.3 | 4.4 | 4.49 | 4.59 | 4.69 | 4.78 | 4.88 | 4.98 | 5.07 | 5.17 | 5.26 | 5.36 | 5.46 | 5.55 | 5.65 | 5.75 | 5.84 | 5.94 | 6.03 | 6.13 | 6.23 | 6.32 | 6.42 | 6.52 | 6.61 | 6.71 | 6.8 | 6.9 | 7 | 7.09 | 7.19 | 7.29
35.44 | 3.44 | 3.53 | 3.63 | 3.72 | 3.82 | 3.92 | 4.01 | 4.11 | 4.21 | 4.3 | 4.4 | 4.49 | 4.59 | 4.69 | 4.78 | 4.88 | 4.98 | 5.07 | 5.17 | 5.26 | 5.36 | 5.46 | 5.55 | 5.65 | 5.74 | 5.84 | 5.94 | 6.03 | 6.13 | 6.23 | 6.32 | 6.42 | 6.51 | 6.61 | 6.71 | 6.8 | 6.9 | 7 | 7.09 | 7.19 | 7.28
35.45 | 3.43 | 3.53 | 3.63 | 3.72 | 3.82 | 3.92 | 4.01 | 4.11 | 4.2 | 4.3 | 4.4 | 4.49 | 4.59 | 4.69 | 4.78 | 4.88 | 4.97 | 5.07 | 5.17 | 5.26 | 5.36 | 5.46 | 5.55 | 5.65 | 5.74 | 5.84 | 5.94 | 6.03 | 6.13 | 6.22 | 6.32 | 6.42 | 6.51 | 6.61 | 6.71 | 6.8 | 6.9 | 6.99 | 7.09 | 7.19 | 7.28
35.46 | 3.43 | 3.53 | 3.63 | 3.72 | 3.82 | 3.91 | 4.01 | 4.11 | 4.2 | 4.3 | 4.4 | 4.49 | 4.59 | 4.68 | 4.78 | 4.88 | 4.97 | 5.07 | 5.17 | 5.26 | 5.36 | 5.45 | 5.55 | 5.65 | 5.74 | 5.84 | 5.94 | 6.03 | 6.13 | 6.22 | 6.32 | 6.42 | 6.51 | 6.61 | 6.7 | 6.8 | 6.9 | 6.99 | 7.09 | 7.19 | 7.28
35.47 | 3.43 | 3.53 | 3.63 | 3.72 | 3.82 | 3.91 | 4.01 | 4.11 | 4.2 | 4.3 | 4.4 | 4.49 | 4.59 | 4.68 | 4.78 | 4.88 | 4.97 | 5.07 | 5.16 | 5.26 | 5.36 | 5.45 | 5.55 | 5.65 | 5.74 | 5.84 | 5.93 | 6.03 | 6.13 | 6.22 | 6.32 | 6.42 | 6.51 | 6.61 | 6.7 | 6.8 | 6.9 | 6.99 | 7.09 | 7.18 | 7.28
35.48 | 3.43 | 3.53 | 3.62 | 3.72 | 3.82 | 3.91 | 4.01 | 4.11 | 4.2 | 4.3 | 4.39 | 4.49 | 4.59 | 4.68 | 4.78 | 4.88 | 4.97 | 5.07 | 5.16 | 5.26 | 5.36 | 5.45 | 5.55 | 5.64 | 5.74 | 5.84 | 5.93 | 6.03 | 6.13 | 6.22 | 6.32 | 6.41 | 6.51 | 6.61 | 6.7 | 6.8 | 6.89 | 6.99 | 7.09 | 7.18 | 7.28
35.49 | 3.43 | 3.53 | 3.62 | 3.72 | 3.82 | 3.91 | 4.01 | 4.11 | 4.2 | 4.3 | 4.39 | 4.49 | 4.59 | 4.68 | 4.78 | 4.87 | 4.97 | 5.07 | 5.16 | 5.26 | 5.36 | 5.45 | 5.55 | 5.64 | 5.74 | 5.84 | 5.93 | 6.03 | 6.12 | 6.22 | 6.32 | 6.41 | 6.51 | 6.61 | 6.7 | 6.8 | 6.89 | 6.99 | 7.09 | 7.18 | 7.28
35.5 | 3.43 | 3.53 | 3.62 | 3.72 | 3.82 | 3.91 | 4.01 | 4.1 | 4.2 | 4.3 | 4.39 | 4.49 | 4.59 | 4.68 | 4.78 | 4.87 | 4.97 | 5.07 | 5.16 | 5.26 | 5.35 | 5.45 | 5.55 | 5.64 | 5.74 | 5.84 | 5.93 | 6.03 | 6.12 | 6.22 | 6.32 | 6.41 | 6.51 | 6.6 | 6.7 | 6.8 | 6.89 | 6.99 | 7.08 | 7.18 | 7.28
35.51 | 3.43 | 3.53 | 3.62 | 3.72 | 3.82 | 3.91 | 4.01 | 4.1 | 4.2 | 4.3 | 4.39 | 4.49 | 4.58 | 4.68 | 4.78 | 4.87 | 4.97 | 5.07 | 5.16 | 5.26 | 5.35 | 5.45 | 5.55 | 5.64 | 5.74 | 5.83 | 5.93 | 6.03 | 6.12 | 6.22 | 6.31 | 6.41 | 6.51 | 6.6 | 6.7 | 6.8 | 6.89 | 6.99 | 7.08 | 7.18 | 7.28
35.52 | 3.43 | 3.53 | 3.62 | 3.72 | 3.81 | 3.91 | 4.01 | 4.1 | 4.2 | 4.3 | 4.39 | 4.49 | 4.58 | 4.68 | 4.78 | 4.87 | 4.97 | 5.06 | 5.16 | 5.26 | 5.35 | 5.45 | 5.54 | 5.64 | 5.74 | 5.83 | 5.93 | 6.03 | 6.12 | 6.22 | 6.31 | 6.41 | 6.51 | 6.6 | 6.7 | 6.79 | 6.89 | 6.99 | 7.08 | 7.18 | 7.27
35.53 | 3.43 | 3.53 | 3.62 | 3.72 | 3.81 | 3.91 | 4.01 | 4.1 | 4.2 | 4.29 | 4.39 | 4.49 | 4.58 | 4.68 | 4.77 | 4.87 | 4.97 | 5.06 | 5.16 | 5.26 | 5.35 | 5.45 | 5.54 | 5.64 | 5.74 | 5.83 | 5.93 | 6.02 | 6.12 | 6.22 | 6.31 | 6.41 | 6.5 | 6.6 | 6.7 | 6.79 | 6.89 | 6.99 | 7.08 | 7.18 | 7.27
35.54 | 3.43 | 3.52 | 3.62 | 3.72 | 3.81 | 3.91 | 4.01 | 4.1 | 4.2 | 4.29 | 4.39 | 4.49 | 4.58 | 4.68 | 4.77 | 4.87 | 4.97 | 5.06 | 5.16 | 5.25 | 5.35 | 5.45 | 5.54 | 5.64 | 5.74 | 5.83 | 5.93 | 6.02 | 6.12 | 6.22 | 6.31 | 6.41 | 6.5 | 6.6 | 6.7 | 6.79 | 6.89 | 6.98 | 7.08 | 7.18 | 7.27
35.55 | 3.43 | 3.52 | 3.62 | 3.72 | 3.81 | 3.91 | 4 | 4.1 | 4.2 | 4.29 | 4.39 | 4.48 | 4.58 | 4.68 | 4.77 | 4.87 | 4.97 | 5.06 | 5.16 | 5.25 | 5.35 | 5.45 | 5.54 | 5.64 | 5.73 | 5.83 | 5.93 | 6.02 | 6.12 | 6.21 | 6.31 | 6.41 | 6.5 | 6.6 | 6.69 | 6.79 | 6.89 | 6.98 | 7.08 | 7.18 | 7.27
35.56 | 3.43 | 3.52 | 3.62 | 3.72 | 3.81 | 3.91 | 4 | 4.1 | 4.2 | 4.29 | 4.39 | 4.48 | 4.58 | 4.68 | 4.77 | 4.87 | 4.96 | 5.06 | 5.16 | 5.25 | 5.35 | 5.44 | 5.54 | 5.64 | 5.73 | 5.83 | 5.93 | 6.02 | 6.12 | 6.21 | 6.31 | 6.41 | 6.5 | 6.6 | 6.69 | 6.79 | 6.89 | 6.98 | 7.08 | 7.17 | 7.27
35.57 | 3.43 | 3.52 | 3.62 | 3.71 | 3.81 | 3.91 | 4 | 4.1 | 4.2 | 4.29 | 4.39 | 4.48 | 4.58 | 4.68 | 4.77 | 4.87 | 4.96 | 5.06 | 5.16 | 5.25 | 5.35 | 5.44 | 5.54 | 5.64 | 5.73 | 5.83 | 5.92 | 6.02 | 6.12 | 6.21 | 6.31 | 6.4 | 6.5 | 6.6 | 6.69 | 6.79 | 6.88 | 6.98 | 7.08 | 7.17 | 7.27
35.58 | 3.43 | 3.52 | 3.62 | 3.71 | 3.81 | 3.91 | 4 | 4.1 | 4.19 | 4.29 | 4.39 | 4.48 | 4.58 | 4.67 | 4.77 | 4.87 | 4.96 | 5.06 | 5.15 | 5.25 | 5.35 | 5.44 | 5.54 | 5.64 | 5.73 | 5.83 | 5.92 | 6.02 | 6.12 | 6.21 | 6.31 | 6.4 | 6.5 | 6.6 | 6.69 | 6.79 | 6.88 | 6.98 | 7.08 | 7.17 | 7.27
35.59 | 3.43 | 3.52 | 3.62 | 3.71 | 3.81 | 3.91 | 4 | 4.1 | 4.19 | 4.29 | 4.39 | 4.48 | 4.58 | 4.67 | 4.77 | 4.87 | 4.96 | 5.06 | 5.15 | 5.25 | 5.35 | 5.44 | 5.54 | 5.63 | 5.73 | 5.83 | 5.92 | 6.02 | 6.11 | 6.21 | 6.31 | 6.4 | 6.5 | 6.59 | 6.69 | 6.79 | 6.88 | 6.98 | 7.07 | 7.17 | 7.27
35.6 | 3.42 | 3.52 | 3.62 | 3.71 | 3.81 | 3.9 | 4 | 4.1 | 4.19 | 4.29 | 4.38 | 4.48 | 4.58 | 4.67 | 4.77 | 4.87 | 4.96 | 5.06 | 5.15 | 5.25 | 5.35 | 5.44 | 5.54 | 5.63 | 5.73 | 5.83 | 5.92 | 6.02 | 6.11 | 6.21 | 6.31 | 6.4 | 6.5 | 6.59 | 6.69 | 6.79 | 6.88 | 6.98 | 7.07 | 7.17 | 7.27
35.61 | 3.42 | 3.52 | 3.62 | 3.71 | 3.81 | 3.9 | 4 | 4.1 | 4.19 | 4.29 | 4.38 | 4.48 | 4.58 | 4.67 | 4.77 | 4.86 | 4.96 | 5.06 | 5.15 | 5.25 | 5.34 | 5.44 | 5.54 | 5.63 | 5.73 | 5.82 | 5.92 | 6.02 | 6.11 | 6.21 | 6.3 | 6.4 | 6.5 | 6.59 | 6.69 | 6.78 | 6.88 | 6.98 | 7.07 | 7.17 | 7.26
35.62 | 3.42 | 3.52 | 3.62 | 3.71 | 3.81 | 3.9 | 4 | 4.1 | 4.19 | 4.29 | 4.38 | 4.48 | 4.58 | 4.67 | 4.77 | 4.86 | 4.96 | 5.06 | 5.15 | 5.25 | 5.34 | 5.44 | 5.54 | 5.63 | 5.73 | 5.82 | 5.92 | 6.02 | 6.11 | 6.21 | 6.3 | 6.4 | 6.5 | 6.59 | 6.69 | 6.78 | 6.88 | 6.98 | 7.07 | 7.17 | 7.26
35.63 | 3.42 | 3.52 | 3.61 | 3.71 | 3.81 | 3.9 | 4 | 4.09 | 4.19 | 4.29 | 4.38 | 4.48 | 4.57 | 4.67 | 4.77 | 4.86 | 4.96 | 5.05 | 5.15 | 5.25 | 5.34 | 5.44 | 5.53 | 5.63 | 5.73 | 5.82 | 5.92 | 6.01 | 6.11 | 6.21 | 6.3 | 6.4 | 6.49 | 6.59 | 6.69 | 6.78 | 6.88 | 6.97 | 7.07 | 7.17 | 7.26
35.64 | 3.42 | 3.52 | 3.61 | 3.71 | 3.81 | 3.9 | 4 | 4.09 | 4.19 | 4.29 | 4.38 | 4.48 | 4.57 | 4.67 | 4.77 | 4.86 | 4.96 | 5.05 | 5.15 | 5.25 | 5.34 | 5.44 | 5.53 | 5.63 | 5.73 | 5.82 | 5.92 | 6.01 | 6.11 | 6.21 | 6.3 | 6.4 | 6.49 | 6.59 | 6.69 | 6.78 | 6.88 | 6.97 | 7.07 | 7.16 | 7.26
35.65 | 3.42 | 3.52 | 3.61 | 3.71 | 3.8 | 3.9 | 4 | 4.09 | 4.19 | 4.28 | 4.38 | 4.48 | 4.57 | 4.67 | 4.76 | 4.86 | 4.96 | 5.05 | 5.15 | 5.24 | 5.34 | 5.44 | 5.53 | 5.63 | 5.72 | 5.82 | 5.92 | 6.01 | 6.11 | 6.2 | 6.3 | 6.4 | 6.49 | 6.59 | 6.68 | 6.78 | 6.88 | 6.97 | 7.07 | 7.16 | 7.26
35.66 | 3.42 | 3.52 | 3.61 | 3.71 | 3.8 | 3.9 | 4 | 4.09 | 4.19 | 4.28 | 4.38 | 4.48 | 4.57 | 4.67 | 4.76 | 4.86 | 4.96 | 5.05 | 5.15 | 5.24 | 5.34 | 5.44 | 5.53 | 5.63 | 5.72 | 5.82 | 5.92 | 6.01 | 6.11 | 6.2 | 6.3 | 6.4 | 6.49 | 6.59 | 6.68 | 6.78 | 6.87 | 6.97 | 7.07 | 7.16 | 7.26
35.67 | 3.42 | 3.52 | 3.61 | 3.71 | 3.8 | 3.9 | 4 | 4.09 | 4.19 | 4.28 | 4.38 | 4.48 | 4.57 | 4.67 | 4.76 | 4.86 | 4.95 | 5.05 | 5.15 | 5.24 | 5.34 | 5.43 | 5.53 | 5.63 | 5.72 | 5.82 | 5.91 | 6.01 | 6.11 | 6.2 | 6.3 | 6.39 | 6.49 | 6.59 | 6.68 | 6.78 | 6.87 | 6.97 | 7.07 | 7.16 | 7.26
35.68 | 3.42 | 3.51 | 3.61 | 3.71 | 3.8 | 3.9 | 3.99 | 4.09 | 4.19 | 4.28 | 4.38 | 4.47 | 4.57 | 4.67 | 4.76 | 4.86 | 4.95 | 5.05 | 5.15 | 5.24 | 5.34 | 5.43 | 5.53 | 5.63 | 5.72 | 5.82 | 5.91 | 6.01 | 6.11 | 6.2 | 6.3 | 6.39 | 6.49 | 6.58 | 6.68 | 6.78 | 6.87 | 6.97 | 7.06 | 7.16 | 7.26
35.69 | 3.42 | 3.51 | 3.61 | 3.71 | 3.8 | 3.9 | 3.99 | 4.09 | 4.19 | 4.28 | 4.38 | 4.47 | 4.57 | 4.67 | 4.76 | 4.86 | 4.95 | 5.05 | 5.15 | 5.24 | 5.34 | 5.43 | 5.53 | 5.62 | 5.72 | 5.82 | 5.91 | 6.01 | 6.1 | 6.2 | 6.3 | 6.39 | 6.49 | 6.58 | 6.68 | 6.78 | 6.87 | 6.97 | 7.06 | 7.16 | 7.25
35.7 | 3.42 | 3.51 | 3.61 | 3.71 | 3.8 | 3.9 | 3.99 | 4.09 | 4.19 | 4.28 | 4.38 | 4.47 | 4.57 | 4.66 | 4.76 | 4.86 | 4.95 | 5.05 | 5.14 | 5.24 | 5.34 | 5.43 | 5.53 | 5.62 | 5.72 | 5.82 | 5.91 | 6.01 | 6.1 | 6.2 | 6.29 | 6.39 | 6.49 | 6.58 | 6.68 | 6.77 | 6.87 | 6.97 | 7.06 | 7.16 | 7.25
35.71 | 3.42 | 3.51 | 3.61 | 3.7 | 3.8 | 3.9 | 3.99 | 4.09 | 4.18 | 4.28 | 4.38 | 4.47 | 4.57 | 4.66 | 4.76 | 4.86 | 4.95 | 5.05 | 5.14 | 5.24 | 5.34 | 5.43 | 5.53 | 5.62 | 5.72 | 5.81 | 5.91 | 6.01 | 6.1 | 6.2 | 6.29 | 6.39 | 6.49 | 6.58 | 6.68 | 6.77 | 6.87 | 6.96 | 7.06 | 7.16 | 7.25
35.72 | 3.42 | 3.51 | 3.61 | 3.7 | 3.8 | 3.9 | 3.99 | 4.09 | 4.18 | 4.28 | 4.38 | 4.47 | 4.57 | 4.66 | 4.76 | 4.85 | 4.95 | 5.05 | 5.14 | 5.24 | 5.33 | 5.43 | 5.53 | 5.62 | 5.72 | 5.81 | 5.91 | 6.01 | 6.1 | 6.2 | 6.29 | 6.39 | 6.48 | 6.58 | 6.68 | 6.77 | 6.87 | 6.96 | 7.06 | 7.16 | 7.25
35.73 | 3.42 | 3.51 | 3.61 | 3.7 | 3.8 | 3.9 | 3.99 | 4.09 | 4.18 | 4.28 | 4.37 | 4.47 | 4.57 | 4.66 | 4.76 | 4.85 | 4.95 | 5.05 | 5.14 | 5.24 | 5.33 | 5.43 | 5.52 | 5.62 | 5.72 | 5.81 | 5.91 | 6 | 6.1 | 6.2 | 6.29 | 6.39 | 6.48 | 6.58 | 6.68 | 6.77 | 6.87 | 6.96 | 7.06 | 7.15 | 7.25
35.74 | 3.41 | 3.51 | 3.61 | 3.7 | 3.8 | 3.89 | 3.99 | 4.09 | 4.18 | 4.28 | 4.37 | 4.47 | 4.57 | 4.66 | 4.76 | 4.85 | 4.95 | 5.04 | 5.14 | 5.24 | 5.33 | 5.43 | 5.52 | 5.62 | 5.72 | 5.81 | 5.91 | 6 | 6.1 | 6.19 | 6.29 | 6.39 | 6.48 | 6.58 | 6.67 | 6.77 | 6.87 | 6.96 | 7.06 | 7.15 | 7.25
35.75 | 3.41 | 3.51 | 3.61 | 3.7 | 3.8 | 3.89 | 3.99 | 4.09 | 4.18 | 4.28 | 4.37 | 4.47 | 4.56 | 4.66 | 4.76 | 4.85 | 4.95 | 5.04 | 5.14 | 5.24 | 5.33 | 5.43 | 5.52 | 5.62 | 5.71 | 5.81 | 5.91 | 6 | 6.1 | 6.19 | 6.29 | 6.39 | 6.48 | 6.58 | 6.67 | 6.77 | 6.86 | 6.96 | 7.06 | 7.15 | 7.25
35.76 | 3.41 | 3.51 | 3.61 | 3.7 | 3.8 | 3.89 | 3.99 | 4.08 | 4.18 | 4.28 | 4.37 | 4.47 | 4.56 | 4.66 | 4.76 | 4.85 | 4.95 | 5.04 | 5.14 | 5.23 | 5.33 | 5.43 | 5.52 | 5.62 | 5.71 | 5.81 | 5.91 | 6 | 6.1 | 6.19 | 6.29 | 6.38 | 6.48 | 6.58 | 6.67 | 6.77 | 6.86 | 6.96 | 7.06 | 7.15 | 7.25
35.77 | 3.41 | 3.51 | 3.6 | 3.7 | 3.8 | 3.89 | 3.99 | 4.08 | 4.18 | 4.28 | 4.37 | 4.47 | 4.56 | 4.66 | 4.75 | 4.85 | 4.95 | 5.04 | 5.14 | 5.23 | 5.33 | 5.43 | 5.52 | 5.62 | 5.71 | 5.81 | 5.9 | 6 | 6.1 | 6.19 | 6.29 | 6.38 | 6.48 | 6.58 | 6.67 | 6.77 | 6.86 | 6.96 | 7.05 | 7.15 | 7.25
35.78 | 3.41 | 3.51 | 3.6 | 3.7 | 3.8 | 3.89 | 3.99 | 4.08 | 4.18 | 4.27 | 4.37 | 4.47 | 4.56 | 4.66 | 4.75 | 4.85 | 4.95 | 5.04 | 5.14 | 5.23 | 5.33 | 5.42 | 5.52 | 5.62 | 5.71 | 5.81 | 5.9 | 6 | 6.09 | 6.19 | 6.29 | 6.38 | 6.48 | 6.57 | 6.67 | 6.77 | 6.86 | 6.96 | 7.05 | 7.15 | 7.24
35.79 | 3.41 | 3.51 | 3.6 | 3.7 | 3.79 | 3.89 | 3.99 | 4.08 | 4.18 | 4.27 | 4.37 | 4.47 | 4.56 | 4.66 | 4.75 | 4.85 | 4.94 | 5.04 | 5.14 | 5.23 | 5.33 | 5.42 | 5.52 | 5.62 | 5.71 | 5.81 | 5.9 | 6 | 6.09 | 6.19 | 6.29 | 6.38 | 6.48 | 6.57 | 6.67 | 6.76 | 6.86 | 6.96 | 7.05 | 7.15 | 7.24
35.8 | 3.41 | 3.51 | 3.6 | 3.7 | 3.79 | 3.89 | 3.99 | 4.08 | 4.18 | 4.27 | 4.37 | 4.46 | 4.56 | 4.66 | 4.75 | 4.85 | 4.94 | 5.04 | 5.14 | 5.23 | 5.33 | 5.42 | 5.52 | 5.61 | 5.71 | 5.81 | 5.9 | 6 | 6.09 | 6.19 | 6.28 | 6.38 | 6.48 | 6.57 | 6.67 | 6.76 | 6.86 | 6.95 | 7.05 | 7.15 | 7.24
35.81 | 3.41 | 3.51 | 3.6 | 3.7 | 3.79 | 3.89 | 3.98 | 4.08 | 4.18 | 4.27 | 4.37 | 4.46 | 4.56 | 4.66 | 4.75 | 4.85 | 4.94 | 5.04 | 5.13 | 5.23 | 5.33 | 5.42 | 5.52 | 5.61 | 5.71 | 5.8 | 5.9 | 6 | 6.09 | 6.19 | 6.28 | 6.38 | 6.47 | 6.57 | 6.67 | 6.76 | 6.86 | 6.95 | 7.05 | 7.15 | 7.24
35.82 | 3.41 | 3.51 | 3.6 | 3.7 | 3.79 | 3.89 | 3.98 | 4.08 | 4.18 | 4.27 | 4.37 | 4.46 | 4.56 | 4.65 | 4.75 | 4.85 | 4.94 | 5.04 | 5.13 | 5.23 | 5.32 | 5.42 | 5.52 | 5.61 | 5.71 | 5.8 | 5.9 | 6 | 6.09 | 6.19 | 6.28 | 6.38 | 6.47 | 6.57 | 6.67 | 6.76 | 6.86 | 6.95 | 7.05 | 7.14 | 7.24
35.83 | 3.41 | 3.5 | 3.6 | 3.7 | 3.79 | 3.89 | 3.98 | 4.08 | 4.17 | 4.27 | 4.37 | 4.46 | 4.56 | 4.65 | 4.75 | 4.85 | 4.94 | 5.04 | 5.13 | 5.23 | 5.32 | 5.42 | 5.52 | 5.61 | 5.71 | 5.8 | 5.9 | 5.99 | 6.09 | 6.19 | 6.28 | 6.38 | 6.47 | 6.57 | 6.66 | 6.76 | 6.86 | 6.95 | 7.05 | 7.14 | 7.24
35.84 | 3.41 | 3.5 | 3.6 | 3.7 | 3.79 | 3.89 | 3.98 | 4.08 | 4.17 | 4.27 | 4.37 | 4.46 | 4.56 | 4.65 | 4.75 | 4.84 | 4.94 | 5.04 | 5.13 | 5.23 | 5.32 | 5.42 | 5.51 | 5.61 | 5.71 | 5.8 | 5.9 | 5.99 | 6.09 | 6.18 | 6.28 | 6.38 | 6.47 | 6.57 | 6.66 | 6.76 | 6.85 | 6.95 | 7.05 | 7.14 | 7.24
35.85 | 3.41 | 3.5 | 3.6 | 3.69 | 3.79 | 3.89 | 3.98 | 4.08 | 4.17 | 4.27 | 4.36 | 4.46 | 4.56 | 4.65 | 4.75 | 4.84 | 4.94 | 5.04 | 5.13 | 5.23 | 5.32 | 5.42 | 5.51 | 5.61 | 5.71 | 5.8 | 5.9 | 5.99 | 6.09 | 6.18 | 6.28 | 6.37 | 6.47 | 6.57 | 6.66 | 6.76 | 6.85 | 6.95 | 7.04 | 7.14 | 7.24
35.86 | 3.41 | 3.5 | 3.6 | 3.69 | 3.79 | 3.89 | 3.98 | 4.08 | 4.17 | 4.27 | 4.36 | 4.46 | 4.56 | 4.65 | 4.75 | 4.84 | 4.94 | 5.03 | 5.13 | 5.23 | 5.32 | 5.42 | 5.51 | 5.61 | 5.7 | 5.8 | 5.9 | 5.99 | 6.09 | 6.18 | 6.28 | 6.37 | 6.47 | 6.57 | 6.66 | 6.76 | 6.85 | 6.95 | 7.04 | 7.14 | 7.24
35.87 | 3.41 | 3.5 | 3.6 | 3.69 | 3.79 | 3.88 | 3.98 | 4.08 | 4.17 | 4.27 | 4.36 | 4.46 | 4.55 | 4.65 | 4.75 | 4.84 | 4.94 | 5.03 | 5.13 | 5.22 | 5.32 | 5.42 | 5.51 | 5.61 | 5.7 | 5.8 | 5.89 | 5.99 | 6.09 | 6.18 | 6.28 | 6.37 | 6.47 | 6.56 | 6.66 | 6.76 | 6.85 | 6.95 | 7.04 | 7.14 | 7.23
35.88 | 3.41 | 3.5 | 3.6 | 3.69 | 3.79 | 3.88 | 3.98 | 4.08 | 4.17 | 4.27 | 4.36 | 4.46 | 4.55 | 4.65 | 4.75 | 4.84 | 4.94 | 5.03 | 5.13 | 5.22 | 5.32 | 5.42 | 5.51 | 5.61 | 5.7 | 5.8 | 5.89 | 5.99 | 6.08 | 6.18 | 6.28 | 6.37 | 6.47 | 6.56 | 6.66 | 6.75 | 6.85 | 6.95 | 7.04 | 7.14 | 7.23
35.89 | 3.4 | 3.5 | 3.6 | 3.69 | 3.79 | 3.88 | 3.98 | 4.07 | 4.17 | 4.27 | 4.36 | 4.46 | 4.55 | 4.65 | 4.74 | 4.84 | 4.94 | 5.03 | 5.13 | 5.22 | 5.32 | 5.41 | 5.51 | 5.61 | 5.7 | 5.8 | 5.89 | 5.99 | 6.08 | 6.18 | 6.28 | 6.37 | 6.47 | 6.56 | 6.66 | 6.75 | 6.85 | 6.94 | 7.04 | 7.14 | 7.23
35.9 | 3.4 | 3.5 | 3.6 | 3.69 | 3.79 | 3.88 | 3.98 | 4.07 | 4.17 | 4.27 | 4.36 | 4.46 | 4.55 | 4.65 | 4.74 | 4.84 | 4.93 | 5.03 | 5.13 | 5.22 | 5.32 | 5.41 | 5.51 | 5.6 | 5.7 | 5.8 | 5.89 | 5.99 | 6.08 | 6.18 | 6.27 | 6.37 | 6.47 | 6.56 | 6.66 | 6.75 | 6.85 | 6.94 | 7.04 | 7.13 | 7.23
35.91 | 3.4 | 3.5 | 3.59 | 3.69 | 3.79 | 3.88 | 3.98 | 4.07 | 4.17 | 4.26 | 4.36 | 4.46 | 4.55 | 4.65 | 4.74 | 4.84 | 4.93 | 5.03 | 5.13 | 5.22 | 5.32 | 5.41 | 5.51 | 5.6 | 5.7 | 5.79 | 5.89 | 5.99 | 6.08 | 6.18 | 6.27 | 6.37 | 6.46 | 6.56 | 6.66 | 6.75 | 6.85 | 6.94 | 7.04 | 7.13 | 7.23
35.92 | 3.4 | 3.5 | 3.59 | 3.69 | 3.79 | 3.88 | 3.98 | 4.07 | 4.17 | 4.26 | 4.36 | 4.46 | 4.55 | 4.65 | 4.74 | 4.84 | 4.93 | 5.03 | 5.12 | 5.22 | 5.32 | 5.41 | 5.51 | 5.6 | 5.7 | 5.79 | 5.89 | 5.99 | 6.08 | 6.18 | 6.27 | 6.37 | 6.46 | 6.56 | 6.65 | 6.75 | 6.85 | 6.94 | 7.04 | 7.13 | 7.23
35.93 | 3.4 | 3.5 | 3.59 | 3.69 | 3.78 | 3.88 | 3.98 | 4.07 | 4.17 | 4.26 | 4.36 | 4.45 | 4.55 | 4.65 | 4.74 | 4.84 | 4.93 | 5.03 | 5.12 | 5.22 | 5.31 | 5.41 | 5.51 | 5.6 | 5.7 | 5.79 | 5.89 | 5.98 | 6.08 | 6.18 | 6.27 | 6.37 | 6.46 | 6.56 | 6.65 | 6.75 | 6.84 | 6.94 | 7.04 | 7.13 | 7.23
35.94 | 3.4 | 3.5 | 3.59 | 3.69 | 3.78 | 3.88 | 3.98 | 4.07 | 4.17 | 4.26 | 4.36 | 4.45 | 4.55 | 4.64 | 4.74 | 4.84 | 4.93 | 5.03 | 5.12 | 5.22 | 5.31 | 5.41 | 5.51 | 5.6 | 5.7 | 5.79 | 5.89 | 5.98 | 6.08 | 6.17 | 6.27 | 6.37 | 6.46 | 6.56 | 6.65 | 6.75 | 6.84 | 6.94 | 7.03 | 7.13 | 7.23
35.95 | 3.4 | 3.5 | 3.59 | 3.69 | 3.78 | 3.88 | 3.97 | 4.07 | 4.17 | 4.26 | 4.36 | 4.45 | 4.55 | 4.64 | 4.74 | 4.84 | 4.93 | 5.03 | 5.12 | 5.22 | 5.31 | 5.41 | 5.5 | 5.6 | 5.7 | 5.79 | 5.89 | 5.98 | 6.08 | 6.17 | 6.27 | 6.36 | 6.46 | 6.56 | 6.65 | 6.75 | 6.84 | 6.94 | 7.03 | 7.13 | 7.22
35.96 | 3.4 | 3.5 | 3.59 | 3.69 | 3.78 | 3.88 | 3.97 | 4.07 | 4.16 | 4.26 | 4.36 | 4.45 | 4.55 | 4.64 | 4.74 | 4.83 | 4.93 | 5.03 | 5.12 | 5.22 | 5.31 | 5.41 | 5.5 | 5.6 | 5.69 | 5.79 | 5.89 | 5.98 | 6.08 | 6.17 | 6.27 | 6.36 | 6.46 | 6.55 | 6.65 | 6.75 | 6.84 | 6.94 | 7.03 | 7.13 | 7.22
35.97 | 3.4 | 3.5 | 3.59 | 3.69 | 3.78 | 3.88 | 3.97 | 4.07 | 4.16 | 4.26 | 4.36 | 4.45 | 4.55 | 4.64 | 4.74 | 4.83 | 4.93 | 5.02 | 5.12 | 5.22 | 5.31 | 5.41 | 5.5 | 5.6 | 5.69 | 5.79 | 5.88 | 5.98 | 6.08 | 6.17 | 6.27 | 6.36 | 6.46 | 6.55 | 6.65 | 6.74 | 6.84 | 6.94 | 7.03 | 7.13 | 7.22
35.98 | 3.4 | 3.49 | 3.59 | 3.69 | 3.78 | 3.88 | 3.97 | 4.07 | 4.16 | 4.26 | 4.35 | 4.45 | 4.55 | 4.64 | 4.74 | 4.83 | 4.93 | 5.02 | 5.12 | 5.21 | 5.31 | 5.41 | 5.5 | 5.6 | 5.69 | 5.79 | 5.88 | 5.98 | 6.07 | 6.17 | 6.27 | 6.36 | 6.46 | 6.55 | 6.65 | 6.74 | 6.84 | 6.93 | 7.03 | 7.13 | 7.22
35.99 | 3.4 | 3.49 | 3.59 | 3.68 | 3.78 | 3.88 | 3.97 | 4.07 | 4.16 | 4.26 | 4.35 | 4.45 | 4.54 | 4.64 | 4.74 | 4.83 | 4.93 | 5.02 | 5.12 | 5.21 | 5.31 | 5.4 | 5.5 | 5.6 | 5.69 | 5.79 | 5.88 | 5.98 | 6.07 | 6.17 | 6.26 | 6.36 | 6.46 | 6.55 | 6.65 | 6.74 | 6.84 | 6.93 | 7.03 | 7.12 | 7.22
36 | 3.4 | 3.49 | 3.59 | 3.68 | 3.78 | 3.88 | 3.97 | 4.07 | 4.16 | 4.26 | 4.35 | 4.45 | 4.54 | 4.64 | 4.74 | 4.83 | 4.93 | 5.02 | 5.12 | 5.21 | 5.31 | 5.4 | 5.5 | 5.6 | 5.69 | 5.79 | 5.88 | 5.98 | 6.07 | 6.17 | 6.26 | 6.36 | 6.45 | 6.55 | 6.65 | 6.74 | 6.84 | 6.93 | 7.03 | 7.12 | 7.22
36.01 | 3.4 | 3.49 | 3.59 | 3.68 | 3.78 | 3.87 | 3.97 | 4.07 | 4.16 | 4.26 | 4.35 | 4.45 | 4.54 | 4.64 | 4.73 | 4.83 | 4.93 | 5.02 | 5.12 | 5.21 | 5.31 | 5.4 | 5.5 | 5.59 | 5.69 | 5.79 | 5.88 | 5.98 | 6.07 | 6.17 | 6.26 | 6.36 | 6.45 | 6.55 | 6.64 | 6.74 | 6.84 | 6.93 | 7.03 | 7.12 | 7.22
36.02 | 3.4 | 3.49 | 3.59 | 3.68 | 3.78 | 3.87 | 3.97 | 4.06 | 4.16 | 4.26 | 4.35 | 4.45 | 4.54 | 4.64 | 4.73 | 4.83 | 4.92 | 5.02 | 5.12 | 5.21 | 5.31 | 5.4 | 5.5 | 5.59 | 5.69 | 5.78 | 5.88 | 5.98 | 6.07 | 6.17 | 6.26 | 6.36 | 6.45 | 6.55 | 6.64 | 6.74 | 6.83 | 6.93 | 7.03 | 7.12 | 7.22
36.03 | 3.4 | 3.49 | 3.59 | 3.68 | 3.78 | 3.87 | 3.97 | 4.06 | 4.16 | 4.26 | 4.35 | 4.45 | 4.54 | 4.64 | 4.73 | 4.83 | 4.92 | 5.02 | 5.11 | 5.21 | 5.31 | 5.4 | 5.5 | 5.59 | 5.69 | 5.78 | 5.88 | 5.97 | 6.07 | 6.17 | 6.26 | 6.36 | 6.45 | 6.55 | 6.64 | 6.74 | 6.83 | 6.93 | 7.02 | 7.12 | 7.22
36.04 | 3.39 | 3.49 | 3.59 | 3.68 | 3.78 | 3.87 | 3.97 | 4.06 | 4.16 | 4.25 | 4.35 | 4.45 | 4.54 | 4.64 | 4.73 | 4.83 | 4.92 | 5.02 | 5.11 | 5.21 | 5.3 | 5.4 | 5.5 | 5.59 | 5.69 | 5.78 | 5.88 | 5.97 | 6.07 | 6.16 | 6.26 | 6.36 | 6.45 | 6.55 | 6.64 | 6.74 | 6.83 | 6.93 | 7.02 | 7.12 | 7.21
36.05 | 3.39 | 3.49 | 3.59 | 3.68 | 3.78 | 3.87 | 3.97 | 4.06 | 4.16 | 4.25 | 4.35 | 4.44 | 4.54 | 4.64 | 4.73 | 4.83 | 4.92 | 5.02 | 5.11 | 5.21 | 5.3 | 5.4 | 5.49 | 5.59 | 5.69 | 5.78 | 5.88 | 5.97 | 6.07 | 6.16 | 6.26 | 6.35 | 6.45 | 6.54 | 6.64 | 6.74 | 6.83 | 6.93 | 7.02 | 7.12 | 7.21
36.06 | 3.39 | 3.49 | 3.58 | 3.68 | 3.78 | 3.87 | 3.97 | 4.06 | 4.16 | 4.25 | 4.35 | 4.44 | 4.54 | 4.63 | 4.73 | 4.83 | 4.92 | 5.02 | 5.11 | 5.21 | 5.3 | 5.4 | 5.49 | 5.59 | 5.68 | 5.78 | 5.88 | 5.97 | 6.07 | 6.16 | 6.26 | 6.35 | 6.45 | 6.54 | 6.64 | 6.73 | 6.83 | 6.93 | 7.02 | 7.12 | 7.21
36.07 | 3.39 | 3.49 | 3.58 | 3.68 | 3.77 | 3.87 | 3.97 | 4.06 | 4.16 | 4.25 | 4.35 | 4.44 | 4.54 | 4.63 | 4.73 | 4.82 | 4.92 | 5.02 | 5.11 | 5.21 | 5.3 | 5.4 | 5.49 | 5.59 | 5.68 | 5.78 | 5.87 | 5.97 | 6.07 | 6.16 | 6.26 | 6.35 | 6.45 | 6.54 | 6.64 | 6.73 | 6.83 | 6.92 | 7.02 | 7.12 | 7.21
36.08 | 3.39 | 3.49 | 3.58 | 3.68 | 3.77 | 3.87 | 3.96 | 4.06 | 4.16 | 4.25 | 4.35 | 4.44 | 4.54 | 4.63 | 4.73 | 4.82 | 4.92 | 5.01 | 5.11 | 5.21 | 5.3 | 5.4 | 5.49 | 5.59 | 5.68 | 5.78 | 5.87 | 5.97 | 6.06 | 6.16 | 6.26 | 6.35 | 6.45 | 6.54 | 6.64 | 6.73 | 6.83 | 6.92 | 7.02 | 7.11 | 7.21
36.09 | 3.39 | 3.49 | 3.58 | 3.68 | 3.77 | 3.87 | 3.96 | 4.06 | 4.16 | 4.25 | 4.35 | 4.44 | 4.54 | 4.63 | 4.73 | 4.82 | 4.92 | 5.01 | 5.11 | 5.2 | 5.3 | 5.4 | 5.49 | 5.59 | 5.68 | 5.78 | 5.87 | 5.97 | 6.06 | 6.16 | 6.25 | 6.35 | 6.45 | 6.54 | 6.64 | 6.73 | 6.83 | 6.92 | 7.02 | 7.11 | 7.21
36.1 | 3.39 | 3.49 | 3.58 | 3.68 | 3.77 | 3.87 | 3.96 | 4.06 | 4.15 | 4.25 | 4.35 | 4.44 | 4.54 | 4.63 | 4.73 | 4.82 | 4.92 | 5.01 | 5.11 | 5.2 | 5.3 | 5.39 | 5.49 | 5.59 | 5.68 | 5.78 | 5.87 | 5.97 | 6.06 | 6.16 | 6.25 | 6.35 | 6.44 | 6.54 | 6.63 | 6.73 | 6.83 | 6.92 | 7.02 | 7.11 | 7.21
36.11 | 3.39 | 3.49 | 3.58 | 3.68 | 3.77 | 3.87 | 3.96 | 4.06 | 4.15 | 4.25 | 4.34 | 4.44 | 4.54 | 4.63 | 4.73 | 4.82 | 4.92 | 5.01 | 5.11 | 5.2 | 5.3 | 5.39 | 5.49 | 5.58 | 5.68 | 5.78 | 5.87 | 5.97 | 6.06 | 6.16 | 6.25 | 6.35 | 6.44 | 6.54 | 6.63 | 6.73 | 6.82 | 6.92 | 7.02 | 7.11 | 7.21
36.12 | 3.39 | 3.48 | 3.58 | 3.68 | 3.77 | 3.87 | 3.96 | 4.06 | 4.15 | 4.25 | 4.34 | 4.44 | 4.53 | 4.63 | 4.73 | 4.82 | 4.92 | 5.01 | 5.11 | 5.2 | 5.3 | 5.39 | 5.49 | 5.58 | 5.68 | 5.77 | 5.87 | 5.97 | 6.06 | 6.16 | 6.25 | 6.35 | 6.44 | 6.54 | 6.63 | 6.73 | 6.82 | 6.92 | 7.01 | 7.11 | 7.2
36.13 | 3.39 | 3.48 | 3.58 | 3.67 | 3.77 | 3.87 | 3.96 | 4.06 | 4.15 | 4.25 | 4.34 | 4.44 | 4.53 | 4.63 | 4.72 | 4.82 | 4.92 | 5.01 | 5.11 | 5.2 | 5.3 | 5.39 | 5.49 | 5.58 | 5.68 | 5.77 | 5.87 | 5.96 | 6.06 | 6.15 | 6.25 | 6.35 | 6.44 | 6.54 | 6.63 | 6.73 | 6.82 | 6.92 | 7.01 | 7.11 | 7.2
36.14 | 3.39 | 3.48 | 3.58 | 3.67 | 3.77 | 3.87 | 3.96 | 4.06 | 4.15 | 4.25 | 4.34 | 4.44 | 4.53 | 4.63 | 4.72 | 4.82 | 4.91 | 5.01 | 5.1 | 5.2 | 5.3 | 5.39 | 5.49 | 5.58 | 5.68 | 5.77 | 5.87 | 5.96 | 6.06 | 6.15 | 6.25 | 6.34 | 6.44 | 6.54 | 6.63 | 6.73 | 6.82 | 6.92 | 7.01 | 7.11 | 7.2
36.15 | 3.39 | 3.48 | 3.58 | 3.67 | 3.77 | 3.86 | 3.96 | 4.06 | 4.15 | 4.25 | 4.34 | 4.44 | 4.53 | 4.63 | 4.72 | 4.82 | 4.91 | 5.01 | 5.1 | 5.2 | 5.29 | 5.39 | 5.49 | 5.58 | 5.68 | 5.77 | 5.87 | 5.96 | 6.06 | 6.15 | 6.25 | 6.34 | 6.44 | 6.53 | 6.63 | 6.72 | 6.82 | 6.92 | 7.01 | 7.11 | 7.2
36.16 | 3.39 | 3.48 | 3.58 | 3.67 | 3.77 | 3.86 | 3.96 | 4.05 | 4.15 | 4.24 | 4.34 | 4.44 | 4.53 | 4.63 | 4.72 | 4.82 | 4.91 | 5.01 | 5.1 | 5.2 | 5.29 | 5.39 | 5.48 | 5.58 | 5.68 | 5.77 | 5.87 | 5.96 | 6.06 | 6.15 | 6.25 | 6.34 | 6.44 | 6.53 | 6.63 | 6.72 | 6.82 | 6.91 | 7.01 | 7.1 | 7.2
36.17 | 3.39 | 3.48 | 3.58 | 3.67 | 3.77 | 3.86 | 3.96 | 4.05 | 4.15 | 4.24 | 4.34 | 4.43 | 4.53 | 4.63 | 4.72 | 4.82 | 4.91 | 5.01 | 5.1 | 5.2 | 5.29 | 5.39 | 5.48 | 5.58 | 5.67 | 5.77 | 5.86 | 5.96 | 6.06 | 6.15 | 6.25 | 6.34 | 6.44 | 6.53 | 6.63 | 6.72 | 6.82 | 6.91 | 7.01 | 7.1 | 7.2
36.18 | 3.39 | 3.48 | 3.58 | 3.67 | 3.77 | 3.86 | 3.96 | 4.05 | 4.15 | 4.24 | 4.34 | 4.43 | 4.53 | 4.62 | 4.72 | 4.82 | 4.91 | 5.01 | 5.1 | 5.2 | 5.29 | 5.39 | 5.48 | 5.58 | 5.67 | 5.77 | 5.86 | 5.96 | 6.05 | 6.15 | 6.25 | 6.34 | 6.44 | 6.53 | 6.63 | 6.72 | 6.82 | 6.91 | 7.01 | 7.1 | 7.2
36.19 | 3.38 | 3.48 | 3.58 | 3.67 | 3.77 | 3.86 | 3.96 | 4.05 | 4.15 | 4.24 | 4.34 | 4.43 | 4.53 | 4.62 | 4.72 | 4.81 | 4.91 | 5.01 | 5.1 | 5.2 | 5.29 | 5.39 | 5.48 | 5.58 | 5.67 | 5.77 | 5.86 | 5.96 | 6.05 | 6.15 | 6.24 | 6.34 | 6.43 | 6.53 | 6.63 | 6.72 | 6.82 | 6.91 | 7.01 | 7.1 | 7.2
36.2 | 3.38 | 3.48 | 3.57 | 3.67 | 3.77 | 3.86 | 3.96 | 4.05 | 4.15 | 4.24 | 4.34 | 4.43 | 4.53 | 4.62 | 4.72 | 4.81 | 4.91 | 5 | 5.1 | 5.19 | 5.29 | 5.39 | 5.48 | 5.58 | 5.67 | 5.77 | 5.86 | 5.96 | 6.05 | 6.15 | 6.24 | 6.34 | 6.43 | 6.53 | 6.62 | 6.72 | 6.81 | 6.91 | 7.01 | 7.1 | 7.2
36.21 | 3.38 | 3.48 | 3.57 | 3.67 | 3.76 | 3.86 | 3.96 | 4.05 | 4.15 | 4.24 | 4.34 | 4.43 | 4.53 | 4.62 | 4.72 | 4.81 | 4.91 | 5 | 5.1 | 5.19 | 5.29 | 5.38 | 5.48 | 5.58 | 5.67 | 5.77 | 5.86 | 5.96 | 6.05 | 6.15 | 6.24 | 6.34 | 6.43 | 6.53 | 6.62 | 6.72 | 6.81 | 6.91 | 7 | 7.1 | 7.19
36.22 | 3.38 | 3.48 | 3.57 | 3.67 | 3.76 | 3.86 | 3.95 | 4.05 | 4.15 | 4.24 | 4.34 | 4.43 | 4.53 | 4.62 | 4.72 | 4.81 | 4.91 | 5 | 5.1 | 5.19 | 5.29 | 5.38 | 5.48 | 5.57 | 5.67 | 5.76 | 5.86 | 5.96 | 6.05 | 6.15 | 6.24 | 6.34 | 6.43 | 6.53 | 6.62 | 6.72 | 6.81 | 6.91 | 7 | 7.1 | 7.19
36.23 | 3.38 | 3.48 | 3.57 | 3.67 | 3.76 | 3.86 | 3.95 | 4.05 | 4.14 | 4.24 | 4.33 | 4.43 | 4.53 | 4.62 | 4.72 | 4.81 | 4.91 | 5 | 5.1 | 5.19 | 5.29 | 5.38 | 5.48 | 5.57 | 5.67 | 5.76 | 5.86 | 5.95 | 6.05 | 6.14 | 6.24 | 6.34 | 6.43 | 6.53 | 6.62 | 6.72 | 6.81 | 6.91 | 7 | 7.1 | 7.19
36.24 | 3.38 | 3.48 | 3.57 | 3.67 | 3.76 | 3.86 | 3.95 | 4.05 | 4.14 | 4.24 | 4.33 | 4.43 | 4.52 | 4.62 | 4.72 | 4.81 | 4.91 | 5 | 5.1 | 5.19 | 5.29 | 5.38 | 5.48 | 5.57 | 5.67 | 5.76 | 5.86 | 5.95 | 6.05 | 6.14 | 6.24 | 6.33 | 6.43 | 6.52 | 6.62 | 6.71 | 6.81 | 6.91 | 7 | 7.1 | 7.19
36.25 | 3.38 | 3.48 | 3.57 | 3.67 | 3.76 | 3.86 | 3.95 | 4.05 | 4.14 | 4.24 | 4.33 | 4.43 | 4.52 | 4.62 | 4.71 | 4.81 | 4.9 | 5 | 5.1 | 5.19 | 5.29 | 5.38 | 5.48 | 5.57 | 5.67 | 5.76 | 5.86 | 5.95 | 6.05 | 6.14 | 6.24 | 6.33 | 6.43 | 6.52 | 6.62 | 6.71 | 6.81 | 6.9 | 7 | 7.09 | 7.19
36.26 | 3.38 | 3.48 | 3.57 | 3.67 | 3.76 | 3.86 | 3.95 | 4.05 | 4.14 | 4.24 | 4.33 | 4.43 | 4.52 | 4.62 | 4.71 | 4.81 | 4.9 | 5 | 5.09 | 5.19 | 5.28 | 5.38 | 5.48 | 5.57 | 5.67 | 5.76 | 5.86 | 5.95 | 6.05 | 6.14 | 6.24 | 6.33 | 6.43 | 6.52 | 6.62 | 6.71 | 6.81 | 6.9 | 7 | 7.09 | 7.19
36.27 | 3.38 | 3.47 | 3.57 | 3.67 | 3.76 | 3.86 | 3.95 | 4.05 | 4.14 | 4.24 | 4.33 | 4.43 | 4.52 | 4.62 | 4.71 | 4.81 | 4.9 | 5 | 5.09 | 5.19 | 5.28 | 5.38 | 5.47 | 5.57 | 5.66 | 5.76 | 5.85 | 5.95 | 6.05 | 6.14 | 6.24 | 6.33 | 6.43 | 6.52 | 6.62 | 6.71 | 6.81 | 6.9 | 7 | 7.09 | 7.19
36.28 | 3.38 | 3.47 | 3.57 | 3.66 | 3.76 | 3.85 | 3.95 | 4.05 | 4.14 | 4.24 | 4.33 | 4.43 | 4.52 | 4.62 | 4.71 | 4.81 | 4.9 | 5 | 5.09 | 5.19 | 5.28 | 5.38 | 5.47 | 5.57 | 5.66 | 5.76 | 5.85 | 5.95 | 6.04 | 6.14 | 6.23 | 6.33 | 6.43 | 6.52 | 6.62 | 6.71 | 6.81 | 6.9 | 7 | 7.09 | 7.19
36.29 | 3.38 | 3.47 | 3.57 | 3.66 | 3.76 | 3.85 | 3.95 | 4.04 | 4.14 | 4.23 | 4.33 | 4.43 | 4.52 | 4.62 | 4.71 | 4.81 | 4.9 | 5 | 5.09 | 5.19 | 5.28 | 5.38 | 5.47 | 5.57 | 5.66 | 5.76 | 5.85 | 5.95 | 6.04 | 6.14 | 6.23 | 6.33 | 6.42 | 6.52 | 6.61 | 6.71 | 6.8 | 6.9 | 6.99 | 7.09 | 7.19
36.3 | 3.38 | 3.47 | 3.57 | 3.66 | 3.76 | 3.85 | 3.95 | 4.04 | 4.14 | 4.23 | 4.33 | 4.42 | 4.52 | 4.61 | 4.71 | 4.81 | 4.9 | 5 | 5.09 | 5.19 | 5.28 | 5.38 | 5.47 | 5.57 | 5.66 | 5.76 | 5.85 | 5.95 | 6.04 | 6.14 | 6.23 | 6.33 | 6.42 | 6.52 | 6.61 | 6.71 | 6.8 | 6.9 | 6.99 | 7.09 | 7.18
36.31 | 3.38 | 3.47 | 3.57 | 3.66 | 3.76 | 3.85 | 3.95 | 4.04 | 4.14 | 4.23 | 4.33 | 4.42 | 4.52 | 4.61 | 4.71 | 4.8 | 4.9 | 4.99 | 5.09 | 5.18 | 5.28 | 5.38 | 5.47 | 5.57 | 5.66 | 5.76 | 5.85 | 5.95 | 6.04 | 6.14 | 6.23 | 6.33 | 6.42 | 6.52 | 6.61 | 6.71 | 6.8 | 6.9 | 6.99 | 7.09 | 7.18
36.32 | 3.38 | 3.47 | 3.57 | 3.66 | 3.76 | 3.85 | 3.95 | 4.04 | 4.14 | 4.23 | 4.33 | 4.42 | 4.52 | 4.61 | 4.71 | 4.8 | 4.9 | 4.99 | 5.09 | 5.18 | 5.28 | 5.37 | 5.47 | 5.56 | 5.66 | 5.75 | 5.85 | 5.95 | 6.04 | 6.14 | 6.23 | 6.33 | 6.42 | 6.52 | 6.61 | 6.71 | 6.8 | 6.9 | 6.99 | 7.09 | 7.18
36.33 | 3.38 | 3.47 | 3.57 | 3.66 | 3.76 | 3.85 | 3.95 | 4.04 | 4.14 | 4.23 | 4.33 | 4.42 | 4.52 | 4.61 | 4.71 | 4.8 | 4.9 | 4.99 | 5.09 | 5.18 | 5.28 | 5.37 | 5.47 | 5.56 | 5.66 | 5.75 | 5.85 | 5.94 | 6.04 | 6.13 | 6.23 | 6.32 | 6.42 | 6.51 | 6.61 | 6.71 | 6.8 | 6.9 | 6.99 | 7.09 | 7.18
36.34 | 3.37 | 3.47 | 3.57 | 3.66 | 3.76 | 3.85 | 3.95 | 4.04 | 4.14 | 4.23 | 4.33 | 4.42 | 4.52 | 4.61 | 4.71 | 4.8 | 4.9 | 4.99 | 5.09 | 5.18 | 5.28 | 5.37 | 5.47 | 5.56 | 5.66 | 5.75 | 5.85 | 5.94 | 6.04 | 6.13 | 6.23 | 6.32 | 6.42 | 6.51 | 6.61 | 6.7 | 6.8 | 6.89 | 6.99 | 7.08 | 7.18
36.35 | 3.37 | 3.47 | 3.56 | 3.66 | 3.75 | 3.85 | 3.94 | 4.04 | 4.14 | 4.23 | 4.33 | 4.42 | 4.52 | 4.61 | 4.71 | 4.8 | 4.9 | 4.99 | 5.09 | 5.18 | 5.28 | 5.37 | 5.47 | 5.56 | 5.66 | 5.75 | 5.85 | 5.94 | 6.04 | 6.13 | 6.23 | 6.32 | 6.42 | 6.51 | 6.61 | 6.7 | 6.8 | 6.89 | 6.99 | 7.08 | 7.18
36.36 | 3.37 | 3.47 | 3.56 | 3.66 | 3.75 | 3.85 | 3.94 | 4.04 | 4.13 | 4.23 | 4.32 | 4.42 | 4.51 | 4.61 | 4.7 | 4.8 | 4.9 | 4.99 | 5.09 | 5.18 | 5.28 | 5.37 | 5.47 | 5.56 | 5.66 | 5.75 | 5.85 | 5.94 | 6.04 | 6.13 | 6.23 | 6.32 | 6.42 | 6.51 | 6.61 | 6.7 | 6.8 | 6.89 | 6.99 | 7.08 | 7.18
36.37 | 3.37 | 3.47 | 3.56 | 3.66 | 3.75 | 3.85 | 3.94 | 4.04 | 4.13 | 4.23 | 4.32 | 4.42 | 4.51 | 4.61 | 4.7 | 4.8 | 4.89 | 4.99 | 5.08 | 5.18 | 5.27 | 5.37 | 5.46 | 5.56 | 5.66 | 5.75 | 5.85 | 5.94 | 6.04 | 6.13 | 6.23 | 6.32 | 6.42 | 6.51 | 6.61 | 6.7 | 6.8 | 6.89 | 6.99 | 7.08 | 7.18
36.38 | 3.37 | 3.47 | 3.56 | 3.66 | 3.75 | 3.85 | 3.94 | 4.04 | 4.13 | 4.23 | 4.32 | 4.42 | 4.51 | 4.61 | 4.7 | 4.8 | 4.89 | 4.99 | 5.08 | 5.18 | 5.27 | 5.37 | 5.46 | 5.56 | 5.65 | 5.75 | 5.84 | 5.94 | 6.03 | 6.13 | 6.22 | 6.32 | 6.41 | 6.51 | 6.6 | 6.7 | 6.79 | 6.89 | 6.98 | 7.08 | 7.17
36.39 | 3.37 | 3.47 | 3.56 | 3.66 | 3.75 | 3.85 | 3.94 | 4.04 | 4.13 | 4.23 | 4.32 | 4.42 | 4.51 | 4.61 | 4.7 | 4.8 | 4.89 | 4.99 | 5.08 | 5.18 | 5.27 | 5.37 | 5.46 | 5.56 | 5.65 | 5.75 | 5.84 | 5.94 | 6.03 | 6.13 | 6.22 | 6.32 | 6.41 | 6.51 | 6.6 | 6.7 | 6.79 | 6.89 | 6.98 | 7.08 | 7.17
36.4 | 3.37 | 3.47 | 3.56 | 3.66 | 3.75 | 3.85 | 3.94 | 4.04 | 4.13 | 4.23 | 4.32 | 4.42 | 4.51 | 4.61 | 4.7 | 4.8 | 4.89 | 4.99 | 5.08 | 5.18 | 5.27 | 5.37 | 5.46 | 5.56 | 5.65 | 5.75 | 5.84 | 5.94 | 6.03 | 6.13 | 6.22 | 6.32 | 6.41 | 6.51 | 6.6 | 6.7 | 6.79 | 6.89 | 6.98 | 7.08 | 7.17
36.41 | 3.37 | 3.47 | 3.56 | 3.66 | 3.75 | 3.85 | 3.94 | 4.04 | 4.13 | 4.23 | 4.32 | 4.42 | 4.51 | 4.61 | 4.7 | 4.8 | 4.89 | 4.99 | 5.08 | 5.18 | 5.27 | 5.37 | 5.46 | 5.56 | 5.65 | 5.75 | 5.84 | 5.94 | 6.03 | 6.13 | 6.22 | 6.32 | 6.41 | 6.51 | 6.6 | 6.7 | 6.79 | 6.89 | 6.98 | 7.08 | 7.17
36.42 | 3.37 | 3.46 | 3.56 | 3.65 | 3.75 | 3.84 | 3.94 | 4.03 | 4.13 | 4.22 | 4.32 | 4.41 | 4.51 | 4.6 | 4.7 | 4.8 | 4.89 | 4.99 | 5.08 | 5.18 | 5.27 | 5.37 | 5.46 | 5.56 | 5.65 | 5.75 | 5.84 | 5.94 | 6.03 | 6.13 | 6.22 | 6.32 | 6.41 | 6.51 | 6.6 | 6.7 | 6.79 | 6.89 | 6.98 | 7.08 | 7.17
36.43 | 3.37 | 3.46 | 3.56 | 3.65 | 3.75 | 3.84 | 3.94 | 4.03 | 4.13 | 4.22 | 4.32 | 4.41 | 4.51 | 4.6 | 4.7 | 4.79 | 4.89 | 4.98 | 5.08 | 5.17 | 5.27 | 5.36 | 5.46 | 5.55 | 5.65 | 5.74 | 5.84 | 5.93 | 6.03 | 6.12 | 6.22 | 6.31 | 6.41 | 6.5 | 6.6 | 6.69 | 6.79 | 6.88 | 6.98 | 7.07 | 7.17
36.44 | 3.37 | 3.46 | 3.56 | 3.65 | 3.75 | 3.84 | 3.94 | 4.03 | 4.13 | 4.22 | 4.32 | 4.41 | 4.51 | 4.6 | 4.7 | 4.79 | 4.89 | 4.98 | 5.08 | 5.17 | 5.27 | 5.36 | 5.46 | 5.55 | 5.65 | 5.74 | 5.84 | 5.93 | 6.03 | 6.12 | 6.22 | 6.31 | 6.41 | 6.5 | 6.6 | 6.69 | 6.79 | 6.88 | 6.98 | 7.07 | 7.17
36.45 | 3.37 | 3.46 | 3.56 | 3.65 | 3.75 | 3.84 | 3.94 | 4.03 | 4.13 | 4.22 | 4.32 | 4.41 | 4.51 | 4.6 | 4.7 | 4.79 | 4.89 | 4.98 | 5.08 | 5.17 | 5.27 | 5.36 | 5.46 | 5.55 | 5.65 | 5.74 | 5.84 | 5.93 | 6.03 | 6.12 | 6.22 | 6.31 | 6.41 | 6.5 | 6.6 | 6.69 | 6.79 | 6.88 | 6.98 | 7.07 | 7.17
36.46 | 3.37 | 3.46 | 3.56 | 3.65 | 3.75 | 3.84 | 3.94 | 4.03 | 4.13 | 4.22 | 4.32 | 4.41 | 4.51 | 4.6 | 4.7 | 4.79 | 4.89 | 4.98 | 5.08 | 5.17 | 5.27 | 5.36 | 5.46 | 5.55 | 5.65 | 5.74 | 5.84 | 5.93 | 6.03 | 6.12 | 6.22 | 6.31 | 6.41 | 6.5 | 6.6 | 6.69 | 6.79 | 6.88 | 6.98 | 7.07 | 7.17
36.47 | 3.37 | 3.46 | 3.56 | 3.65 | 3.75 | 3.84 | 3.94 | 4.03 | 4.13 | 4.22 | 4.32 | 4.41 | 4.51 | 4.6 | 4.7 | 4.79 | 4.89 | 4.98 | 5.08 | 5.17 | 5.27 | 5.36 | 5.46 | 5.55 | 5.65 | 5.74 | 5.84 | 5.93 | 6.03 | 6.12 | 6.22 | 6.31 | 6.41 | 6.5 | 6.59 | 6.69 | 6.78 | 6.88 | 6.97 | 7.07 | 7.16
36.48 | 3.37 | 3.46 | 3.56 | 3.65 | 3.75 | 3.84 | 3.94 | 4.03 | 4.13 | 4.22 | 4.32 | 4.41 | 4.51 | 4.6 | 4.69 | 4.79 | 4.88 | 4.98 | 5.07 | 5.17 | 5.26 | 5.36 | 5.45 | 5.55 | 5.64 | 5.74 | 5.83 | 5.93 | 6.02 | 6.12 | 6.21 | 6.31 | 6.4 | 6.5 | 6.59 | 6.69 | 6.78 | 6.88 | 6.97 | 7.07 | 7.16
36.49 | 3.36 | 3.46 | 3.55 | 3.65 | 3.74 | 3.84 | 3.93 | 4.03 | 4.12 | 4.22 | 4.31 | 4.41 | 4.5 | 4.6 | 4.69 | 4.79 | 4.88 | 4.98 | 5.07 | 5.17 | 5.26 | 5.36 | 5.45 | 5.55 | 5.64 | 5.74 | 5.83 | 5.93 | 6.02 | 6.12 | 6.21 | 6.31 | 6.4 | 6.5 | 6.59 | 6.69 | 6.78 | 6.88 | 6.97 | 7.07 | 7.16
36.5 | 3.36 | 3.46 | 3.55 | 3.65 | 3.74 | 3.84 | 3.93 | 4.03 | 4.12 | 4.22 | 4.31 | 4.41 | 4.5 | 4.6 | 4.69 | 4.79 | 4.88 | 4.98 | 5.07 | 5.17 | 5.26 | 5.36 | 5.45 | 5.55 | 5.64 | 5.74 | 5.83 | 5.93 | 6.02 | 6.12 | 6.21 | 6.31 | 6.4 | 6.5 | 6.59 | 6.69 | 6.78 | 6.88 | 6.97 | 7.07 | 7.16
36.51 | 3.36 | 3.46 | 3.55 | 3.65 | 3.74 | 3.84 | 3.93 | 4.03 | 4.12 | 4.22 | 4.31 | 4.41 | 4.5 | 4.6 | 4.69 | 4.79 | 4.88 | 4.98 | 5.07 | 5.17 | 5.26 | 5.36 | 5.45 | 5.55 | 5.64 | 5.74 | 5.83 | 5.93 | 6.02 | 6.12 | 6.21 | 6.31 | 6.4 | 6.5 | 6.59 | 6.69 | 6.78 | 6.88 | 6.97 | 7.07 | 7.16
36.52 | 3.36 | 3.46 | 3.55 | 3.65 | 3.74 | 3.84 | 3.93 | 4.03 | 4.12 | 4.22 | 4.31 | 4.41 | 4.5 | 4.6 | 4.69 | 4.79 | 4.88 | 4.98 | 5.07 | 5.17 | 5.26 | 5.36 | 5.45 | 5.55 | 5.64 | 5.74 | 5.83 | 5.93 | 6.02 | 6.12 | 6.21 | 6.3 | 6.4 | 6.49 | 6.59 | 6.68 | 6.78 | 6.87 | 6.97 | 7.06 | 7.16
36.53 | 3.36 | 3.46 | 3.55 | 3.65 | 3.74 | 3.84 | 3.93 | 4.03 | 4.12 | 4.22 | 4.31 | 4.41 | 4.5 | 4.6 | 4.69 | 4.79 | 4.88 | 4.98 | 5.07 | 5.17 | 5.26 | 5.36 | 5.45 | 5.54 | 5.64 | 5.73 | 5.83 | 5.92 | 6.02 | 6.11 | 6.21 | 6.3 | 6.4 | 6.49 | 6.59 | 6.68 | 6.78 | 6.87 | 6.97 | 7.06 | 7.16
36.54 | 3.36 | 3.46 | 3.55 | 3.65 | 3.74 | 3.84 | 3.93 | 4.03 | 4.12 | 4.22 | 4.31 | 4.41 | 4.5 | 4.6 | 4.69 | 4.78 | 4.88 | 4.97 | 5.07 | 5.16 | 5.26 | 5.35 | 5.45 | 5.54 | 5.64 | 5.73 | 5.83 | 5.92 | 6.02 | 6.11 | 6.21 | 6.3 | 6.4 | 6.49 | 6.59 | 6.68 | 6.78 | 6.87 | 6.97 | 7.06 | 7.16
36.55 | 3.36 | 3.46 | 3.55 | 3.65 | 3.74 | 3.84 | 3.93 | 4.02 | 4.12 | 4.21 | 4.31 | 4.4 | 4.5 | 4.59 | 4.69 | 4.78 | 4.88 | 4.97 | 5.07 | 5.16 | 5.26 | 5.35 | 5.45 | 5.54 | 5.64 | 5.73 | 5.83 | 5.92 | 6.02 | 6.11 | 6.21 | 6.3 | 6.4 | 6.49 | 6.59 | 6.68 | 6.78 | 6.87 | 6.97 | 7.06 | 7.16
36.56 | 3.36 | 3.45 | 3.55 | 3.64 | 3.74 | 3.83 | 3.93 | 4.02 | 4.12 | 4.21 | 4.31 | 4.4 | 4.5 | 4.59 | 4.69 | 4.78 | 4.88 | 4.97 | 5.07 | 5.16 | 5.26 | 5.35 | 5.45 | 5.54 | 5.64 | 5.73 | 5.83 | 5.92 | 6.02 | 6.11 | 6.21 | 6.3 | 6.4 | 6.49 | 6.59 | 6.68 | 6.77 | 6.87 | 6.96 | 7.06 | 7.15
36.57 | 3.36 | 3.45 | 3.55 | 3.64 | 3.74 | 3.83 | 3.93 | 4.02 | 4.12 | 4.21 | 4.31 | 4.4 | 4.5 | 4.59 | 4.69 | 4.78 | 4.88 | 4.97 | 5.07 | 5.16 | 5.26 | 5.35 | 5.45 | 5.54 | 5.64 | 5.73 | 5.83 | 5.92 | 6.02 | 6.11 | 6.2 | 6.3 | 6.39 | 6.49 | 6.58 | 6.68 | 6.77 | 6.87 | 6.96 | 7.06 | 7.15
36.58 | 3.36 | 3.45 | 3.55 | 3.64 | 3.74 | 3.83 | 3.93 | 4.02 | 4.12 | 4.21 | 4.31 | 4.4 | 4.5 | 4.59 | 4.69 | 4.78 | 4.88 | 4.97 | 5.07 | 5.16 | 5.26 | 5.35 | 5.45 | 5.54 | 5.63 | 5.73 | 5.82 | 5.92 | 6.01 | 6.11 | 6.2 | 6.3 | 6.39 | 6.49 | 6.58 | 6.68 | 6.77 | 6.87 | 6.96 | 7.06 | 7.15
36.59 | 3.36 | 3.45 | 3.55 | 3.64 | 3.74 | 3.83 | 3.93 | 4.02 | 4.12 | 4.21 | 4.31 | 4.4 | 4.5 | 4.59 | 4.69 | 4.78 | 4.88 | 4.97 | 5.07 | 5.16 | 5.25 | 5.35 | 5.44 | 5.54 | 5.63 | 5.73 | 5.82 | 5.92 | 6.01 | 6.11 | 6.2 | 6.3 | 6.39 | 6.49 | 6.58 | 6.68 | 6.77 | 6.87 | 6.96 | 7.06 | 7.15
36.6 | 3.36 | 3.45 | 3.55 | 3.64 | 3.74 | 3.83 | 3.93 | 4.02 | 4.12 | 4.21 | 4.31 | 4.4 | 4.5 | 4.59 | 4.68 | 4.78 | 4.87 | 4.97 | 5.06 | 5.16 | 5.25 | 5.35 | 5.44 | 5.54 | 5.63 | 5.73 | 5.82 | 5.92 | 6.01 | 6.11 | 6.2 | 6.3 | 6.39 | 6.49 | 6.58 | 6.68 | 6.77 | 6.87 | 6.96 | 7.05 | 7.15
36.61 | 3.36 | 3.45 | 3.55 | 3.64 | 3.74 | 3.83 | 3.93 | 4.02 | 4.12 | 4.21 | 4.3 | 4.4 | 4.49 | 4.59 | 4.68 | 4.78 | 4.87 | 4.97 | 5.06 | 5.16 | 5.25 | 5.35 | 5.44 | 5.54 | 5.63 | 5.73 | 5.82 | 5.92 | 6.01 | 6.11 | 6.2 | 6.3 | 6.39 | 6.49 | 6.58 | 6.67 | 6.77 | 6.86 | 6.96 | 7.05 | 7.15
36.62 | 3.36 | 3.45 | 3.55 | 3.64 | 3.74 | 3.83 | 3.92 | 4.02 | 4.11 | 4.21 | 4.3 | 4.4 | 4.49 | 4.59 | 4.68 | 4.78 | 4.87 | 4.97 | 5.06 | 5.16 | 5.25 | 5.35 | 5.44 | 5.54 | 5.63 | 5.73 | 5.82 | 5.92 | 6.01 | 6.11 | 6.2 | 6.29 | 6.39 | 6.48 | 6.58 | 6.67 | 6.77 | 6.86 | 6.96 | 7.05 | 7.15
36.63 | 3.36 | 3.45 | 3.55 | 3.64 | 3.73 | 3.83 | 3.92 | 4.02 | 4.11 | 4.21 | 4.3 | 4.4 | 4.49 | 4.59 | 4.68 | 4.78 | 4.87 | 4.97 | 5.06 | 5.16 | 5.25 | 5.35 | 5.44 | 5.54 | 5.63 | 5.72 | 5.82 | 5.91 | 6.01 | 6.1 | 6.2 | 6.29 | 6.39 | 6.48 | 6.58 | 6.67 | 6.77 | 6.86 | 6.96 | 7.05 | 7.15
36.64 | 3.35 | 3.45 | 3.54 | 3.64 | 3.73 | 3.83 | 3.92 | 4.02 | 4.11 | 4.21 | 4.3 | 4.4 | 4.49 | 4.59 | 4.68 | 4.78 | 4.87 | 4.97 | 5.06 | 5.16 | 5.25 | 5.35 | 5.44 | 5.53 | 5.63 | 5.72 | 5.82 | 5.91 | 6.01 | 6.1 | 6.2 | 6.29 | 6.39 | 6.48 | 6.58 | 6.67 | 6.77 | 6.86 | 6.96 | 7.05 | 7.15
36.65 | 3.35 | 3.45 | 3.54 | 3.64 | 3.73 | 3.83 | 3.92 | 4.02 | 4.11 | 4.21 | 4.3 | 4.4 | 4.49 | 4.59 | 4.68 | 4.78 | 4.87 | 4.97 | 5.06 | 5.15 | 5.25 | 5.34 | 5.44 | 5.53 | 5.63 | 5.72 | 5.82 | 5.91 | 6.01 | 6.1 | 6.2 | 6.29 | 6.39 | 6.48 | 6.58 | 6.67 | 6.77 | 6.86 | 6.95 | 7.05 | 7.14
36.66 | 3.35 | 3.45 | 3.54 | 3.64 | 3.73 | 3.83 | 3.92 | 4.02 | 4.11 | 4.21 | 4.3 | 4.4 | 4.49 | 4.59 | 4.68 | 4.77 | 4.87 | 4.96 | 5.06 | 5.15 | 5.25 | 5.34 | 5.44 | 5.53 | 5.63 | 5.72 | 5.82 | 5.91 | 6.01 | 6.1 | 6.2 | 6.29 | 6.39 | 6.48 | 6.57 | 6.67 | 6.76 | 6.86 | 6.95 | 7.05 | 7.14
36.67 | 3.35 | 3.45 | 3.54 | 3.64 | 3.73 | 3.83 | 3.92 | 4.02 | 4.11 | 4.21 | 4.3 | 4.39 | 4.49 | 4.58 | 4.68 | 4.77 | 4.87 | 4.96 | 5.06 | 5.15 | 5.25 | 5.34 | 5.44 | 5.53 | 5.63 | 5.72 | 5.82 | 5.91 | 6.01 | 6.1 | 6.19 | 6.29 | 6.38 | 6.48 | 6.57 | 6.67 | 6.76 | 6.86 | 6.95 | 7.05 | 7.14
36.68 | 3.35 | 3.45 | 3.54 | 3.64 | 3.73 | 3.83 | 3.92 | 4.02 | 4.11 | 4.2 | 4.3 | 4.39 | 4.49 | 4.58 | 4.68 | 4.77 | 4.87 | 4.96 | 5.06 | 5.15 | 5.25 | 5.34 | 5.44 | 5.53 | 5.63 | 5.72 | 5.81 | 5.91 | 6 | 6.1 | 6.19 | 6.29 | 6.38 | 6.48 | 6.57 | 6.67 | 6.76 | 6.86 | 6.95 | 7.05 | 7.14
36.69 | 3.35 | 3.45 | 3.54 | 3.64 | 3.73 | 3.83 | 3.92 | 4.01 | 4.11 | 4.2 | 4.3 | 4.39 | 4.49 | 4.58 | 4.68 | 4.77 | 4.87 | 4.96 | 5.06 | 5.15 | 5.25 | 5.34 | 5.44 | 5.53 | 5.62 | 5.72 | 5.81 | 5.91 | 6 | 6.1 | 6.19 | 6.29 | 6.38 | 6.48 | 6.57 | 6.67 | 6.76 | 6.86 | 6.95 | 7.04 | 7.14
36.7 | 3.35 | 3.45 | 3.54 | 3.63 | 3.73 | 3.82 | 3.92 | 4.01 | 4.11 | 4.2 | 4.3 | 4.39 | 4.49 | 4.58 | 4.68 | 4.77 | 4.87 | 4.96 | 5.06 | 5.15 | 5.24 | 5.34 | 5.43 | 5.53 | 5.62 | 5.72 | 5.81 | 5.91 | 6 | 6.1 | 6.19 | 6.29 | 6.38 | 6.48 | 6.57 | 6.66 | 6.76 | 6.85 | 6.95 | 7.04 | 7.14
36.71 | 3.35 | 3.44 | 3.54 | 3.63 | 3.73 | 3.82 | 3.92 | 4.01 | 4.11 | 4.2 | 4.3 | 4.39 | 4.49 | 4.58 | 4.68 | 4.77 | 4.87 | 4.96 | 5.05 | 5.15 | 5.24 | 5.34 | 5.43 | 5.53 | 5.62 | 5.72 | 5.81 | 5.91 | 6 | 6.1 | 6.19 | 6.29 | 6.38 | 6.47 | 6.57 | 6.66 | 6.76 | 6.85 | 6.95 | 7.04 | 7.14
36.72 | 3.35 | 3.44 | 3.54 | 3.63 | 3.73 | 3.82 | 3.92 | 4.01 | 4.11 | 4.2 | 4.3 | 4.39 | 4.49 | 4.58 | 4.68 | 4.77 | 4.86 | 4.96 | 5.05 | 5.15 | 5.24 | 5.34 | 5.43 | 5.53 | 5.62 | 5.72 | 5.81 | 5.91 | 6 | 6.09 | 6.19 | 6.28 | 6.38 | 6.47 | 6.57 | 6.66 | 6.76 | 6.85 | 6.95 | 7.04 | 7.14
36.73 | 3.35 | 3.44 | 3.54 | 3.63 | 3.73 | 3.82 | 3.92 | 4.01 | 4.11 | 4.2 | 4.3 | 4.39 | 4.48 | 4.58 | 4.67 | 4.77 | 4.86 | 4.96 | 5.05 | 5.15 | 5.24 | 5.34 | 5.43 | 5.53 | 5.62 | 5.72 | 5.81 | 5.9 | 6 | 6.09 | 6.19 | 6.28 | 6.38 | 6.47 | 6.57 | 6.66 | 6.76 | 6.85 | 6.95 | 7.04 | 7.13
36.74 | 3.35 | 3.44 | 3.54 | 3.63 | 3.73 | 3.82 | 3.92 | 4.01 | 4.11 | 4.2 | 4.29 | 4.39 | 4.48 | 4.58 | 4.67 | 4.77 | 4.86 | 4.96 | 5.05 | 5.15 | 5.24 | 5.34 | 5.43 | 5.53 | 5.62 | 5.71 | 5.81 | 5.9 | 6 | 6.09 | 6.19 | 6.28 | 6.38 | 6.47 | 6.57 | 6.66 | 6.76 | 6.85 | 6.94 | 7.04 | 7.13
36.75 | 3.35 | 3.44 | 3.54 | 3.63 | 3.73 | 3.82 | 3.92 | 4.01 | 4.1 | 4.2 | 4.29 | 4.39 | 4.48 | 4.58 | 4.67 | 4.77 | 4.86 | 4.96 | 5.05 | 5.15 | 5.24 | 5.33 | 5.43 | 5.52 | 5.62 | 5.71 | 5.81 | 5.9 | 6 | 6.09 | 6.19 | 6.28 | 6.38 | 6.47 | 6.56 | 6.66 | 6.75 | 6.85 | 6.94 | 7.04 | 7.13
36.76 | 3.35 | 3.44 | 3.54 | 3.63 | 3.73 | 3.82 | 3.91 | 4.01 | 4.1 | 4.2 | 4.29 | 4.39 | 4.48 | 4.58 | 4.67 | 4.77 | 4.86 | 4.96 | 5.05 | 5.14 | 5.24 | 5.33 | 5.43 | 5.52 | 5.62 | 5.71 | 5.81 | 5.9 | 6 | 6.09 | 6.19 | 6.28 | 6.37 | 6.47 | 6.56 | 6.66 | 6.75 | 6.85 | 6.94 | 7.04 | 7.13
36.77 | 3.35 | 3.44 | 3.54 | 3.63 | 3.72 | 3.82 | 3.91 | 4.01 | 4.1 | 4.2 | 4.29 | 4.39 | 4.48 | 4.58 | 4.67 | 4.77 | 4.86 | 4.95 | 5.05 | 5.14 | 5.24 | 5.33 | 5.43 | 5.52 | 5.62 | 5.71 | 5.81 | 5.9 | 6 | 6.09 | 6.18 | 6.28 | 6.37 | 6.47 | 6.56 | 6.66 | 6.75 | 6.85 | 6.94 | 7.04 | 7.13
36.78 | 3.35 | 3.44 | 3.53 | 3.63 | 3.72 | 3.82 | 3.91 | 4.01 | 4.1 | 4.2 | 4.29 | 4.39 | 4.48 | 4.58 | 4.67 | 4.76 | 4.86 | 4.95 | 5.05 | 5.14 | 5.24 | 5.33 | 5.43 | 5.52 | 5.62 | 5.71 | 5.81 | 5.9 | 5.99 | 6.09 | 6.18 | 6.28 | 6.37 | 6.47 | 6.56 | 6.66 | 6.75 | 6.85 | 6.94 | 7.03 | 7.13
36.79 | 3.34 | 3.44 | 3.53 | 3.63 | 3.72 | 3.82 | 3.91 | 4.01 | 4.1 | 4.2 | 4.29 | 4.39 | 4.48 | 4.57 | 4.67 | 4.76 | 4.86 | 4.95 | 5.05 | 5.14 | 5.24 | 5.33 | 5.43 | 5.52 | 5.62 | 5.71 | 5.8 | 5.9 | 5.99 | 6.09 | 6.18 | 6.28 | 6.37 | 6.47 | 6.56 | 6.66 | 6.75 | 6.84 | 6.94 | 7.03 | 7.13
36.8 | 3.34 | 3.44 | 3.53 | 3.63 | 3.72 | 3.82 | 3.91 | 4.01 | 4.1 | 4.2 | 4.29 | 4.38 | 4.48 | 4.57 | 4.67 | 4.76 | 4.86 | 4.95 | 5.05 | 5.14 | 5.24 | 5.33 | 5.42 | 5.52 | 5.61 | 5.71 | 5.8 | 5.9 | 5.99 | 6.09 | 6.18 | 6.28 | 6.37 | 6.47 | 6.56 | 6.65 | 6.75 | 6.84 | 6.94 | 7.03 | 7.13
36.81 | 3.34 | 3.44 | 3.53 | 3.63 | 3.72 | 3.82 | 3.91 | 4.01 | 4.1 | 4.19 | 4.29 | 4.38 | 4.48 | 4.57 | 4.67 | 4.76 | 4.86 | 4.95 | 5.05 | 5.14 | 5.23 | 5.33 | 5.42 | 5.52 | 5.61 | 5.71 | 5.8 | 5.9 | 5.99 | 6.09 | 6.18 | 6.27 | 6.37 | 6.46 | 6.56 | 6.65 | 6.75 | 6.84 | 6.94 | 7.03 | 7.13
36.82 | 3.34 | 3.44 | 3.53 | 3.63 | 3.72 | 3.82 | 3.91 | 4 | 4.1 | 4.19 | 4.29 | 4.38 | 4.48 | 4.57 | 4.67 | 4.76 | 4.86 | 4.95 | 5.04 | 5.14 | 5.23 | 5.33 | 5.42 | 5.52 | 5.61 | 5.71 | 5.8 | 5.9 | 5.99 | 6.08 | 6.18 | 6.27 | 6.37 | 6.46 | 6.56 | 6.65 | 6.75 | 6.84 | 6.94 | 7.03 | 7.12
36.83 | 3.34 | 3.44 | 3.53 | 3.63 | 3.72 | 3.82 | 3.91 | 4 | 4.1 | 4.19 | 4.29 | 4.38 | 4.48 | 4.57 | 4.67 | 4.76 | 4.85 | 4.95 | 5.04 | 5.14 | 5.23 | 5.33 | 5.42 | 5.52 | 5.61 | 5.71 | 5.8 | 5.89 | 5.99 | 6.08 | 6.18 | 6.27 | 6.37 | 6.46 | 6.56 | 6.65 | 6.75 | 6.84 | 6.93 | 7.03 | 7.12
36.84 | 3.34 | 3.44 | 3.53 | 3.63 | 3.72 | 3.81 | 3.91 | 4 | 4.1 | 4.19 | 4.29 | 4.38 | 4.48 | 4.57 | 4.67 | 4.76 | 4.85 | 4.95 | 5.04 | 5.14 | 5.23 | 5.33 | 5.42 | 5.52 | 5.61 | 5.7 | 5.8 | 5.89 | 5.99 | 6.08 | 6.18 | 6.27 | 6.37 | 6.46 | 6.56 | 6.65 | 6.74 | 6.84 | 6.93 | 7.03 | 7.12
36.85 | 3.34 | 3.44 | 3.53 | 3.62 | 3.72 | 3.81 | 3.91 | 4 | 4.1 | 4.19 | 4.29 | 4.38 | 4.48 | 4.57 | 4.66 | 4.76 | 4.85 | 4.95 | 5.04 | 5.14 | 5.23 | 5.33 | 5.42 | 5.51 | 5.61 | 5.7 | 5.8 | 5.89 | 5.99 | 6.08 | 6.18 | 6.27 | 6.37 | 6.46 | 6.55 | 6.65 | 6.74 | 6.84 | 6.93 | 7.03 | 7.12
36.86 | 3.34 | 3.43 | 3.53 | 3.62 | 3.72 | 3.81 | 3.91 | 4 | 4.1 | 4.19 | 4.29 | 4.38 | 4.47 | 4.57 | 4.66 | 4.76 | 4.85 | 4.95 | 5.04 | 5.14 | 5.23 | 5.32 | 5.42 | 5.51 | 5.61 | 5.7 | 5.8 | 5.89 | 5.99 | 6.08 | 6.18 | 6.27 | 6.36 | 6.46 | 6.55 | 6.65 | 6.74 | 6.84 | 6.93 | 7.03 | 7.12
36.87 | 3.34 | 3.43 | 3.53 | 3.62 | 3.72 | 3.81 | 3.91 | 4 | 4.1 | 4.19 | 4.28 | 4.38 | 4.47 | 4.57 | 4.66 | 4.76 | 4.85 | 4.95 | 5.04 | 5.14 | 5.23 | 5.32 | 5.42 | 5.51 | 5.61 | 5.7 | 5.8 | 5.89 | 5.99 | 6.08 | 6.17 | 6.27 | 6.36 | 6.46 | 6.55 | 6.65 | 6.74 | 6.84 | 6.93 | 7.02 | 7.12
36.88 | 3.34 | 3.43 | 3.53 | 3.62 | 3.72 | 3.81 | 3.91 | 4 | 4.09 | 4.19 | 4.28 | 4.38 | 4.47 | 4.57 | 4.66 | 4.76 | 4.85 | 4.95 | 5.04 | 5.13 | 5.23 | 5.32 | 5.42 | 5.51 | 5.61 | 5.7 | 5.8 | 5.89 | 5.98 | 6.08 | 6.17 | 6.27 | 6.36 | 6.46 | 6.55 | 6.65 | 6.74 | 6.83 | 6.93 | 7.02 | 7.12
36.89 | 3.34 | 3.43 | 3.53 | 3.62 | 3.72 | 3.81 | 3.91 | 4 | 4.09 | 4.19 | 4.28 | 4.38 | 4.47 | 4.57 | 4.66 | 4.76 | 4.85 | 4.94 | 5.04 | 5.13 | 5.23 | 5.32 | 5.42 | 5.51 | 5.61 | 5.7 | 5.79 | 5.89 | 5.98 | 6.08 | 6.17 | 6.27 | 6.36 | 6.46 | 6.55 | 6.64 | 6.74 | 6.83 | 6.93 | 7.02 | 7.12
36.9 | 3.34 | 3.43 | 3.53 | 3.62 | 3.72 | 3.81 | 3.9 | 4 | 4.09 | 4.19 | 4.28 | 4.38 | 4.47 | 4.57 | 4.66 | 4.75 | 4.85 | 4.94 | 5.04 | 5.13 | 5.23 | 5.32 | 5.42 | 5.51 | 5.6 | 5.7 | 5.79 | 5.89 | 5.98 | 6.08 | 6.17 | 6.27 | 6.36 | 6.45 | 6.55 | 6.64 | 6.74 | 6.83 | 6.93 | 7.02 | 7.12
36.91 | 3.34 | 3.43 | 3.53 | 3.62 | 3.71 | 3.81 | 3.9 | 4 | 4.09 | 4.19 | 4.28 | 4.38 | 4.47 | 4.56 | 4.66 | 4.75 | 4.85 | 4.94 | 5.04 | 5.13 | 5.23 | 5.32 | 5.41 | 5.51 | 5.6 | 5.7 | 5.79 | 5.89 | 5.98 | 6.08 | 6.17 | 6.26 | 6.36 | 6.45 | 6.55 | 6.64 | 6.74 | 6.83 | 6.93 | 7.02 | 7.11
36.92 | 3.34 | 3.43 | 3.53 | 3.62 | 3.71 | 3.81 | 3.9 | 4 | 4.09 | 4.19 | 4.28 | 4.38 | 4.47 | 4.56 | 4.66 | 4.75 | 4.85 | 4.94 | 5.04 | 5.13 | 5.23 | 5.32 | 5.41 | 5.51 | 5.6 | 5.7 | 5.79 | 5.89 | 5.98 | 6.07 | 6.17 | 6.26 | 6.36 | 6.45 | 6.55 | 6.64 | 6.74 | 6.83 | 6.92 | 7.02 | 7.11
36.93 | 3.34 | 3.43 | 3.52 | 3.62 | 3.71 | 3.81 | 3.9 | 4 | 4.09 | 4.19 | 4.28 | 4.37 | 4.47 | 4.56 | 4.66 | 4.75 | 4.85 | 4.94 | 5.04 | 5.13 | 5.22 | 5.32 | 5.41 | 5.51 | 5.6 | 5.7 | 5.79 | 5.89 | 5.98 | 6.07 | 6.17 | 6.26 | 6.36 | 6.45 | 6.55 | 6.64 | 6.73 | 6.83 | 6.92 | 7.02 | 7.11
36.94 | 3.33 | 3.43 | 3.52 | 3.62 | 3.71 | 3.81 | 3.9 | 4 | 4.09 | 4.18 | 4.28 | 4.37 | 4.47 | 4.56 | 4.66 | 4.75 | 4.85 | 4.94 | 5.03 | 5.13 | 5.22 | 5.32 | 5.41 | 5.51 | 5.6 | 5.7 | 5.79 | 5.88 | 5.98 | 6.07 | 6.17 | 6.26 | 6.36 | 6.45 | 6.54 | 6.64 | 6.73 | 6.83 | 6.92 | 7.02 | 7.11
36.95 | 3.33 | 3.43 | 3.52 | 3.62 | 3.71 | 3.81 | 3.9 | 4 | 4.09 | 4.18 | 4.28 | 4.37 | 4.47 | 4.56 | 4.66 | 4.75 | 4.84 | 4.94 | 5.03 | 5.13 | 5.22 | 5.32 | 5.41 | 5.51 | 5.6 | 5.69 | 5.79 | 5.88 | 5.98 | 6.07 | 6.17 | 6.26 | 6.35 | 6.45 | 6.54 | 6.64 | 6.73 | 6.83 | 6.92 | 7.02 | 7.11
36.96 | 3.33 | 3.43 | 3.52 | 3.62 | 3.71 | 3.81 | 3.9 | 3.99 | 4.09 | 4.18 | 4.28 | 4.37 | 4.47 | 4.56 | 4.66 | 4.75 | 4.84 | 4.94 | 5.03 | 5.13 | 5.22 | 5.32 | 5.41 | 5.5 | 5.6 | 5.69 | 5.79 | 5.88 | 5.98 | 6.07 | 6.17 | 6.26 | 6.35 | 6.45 | 6.54 | 6.64 | 6.73 | 6.83 | 6.92 | 7.01 | 7.11
36.97 | 3.33 | 3.43 | 3.52 | 3.62 | 3.71 | 3.8 | 3.9 | 3.99 | 4.09 | 4.18 | 4.28 | 4.37 | 4.47 | 4.56 | 4.65 | 4.75 | 4.84 | 4.94 | 5.03 | 5.13 | 5.22 | 5.31 | 5.41 | 5.5 | 5.6 | 5.69 | 5.79 | 5.88 | 5.98 | 6.07 | 6.16 | 6.26 | 6.35 | 6.45 | 6.54 | 6.64 | 6.73 | 6.82 | 6.92 | 7.01 | 7.11
36.98 | 3.33 | 3.43 | 3.52 | 3.62 | 3.71 | 3.8 | 3.9 | 3.99 | 4.09 | 4.18 | 4.28 | 4.37 | 4.46 | 4.56 | 4.65 | 4.75 | 4.84 | 4.94 | 5.03 | 5.13 | 5.22 | 5.31 | 5.41 | 5.5 | 5.6 | 5.69 | 5.79 | 5.88 | 5.97 | 6.07 | 6.16 | 6.26 | 6.35 | 6.45 | 6.54 | 6.63 | 6.73 | 6.82 | 6.92 | 7.01 | 7.11
36.99 | 3.33 | 3.43 | 3.52 | 3.61 | 3.71 | 3.8 | 3.9 | 3.99 | 4.09 | 4.18 | 4.28 | 4.37 | 4.46 | 4.56 | 4.65 | 4.75 | 4.84 | 4.94 | 5.03 | 5.12 | 5.22 | 5.31 | 5.41 | 5.5 | 5.6 | 5.69 | 5.78 | 5.88 | 5.97 | 6.07 | 6.16 | 6.26 | 6.35 | 6.45 | 6.54 | 6.63 | 6.73 | 6.82 | 6.92 | 7.01 | 7.11
37 | 3.33 | 3.43 | 3.52 | 3.61 | 3.71 | 3.8 | 3.9 | 3.99 | 4.09 | 4.18 | 4.27 | 4.37 | 4.46 | 4.56 | 4.65 | 4.75 | 4.84 | 4.93 | 5.03 | 5.12 | 5.22 | 5.31 | 5.41 | 5.5 | 5.6 | 5.69 | 5.78 | 5.88 | 5.97 | 6.07 | 6.16 | 6.26 | 6.35 | 6.44 | 6.54 | 6.63 | 6.73 | 6.82 | 6.92 | 7.01 | 7.1
37.01 | 3.33 | 3.42 | 3.52 | 3.61 | 3.71 | 3.8 | 3.9 | 3.99 | 4.09 | 4.18 | 4.27 | 4.37 | 4.46 | 4.56 | 4.65 | 4.75 | 4.84 | 4.93 | 5.03 | 5.12 | 5.22 | 5.31 | 5.41 | 5.5 | 5.59 | 5.69 | 5.78 | 5.88 | 5.97 | 6.07 | 6.16 | 6.25 | 6.35 | 6.44 | 6.54 | 6.63 | 6.73 | 6.82 | 6.91 | 7.01 | 7.1
37.02 | 3.33 | 3.42 | 3.52 | 3.61 | 3.71 | 3.8 | 3.9 | 3.99 | 4.08 | 4.18 | 4.27 | 4.37 | 4.46 | 4.56 | 4.65 | 4.74 | 4.84 | 4.93 | 5.03 | 5.12 | 5.22 | 5.31 | 5.4 | 5.5 | 5.59 | 5.69 | 5.78 | 5.88 | 5.97 | 6.06 | 6.16 | 6.25 | 6.35 | 6.44 | 6.54 | 6.63 | 6.72 | 6.82 | 6.91 | 7.01 | 7.1
37.03 | 3.33 | 3.42 | 3.52 | 3.61 | 3.71 | 3.8 | 3.89 | 3.99 | 4.08 | 4.18 | 4.27 | 4.37 | 4.46 | 4.56 | 4.65 | 4.74 | 4.84 | 4.93 | 5.03 | 5.12 | 5.22 | 5.31 | 5.4 | 5.5 | 5.59 | 5.69 | 5.78 | 5.88 | 5.97 | 6.06 | 6.16 | 6.25 | 6.35 | 6.44 | 6.54 | 6.63 | 6.72 | 6.82 | 6.91 | 7.01 | 7.1
37.04 | 3.33 | 3.42 | 3.52 | 3.61 | 3.71 | 3.8 | 3.89 | 3.99 | 4.08 | 4.18 | 4.27 | 4.37 | 4.46 | 4.55 | 4.65 | 4.74 | 4.84 | 4.93 | 5.03 | 5.12 | 5.21 | 5.31 | 5.4 | 5.5 | 5.59 | 5.69 | 5.78 | 5.87 | 5.97 | 6.06 | 6.16 | 6.25 | 6.35 | 6.44 | 6.53 | 6.63 | 6.72 | 6.82 | 6.91 | 7.01 | 7.1
37.05 | 3.33 | 3.42 | 3.52 | 3.61 | 3.7 | 3.8 | 3.89 | 3.99 | 4.08 | 4.18 | 4.27 | 4.36 | 4.46 | 4.55 | 4.65 | 4.74 | 4.84 | 4.93 | 5.02 | 5.12 | 5.21 | 5.31 | 5.4 | 5.5 | 5.59 | 5.68 | 5.78 | 5.87 | 5.97 | 6.06 | 6.16 | 6.25 | 6.34 | 6.44 | 6.53 | 6.63 | 6.72 | 6.82 | 6.91 | 7 | 7.1
37.06 | 3.33 | 3.42 | 3.52 | 3.61 | 3.7 | 3.8 | 3.89 | 3.99 | 4.08 | 4.18 | 4.27 | 4.36 | 4.46 | 4.55 | 4.65 | 4.74 | 4.84 | 4.93 | 5.02 | 5.12 | 5.21 | 5.31 | 5.4 | 5.5 | 5.59 | 5.68 | 5.78 | 5.87 | 5.97 | 6.06 | 6.16 | 6.25 | 6.34 | 6.44 | 6.53 | 6.63 | 6.72 | 6.81 | 6.91 | 7 | 7.1
37.07 | 3.33 | 3.42 | 3.51 | 3.61 | 3.7 | 3.8 | 3.89 | 3.99 | 4.08 | 4.17 | 4.27 | 4.36 | 4.46 | 4.55 | 4.65 | 4.74 | 4.83 | 4.93 | 5.02 | 5.12 | 5.21 | 5.31 | 5.4 | 5.49 | 5.59 | 5.68 | 5.78 | 5.87 | 5.97 | 6.06 | 6.15 | 6.25 | 6.34 | 6.44 | 6.53 | 6.63 | 6.72 | 6.81 | 6.91 | 7 | 7.1
37.08 | 3.33 | 3.42 | 3.51 | 3.61 | 3.7 | 3.8 | 3.89 | 3.99 | 4.08 | 4.17 | 4.27 | 4.36 | 4.46 | 4.55 | 4.65 | 4.74 | 4.83 | 4.93 | 5.02 | 5.12 | 5.21 | 5.3 | 5.4 | 5.49 | 5.59 | 5.68 | 5.78 | 5.87 | 5.96 | 6.06 | 6.15 | 6.25 | 6.34 | 6.44 | 6.53 | 6.62 | 6.72 | 6.81 | 6.91 | 7 | 7.1
37.09 | 3.33 | 3.42 | 3.51 | 3.61 | 3.7 | 3.8 | 3.89 | 3.98 | 4.08 | 4.17 | 4.27 | 4.36 | 4.46 | 4.55 | 4.64 | 4.74 | 4.83 | 4.93 | 5.02 | 5.12 | 5.21 | 5.3 | 5.4 | 5.49 | 5.59 | 5.68 | 5.78 | 5.87 | 5.96 | 6.06 | 6.15 | 6.25 | 6.34 | 6.43 | 6.53 | 6.62 | 6.72 | 6.81 | 6.91 | 7 | 7.09
37.1 | 3.32 | 3.42 | 3.51 | 3.61 | 3.7 | 3.8 | 3.89 | 3.98 | 4.08 | 4.17 | 4.27 | 4.36 | 4.46 | 4.55 | 4.64 | 4.74 | 4.83 | 4.93 | 5.02 | 5.11 | 5.21 | 5.3 | 5.4 | 5.49 | 5.59 | 5.68 | 5.77 | 5.87 | 5.96 | 6.06 | 6.15 | 6.25 | 6.34 | 6.43 | 6.53 | 6.62 | 6.72 | 6.81 | 6.9 | 7 | 7.09
37.11 | 3.32 | 3.42 | 3.51 | 3.61 | 3.7 | 3.79 | 3.89 | 3.98 | 4.08 | 4.17 | 4.27 | 4.36 | 4.45 | 4.55 | 4.64 | 4.74 | 4.83 | 4.93 | 5.02 | 5.11 | 5.21 | 5.3 | 5.4 | 5.49 | 5.58 | 5.68 | 5.77 | 5.87 | 5.96 | 6.06 | 6.15 | 6.24 | 6.34 | 6.43 | 6.53 | 6.62 | 6.72 | 6.81 | 6.9 | 7 | 7.09
37.12 | 3.32 | 3.42 | 3.51 | 3.61 | 3.7 | 3.79 | 3.89 | 3.98 | 4.08 | 4.17 | 4.27 | 4.36 | 4.45 | 4.55 | 4.64 | 4.74 | 4.83 | 4.92 | 5.02 | 5.11 | 5.21 | 5.3 | 5.4 | 5.49 | 5.58 | 5.68 | 5.77 | 5.87 | 5.96 | 6.05 | 6.15 | 6.24 | 6.34 | 6.43 | 6.53 | 6.62 | 6.71 | 6.81 | 6.9 | 7 | 7.09
37.13 | 3.32 | 3.42 | 3.51 | 3.61 | 3.7 | 3.79 | 3.89 | 3.98 | 4.08 | 4.17 | 4.26 | 4.36 | 4.45 | 4.55 | 4.64 | 4.74 | 4.83 | 4.92 | 5.02 | 5.11 | 5.21 | 5.3 | 5.39 | 5.49 | 5.58 | 5.68 | 5.77 | 5.87 | 5.96 | 6.05 | 6.15 | 6.24 | 6.34 | 6.43 | 6.52 | 6.62 | 6.71 | 6.81 | 6.9 | 7 | 7.09
37.14 | 3.32 | 3.42 | 3.51 | 3.6 | 3.7 | 3.79 | 3.89 | 3.98 | 4.08 | 4.17 | 4.26 | 4.36 | 4.45 | 4.55 | 4.64 | 4.73 | 4.83 | 4.92 | 5.02 | 5.11 | 5.21 | 5.3 | 5.39 | 5.49 | 5.58 | 5.68 | 5.77 | 5.86 | 5.96 | 6.05 | 6.15 | 6.24 | 6.34 | 6.43 | 6.52 | 6.62 | 6.71 | 6.81 | 6.9 | 6.99 | 7.09
37.15 | 3.32 | 3.42 | 3.51 | 3.6 | 3.7 | 3.79 | 3.89 | 3.98 | 4.07 | 4.17 | 4.26 | 4.36 | 4.45 | 4.55 | 4.64 | 4.73 | 4.83 | 4.92 | 5.02 | 5.11 | 5.2 | 5.3 | 5.39 | 5.49 | 5.58 | 5.68 | 5.77 | 5.86 | 5.96 | 6.05 | 6.15 | 6.24 | 6.33 | 6.43 | 6.52 | 6.62 | 6.71 | 6.8 | 6.9 | 6.99 | 7.09
37.16 | 3.32 | 3.41 | 3.51 | 3.6 | 3.7 | 3.79 | 3.89 | 3.98 | 4.07 | 4.17 | 4.26 | 4.36 | 4.45 | 4.54 | 4.64 | 4.73 | 4.83 | 4.92 | 5.02 | 5.11 | 5.2 | 5.3 | 5.39 | 5.49 | 5.58 | 5.67 | 5.77 | 5.86 | 5.96 | 6.05 | 6.14 | 6.24 | 6.33 | 6.43 | 6.52 | 6.62 | 6.71 | 6.8 | 6.9 | 6.99 | 7.09
37.17 | 3.32 | 3.41 | 3.51 | 3.6 | 3.7 | 3.79 | 3.88 | 3.98 | 4.07 | 4.17 | 4.26 | 4.36 | 4.45 | 4.54 | 4.64 | 4.73 | 4.83 | 4.92 | 5.01 | 5.11 | 5.2 | 5.3 | 5.39 | 5.49 | 5.58 | 5.67 | 5.77 | 5.86 | 5.96 | 6.05 | 6.14 | 6.24 | 6.33 | 6.43 | 6.52 | 6.61 | 6.71 | 6.8 | 6.9 | 6.99 | 7.09
37.18 | 3.32 | 3.41 | 3.51 | 3.6 | 3.7 | 3.79 | 3.88 | 3.98 | 4.07 | 4.17 | 4.26 | 4.35 | 4.45 | 4.54 | 4.64 | 4.73 | 4.83 | 4.92 | 5.01 | 5.11 | 5.2 | 5.3 | 5.39 | 5.48 | 5.58 | 5.67 | 5.77 | 5.86 | 5.95 | 6.05 | 6.14 | 6.24 | 6.33 | 6.43 | 6.52 | 6.61 | 6.71 | 6.8 | 6.9 | 6.99 | 7.08
37.19 | 3.32 | 3.41 | 3.51 | 3.6 | 3.7 | 3.79 | 3.88 | 3.98 | 4.07 | 4.17 | 4.26 | 4.35 | 4.45 | 4.54 | 4.64 | 4.73 | 4.82 | 4.92 | 5.01 | 5.11 | 5.2 | 5.29 | 5.39 | 5.48 | 5.58 | 5.67 | 5.77 | 5.86 | 5.95 | 6.05 | 6.14 | 6.24 | 6.33 | 6.42 | 6.52 | 6.61 | 6.71 | 6.8 | 6.89 | 6.99 | 7.08
37.2 | 3.32 | 3.41 | 3.51 | 3.6 | 3.69 | 3.79 | 3.88 | 3.98 | 4.07 | 4.16 | 4.26 | 4.35 | 4.45 | 4.54 | 4.64 | 4.73 | 4.82 | 4.92 | 5.01 | 5.11 | 5.2 | 5.29 | 5.39 | 5.48 | 5.58 | 5.67 | 5.76 | 5.86 | 5.95 | 6.05 | 6.14 | 6.23 | 6.33 | 6.42 | 6.52 | 6.61 | 6.71 | 6.8 | 6.89 | 6.99 | 7.08
37.21 | 3.32 | 3.41 | 3.51 | 3.6 | 3.69 | 3.79 | 3.88 | 3.98 | 4.07 | 4.16 | 4.26 | 4.35 | 4.45 | 4.54 | 4.63 | 4.73 | 4.82 | 4.92 | 5.01 | 5.11 | 5.2 | 5.29 | 5.39 | 5.48 | 5.58 | 5.67 | 5.76 | 5.86 | 5.95 | 6.05 | 6.14 | 6.23 | 6.33 | 6.42 | 6.52 | 6.61 | 6.7 | 6.8 | 6.89 | 6.99 | 7.08
37.22 | 3.32 | 3.41 | 3.5 | 3.6 | 3.69 | 3.79 | 3.88 | 3.98 | 4.07 | 4.16 | 4.26 | 4.35 | 4.45 | 4.54 | 4.63 | 4.73 | 4.82 | 4.92 | 5.01 | 5.1 | 5.2 | 5.29 | 5.39 | 5.48 | 5.57 | 5.67 | 5.76 | 5.86 | 5.95 | 6.04 | 6.14 | 6.23 | 6.33 | 6.42 | 6.52 | 6.61 | 6.7 | 6.8 | 6.89 | 6.99 | 7.08
37.23 | 3.32 | 3.41 | 3.5 | 3.6 | 3.69 | 3.79 | 3.88 | 3.97 | 4.07 | 4.16 | 4.26 | 4.35 | 4.44 | 4.54 | 4.63 | 4.73 | 4.82 | 4.92 | 5.01 | 5.1 | 5.2 | 5.29 | 5.39 | 5.48 | 5.57 | 5.67 | 5.76 | 5.86 | 5.95 | 6.04 | 6.14 | 6.23 | 6.33 | 6.42 | 6.51 | 6.61 | 6.7 | 6.8 | 6.89 | 6.98 | 7.08
37.24 | 3.32 | 3.41 | 3.5 | 3.6 | 3.69 | 3.79 | 3.88 | 3.97 | 4.07 | 4.16 | 4.26 | 4.35 | 4.44 | 4.54 | 4.63 | 4.73 | 4.82 | 4.91 | 5.01 | 5.1 | 5.2 | 5.29 | 5.38 | 5.48 | 5.57 | 5.67 | 5.76 | 5.85 | 5.95 | 6.04 | 6.14 | 6.23 | 6.32 | 6.42 | 6.51 | 6.61 | 6.7 | 6.8 | 6.89 | 6.98 | 7.08
37.25 | 3.31 | 3.41 | 3.5 | 3.6 | 3.69 | 3.78 | 3.88 | 3.97 | 4.07 | 4.16 | 4.26 | 4.35 | 4.44 | 4.54 | 4.63 | 4.73 | 4.82 | 4.91 | 5.01 | 5.1 | 5.2 | 5.29 | 5.38 | 5.48 | 5.57 | 5.67 | 5.76 | 5.85 | 5.95 | 6.04 | 6.14 | 6.23 | 6.32 | 6.42 | 6.51 | 6.61 | 6.7 | 6.79 | 6.89 | 6.98 | 7.08
37.26 | 3.31 | 3.41 | 3.5 | 3.6 | 3.69 | 3.78 | 3.88 | 3.97 | 4.07 | 4.16 | 4.25 | 4.35 | 4.44 | 4.54 | 4.63 | 4.72 | 4.82 | 4.91 | 5.01 | 5.1 | 5.19 | 5.29 | 5.38 | 5.48 | 5.57 | 5.66 | 5.76 | 5.85 | 5.95 | 6.04 | 6.13 | 6.23 | 6.32 | 6.42 | 6.51 | 6.6 | 6.7 | 6.79 | 6.89 | 6.98 | 7.07
37.27 | 3.31 | 3.41 | 3.5 | 3.6 | 3.69 | 3.78 | 3.88 | 3.97 | 4.07 | 4.16 | 4.25 | 4.35 | 4.44 | 4.54 | 4.63 | 4.72 | 4.82 | 4.91 | 5.01 | 5.1 | 5.19 | 5.29 | 5.38 | 5.48 | 5.57 | 5.66 | 5.76 | 5.85 | 5.95 | 6.04 | 6.13 | 6.23 | 6.32 | 6.42 | 6.51 | 6.6 | 6.7 | 6.79 | 6.89 | 6.98 | 7.07
37.28 | 3.31 | 3.41 | 3.5 | 3.59 | 3.69 | 3.78 | 3.88 | 3.97 | 4.06 | 4.16 | 4.25 | 4.35 | 4.44 | 4.53 | 4.63 | 4.72 | 4.82 | 4.91 | 5 | 5.1 | 5.19 | 5.29 | 5.38 | 5.47 | 5.57 | 5.66 | 5.76 | 5.85 | 5.94 | 6.04 | 6.13 | 6.23 | 6.32 | 6.41 | 6.51 | 6.6 | 6.7 | 6.79 | 6.88 | 6.98 | 7.07
37.29 | 3.31 | 3.41 | 3.5 | 3.59 | 3.69 | 3.78 | 3.88 | 3.97 | 4.06 | 4.16 | 4.25 | 4.35 | 4.44 | 4.53 | 4.63 | 4.72 | 4.82 | 4.91 | 5 | 5.1 | 5.19 | 5.29 | 5.38 | 5.47 | 5.57 | 5.66 | 5.76 | 5.85 | 5.94 | 6.04 | 6.13 | 6.23 | 6.32 | 6.41 | 6.51 | 6.6 | 6.7 | 6.79 | 6.88 | 6.98 | 7.07
37.3 | 3.31 | 3.41 | 3.5 | 3.59 | 3.69 | 3.78 | 3.88 | 3.97 | 4.06 | 4.16 | 4.25 | 4.35 | 4.44 | 4.53 | 4.63 | 4.72 | 4.82 | 4.91 | 5 | 5.1 | 5.19 | 5.29 | 5.38 | 5.47 | 5.57 | 5.66 | 5.75 | 5.85 | 5.94 | 6.04 | 6.13 | 6.22 | 6.32 | 6.41 | 6.51 | 6.6 | 6.69 | 6.79 | 6.88 | 6.98 | 7.07
37.31 | 3.31 | 3.4 | 3.5 | 3.59 | 3.69 | 3.78 | 3.87 | 3.97 | 4.06 | 4.16 | 4.25 | 4.34 | 4.44 | 4.53 | 4.63 | 4.72 | 4.81 | 4.91 | 5 | 5.1 | 5.19 | 5.28 | 5.38 | 5.47 | 5.57 | 5.66 | 5.75 | 5.85 | 5.94 | 6.04 | 6.13 | 6.22 | 6.32 | 6.41 | 6.51 | 6.6 | 6.69 | 6.79 | 6.88 | 6.98 | 7.07
37.32 | 3.31 | 3.4 | 3.5 | 3.59 | 3.69 | 3.78 | 3.87 | 3.97 | 4.06 | 4.16 | 4.25 | 4.34 | 4.44 | 4.53 | 4.63 | 4.72 | 4.81 | 4.91 | 5 | 5.1 | 5.19 | 5.28 | 5.38 | 5.47 | 5.57 | 5.66 | 5.75 | 5.85 | 5.94 | 6.03 | 6.13 | 6.22 | 6.32 | 6.41 | 6.5 | 6.6 | 6.69 | 6.79 | 6.88 | 6.97 | 7.07
37.33 | 3.31 | 3.4 | 3.5 | 3.59 | 3.69 | 3.78 | 3.87 | 3.97 | 4.06 | 4.15 | 4.25 | 4.34 | 4.44 | 4.53 | 4.62 | 4.72 | 4.81 | 4.91 | 5 | 5.09 | 5.19 | 5.28 | 5.38 | 5.47 | 5.56 | 5.66 | 5.75 | 5.85 | 5.94 | 6.03 | 6.13 | 6.22 | 6.32 | 6.41 | 6.5 | 6.6 | 6.69 | 6.79 | 6.88 | 6.97 | 7.07
37.34 | 3.31 | 3.4 | 3.5 | 3.59 | 3.68 | 3.78 | 3.87 | 3.97 | 4.06 | 4.15 | 4.25 | 4.34 | 4.44 | 4.53 | 4.62 | 4.72 | 4.81 | 4.91 | 5 | 5.09 | 5.19 | 5.28 | 5.38 | 5.47 | 5.56 | 5.66 | 5.75 | 5.85 | 5.94 | 6.03 | 6.13 | 6.22 | 6.31 | 6.41 | 6.5 | 6.6 | 6.69 | 6.78 | 6.88 | 6.97 | 7.07
37.35 | 3.31 | 3.4 | 3.5 | 3.59 | 3.68 | 3.78 | 3.87 | 3.97 | 4.06 | 4.15 | 4.25 | 4.34 | 4.44 | 4.53 | 4.62 | 4.72 | 4.81 | 4.9 | 5 | 5.09 | 5.19 | 5.28 | 5.37 | 5.47 | 5.56 | 5.66 | 5.75 | 5.84 | 5.94 | 6.03 | 6.13 | 6.22 | 6.31 | 6.41 | 6.5 | 6.6 | 6.69 | 6.78 | 6.88 | 6.97 | 7.06
37.36 | 3.31 | 3.4 | 3.5 | 3.59 | 3.68 | 3.78 | 3.87 | 3.96 | 4.06 | 4.15 | 4.25 | 4.34 | 4.43 | 4.53 | 4.62 | 4.72 | 4.81 | 4.9 | 5 | 5.09 | 5.19 | 5.28 | 5.37 | 5.47 | 5.56 | 5.66 | 5.75 | 5.84 | 5.94 | 6.03 | 6.12 | 6.22 | 6.31 | 6.41 | 6.5 | 6.59 | 6.69 | 6.78 | 6.88 | 6.97 | 7.06
37.37 | 3.31 | 3.4 | 3.49 | 3.59 | 3.68 | 3.78 | 3.87 | 3.96 | 4.06 | 4.15 | 4.25 | 4.34 | 4.43 | 4.53 | 4.62 | 4.72 | 4.81 | 4.9 | 5 | 5.09 | 5.18 | 5.28 | 5.37 | 5.47 | 5.56 | 5.65 | 5.75 | 5.84 | 5.94 | 6.03 | 6.12 | 6.22 | 6.31 | 6.41 | 6.5 | 6.59 | 6.69 | 6.78 | 6.87 | 6.97 | 7.06
37.38 | 3.31 | 3.4 | 3.49 | 3.59 | 3.68 | 3.78 | 3.87 | 3.96 | 4.06 | 4.15 | 4.25 | 4.34 | 4.43 | 4.53 | 4.62 | 4.71 | 4.81 | 4.9 | 5 | 5.09 | 5.18 | 5.28 | 5.37 | 5.47 | 5.56 | 5.65 | 5.75 | 5.84 | 5.94 | 6.03 | 6.12 | 6.22 | 6.31 | 6.4 | 6.5 | 6.59 | 6.69 | 6.78 | 6.87 | 6.97 | 7.06
37.39 | 3.31 | 3.4 | 3.49 | 3.59 | 3.68 | 3.77 | 3.87 | 3.96 | 4.06 | 4.15 | 4.24 | 4.34 | 4.43 | 4.53 | 4.62 | 4.71 | 4.81 | 4.9 | 5 | 5.09 | 5.18 | 5.28 | 5.37 | 5.46 | 5.56 | 5.65 | 5.75 | 5.84 | 5.93 | 6.03 | 6.12 | 6.22 | 6.31 | 6.4 | 6.5 | 6.59 | 6.68 | 6.78 | 6.87 | 6.97 | 7.06
37.4 | 3.3 | 3.4 | 3.49 | 3.59 | 3.68 | 3.77 | 3.87 | 3.96 | 4.06 | 4.15 | 4.24 | 4.34 | 4.43 | 4.53 | 4.62 | 4.71 | 4.81 | 4.9 | 4.99 | 5.09 | 5.18 | 5.28 | 5.37 | 5.46 | 5.56 | 5.65 | 5.75 | 5.84 | 5.93 | 6.03 | 6.12 | 6.21 | 6.31 | 6.4 | 6.5 | 6.59 | 6.68 | 6.78 | 6.87 | 6.97 | 7.06
37.41 | 3.3 | 3.4 | 3.49 | 3.59 | 3.68 | 3.77 | 3.87 | 3.96 | 4.06 | 4.15 | 4.24 | 4.34 | 4.43 | 4.52 | 4.62 | 4.71 | 4.81 | 4.9 | 4.99 | 5.09 | 5.18 | 5.28 | 5.37 | 5.46 | 5.56 | 5.65 | 5.74 | 5.84 | 5.93 | 6.03 | 6.12 | 6.21 | 6.31 | 6.4 | 6.5 | 6.59 | 6.68 | 6.78 | 6.87 | 6.96 | 7.06
37.42 | 3.3 | 3.4 | 3.49 | 3.58 | 3.68 | 3.77 | 3.87 | 3.96 | 4.05 | 4.15 | 4.24 | 4.34 | 4.43 | 4.52 | 4.62 | 4.71 | 4.81 | 4.9 | 4.99 | 5.09 | 5.18 | 5.27 | 5.37 | 5.46 | 5.56 | 5.65 | 5.74 | 5.84 | 5.93 | 6.02 | 6.12 | 6.21 | 6.31 | 6.4 | 6.49 | 6.59 | 6.68 | 6.78 | 6.87 | 6.96 | 7.06
37.43 | 3.3 | 3.4 | 3.49 | 3.58 | 3.68 | 3.77 | 3.87 | 3.96 | 4.05 | 4.15 | 4.24 | 4.34 | 4.43 | 4.52 | 4.62 | 4.71 | 4.8 | 4.9 | 4.99 | 5.09 | 5.18 | 5.27 | 5.37 | 5.46 | 5.55 | 5.65 | 5.74 | 5.84 | 5.93 | 6.02 | 6.12 | 6.21 | 6.31 | 6.4 | 6.49 | 6.59 | 6.68 | 6.77 | 6.87 | 6.96 | 7.06
37.44 | 3.3 | 3.4 | 3.49 | 3.58 | 3.68 | 3.77 | 3.87 | 3.96 | 4.05 | 4.15 | 4.24 | 4.33 | 4.43 | 4.52 | 4.62 | 4.71 | 4.8 | 4.9 | 4.99 | 5.08 | 5.18 | 5.27 | 5.37 | 5.46 | 5.55 | 5.65 | 5.74 | 5.84 | 5.93 | 6.02 | 6.12 | 6.21 | 6.3 | 6.4 | 6.49 | 6.59 | 6.68 | 6.77 | 6.87 | 6.96 | 7.05
37.45 | 3.3 | 3.4 | 3.49 | 3.58 | 3.68 | 3.77 | 3.86 | 3.96 | 4.05 | 4.15 | 4.24 | 4.33 | 4.43 | 4.52 | 4.61 | 4.71 | 4.8 | 4.9 | 4.99 | 5.08 | 5.18 | 5.27 | 5.37 | 5.46 | 5.55 | 5.65 | 5.74 | 5.83 | 5.93 | 6.02 | 6.12 | 6.21 | 6.3 | 6.4 | 6.49 | 6.58 | 6.68 | 6.77 | 6.87 | 6.96 | 7.05
37.46 | 3.3 | 3.39 | 3.49 | 3.58 | 3.68 | 3.77 | 3.86 | 3.96 | 4.05 | 4.15 | 4.24 | 4.33 | 4.43 | 4.52 | 4.61 | 4.71 | 4.8 | 4.9 | 4.99 | 5.08 | 5.18 | 5.27 | 5.36 | 5.46 | 5.55 | 5.65 | 5.74 | 5.83 | 5.93 | 6.02 | 6.11 | 6.21 | 6.3 | 6.4 | 6.49 | 6.58 | 6.68 | 6.77 | 6.86 | 6.96 | 7.05
37.47 | 3.3 | 3.39 | 3.49 | 3.58 | 3.68 | 3.77 | 3.86 | 3.96 | 4.05 | 4.14 | 4.24 | 4.33 | 4.43 | 4.52 | 4.61 | 4.71 | 4.8 | 4.89 | 4.99 | 5.08 | 5.18 | 5.27 | 5.36 | 5.46 | 5.55 | 5.64 | 5.74 | 5.83 | 5.93 | 6.02 | 6.11 | 6.21 | 6.3 | 6.4 | 6.49 | 6.58 | 6.68 | 6.77 | 6.86 | 6.96 | 7.05
37.48 | 3.3 | 3.39 | 3.49 | 3.58 | 3.67 | 3.77 | 3.86 | 3.96 | 4.05 | 4.14 | 4.24 | 4.33 | 4.42 | 4.52 | 4.61 | 4.71 | 4.8 | 4.89 | 4.99 | 5.08 | 5.18 | 5.27 | 5.36 | 5.46 | 5.55 | 5.64 | 5.74 | 5.83 | 5.93 | 6.02 | 6.11 | 6.21 | 6.3 | 6.39 | 6.49 | 6.58 | 6.68 | 6.77 | 6.86 | 6.96 | 7.05
37.49 | 3.3 | 3.39 | 3.49 | 3.58 | 3.67 | 3.77 | 3.86 | 3.96 | 4.05 | 4.14 | 4.24 | 4.33 | 4.42 | 4.52 | 4.61 | 4.71 | 4.8 | 4.89 | 4.99 | 5.08 | 5.17 | 5.27 | 5.36 | 5.46 | 5.55 | 5.64 | 5.74 | 5.83 | 5.92 | 6.02 | 6.11 | 6.21 | 6.3 | 6.39 | 6.49 | 6.58 | 6.67 | 6.77 | 6.86 | 6.96 | 7.05
37.5 | 3.3 | 3.39 | 3.49 | 3.58 | 3.67 | 3.77 | 3.86 | 3.95 | 4.05 | 4.14 | 4.24 | 4.33 | 4.42 | 4.52 | 4.61 | 4.7 | 4.8 | 4.89 | 4.99 | 5.08 | 5.17 | 5.27 | 5.36 | 5.45 | 5.55 | 5.64 | 5.74 | 5.83 | 5.92 | 6.02 | 6.11 | 6.2 | 6.3 | 6.39 | 6.49 | 6.58 | 6.67 | 6.77 | 6.86 | 6.95 | 7.05
37.51 | 3.3 | 3.39 | 3.49 | 3.58 | 3.67 | 3.77 | 3.86 | 3.95 | 4.05 | 4.14 | 4.24 | 4.33 | 4.42 | 4.52 | 4.61 | 4.7 | 4.8 | 4.89 | 4.98 | 5.08 | 5.17 | 5.27 | 5.36 | 5.45 | 5.55 | 5.64 | 5.73 | 5.83 | 5.92 | 6.02 | 6.11 | 6.2 | 6.3 | 6.39 | 6.48 | 6.58 | 6.67 | 6.77 | 6.86 | 6.95 | 7.05
37.52 | 3.3 | 3.39 | 3.48 | 3.58 | 3.67 | 3.77 | 3.86 | 3.95 | 4.05 | 4.14 | 4.23 | 4.33 | 4.42 | 4.52 | 4.61 | 4.7 | 4.8 | 4.89 | 4.98 | 5.08 | 5.17 | 5.27 | 5.36 | 5.45 | 5.55 | 5.64 | 5.73 | 5.83 | 5.92 | 6.01 | 6.11 | 6.2 | 6.3 | 6.39 | 6.48 | 6.58 | 6.67 | 6.76 | 6.86 | 6.95 | 7.05
37.53 | 3.3 | 3.39 | 3.48 | 3.58 | 3.67 | 3.76 | 3.86 | 3.95 | 4.05 | 4.14 | 4.23 | 4.33 | 4.42 | 4.51 | 4.61 | 4.7 | 4.8 | 4.89 | 4.98 | 5.08 | 5.17 | 5.26 | 5.36 | 5.45 | 5.55 | 5.64 | 5.73 | 5.83 | 5.92 | 6.01 | 6.11 | 6.2 | 6.3 | 6.39 | 6.48 | 6.58 | 6.67 | 6.76 | 6.86 | 6.95 | 7.04
37.54 | 3.3 | 3.39 | 3.48 | 3.58 | 3.67 | 3.76 | 3.86 | 3.95 | 4.05 | 4.14 | 4.23 | 4.33 | 4.42 | 4.51 | 4.61 | 4.7 | 4.79 | 4.89 | 4.98 | 5.08 | 5.17 | 5.26 | 5.36 | 5.45 | 5.54 | 5.64 | 5.73 | 5.83 | 5.92 | 6.01 | 6.11 | 6.2 | 6.29 | 6.39 | 6.48 | 6.58 | 6.67 | 6.76 | 6.86 | 6.95 | 7.04
37.55 | 3.29 | 3.39 | 3.48 | 3.58 | 3.67 | 3.76 | 3.86 | 3.95 | 4.04 | 4.14 | 4.23 | 4.33 | 4.42 | 4.51 | 4.61 | 4.7 | 4.79 | 4.89 | 4.98 | 5.08 | 5.17 | 5.26 | 5.36 | 5.45 | 5.54 | 5.64 | 5.73 | 5.82 | 5.92 | 6.01 | 6.11 | 6.2 | 6.29 | 6.39 | 6.48 | 6.57 | 6.67 | 6.76 | 6.86 | 6.95 | 7.04
37.56 | 3.29 | 3.39 | 3.48 | 3.58 | 3.67 | 3.76 | 3.86 | 3.95 | 4.04 | 4.14 | 4.23 | 4.32 | 4.42 | 4.51 | 4.61 | 4.7 | 4.79 | 4.89 | 4.98 | 5.07 | 5.17 | 5.26 | 5.36 | 5.45 | 5.54 | 5.64 | 5.73 | 5.82 | 5.92 | 6.01 | 6.1 | 6.2 | 6.29 | 6.39 | 6.48 | 6.57 | 6.67 | 6.76 | 6.85 | 6.95 | 7.04
37.57 | 3.29 | 3.39 | 3.48 | 3.57 | 3.67 | 3.76 | 3.86 | 3.95 | 4.04 | 4.14 | 4.23 | 4.32 | 4.42 | 4.51 | 4.61 | 4.7 | 4.79 | 4.89 | 4.98 | 5.07 | 5.17 | 5.26 | 5.35 | 5.45 | 5.54 | 5.64 | 5.73 | 5.82 | 5.92 | 6.01 | 6.1 | 6.2 | 6.29 | 6.38 | 6.48 | 6.57 | 6.67 | 6.76 | 6.85 | 6.95 | 7.04
37.58 | 3.29 | 3.39 | 3.48 | 3.57 | 3.67 | 3.76 | 3.85 | 3.95 | 4.04 | 4.14 | 4.23 | 4.32 | 4.42 | 4.51 | 4.6 | 4.7 | 4.79 | 4.89 | 4.98 | 5.07 | 5.17 | 5.26 | 5.35 | 5.45 | 5.54 | 5.63 | 5.73 | 5.82 | 5.92 | 6.01 | 6.1 | 6.2 | 6.29 | 6.38 | 6.48 | 6.57 | 6.66 | 6.76 | 6.85 | 6.95 | 7.04
37.59 | 3.29 | 3.39 | 3.48 | 3.57 | 3.67 | 3.76 | 3.85 | 3.95 | 4.04 | 4.14 | 4.23 | 4.32 | 4.42 | 4.51 | 4.6 | 4.7 | 4.79 | 4.88 | 4.98 | 5.07 | 5.17 | 5.26 | 5.35 | 5.45 | 5.54 | 5.63 | 5.73 | 5.82 | 5.91 | 6.01 | 6.1 | 6.2 | 6.29 | 6.38 | 6.48 | 6.57 | 6.66 | 6.76 | 6.85 | 6.94 | 7.04
37.6 | 3.29 | 3.39 | 3.48 | 3.57 | 3.67 | 3.76 | 3.85 | 3.95 | 4.04 | 4.13 | 4.23 | 4.32 | 4.42 | 4.51 | 4.6 | 4.7 | 4.79 | 4.88 | 4.98 | 5.07 | 5.16 | 5.26 | 5.35 | 5.45 | 5.54 | 5.63 | 5.73 | 5.82 | 5.91 | 6.01 | 6.1 | 6.19 | 6.29 | 6.38 | 6.48 | 6.57 | 6.66 | 6.76 | 6.85 | 6.94 | 7.04
37.61 | 3.29 | 3.38 | 3.48 | 3.57 | 3.67 | 3.76 | 3.85 | 3.95 | 4.04 | 4.13 | 4.23 | 4.32 | 4.41 | 4.51 | 4.6 | 4.7 | 4.79 | 4.88 | 4.98 | 5.07 | 5.16 | 5.26 | 5.35 | 5.44 | 5.54 | 5.63 | 5.73 | 5.82 | 5.91 | 6.01 | 6.1 | 6.19 | 6.29 | 6.38 | 6.47 | 6.57 | 6.66 | 6.75 | 6.85 | 6.94 | 7.04
37.62 | 3.29 | 3.38 | 3.48 | 3.57 | 3.66 | 3.76 | 3.85 | 3.95 | 4.04 | 4.13 | 4.23 | 4.32 | 4.41 | 4.51 | 4.6 | 4.69 | 4.79 | 4.88 | 4.98 | 5.07 | 5.16 | 5.26 | 5.35 | 5.44 | 5.54 | 5.63 | 5.72 | 5.82 | 5.91 | 6.01 | 6.1 | 6.19 | 6.29 | 6.38 | 6.47 | 6.57 | 6.66 | 6.75 | 6.85 | 6.94 | 7.03
37.63 | 3.29 | 3.38 | 3.48 | 3.57 | 3.66 | 3.76 | 3.85 | 3.94 | 4.04 | 4.13 | 4.23 | 4.32 | 4.41 | 4.51 | 4.6 | 4.69 | 4.79 | 4.88 | 4.97 | 5.07 | 5.16 | 5.26 | 5.35 | 5.44 | 5.54 | 5.63 | 5.72 | 5.82 | 5.91 | 6 | 6.1 | 6.19 | 6.28 | 6.38 | 6.47 | 6.57 | 6.66 | 6.75 | 6.85 | 6.94 | 7.03
37.64 | 3.29 | 3.38 | 3.48 | 3.57 | 3.66 | 3.76 | 3.85 | 3.94 | 4.04 | 4.13 | 4.23 | 4.32 | 4.41 | 4.51 | 4.6 | 4.69 | 4.79 | 4.88 | 4.97 | 5.07 | 5.16 | 5.25 | 5.35 | 5.44 | 5.54 | 5.63 | 5.72 | 5.82 | 5.91 | 6 | 6.1 | 6.19 | 6.28 | 6.38 | 6.47 | 6.56 | 6.66 | 6.75 | 6.85 | 6.94 | 7.03
37.65 | 3.29 | 3.38 | 3.48 | 3.57 | 3.66 | 3.76 | 3.85 | 3.94 | 4.04 | 4.13 | 4.22 | 4.32 | 4.41 | 4.51 | 4.6 | 4.69 | 4.79 | 4.88 | 4.97 | 5.07 | 5.16 | 5.25 | 5.35 | 5.44 | 5.53 | 5.63 | 5.72 | 5.82 | 5.91 | 6 | 6.1 | 6.19 | 6.28 | 6.38 | 6.47 | 6.56 | 6.66 | 6.75 | 6.84 | 6.94 | 7.03
37.66 | 3.29 | 3.38 | 3.47 | 3.57 | 3.66 | 3.76 | 3.85 | 3.94 | 4.04 | 4.13 | 4.22 | 4.32 | 4.41 | 4.5 | 4.6 | 4.69 | 4.78 | 4.88 | 4.97 | 5.07 | 5.16 | 5.25 | 5.35 | 5.44 | 5.53 | 5.63 | 5.72 | 5.81 | 5.91 | 6 | 6.09 | 6.19 | 6.28 | 6.38 | 6.47 | 6.56 | 6.66 | 6.75 | 6.84 | 6.94 | 7.03
37.67 | 3.29 | 3.38 | 3.47 | 3.57 | 3.66 | 3.75 | 3.85 | 3.94 | 4.04 | 4.13 | 4.22 | 4.32 | 4.41 | 4.5 | 4.6 | 4.69 | 4.78 | 4.88 | 4.97 | 5.06 | 5.16 | 5.25 | 5.35 | 5.44 | 5.53 | 5.63 | 5.72 | 5.81 | 5.91 | 6 | 6.09 | 6.19 | 6.28 | 6.37 | 6.47 | 6.56 | 6.65 | 6.75 | 6.84 | 6.94 | 7.03
37.68 | 3.29 | 3.38 | 3.47 | 3.57 | 3.66 | 3.75 | 3.85 | 3.94 | 4.03 | 4.13 | 4.22 | 4.32 | 4.41 | 4.5 | 4.6 | 4.69 | 4.78 | 4.88 | 4.97 | 5.06 | 5.16 | 5.25 | 5.34 | 5.44 | 5.53 | 5.63 | 5.72 | 5.81 | 5.91 | 6 | 6.09 | 6.19 | 6.28 | 6.37 | 6.47 | 6.56 | 6.65 | 6.75 | 6.84 | 6.93 | 7.03
37.69 | 3.29 | 3.38 | 3.47 | 3.57 | 3.66 | 3.75 | 3.85 | 3.94 | 4.03 | 4.13 | 4.22 | 4.31 | 4.41 | 4.5 | 4.6 | 4.69 | 4.78 | 4.88 | 4.97 | 5.06 | 5.16 | 5.25 | 5.34 | 5.44 | 5.53 | 5.62 | 5.72 | 5.81 | 5.9 | 6 | 6.09 | 6.19 | 6.28 | 6.37 | 6.47 | 6.56 | 6.65 | 6.75 | 6.84 | 6.93 | 7.03
37.7 | 3.29 | 3.38 | 3.47 | 3.57 | 3.66 | 3.75 | 3.85 | 3.94 | 4.03 | 4.13 | 4.22 | 4.31 | 4.41 | 4.5 | 4.59 | 4.69 | 4.78 | 4.88 | 4.97 | 5.06 | 5.16 | 5.25 | 5.34 | 5.44 | 5.53 | 5.62 | 5.72 | 5.81 | 5.9 | 6 | 6.09 | 6.18 | 6.28 | 6.37 | 6.46 | 6.56 | 6.65 | 6.75 | 6.84 | 6.93 | 7.03
37.71 | 3.28 | 3.38 | 3.47 | 3.57 | 3.66 | 3.75 | 3.85 | 3.94 | 4.03 | 4.13 | 4.22 | 4.31 | 4.41 | 4.5 | 4.59 | 4.69 | 4.78 | 4.87 | 4.97 | 5.06 | 5.15 | 5.25 | 5.34 | 5.44 | 5.53 | 5.62 | 5.72 | 5.81 | 5.9 | 6 | 6.09 | 6.18 | 6.28 | 6.37 | 6.46 | 6.56 | 6.65 | 6.74 | 6.84 | 6.93 | 7.02
37.72 | 3.28 | 3.38 | 3.47 | 3.56 | 3.66 | 3.75 | 3.84 | 3.94 | 4.03 | 4.13 | 4.22 | 4.31 | 4.41 | 4.5 | 4.59 | 4.69 | 4.78 | 4.87 | 4.97 | 5.06 | 5.15 | 5.25 | 5.34 | 5.43 | 5.53 | 5.62 | 5.71 | 5.81 | 5.9 | 6 | 6.09 | 6.18 | 6.28 | 6.37 | 6.46 | 6.56 | 6.65 | 6.74 | 6.84 | 6.93 | 7.02
37.73 | 3.28 | 3.38 | 3.47 | 3.56 | 3.66 | 3.75 | 3.84 | 3.94 | 4.03 | 4.12 | 4.22 | 4.31 | 4.41 | 4.5 | 4.59 | 4.69 | 4.78 | 4.87 | 4.97 | 5.06 | 5.15 | 5.25 | 5.34 | 5.43 | 5.53 | 5.62 | 5.71 | 5.81 | 5.9 | 5.99 | 6.09 | 6.18 | 6.27 | 6.37 | 6.46 | 6.56 | 6.65 | 6.74 | 6.84 | 6.93 | 7.02
37.74 | 3.28 | 3.38 | 3.47 | 3.56 | 3.66 | 3.75 | 3.84 | 3.94 | 4.03 | 4.12 | 4.22 | 4.31 | 4.4 | 4.5 | 4.59 | 4.68 | 4.78 | 4.87 | 4.97 | 5.06 | 5.15 | 5.25 | 5.34 | 5.43 | 5.53 | 5.62 | 5.71 | 5.81 | 5.9 | 5.99 | 6.09 | 6.18 | 6.27 | 6.37 | 6.46 | 6.55 | 6.65 | 6.74 | 6.83 | 6.93 | 7.02
37.75 | 3.28 | 3.38 | 3.47 | 3.56 | 3.66 | 3.75 | 3.84 | 3.94 | 4.03 | 4.12 | 4.22 | 4.31 | 4.4 | 4.5 | 4.59 | 4.68 | 4.78 | 4.87 | 4.96 | 5.06 | 5.15 | 5.24 | 5.34 | 5.43 | 5.53 | 5.62 | 5.71 | 5.81 | 5.9 | 5.99 | 6.09 | 6.18 | 6.27 | 6.37 | 6.46 | 6.55 | 6.65 | 6.74 | 6.83 | 6.93 | 7.02
37.76 | 3.28 | 3.37 | 3.47 | 3.56 | 3.66 | 3.75 | 3.84 | 3.94 | 4.03 | 4.12 | 4.22 | 4.31 | 4.4 | 4.5 | 4.59 | 4.68 | 4.78 | 4.87 | 4.96 | 5.06 | 5.15 | 5.24 | 5.34 | 5.43 | 5.52 | 5.62 | 5.71 | 5.8 | 5.9 | 5.99 | 6.08 | 6.18 | 6.27 | 6.37 | 6.46 | 6.55 | 6.65 | 6.74 | 6.83 | 6.93 | 7.02
37.77 | 3.28 | 3.37 | 3.47 | 3.56 | 3.65 | 3.75 | 3.84 | 3.93 | 4.03 | 4.12 | 4.22 | 4.31 | 4.4 | 4.5 | 4.59 | 4.68 | 4.78 | 4.87 | 4.96 | 5.06 | 5.15 | 5.24 | 5.34 | 5.43 | 5.52 | 5.62 | 5.71 | 5.8 | 5.9 | 5.99 | 6.08 | 6.18 | 6.27 | 6.36 | 6.46 | 6.55 | 6.64 | 6.74 | 6.83 | 6.92 | 7.02
37.78 | 3.28 | 3.37 | 3.47 | 3.56 | 3.65 | 3.75 | 3.84 | 3.93 | 4.03 | 4.12 | 4.21 | 4.31 | 4.4 | 4.49 | 4.59 | 4.68 | 4.77 | 4.87 | 4.96 | 5.06 | 5.15 | 5.24 | 5.34 | 5.43 | 5.52 | 5.62 | 5.71 | 5.8 | 5.9 | 5.99 | 6.08 | 6.18 | 6.27 | 6.36 | 6.46 | 6.55 | 6.64 | 6.74 | 6.83 | 6.92 | 7.02
37.79 | 3.28 | 3.37 | 3.47 | 3.56 | 3.65 | 3.75 | 3.84 | 3.93 | 4.03 | 4.12 | 4.21 | 4.31 | 4.4 | 4.49 | 4.59 | 4.68 | 4.77 | 4.87 | 4.96 | 5.05 | 5.15 | 5.24 | 5.33 | 5.43 | 5.52 | 5.61 | 5.71 | 5.8 | 5.89 | 5.99 | 6.08 | 6.18 | 6.27 | 6.36 | 6.46 | 6.55 | 6.64 | 6.74 | 6.83 | 6.92 | 7.02
37.8 | 3.28 | 3.37 | 3.47 | 3.56 | 3.65 | 3.75 | 3.84 | 3.93 | 4.03 | 4.12 | 4.21 | 4.31 | 4.4 | 4.49 | 4.59 | 4.68 | 4.77 | 4.87 | 4.96 | 5.05 | 5.15 | 5.24 | 5.33 | 5.43 | 5.52 | 5.61 | 5.71 | 5.8 | 5.89 | 5.99 | 6.08 | 6.17 | 6.27 | 6.36 | 6.45 | 6.55 | 6.64 | 6.73 | 6.83 | 6.92 | 7.01
37.81 | 3.28 | 3.37 | 3.46 | 3.56 | 3.65 | 3.74 | 3.84 | 3.93 | 4.03 | 4.12 | 4.21 | 4.31 | 4.4 | 4.49 | 4.59 | 4.68 | 4.77 | 4.87 | 4.96 | 5.05 | 5.15 | 5.24 | 5.33 | 5.43 | 5.52 | 5.61 | 5.71 | 5.8 | 5.89 | 5.99 | 6.08 | 6.17 | 6.27 | 6.36 | 6.45 | 6.55 | 6.64 | 6.73 | 6.83 | 6.92 | 7.01
37.82 | 3.28 | 3.37 | 3.46 | 3.56 | 3.65 | 3.74 | 3.84 | 3.93 | 4.02 | 4.12 | 4.21 | 4.3 | 4.4 | 4.49 | 4.58 | 4.68 | 4.77 | 4.86 | 4.96 | 5.05 | 5.15 | 5.24 | 5.33 | 5.43 | 5.52 | 5.61 | 5.71 | 5.8 | 5.89 | 5.99 | 6.08 | 6.17 | 6.27 | 6.36 | 6.45 | 6.55 | 6.64 | 6.73 | 6.83 | 6.92 | 7.01
37.83 | 3.28 | 3.37 | 3.46 | 3.56 | 3.65 | 3.74 | 3.84 | 3.93 | 4.02 | 4.12 | 4.21 | 4.3 | 4.4 | 4.49 | 4.58 | 4.68 | 4.77 | 4.86 | 4.96 | 5.05 | 5.14 | 5.24 | 5.33 | 5.42 | 5.52 | 5.61 | 5.7 | 5.8 | 5.89 | 5.98 | 6.08 | 6.17 | 6.26 | 6.36 | 6.45 | 6.54 | 6.64 | 6.73 | 6.82 | 6.92 | 7.01
37.84 | 3.28 | 3.37 | 3.46 | 3.56 | 3.65 | 3.74 | 3.84 | 3.93 | 4.02 | 4.12 | 4.21 | 4.3 | 4.4 | 4.49 | 4.58 | 4.68 | 4.77 | 4.86 | 4.96 | 5.05 | 5.14 | 5.24 | 5.33 | 5.42 | 5.52 | 5.61 | 5.7 | 5.8 | 5.89 | 5.98 | 6.08 | 6.17 | 6.26 | 6.36 | 6.45 | 6.54 | 6.64 | 6.73 | 6.82 | 6.92 | 7.01
37.85 | 3.28 | 3.37 | 3.46 | 3.56 | 3.65 | 3.74 | 3.84 | 3.93 | 4.02 | 4.12 | 4.21 | 4.3 | 4.4 | 4.49 | 4.58 | 4.68 | 4.77 | 4.86 | 4.96 | 5.05 | 5.14 | 5.24 | 5.33 | 5.42 | 5.52 | 5.61 | 5.7 | 5.8 | 5.89 | 5.98 | 6.08 | 6.17 | 6.26 | 6.36 | 6.45 | 6.54 | 6.64 | 6.73 | 6.82 | 6.92 | 7.01
37.86 | 3.27 | 3.37 | 3.46 | 3.55 | 3.65 | 3.74 | 3.83 | 3.93 | 4.02 | 4.11 | 4.21 | 4.3 | 4.39 | 4.49 | 4.58 | 4.67 | 4.77 | 4.86 | 4.95 | 5.05 | 5.14 | 5.23 | 5.33 | 5.42 | 5.51 | 5.61 | 5.7 | 5.79 | 5.89 | 5.98 | 6.07 | 6.17 | 6.26 | 6.35 | 6.45 | 6.54 | 6.63 | 6.73 | 6.82 | 6.91 | 7.01
37.87 | 3.27 | 3.37 | 3.46 | 3.55 | 3.65 | 3.74 | 3.83 | 3.93 | 4.02 | 4.11 | 4.21 | 4.3 | 4.39 | 4.49 | 4.58 | 4.67 | 4.77 | 4.86 | 4.95 | 5.05 | 5.14 | 5.23 | 5.33 | 5.42 | 5.51 | 5.61 | 5.7 | 5.79 | 5.89 | 5.98 | 6.07 | 6.17 | 6.26 | 6.35 | 6.45 | 6.54 | 6.63 | 6.73 | 6.82 | 6.91 | 7.01
37.88 | 3.27 | 3.37 | 3.46 | 3.55 | 3.65 | 3.74 | 3.83 | 3.93 | 4.02 | 4.11 | 4.21 | 4.3 | 4.39 | 4.49 | 4.58 | 4.67 | 4.77 | 4.86 | 4.95 | 5.05 | 5.14 | 5.23 | 5.33 | 5.42 | 5.51 | 5.61 | 5.7 | 5.79 | 5.89 | 5.98 | 6.07 | 6.17 | 6.26 | 6.35 | 6.45 | 6.54 | 6.63 | 6.73 | 6.82 | 6.91 | 7.01
37.89 | 3.27 | 3.37 | 3.46 | 3.55 | 3.65 | 3.74 | 3.83 | 3.93 | 4.02 | 4.11 | 4.21 | 4.3 | 4.39 | 4.49 | 4.58 | 4.67 | 4.77 | 4.86 | 4.95 | 5.05 | 5.14 | 5.23 | 5.33 | 5.42 | 5.51 | 5.61 | 5.7 | 5.79 | 5.89 | 5.98 | 6.07 | 6.17 | 6.26 | 6.35 | 6.44 | 6.54 | 6.63 | 6.72 | 6.82 | 6.91 | 7
37.9 | 3.27 | 3.37 | 3.46 | 3.55 | 3.65 | 3.74 | 3.83 | 3.93 | 4.02 | 4.11 | 4.21 | 4.3 | 4.39 | 4.48 | 4.58 | 4.67 | 4.76 | 4.86 | 4.95 | 5.04 | 5.14 | 5.23 | 5.32 | 5.42 | 5.51 | 5.6 | 5.7 | 5.79 | 5.88 | 5.98 | 6.07 | 6.16 | 6.26 | 6.35 | 6.44 | 6.54 | 6.63 | 6.72 | 6.82 | 6.91 | 7
37.91 | 3.27 | 3.36 | 3.46 | 3.55 | 3.64 | 3.74 | 3.83 | 3.92 | 4.02 | 4.11 | 4.2 | 4.3 | 4.39 | 4.48 | 4.58 | 4.67 | 4.76 | 4.86 | 4.95 | 5.04 | 5.14 | 5.23 | 5.32 | 5.42 | 5.51 | 5.6 | 5.7 | 5.79 | 5.88 | 5.98 | 6.07 | 6.16 | 6.26 | 6.35 | 6.44 | 6.54 | 6.63 | 6.72 | 6.82 | 6.91 | 7
37.92 | 3.27 | 3.36 | 3.46 | 3.55 | 3.64 | 3.74 | 3.83 | 3.92 | 4.02 | 4.11 | 4.2 | 4.3 | 4.39 | 4.48 | 4.58 | 4.67 | 4.76 | 4.86 | 4.95 | 5.04 | 5.14 | 5.23 | 5.32 | 5.42 | 5.51 | 5.6 | 5.7 | 5.79 | 5.88 | 5.98 | 6.07 | 6.16 | 6.26 | 6.35 | 6.44 | 6.54 | 6.63 | 6.72 | 6.81 | 6.91 | 7
37.93 | 3.27 | 3.36 | 3.46 | 3.55 | 3.64 | 3.74 | 3.83 | 3.92 | 4.02 | 4.11 | 4.2 | 4.3 | 4.39 | 4.48 | 4.58 | 4.67 | 4.76 | 4.86 | 4.95 | 5.04 | 5.14 | 5.23 | 5.32 | 5.42 | 5.51 | 5.6 | 5.69 | 5.79 | 5.88 | 5.97 | 6.07 | 6.16 | 6.25 | 6.35 | 6.44 | 6.53 | 6.63 | 6.72 | 6.81 | 6.91 | 7
37.94 | 3.27 | 3.36 | 3.46 | 3.55 | 3.64 | 3.74 | 3.83 | 3.92 | 4.02 | 4.11 | 4.2 | 4.3 | 4.39 | 4.48 | 4.58 | 4.67 | 4.76 | 4.85 | 4.95 | 5.04 | 5.13 | 5.23 | 5.32 | 5.41 | 5.51 | 5.6 | 5.69 | 5.79 | 5.88 | 5.97 | 6.07 | 6.16 | 6.25 | 6.35 | 6.44 | 6.53 | 6.63 | 6.72 | 6.81 | 6.91 | 7
37.95 | 3.27 | 3.36 | 3.46 | 3.55 | 3.64 | 3.74 | 3.83 | 3.92 | 4.01 | 4.11 | 4.2 | 4.29 | 4.39 | 4.48 | 4.57 | 4.67 | 4.76 | 4.85 | 4.95 | 5.04 | 5.13 | 5.23 | 5.32 | 5.41 | 5.51 | 5.6 | 5.69 | 5.79 | 5.88 | 5.97 | 6.07 | 6.16 | 6.25 | 6.35 | 6.44 | 6.53 | 6.63 | 6.72 | 6.81 | 6.9 | 7
37.96 | 3.27 | 3.36 | 3.45 | 3.55 | 3.64 | 3.73 | 3.83 | 3.92 | 4.01 | 4.11 | 4.2 | 4.29 | 4.39 | 4.48 | 4.57 | 4.67 | 4.76 | 4.85 | 4.95 | 5.04 | 5.13 | 5.23 | 5.32 | 5.41 | 5.51 | 5.6 | 5.69 | 5.79 | 5.88 | 5.97 | 6.06 | 6.16 | 6.25 | 6.34 | 6.44 | 6.53 | 6.62 | 6.72 | 6.81 | 6.9 | 7
37.97 | 3.27 | 3.36 | 3.45 | 3.55 | 3.64 | 3.73 | 3.83 | 3.92 | 4.01 | 4.11 | 4.2 | 4.29 | 4.39 | 4.48 | 4.57 | 4.67 | 4.76 | 4.85 | 4.95 | 5.04 | 5.13 | 5.23 | 5.32 | 5.41 | 5.5 | 5.6 | 5.69 | 5.78 | 5.88 | 5.97 | 6.06 | 6.16 | 6.25 | 6.34 | 6.44 | 6.53 | 6.62 | 6.72 | 6.81 | 6.9 | 7
37.98 | 3.27 | 3.36 | 3.45 | 3.55 | 3.64 | 3.73 | 3.83 | 3.92 | 4.01 | 4.11 | 4.2 | 4.29 | 4.39 | 4.48 | 4.57 | 4.67 | 4.76 | 4.85 | 4.94 | 5.04 | 5.13 | 5.22 | 5.32 | 5.41 | 5.5 | 5.6 | 5.69 | 5.78 | 5.88 | 5.97 | 6.06 | 6.16 | 6.25 | 6.34 | 6.44 | 6.53 | 6.62 | 6.72 | 6.81 | 6.9 | 6.99
37.99 | 3.27 | 3.36 | 3.45 | 3.55 | 3.64 | 3.73 | 3.83 | 3.92 | 4.01 | 4.11 | 4.2 | 4.29 | 4.38 | 4.48 | 4.57 | 4.66 | 4.76 | 4.85 | 4.94 | 5.04 | 5.13 | 5.22 | 5.32 | 5.41 | 5.5 | 5.6 | 5.69 | 5.78 | 5.88 | 5.97 | 6.06 | 6.16 | 6.25 | 6.34 | 6.43 | 6.53 | 6.62 | 6.71 | 6.81 | 6.9 | 6.99
38 | 3.27 | 3.36 | 3.45 | 3.55 | 3.64 | 3.73 | 3.82 | 3.92 | 4.01 | 4.1 | 4.2 | 4.29 | 4.38 | 4.48 | 4.57 | 4.66 | 4.76 | 4.85 | 4.94 | 5.04 | 5.13 | 5.22 | 5.32 | 5.41 | 5.5 | 5.6 | 5.69 | 5.78 | 5.87 | 5.97 | 6.06 | 6.15 | 6.25 | 6.34 | 6.43 | 6.53 | 6.62 | 6.71 | 6.81 | 6.9 | 6.99
38.01 | 3.26 | 3.36 | 3.45 | 3.54 | 3.64 | 3.73 | 3.82 | 3.92 | 4.01 | 4.1 | 4.2 | 4.29 | 4.38 | 4.48 | 4.57 | 4.66 | 4.76 | 4.85 | 4.94 | 5.04 | 5.13 | 5.22 | 5.31 | 5.41 | 5.5 | 5.59 | 5.69 | 5.78 | 5.87 | 5.97 | 6.06 | 6.15 | 6.25 | 6.34 | 6.43 | 6.53 | 6.62 | 6.71 | 6.81 | 6.9 | 6.99
38.02 | 3.26 | 3.36 | 3.45 | 3.54 | 3.64 | 3.73 | 3.82 | 3.92 | 4.01 | 4.1 | 4.2 | 4.29 | 4.38 | 4.48 | 4.57 | 4.66 | 4.75 | 4.85 | 4.94 | 5.03 | 5.13 | 5.22 | 5.31 | 5.41 | 5.5 | 5.59 | 5.69 | 5.78 | 5.87 | 5.97 | 6.06 | 6.15 | 6.25 | 6.34 | 6.43 | 6.52 | 6.62 | 6.71 | 6.8 | 6.9 | 6.99
38.03 | 3.26 | 3.36 | 3.45 | 3.54 | 3.64 | 3.73 | 3.82 | 3.92 | 4.01 | 4.1 | 4.2 | 4.29 | 4.38 | 4.47 | 4.57 | 4.66 | 4.75 | 4.85 | 4.94 | 5.03 | 5.13 | 5.22 | 5.31 | 5.41 | 5.5 | 5.59 | 5.69 | 5.78 | 5.87 | 5.96 | 6.06 | 6.15 | 6.24 | 6.34 | 6.43 | 6.52 | 6.62 | 6.71 | 6.8 | 6.9 | 6.99
38.04 | 3.26 | 3.36 | 3.45 | 3.54 | 3.64 | 3.73 | 3.82 | 3.91 | 4.01 | 4.1 | 4.19 | 4.29 | 4.38 | 4.47 | 4.57 | 4.66 | 4.75 | 4.85 | 4.94 | 5.03 | 5.13 | 5.22 | 5.31 | 5.41 | 5.5 | 5.59 | 5.68 | 5.78 | 5.87 | 5.96 | 6.06 | 6.15 | 6.24 | 6.34 | 6.43 | 6.52 | 6.62 | 6.71 | 6.8 | 6.89 | 6.99
38.05 | 3.26 | 3.36 | 3.45 | 3.54 | 3.63 | 3.73 | 3.82 | 3.91 | 4.01 | 4.1 | 4.19 | 4.29 | 4.38 | 4.47 | 4.57 | 4.66 | 4.75 | 4.85 | 4.94 | 5.03 | 5.12 | 5.22 | 5.31 | 5.4 | 5.5 | 5.59 | 5.68 | 5.78 | 5.87 | 5.96 | 6.06 | 6.15 | 6.24 | 6.34 | 6.43 | 6.52 | 6.61 | 6.71 | 6.8 | 6.89 | 6.99
38.06 | 3.26 | 3.35 | 3.45 | 3.54 | 3.63 | 3.73 | 3.82 | 3.91 | 4.01 | 4.1 | 4.19 | 4.29 | 4.38 | 4.47 | 4.57 | 4.66 | 4.75 | 4.84 | 4.94 | 5.03 | 5.12 | 5.22 | 5.31 | 5.4 | 5.5 | 5.59 | 5.68 | 5.78 | 5.87 | 5.96 | 6.05 | 6.15 | 6.24 | 6.33 | 6.43 | 6.52 | 6.61 | 6.71 | 6.8 | 6.89 | 6.99
38.07 | 3.26 | 3.35 | 3.45 | 3.54 | 3.63 | 3.73 | 3.82 | 3.91 | 4.01 | 4.1 | 4.19 | 4.29 | 4.38 | 4.47 | 4.56 | 4.66 | 4.75 | 4.84 | 4.94 | 5.03 | 5.12 | 5.22 | 5.31 | 5.4 | 5.5 | 5.59 | 5.68 | 5.77 | 5.87 | 5.96 | 6.05 | 6.15 | 6.24 | 6.33 | 6.43 | 6.52 | 6.61 | 6.71 | 6.8 | 6.89 | 6.98
38.08 | 3.26 | 3.35 | 3.45 | 3.54 | 3.63 | 3.73 | 3.82 | 3.91 | 4.01 | 4.1 | 4.19 | 4.28 | 4.38 | 4.47 | 4.56 | 4.66 | 4.75 | 4.84 | 4.94 | 5.03 | 5.12 | 5.22 | 5.31 | 5.4 | 5.49 | 5.59 | 5.68 | 5.77 | 5.87 | 5.96 | 6.05 | 6.15 | 6.24 | 6.33 | 6.43 | 6.52 | 6.61 | 6.7 | 6.8 | 6.89 | 6.98
38.09 | 3.26 | 3.35 | 3.45 | 3.54 | 3.63 | 3.73 | 3.82 | 3.91 | 4 | 4.1 | 4.19 | 4.28 | 4.38 | 4.47 | 4.56 | 4.66 | 4.75 | 4.84 | 4.94 | 5.03 | 5.12 | 5.21 | 5.31 | 5.4 | 5.49 | 5.59 | 5.68 | 5.77 | 5.87 | 5.96 | 6.05 | 6.15 | 6.24 | 6.33 | 6.42 | 6.52 | 6.61 | 6.7 | 6.8 | 6.89 | 6.98
38.1 | 3.26 | 3.35 | 3.45 | 3.54 | 3.63 | 3.72 | 3.82 | 3.91 | 4 | 4.1 | 4.19 | 4.28 | 4.38 | 4.47 | 4.56 | 4.66 | 4.75 | 4.84 | 4.93 | 5.03 | 5.12 | 5.21 | 5.31 | 5.4 | 5.49 | 5.59 | 5.68 | 5.77 | 5.86 | 5.96 | 6.05 | 6.14 | 6.24 | 6.33 | 6.42 | 6.52 | 6.61 | 6.7 | 6.8 | 6.89 | 6.98
38.11 | 3.26 | 3.35 | 3.44 | 3.54 | 3.63 | 3.72 | 3.82 | 3.91 | 4 | 4.1 | 4.19 | 4.28 | 4.38 | 4.47 | 4.56 | 4.65 | 4.75 | 4.84 | 4.93 | 5.03 | 5.12 | 5.21 | 5.31 | 5.4 | 5.49 | 5.58 | 5.68 | 5.77 | 5.86 | 5.96 | 6.05 | 6.14 | 6.24 | 6.33 | 6.42 | 6.52 | 6.61 | 6.7 | 6.79 | 6.89 | 6.98
38.12 | 3.26 | 3.35 | 3.44 | 3.54 | 3.63 | 3.72 | 3.82 | 3.91 | 4 | 4.1 | 4.19 | 4.28 | 4.37 | 4.47 | 4.56 | 4.65 | 4.75 | 4.84 | 4.93 | 5.03 | 5.12 | 5.21 | 5.3 | 5.4 | 5.49 | 5.58 | 5.68 | 5.77 | 5.86 | 5.96 | 6.05 | 6.14 | 6.24 | 6.33 | 6.42 | 6.51 | 6.61 | 6.7 | 6.79 | 6.89 | 6.98
38.13 | 3.26 | 3.35 | 3.44 | 3.54 | 3.63 | 3.72 | 3.82 | 3.91 | 4 | 4.09 | 4.19 | 4.28 | 4.37 | 4.47 | 4.56 | 4.65 | 4.75 | 4.84 | 4.93 | 5.02 | 5.12 | 5.21 | 5.3 | 5.4 | 5.49 | 5.58 | 5.68 | 5.77 | 5.86 | 5.96 | 6.05 | 6.14 | 6.23 | 6.33 | 6.42 | 6.51 | 6.61 | 6.7 | 6.79 | 6.89 | 6.98
38.14 | 3.26 | 3.35 | 3.44 | 3.54 | 3.63 | 3.72 | 3.81 | 3.91 | 4 | 4.09 | 4.19 | 4.28 | 4.37 | 4.47 | 4.56 | 4.65 | 4.74 | 4.84 | 4.93 | 5.02 | 5.12 | 5.21 | 5.3 | 5.4 | 5.49 | 5.58 | 5.68 | 5.77 | 5.86 | 5.95 | 6.05 | 6.14 | 6.23 | 6.33 | 6.42 | 6.51 | 6.61 | 6.7 | 6.79 | 6.88 | 6.98
38.15 | 3.26 | 3.35 | 3.44 | 3.53 | 3.63 | 3.72 | 3.81 | 3.91 | 4 | 4.09 | 4.19 | 4.28 | 4.37 | 4.47 | 4.56 | 4.65 | 4.74 | 4.84 | 4.93 | 5.02 | 5.12 | 5.21 | 5.3 | 5.4 | 5.49 | 5.58 | 5.67 | 5.77 | 5.86 | 5.95 | 6.05 | 6.14 | 6.23 | 6.33 | 6.42 | 6.51 | 6.6 | 6.7 | 6.79 | 6.88 | 6.98
38.16 | 3.26 | 3.35 | 3.44 | 3.53 | 3.63 | 3.72 | 3.81 | 3.91 | 4 | 4.09 | 4.19 | 4.28 | 4.37 | 4.46 | 4.56 | 4.65 | 4.74 | 4.84 | 4.93 | 5.02 | 5.12 | 5.21 | 5.3 | 5.39 | 5.49 | 5.58 | 5.67 | 5.77 | 5.86 | 5.95 | 6.05 | 6.14 | 6.23 | 6.32 | 6.42 | 6.51 | 6.6 | 6.7 | 6.79 | 6.88 | 6.97
38.17 | 3.25 | 3.35 | 3.44 | 3.53 | 3.63 | 3.72 | 3.81 | 3.91 | 4 | 4.09 | 4.18 | 4.28 | 4.37 | 4.46 | 4.56 | 4.65 | 4.74 | 4.84 | 4.93 | 5.02 | 5.11 | 5.21 | 5.3 | 5.39 | 5.49 | 5.58 | 5.67 | 5.77 | 5.86 | 5.95 | 6.04 | 6.14 | 6.23 | 6.32 | 6.42 | 6.51 | 6.6 | 6.69 | 6.79 | 6.88 | 6.97
38.18 | 3.25 | 3.35 | 3.44 | 3.53 | 3.63 | 3.72 | 3.81 | 3.9 | 4 | 4.09 | 4.18 | 4.28 | 4.37 | 4.46 | 4.56 | 4.65 | 4.74 | 4.83 | 4.93 | 5.02 | 5.11 | 5.21 | 5.3 | 5.39 | 5.49 | 5.58 | 5.67 | 5.76 | 5.86 | 5.95 | 6.04 | 6.14 | 6.23 | 6.32 | 6.41 | 6.51 | 6.6 | 6.69 | 6.79 | 6.88 | 6.97
38.19 | 3.25 | 3.35 | 3.44 | 3.53 | 3.63 | 3.72 | 3.81 | 3.9 | 4 | 4.09 | 4.18 | 4.28 | 4.37 | 4.46 | 4.55 | 4.65 | 4.74 | 4.83 | 4.93 | 5.02 | 5.11 | 5.21 | 5.3 | 5.39 | 5.48 | 5.58 | 5.67 | 5.76 | 5.86 | 5.95 | 6.04 | 6.14 | 6.23 | 6.32 | 6.41 | 6.51 | 6.6 | 6.69 | 6.79 | 6.88 | 6.97
38.2 | 3.25 | 3.35 | 3.44 | 3.53 | 3.62 | 3.72 | 3.81 | 3.9 | 4 | 4.09 | 4.18 | 4.28 | 4.37 | 4.46 | 4.55 | 4.65 | 4.74 | 4.83 | 4.93 | 5.02 | 5.11 | 5.2 | 5.3 | 5.39 | 5.48 | 5.58 | 5.67 | 5.76 | 5.86 | 5.95 | 6.04 | 6.13 | 6.23 | 6.32 | 6.41 | 6.51 | 6.6 | 6.69 | 6.78 | 6.88 | 6.97
38.21 | 3.25 | 3.34 | 3.44 | 3.53 | 3.62 | 3.72 | 3.81 | 3.9 | 4 | 4.09 | 4.18 | 4.27 | 4.37 | 4.46 | 4.55 | 4.65 | 4.74 | 4.83 | 4.92 | 5.02 | 5.11 | 5.2 | 5.3 | 5.39 | 5.48 | 5.58 | 5.67 | 5.76 | 5.85 | 5.95 | 6.04 | 6.13 | 6.23 | 6.32 | 6.41 | 6.5 | 6.6 | 6.69 | 6.78 | 6.88 | 6.97
38.22 | 3.25 | 3.34 | 3.44 | 3.53 | 3.62 | 3.72 | 3.81 | 3.9 | 3.99 | 4.09 | 4.18 | 4.27 | 4.37 | 4.46 | 4.55 | 4.65 | 4.74 | 4.83 | 4.92 | 5.02 | 5.11 | 5.2 | 5.3 | 5.39 | 5.48 | 5.57 | 5.67 | 5.76 | 5.85 | 5.95 | 6.04 | 6.13 | 6.22 | 6.32 | 6.41 | 6.5 | 6.6 | 6.69 | 6.78 | 6.88 | 6.97
38.23 | 3.25 | 3.34 | 3.44 | 3.53 | 3.62 | 3.72 | 3.81 | 3.9 | 3.99 | 4.09 | 4.18 | 4.27 | 4.37 | 4.46 | 4.55 | 4.64 | 4.74 | 4.83 | 4.92 | 5.02 | 5.11 | 5.2 | 5.29 | 5.39 | 5.48 | 5.57 | 5.67 | 5.76 | 5.85 | 5.95 | 6.04 | 6.13 | 6.22 | 6.32 | 6.41 | 6.5 | 6.6 | 6.69 | 6.78 | 6.87 | 6.97
38.24 | 3.25 | 3.34 | 3.44 | 3.53 | 3.62 | 3.71 | 3.81 | 3.9 | 3.99 | 4.09 | 4.18 | 4.27 | 4.36 | 4.46 | 4.55 | 4.64 | 4.74 | 4.83 | 4.92 | 5.02 | 5.11 | 5.2 | 5.29 | 5.39 | 5.48 | 5.57 | 5.67 | 5.76 | 5.85 | 5.94 | 6.04 | 6.13 | 6.22 | 6.32 | 6.41 | 6.5 | 6.59 | 6.69 | 6.78 | 6.87 | 6.97
38.25 | 3.25 | 3.34 | 3.44 | 3.53 | 3.62 | 3.71 | 3.81 | 3.9 | 3.99 | 4.09 | 4.18 | 4.27 | 4.36 | 4.46 | 4.55 | 4.64 | 4.74 | 4.83 | 4.92 | 5.01 | 5.11 | 5.2 | 5.29 | 5.39 | 5.48 | 5.57 | 5.66 | 5.76 | 5.85 | 5.94 | 6.04 | 6.13 | 6.22 | 6.31 | 6.41 | 6.5 | 6.59 | 6.69 | 6.78 | 6.87 | 6.96
38.26 | 3.25 | 3.34 | 3.43 | 3.53 | 3.62 | 3.71 | 3.81 | 3.9 | 3.99 | 4.08 | 4.18 | 4.27 | 4.36 | 4.46 | 4.55 | 4.64 | 4.73 | 4.83 | 4.92 | 5.01 | 5.11 | 5.2 | 5.29 | 5.39 | 5.48 | 5.57 | 5.66 | 5.76 | 5.85 | 5.94 | 6.04 | 6.13 | 6.22 | 6.31 | 6.41 | 6.5 | 6.59 | 6.69 | 6.78 | 6.87 | 6.96
38.27 | 3.25 | 3.34 | 3.43 | 3.53 | 3.62 | 3.71 | 3.81 | 3.9 | 3.99 | 4.08 | 4.18 | 4.27 | 4.36 | 4.46 | 4.55 | 4.64 | 4.73 | 4.83 | 4.92 | 5.01 | 5.11 | 5.2 | 5.29 | 5.38 | 5.48 | 5.57 | 5.66 | 5.76 | 5.85 | 5.94 | 6.03 | 6.13 | 6.22 | 6.31 | 6.41 | 6.5 | 6.59 | 6.68 | 6.78 | 6.87 | 6.96
38.28 | 3.25 | 3.34 | 3.43 | 3.53 | 3.62 | 3.71 | 3.8 | 3.9 | 3.99 | 4.08 | 4.18 | 4.27 | 4.36 | 4.45 | 4.55 | 4.64 | 4.73 | 4.83 | 4.92 | 5.01 | 5.1 | 5.2 | 5.29 | 5.38 | 5.48 | 5.57 | 5.66 | 5.75 | 5.85 | 5.94 | 6.03 | 6.13 | 6.22 | 6.31 | 6.4 | 6.5 | 6.59 | 6.68 | 6.78 | 6.87 | 6.96
38.29 | 3.25 | 3.34 | 3.43 | 3.53 | 3.62 | 3.71 | 3.8 | 3.9 | 3.99 | 4.08 | 4.18 | 4.27 | 4.36 | 4.45 | 4.55 | 4.64 | 4.73 | 4.83 | 4.92 | 5.01 | 5.1 | 5.2 | 5.29 | 5.38 | 5.48 | 5.57 | 5.66 | 5.75 | 5.85 | 5.94 | 6.03 | 6.13 | 6.22 | 6.31 | 6.4 | 6.5 | 6.59 | 6.68 | 6.77 | 6.87 | 6.96
38.3 | 3.25 | 3.34 | 3.43 | 3.52 | 3.62 | 3.71 | 3.8 | 3.9 | 3.99 | 4.08 | 4.17 | 4.27 | 4.36 | 4.45 | 4.55 | 4.64 | 4.73 | 4.82 | 4.92 | 5.01 | 5.1 | 5.2 | 5.29 | 5.38 | 5.47 | 5.57 | 5.66 | 5.75 | 5.85 | 5.94 | 6.03 | 6.12 | 6.22 | 6.31 | 6.4 | 6.5 | 6.59 | 6.68 | 6.77 | 6.87 | 6.96
38.31 | 3.25 | 3.34 | 3.43 | 3.52 | 3.62 | 3.71 | 3.8 | 3.9 | 3.99 | 4.08 | 4.17 | 4.27 | 4.36 | 4.45 | 4.55 | 4.64 | 4.73 | 4.82 | 4.92 | 5.01 | 5.1 | 5.19 | 5.29 | 5.38 | 5.47 | 5.57 | 5.66 | 5.75 | 5.84 | 5.94 | 6.03 | 6.12 | 6.22 | 6.31 | 6.4 | 6.49 | 6.59 | 6.68 | 6.77 | 6.87 | 6.96
38.32 | 3.24 | 3.34 | 3.43 | 3.52 | 3.62 | 3.71 | 3.8 | 3.89 | 3.99 | 4.08 | 4.17 | 4.27 | 4.36 | 4.45 | 4.54 | 4.64 | 4.73 | 4.82 | 4.92 | 5.01 | 5.1 | 5.19 | 5.29 | 5.38 | 5.47 | 5.57 | 5.66 | 5.75 | 5.84 | 5.94 | 6.03 | 6.12 | 6.21 | 6.31 | 6.4 | 6.49 | 6.59 | 6.68 | 6.77 | 6.86 | 6.96
38.33 | 3.24 | 3.34 | 3.43 | 3.52 | 3.62 | 3.71 | 3.8 | 3.89 | 3.99 | 4.08 | 4.17 | 4.27 | 4.36 | 4.45 | 4.54 | 4.64 | 4.73 | 4.82 | 4.91 | 5.01 | 5.1 | 5.19 | 5.29 | 5.38 | 5.47 | 5.56 | 5.66 | 5.75 | 5.84 | 5.94 | 6.03 | 6.12 | 6.21 | 6.31 | 6.4 | 6.49 | 6.58 | 6.68 | 6.77 | 6.86 | 6.96
38.34 | 3.24 | 3.34 | 3.43 | 3.52 | 3.61 | 3.71 | 3.8 | 3.89 | 3.99 | 4.08 | 4.17 | 4.26 | 4.36 | 4.45 | 4.54 | 4.64 | 4.73 | 4.82 | 4.91 | 5.01 | 5.1 | 5.19 | 5.29 | 5.38 | 5.47 | 5.56 | 5.66 | 5.75 | 5.84 | 5.93 | 6.03 | 6.12 | 6.21 | 6.31 | 6.4 | 6.49 | 6.58 | 6.68 | 6.77 | 6.86 | 6.95
38.35 | 3.24 | 3.34 | 3.43 | 3.52 | 3.61 | 3.71 | 3.8 | 3.89 | 3.99 | 4.08 | 4.17 | 4.26 | 4.36 | 4.45 | 4.54 | 4.63 | 4.73 | 4.82 | 4.91 | 5.01 | 5.1 | 5.19 | 5.28 | 5.38 | 5.47 | 5.56 | 5.66 | 5.75 | 5.84 | 5.93 | 6.03 | 6.12 | 6.21 | 6.3 | 6.4 | 6.49 | 6.58 | 6.68 | 6.77 | 6.86 | 6.95
38.36 | 3.24 | 3.34 | 3.43 | 3.52 | 3.61 | 3.71 | 3.8 | 3.89 | 3.98 | 4.08 | 4.17 | 4.26 | 4.36 | 4.45 | 4.54 | 4.63 | 4.73 | 4.82 | 4.91 | 5.01 | 5.1 | 5.19 | 5.28 | 5.38 | 5.47 | 5.56 | 5.65 | 5.75 | 5.84 | 5.93 | 6.03 | 6.12 | 6.21 | 6.3 | 6.4 | 6.49 | 6.58 | 6.67 | 6.77 | 6.86 | 6.95
38.37 | 3.24 | 3.33 | 3.43 | 3.52 | 3.61 | 3.71 | 3.8 | 3.89 | 3.98 | 4.08 | 4.17 | 4.26 | 4.35 | 4.45 | 4.54 | 4.63 | 4.73 | 4.82 | 4.91 | 5 | 5.1 | 5.19 | 5.28 | 5.38 | 5.47 | 5.56 | 5.65 | 5.75 | 5.84 | 5.93 | 6.02 | 6.12 | 6.21 | 6.3 | 6.4 | 6.49 | 6.58 | 6.67 | 6.77 | 6.86 | 6.95
38.38 | 3.24 | 3.33 | 3.43 | 3.52 | 3.61 | 3.7 | 3.8 | 3.89 | 3.98 | 4.08 | 4.17 | 4.26 | 4.35 | 4.45 | 4.54 | 4.63 | 4.73 | 4.82 | 4.91 | 5 | 5.1 | 5.19 | 5.28 | 5.37 | 5.47 | 5.56 | 5.65 | 5.75 | 5.84 | 5.93 | 6.02 | 6.12 | 6.21 | 6.3 | 6.39 | 6.49 | 6.58 | 6.67 | 6.77 | 6.86 | 6.95
38.39 | 3.24 | 3.33 | 3.43 | 3.52 | 3.61 | 3.7 | 3.8 | 3.89 | 3.98 | 4.08 | 4.17 | 4.26 | 4.35 | 4.45 | 4.54 | 4.63 | 4.72 | 4.82 | 4.91 | 5 | 5.1 | 5.19 | 5.28 | 5.37 | 5.47 | 5.56 | 5.65 | 5.74 | 5.84 | 5.93 | 6.02 | 6.12 | 6.21 | 6.3 | 6.39 | 6.49 | 6.58 | 6.67 | 6.76 | 6.86 | 6.95
38.4 | 3.24 | 3.33 | 3.43 | 3.52 | 3.61 | 3.7 | 3.8 | 3.89 | 3.98 | 4.07 | 4.17 | 4.26 | 4.35 | 4.45 | 4.54 | 4.63 | 4.72 | 4.82 | 4.91 | 5 | 5.09 | 5.19 | 5.28 | 5.37 | 5.47 | 5.56 | 5.65 | 5.74 | 5.84 | 5.93 | 6.02 | 6.11 | 6.21 | 6.3 | 6.39 | 6.48 | 6.58 | 6.67 | 6.76 | 6.86 | 6.95
38.41 | 3.24 | 3.33 | 3.42 | 3.52 | 3.61 | 3.7 | 3.8 | 3.89 | 3.98 | 4.07 | 4.17 | 4.26 | 4.35 | 4.44 | 4.54 | 4.63 | 4.72 | 4.82 | 4.91 | 5 | 5.09 | 5.19 | 5.28 | 5.37 | 5.46 | 5.56 | 5.65 | 5.74 | 5.84 | 5.93 | 6.02 | 6.11 | 6.21 | 6.3 | 6.39 | 6.48 | 6.58 | 6.67 | 6.76 | 6.85 | 6.95
38.42 | 3.24 | 3.33 | 3.42 | 3.52 | 3.61 | 3.7 | 3.79 | 3.89 | 3.98 | 4.07 | 4.17 | 4.26 | 4.35 | 4.44 | 4.54 | 4.63 | 4.72 | 4.81 | 4.91 | 5 | 5.09 | 5.19 | 5.28 | 5.37 | 5.46 | 5.56 | 5.65 | 5.74 | 5.83 | 5.93 | 6.02 | 6.11 | 6.2 | 6.3 | 6.39 | 6.48 | 6.58 | 6.67 | 6.76 | 6.85 | 6.95
38.43 | 3.24 | 3.33 | 3.42 | 3.52 | 3.61 | 3.7 | 3.79 | 3.89 | 3.98 | 4.07 | 4.16 | 4.26 | 4.35 | 4.44 | 4.54 | 4.63 | 4.72 | 4.81 | 4.91 | 5 | 5.09 | 5.18 | 5.28 | 5.37 | 5.46 | 5.56 | 5.65 | 5.74 | 5.83 | 5.93 | 6.02 | 6.11 | 6.2 | 6.3 | 6.39 | 6.48 | 6.57 | 6.67 | 6.76 | 6.85 | 6.95
38.44 | 3.24 | 3.33 | 3.42 | 3.52 | 3.61 | 3.7 | 3.79 | 3.89 | 3.98 | 4.07 | 4.16 | 4.26 | 4.35 | 4.44 | 4.53 | 4.63 | 4.72 | 4.81 | 4.91 | 5 | 5.09 | 5.18 | 5.28 | 5.37 | 5.46 | 5.55 | 5.65 | 5.74 | 5.83 | 5.92 | 6.02 | 6.11 | 6.2 | 6.3 | 6.39 | 6.48 | 6.57 | 6.67 | 6.76 | 6.85 | 6.94
38.45 | 3.24 | 3.33 | 3.42 | 3.51 | 3.61 | 3.7 | 3.79 | 3.89 | 3.98 | 4.07 | 4.16 | 4.26 | 4.35 | 4.44 | 4.53 | 4.63 | 4.72 | 4.81 | 4.9 | 5 | 5.09 | 5.18 | 5.28 | 5.37 | 5.46 | 5.55 | 5.65 | 5.74 | 5.83 | 5.92 | 6.02 | 6.11 | 6.2 | 6.29 | 6.39 | 6.48 | 6.57 | 6.67 | 6.76 | 6.85 | 6.94
38.46 | 3.24 | 3.33 | 3.42 | 3.51 | 3.61 | 3.7 | 3.79 | 3.88 | 3.98 | 4.07 | 4.16 | 4.26 | 4.35 | 4.44 | 4.53 | 4.63 | 4.72 | 4.81 | 4.9 | 5 | 5.09 | 5.18 | 5.27 | 5.37 | 5.46 | 5.55 | 5.64 | 5.74 | 5.83 | 5.92 | 6.02 | 6.11 | 6.2 | 6.29 | 6.39 | 6.48 | 6.57 | 6.66 | 6.76 | 6.85 | 6.94
38.47 | 3.24 | 3.33 | 3.42 | 3.51 | 3.61 | 3.7 | 3.79 | 3.88 | 3.98 | 4.07 | 4.16 | 4.25 | 4.35 | 4.44 | 4.53 | 4.62 | 4.72 | 4.81 | 4.9 | 5 | 5.09 | 5.18 | 5.27 | 5.37 | 5.46 | 5.55 | 5.64 | 5.74 | 5.83 | 5.92 | 6.01 | 6.11 | 6.2 | 6.29 | 6.39 | 6.48 | 6.57 | 6.66 | 6.76 | 6.85 | 6.94
38.48 | 3.23 | 3.33 | 3.42 | 3.51 | 3.61 | 3.7 | 3.79 | 3.88 | 3.98 | 4.07 | 4.16 | 4.25 | 4.35 | 4.44 | 4.53 | 4.62 | 4.72 | 4.81 | 4.9 | 4.99 | 5.09 | 5.18 | 5.27 | 5.37 | 5.46 | 5.55 | 5.64 | 5.74 | 5.83 | 5.92 | 6.01 | 6.11 | 6.2 | 6.29 | 6.38 | 6.48 | 6.57 | 6.66 | 6.75 | 6.85 | 6.94
38.49 | 3.23 | 3.33 | 3.42 | 3.51 | 3.6 | 3.7 | 3.79 | 3.88 | 3.97 | 4.07 | 4.16 | 4.25 | 4.35 | 4.44 | 4.53 | 4.62 | 4.72 | 4.81 | 4.9 | 4.99 | 5.09 | 5.18 | 5.27 | 5.36 | 5.46 | 5.55 | 5.64 | 5.73 | 5.83 | 5.92 | 6.01 | 6.11 | 6.2 | 6.29 | 6.38 | 6.48 | 6.57 | 6.66 | 6.75 | 6.85 | 6.94
38.5 | 3.23 | 3.33 | 3.42 | 3.51 | 3.6 | 3.7 | 3.79 | 3.88 | 3.97 | 4.07 | 4.16 | 4.25 | 4.34 | 4.44 | 4.53 | 4.62 | 4.72 | 4.81 | 4.9 | 4.99 | 5.09 | 5.18 | 5.27 | 5.36 | 5.46 | 5.55 | 5.64 | 5.73 | 5.83 | 5.92 | 6.01 | 6.1 | 6.2 | 6.29 | 6.38 | 6.47 | 6.57 | 6.66 | 6.75 | 6.84 | 6.94
38.51 | 3.23 | 3.33 | 3.42 | 3.51 | 3.6 | 3.7 | 3.79 | 3.88 | 3.97 | 4.07 | 4.16 | 4.25 | 4.34 | 4.44 | 4.53 | 4.62 | 4.71 | 4.81 | 4.9 | 4.99 | 5.08 | 5.18 | 5.27 | 5.36 | 5.46 | 5.55 | 5.64 | 5.73 | 5.83 | 5.92 | 6.01 | 6.1 | 6.2 | 6.29 | 6.38 | 6.47 | 6.57 | 6.66 | 6.75 | 6.84 | 6.94
38.52 | 3.23 | 3.32 | 3.42 | 3.51 | 3.6 | 3.69 | 3.79 | 3.88 | 3.97 | 4.07 | 4.16 | 4.25 | 4.34 | 4.44 | 4.53 | 4.62 | 4.71 | 4.81 | 4.9 | 4.99 | 5.08 | 5.18 | 5.27 | 5.36 | 5.45 | 5.55 | 5.64 | 5.73 | 5.82 | 5.92 | 6.01 | 6.1 | 6.19 | 6.29 | 6.38 | 6.47 | 6.57 | 6.66 | 6.75 | 6.84 | 6.94
38.53 | 3.23 | 3.32 | 3.42 | 3.51 | 3.6 | 3.69 | 3.79 | 3.88 | 3.97 | 4.06 | 4.16 | 4.25 | 4.34 | 4.43 | 4.53 | 4.62 | 4.71 | 4.81 | 4.9 | 4.99 | 5.08 | 5.18 | 5.27 | 5.36 | 5.45 | 5.55 | 5.64 | 5.73 | 5.82 | 5.92 | 6.01 | 6.1 | 6.19 | 6.29 | 6.38 | 6.47 | 6.56 | 6.66 | 6.75 | 6.84 | 6.93
38.54 | 3.23 | 3.32 | 3.42 | 3.51 | 3.6 | 3.69 | 3.79 | 3.88 | 3.97 | 4.06 | 4.16 | 4.25 | 4.34 | 4.43 | 4.53 | 4.62 | 4.71 | 4.8 | 4.9 | 4.99 | 5.08 | 5.17 | 5.27 | 5.36 | 5.45 | 5.54 | 5.64 | 5.73 | 5.82 | 5.92 | 6.01 | 6.1 | 6.19 | 6.29 | 6.38 | 6.47 | 6.56 | 6.66 | 6.75 | 6.84 | 6.93
38.55 | 3.23 | 3.32 | 3.42 | 3.51 | 3.6 | 3.69 | 3.79 | 3.88 | 3.97 | 4.06 | 4.16 | 4.25 | 4.34 | 4.43 | 4.53 | 4.62 | 4.71 | 4.8 | 4.9 | 4.99 | 5.08 | 5.17 | 5.27 | 5.36 | 5.45 | 5.54 | 5.64 | 5.73 | 5.82 | 5.91 | 6.01 | 6.1 | 6.19 | 6.28 | 6.38 | 6.47 | 6.56 | 6.65 | 6.75 | 6.84 | 6.93
38.56 | 3.23 | 3.32 | 3.41 | 3.51 | 3.6 | 3.69 | 3.78 | 3.88 | 3.97 | 4.06 | 4.15 | 4.25 | 4.34 | 4.43 | 4.53 | 4.62 | 4.71 | 4.8 | 4.9 | 4.99 | 5.08 | 5.17 | 5.27 | 5.36 | 5.45 | 5.54 | 5.64 | 5.73 | 5.82 | 5.91 | 6.01 | 6.1 | 6.19 | 6.28 | 6.38 | 6.47 | 6.56 | 6.65 | 6.75 | 6.84 | 6.93
38.57 | 3.23 | 3.32 | 3.41 | 3.51 | 3.6 | 3.69 | 3.78 | 3.88 | 3.97 | 4.06 | 4.15 | 4.25 | 4.34 | 4.43 | 4.52 | 4.62 | 4.71 | 4.8 | 4.89 | 4.99 | 5.08 | 5.17 | 5.26 | 5.36 | 5.45 | 5.54 | 5.63 | 5.73 | 5.82 | 5.91 | 6 | 6.1 | 6.19 | 6.28 | 6.37 | 6.47 | 6.56 | 6.65 | 6.74 | 6.84 | 6.93
38.58 | 3.23 | 3.32 | 3.41 | 3.51 | 3.6 | 3.69 | 3.78 | 3.88 | 3.97 | 4.06 | 4.15 | 4.25 | 4.34 | 4.43 | 4.52 | 4.62 | 4.71 | 4.8 | 4.89 | 4.99 | 5.08 | 5.17 | 5.26 | 5.36 | 5.45 | 5.54 | 5.63 | 5.73 | 5.82 | 5.91 | 6 | 6.1 | 6.19 | 6.28 | 6.37 | 6.47 | 6.56 | 6.65 | 6.74 | 6.84 | 6.93
38.59 | 3.23 | 3.32 | 3.41 | 3.5 | 3.6 | 3.69 | 3.78 | 3.88 | 3.97 | 4.06 | 4.15 | 4.25 | 4.34 | 4.43 | 4.52 | 4.62 | 4.71 | 4.8 | 4.89 | 4.99 | 5.08 | 5.17 | 5.26 | 5.36 | 5.45 | 5.54 | 5.63 | 5.73 | 5.82 | 5.91 | 6 | 6.1 | 6.19 | 6.28 | 6.37 | 6.47 | 6.56 | 6.65 | 6.74 | 6.84 | 6.93
38.6 | 3.23 | 3.32 | 3.41 | 3.5 | 3.6 | 3.69 | 3.78 | 3.87 | 3.97 | 4.06 | 4.15 | 4.24 | 4.34 | 4.43 | 4.52 | 4.61 | 4.71 | 4.8 | 4.89 | 4.98 | 5.08 | 5.17 | 5.26 | 5.35 | 5.45 | 5.54 | 5.63 | 5.72 | 5.82 | 5.91 | 6 | 6.09 | 6.19 | 6.28 | 6.37 | 6.46 | 6.56 | 6.65 | 6.74 | 6.83 | 6.93
38.61 | 3.23 | 3.32 | 3.41 | 3.5 | 3.6 | 3.69 | 3.78 | 3.87 | 3.97 | 4.06 | 4.15 | 4.24 | 4.34 | 4.43 | 4.52 | 4.61 | 4.71 | 4.8 | 4.89 | 4.98 | 5.08 | 5.17 | 5.26 | 5.35 | 5.45 | 5.54 | 5.63 | 5.72 | 5.82 | 5.91 | 6 | 6.09 | 6.19 | 6.28 | 6.37 | 6.46 | 6.56 | 6.65 | 6.74 | 6.83 | 6.93
38.62 | 3.23 | 3.32 | 3.41 | 3.5 | 3.6 | 3.69 | 3.78 | 3.87 | 3.97 | 4.06 | 4.15 | 4.24 | 4.34 | 4.43 | 4.52 | 4.61 | 4.71 | 4.8 | 4.89 | 4.98 | 5.08 | 5.17 | 5.26 | 5.35 | 5.45 | 5.54 | 5.63 | 5.72 | 5.81 | 5.91 | 6 | 6.09 | 6.18 | 6.28 | 6.37 | 6.46 | 6.55 | 6.65 | 6.74 | 6.83 | 6.92
38.63 | 3.22 | 3.32 | 3.41 | 3.5 | 3.59 | 3.69 | 3.78 | 3.87 | 3.96 | 4.06 | 4.15 | 4.24 | 4.33 | 4.43 | 4.52 | 4.61 | 4.7 | 4.8 | 4.89 | 4.98 | 5.07 | 5.17 | 5.26 | 5.35 | 5.44 | 5.54 | 5.63 | 5.72 | 5.81 | 5.91 | 6 | 6.09 | 6.18 | 6.28 | 6.37 | 6.46 | 6.55 | 6.65 | 6.74 | 6.83 | 6.92
38.64 | 3.22 | 3.32 | 3.41 | 3.5 | 3.59 | 3.69 | 3.78 | 3.87 | 3.96 | 4.06 | 4.15 | 4.24 | 4.33 | 4.43 | 4.52 | 4.61 | 4.7 | 4.8 | 4.89 | 4.98 | 5.07 | 5.17 | 5.26 | 5.35 | 5.44 | 5.54 | 5.63 | 5.72 | 5.81 | 5.91 | 6 | 6.09 | 6.18 | 6.28 | 6.37 | 6.46 | 6.55 | 6.64 | 6.74 | 6.83 | 6.92
38.65 | 3.22 | 3.32 | 3.41 | 3.5 | 3.59 | 3.69 | 3.78 | 3.87 | 3.96 | 4.06 | 4.15 | 4.24 | 4.33 | 4.43 | 4.52 | 4.61 | 4.7 | 4.8 | 4.89 | 4.98 | 5.07 | 5.17 | 5.26 | 5.35 | 5.44 | 5.53 | 5.63 | 5.72 | 5.81 | 5.9 | 6 | 6.09 | 6.18 | 6.27 | 6.37 | 6.46 | 6.55 | 6.64 | 6.74 | 6.83 | 6.92
38.66 | 3.22 | 3.32 | 3.41 | 3.5 | 3.59 | 3.69 | 3.78 | 3.87 | 3.96 | 4.05 | 4.15 | 4.24 | 4.33 | 4.42 | 4.52 | 4.61 | 4.7 | 4.79 | 4.89 | 4.98 | 5.07 | 5.16 | 5.26 | 5.35 | 5.44 | 5.53 | 5.63 | 5.72 | 5.81 | 5.9 | 6 | 6.09 | 6.18 | 6.27 | 6.37 | 6.46 | 6.55 | 6.64 | 6.74 | 6.83 | 6.92
38.67 | 3.22 | 3.31 | 3.41 | 3.5 | 3.59 | 3.68 | 3.78 | 3.87 | 3.96 | 4.05 | 4.15 | 4.24 | 4.33 | 4.42 | 4.52 | 4.61 | 4.7 | 4.79 | 4.89 | 4.98 | 5.07 | 5.16 | 5.26 | 5.35 | 5.44 | 5.53 | 5.63 | 5.72 | 5.81 | 5.9 | 5.99 | 6.09 | 6.18 | 6.27 | 6.36 | 6.46 | 6.55 | 6.64 | 6.73 | 6.83 | 6.92
38.68 | 3.22 | 3.31 | 3.41 | 3.5 | 3.59 | 3.68 | 3.78 | 3.87 | 3.96 | 4.05 | 4.15 | 4.24 | 4.33 | 4.42 | 4.52 | 4.61 | 4.7 | 4.79 | 4.89 | 4.98 | 5.07 | 5.16 | 5.25 | 5.35 | 5.44 | 5.53 | 5.62 | 5.72 | 5.81 | 5.9 | 5.99 | 6.09 | 6.18 | 6.27 | 6.36 | 6.46 | 6.55 | 6.64 | 6.73 | 6.83 | 6.92
38.69 | 3.22 | 3.31 | 3.41 | 3.5 | 3.59 | 3.68 | 3.78 | 3.87 | 3.96 | 4.05 | 4.15 | 4.24 | 4.33 | 4.42 | 4.51 | 4.61 | 4.7 | 4.79 | 4.88 | 4.98 | 5.07 | 5.16 | 5.25 | 5.35 | 5.44 | 5.53 | 5.62 | 5.72 | 5.81 | 5.9 | 5.99 | 6.09 | 6.18 | 6.27 | 6.36 | 6.45 | 6.55 | 6.64 | 6.73 | 6.82 | 6.92
38.7 | 3.22 | 3.31 | 3.41 | 3.5 | 3.59 | 3.68 | 3.77 | 3.87 | 3.96 | 4.05 | 4.14 | 4.24 | 4.33 | 4.42 | 4.51 | 4.61 | 4.7 | 4.79 | 4.88 | 4.98 | 5.07 | 5.16 | 5.25 | 5.35 | 5.44 | 5.53 | 5.62 | 5.71 | 5.81 | 5.9 | 5.99 | 6.08 | 6.18 | 6.27 | 6.36 | 6.45 | 6.55 | 6.64 | 6.73 | 6.82 | 6.92
38.71 | 3.22 | 3.31 | 3.4 | 3.5 | 3.59 | 3.68 | 3.77 | 3.87 | 3.96 | 4.05 | 4.14 | 4.24 | 4.33 | 4.42 | 4.51 | 4.61 | 4.7 | 4.79 | 4.88 | 4.98 | 5.07 | 5.16 | 5.25 | 5.34 | 5.44 | 5.53 | 5.62 | 5.71 | 5.81 | 5.9 | 5.99 | 6.08 | 6.18 | 6.27 | 6.36 | 6.45 | 6.55 | 6.64 | 6.73 | 6.82 | 6.91
38.72 | 3.22 | 3.31 | 3.4 | 3.5 | 3.59 | 3.68 | 3.77 | 3.87 | 3.96 | 4.05 | 4.14 | 4.24 | 4.33 | 4.42 | 4.51 | 4.6 | 4.7 | 4.79 | 4.88 | 4.97 | 5.07 | 5.16 | 5.25 | 5.34 | 5.44 | 5.53 | 5.62 | 5.71 | 5.81 | 5.9 | 5.99 | 6.08 | 6.17 | 6.27 | 6.36 | 6.45 | 6.54 | 6.64 | 6.73 | 6.82 | 6.91
38.73 | 3.22 | 3.31 | 3.4 | 3.5 | 3.59 | 3.68 | 3.77 | 3.86 | 3.96 | 4.05 | 4.14 | 4.23 | 4.33 | 4.42 | 4.51 | 4.6 | 4.7 | 4.79 | 4.88 | 4.97 | 5.07 | 5.16 | 5.25 | 5.34 | 5.44 | 5.53 | 5.62 | 5.71 | 5.8 | 5.9 | 5.99 | 6.08 | 6.17 | 6.27 | 6.36 | 6.45 | 6.54 | 6.64 | 6.73 | 6.82 | 6.91
38.74 | 3.22 | 3.31 | 3.4 | 3.49 | 3.59 | 3.68 | 3.77 | 3.86 | 3.96 | 4.05 | 4.14 | 4.23 | 4.33 | 4.42 | 4.51 | 4.6 | 4.7 | 4.79 | 4.88 | 4.97 | 5.06 | 5.16 | 5.25 | 5.34 | 5.43 | 5.53 | 5.62 | 5.71 | 5.8 | 5.9 | 5.99 | 6.08 | 6.17 | 6.27 | 6.36 | 6.45 | 6.54 | 6.63 | 6.73 | 6.82 | 6.91
38.75 | 3.22 | 3.31 | 3.4 | 3.49 | 3.59 | 3.68 | 3.77 | 3.86 | 3.96 | 4.05 | 4.14 | 4.23 | 4.33 | 4.42 | 4.51 | 4.6 | 4.69 | 4.79 | 4.88 | 4.97 | 5.06 | 5.16 | 5.25 | 5.34 | 5.43 | 5.53 | 5.62 | 5.71 | 5.8 | 5.89 | 5.99 | 6.08 | 6.17 | 6.26 | 6.36 | 6.45 | 6.54 | 6.63 | 6.73 | 6.82 | 6.91
38.76 | 3.22 | 3.31 | 3.4 | 3.49 | 3.59 | 3.68 | 3.77 | 3.86 | 3.96 | 4.05 | 4.14 | 4.23 | 4.32 | 4.42 | 4.51 | 4.6 | 4.69 | 4.79 | 4.88 | 4.97 | 5.06 | 5.16 | 5.25 | 5.34 | 5.43 | 5.52 | 5.62 | 5.71 | 5.8 | 5.89 | 5.99 | 6.08 | 6.17 | 6.26 | 6.36 | 6.45 | 6.54 | 6.63 | 6.72 | 6.82 | 6.91
38.77 | 3.22 | 3.31 | 3.4 | 3.49 | 3.59 | 3.68 | 3.77 | 3.86 | 3.95 | 4.05 | 4.14 | 4.23 | 4.32 | 4.42 | 4.51 | 4.6 | 4.69 | 4.79 | 4.88 | 4.97 | 5.06 | 5.15 | 5.25 | 5.34 | 5.43 | 5.52 | 5.62 | 5.71 | 5.8 | 5.89 | 5.99 | 6.08 | 6.17 | 6.26 | 6.35 | 6.45 | 6.54 | 6.63 | 6.72 | 6.82 | 6.91
38.78 | 3.22 | 3.31 | 3.4 | 3.49 | 3.58 | 3.68 | 3.77 | 3.86 | 3.95 | 4.05 | 4.14 | 4.23 | 4.32 | 4.42 | 4.51 | 4.6 | 4.69 | 4.78 | 4.88 | 4.97 | 5.06 | 5.15 | 5.25 | 5.34 | 5.43 | 5.52 | 5.62 | 5.71 | 5.8 | 5.89 | 5.98 | 6.08 | 6.17 | 6.26 | 6.35 | 6.45 | 6.54 | 6.63 | 6.72 | 6.81 | 6.91
38.79 | 3.21 | 3.31 | 3.4 | 3.49 | 3.58 | 3.68 | 3.77 | 3.86 | 3.95 | 4.05 | 4.14 | 4.23 | 4.32 | 4.41 | 4.51 | 4.6 | 4.69 | 4.78 | 4.88 | 4.97 | 5.06 | 5.15 | 5.25 | 5.34 | 5.43 | 5.52 | 5.61 | 5.71 | 5.8 | 5.89 | 5.98 | 6.08 | 6.17 | 6.26 | 6.35 | 6.44 | 6.54 | 6.63 | 6.72 | 6.81 | 6.91
38.8 | 3.21 | 3.31 | 3.4 | 3.49 | 3.58 | 3.68 | 3.77 | 3.86 | 3.95 | 4.04 | 4.14 | 4.23 | 4.32 | 4.41 | 4.51 | 4.6 | 4.69 | 4.78 | 4.88 | 4.97 | 5.06 | 5.15 | 5.24 | 5.34 | 5.43 | 5.52 | 5.61 | 5.71 | 5.8 | 5.89 | 5.98 | 6.07 | 6.17 | 6.26 | 6.35 | 6.44 | 6.54 | 6.63 | 6.72 | 6.81 | 6.9
38.81 | 3.21 | 3.31 | 3.4 | 3.49 | 3.58 | 3.67 | 3.77 | 3.86 | 3.95 | 4.04 | 4.14 | 4.23 | 4.32 | 4.41 | 4.51 | 4.6 | 4.69 | 4.78 | 4.87 | 4.97 | 5.06 | 5.15 | 5.24 | 5.34 | 5.43 | 5.52 | 5.61 | 5.7 | 5.8 | 5.89 | 5.98 | 6.07 | 6.17 | 6.26 | 6.35 | 6.44 | 6.53 | 6.63 | 6.72 | 6.81 | 6.9
38.82 | 3.21 | 3.3 | 3.4 | 3.49 | 3.58 | 3.67 | 3.77 | 3.86 | 3.95 | 4.04 | 4.14 | 4.23 | 4.32 | 4.41 | 4.5 | 4.6 | 4.69 | 4.78 | 4.87 | 4.97 | 5.06 | 5.15 | 5.24 | 5.33 | 5.43 | 5.52 | 5.61 | 5.7 | 5.8 | 5.89 | 5.98 | 6.07 | 6.16 | 6.26 | 6.35 | 6.44 | 6.53 | 6.63 | 6.72 | 6.81 | 6.9
38.83 | 3.21 | 3.3 | 3.4 | 3.49 | 3.58 | 3.67 | 3.77 | 3.86 | 3.95 | 4.04 | 4.13 | 4.23 | 4.32 | 4.41 | 4.5 | 4.6 | 4.69 | 4.78 | 4.87 | 4.96 | 5.06 | 5.15 | 5.24 | 5.33 | 5.43 | 5.52 | 5.61 | 5.7 | 5.79 | 5.89 | 5.98 | 6.07 | 6.16 | 6.26 | 6.35 | 6.44 | 6.53 | 6.62 | 6.72 | 6.81 | 6.9
38.84 | 3.21 | 3.3 | 3.4 | 3.49 | 3.58 | 3.67 | 3.76 | 3.86 | 3.95 | 4.04 | 4.13 | 4.23 | 4.32 | 4.41 | 4.5 | 4.59 | 4.69 | 4.78 | 4.87 | 4.96 | 5.06 | 5.15 | 5.24 | 5.33 | 5.43 | 5.52 | 5.61 | 5.7 | 5.79 | 5.89 | 5.98 | 6.07 | 6.16 | 6.26 | 6.35 | 6.44 | 6.53 | 6.62 | 6.72 | 6.81 | 6.9
38.85 | 3.21 | 3.3 | 3.4 | 3.49 | 3.58 | 3.67 | 3.76 | 3.86 | 3.95 | 4.04 | 4.13 | 4.23 | 4.32 | 4.41 | 4.5 | 4.59 | 4.69 | 4.78 | 4.87 | 4.96 | 5.06 | 5.15 | 5.24 | 5.33 | 5.42 | 5.52 | 5.61 | 5.7 | 5.79 | 5.89 | 5.98 | 6.07 | 6.16 | 6.25 | 6.35 | 6.44 | 6.53 | 6.62 | 6.72 | 6.81 | 6.9
38.86 | 3.21 | 3.3 | 3.39 | 3.49 | 3.58 | 3.67 | 3.76 | 3.86 | 3.95 | 4.04 | 4.13 | 4.22 | 4.32 | 4.41 | 4.5 | 4.59 | 4.69 | 4.78 | 4.87 | 4.96 | 5.05 | 5.15 | 5.24 | 5.33 | 5.42 | 5.52 | 5.61 | 5.7 | 5.79 | 5.88 | 5.98 | 6.07 | 6.16 | 6.25 | 6.35 | 6.44 | 6.53 | 6.62 | 6.71 | 6.81 | 6.9
38.87 | 3.21 | 3.3 | 3.39 | 3.49 | 3.58 | 3.67 | 3.76 | 3.85 | 3.95 | 4.04 | 4.13 | 4.22 | 4.32 | 4.41 | 4.5 | 4.59 | 4.68 | 4.78 | 4.87 | 4.96 | 5.05 | 5.15 | 5.24 | 5.33 | 5.42 | 5.51 | 5.61 | 5.7 | 5.79 | 5.88 | 5.98 | 6.07 | 6.16 | 6.25 | 6.34 | 6.44 | 6.53 | 6.62 | 6.71 | 6.81 | 6.9
38.88 | 3.21 | 3.3 | 3.39 | 3.49 | 3.58 | 3.67 | 3.76 | 3.85 | 3.95 | 4.04 | 4.13 | 4.22 | 4.32 | 4.41 | 4.5 | 4.59 | 4.68 | 4.78 | 4.87 | 4.96 | 5.05 | 5.14 | 5.24 | 5.33 | 5.42 | 5.51 | 5.61 | 5.7 | 5.79 | 5.88 | 5.97 | 6.07 | 6.16 | 6.25 | 6.34 | 6.44 | 6.53 | 6.62 | 6.71 | 6.8 | 6.9
38.89 | 3.21 | 3.3 | 3.39 | 3.48 | 3.58 | 3.67 | 3.76 | 3.85 | 3.95 | 4.04 | 4.13 | 4.22 | 4.31 | 4.41 | 4.5 | 4.59 | 4.68 | 4.78 | 4.87 | 4.96 | 5.05 | 5.14 | 5.24 | 5.33 | 5.42 | 5.51 | 5.6 | 5.7 | 5.79 | 5.88 | 5.97 | 6.07 | 6.16 | 6.25 | 6.34 | 6.43 | 6.53 | 6.62 | 6.71 | 6.8 | 6.9
38.9 | 3.21 | 3.3 | 3.39 | 3.48 | 3.58 | 3.67 | 3.76 | 3.85 | 3.94 | 4.04 | 4.13 | 4.22 | 4.31 | 4.41 | 4.5 | 4.59 | 4.68 | 4.77 | 4.87 | 4.96 | 5.05 | 5.14 | 5.24 | 5.33 | 5.42 | 5.51 | 5.6 | 5.7 | 5.79 | 5.88 | 5.97 | 6.06 | 6.16 | 6.25 | 6.34 | 6.43 | 6.53 | 6.62 | 6.71 | 6.8 | 6.89
38.91 | 3.21 | 3.3 | 3.39 | 3.48 | 3.58 | 3.67 | 3.76 | 3.85 | 3.94 | 4.04 | 4.13 | 4.22 | 4.31 | 4.41 | 4.5 | 4.59 | 4.68 | 4.77 | 4.87 | 4.96 | 5.05 | 5.14 | 5.23 | 5.33 | 5.42 | 5.51 | 5.6 | 5.7 | 5.79 | 5.88 | 5.97 | 6.06 | 6.16 | 6.25 | 6.34 | 6.43 | 6.52 | 6.62 | 6.71 | 6.8 | 6.89
38.92 | 3.21 | 3.3 | 3.39 | 3.48 | 3.57 | 3.67 | 3.76 | 3.85 | 3.94 | 4.04 | 4.13 | 4.22 | 4.31 | 4.4 | 4.5 | 4.59 | 4.68 | 4.77 | 4.86 | 4.96 | 5.05 | 5.14 | 5.23 | 5.33 | 5.42 | 5.51 | 5.6 | 5.69 | 5.79 | 5.88 | 5.97 | 6.06 | 6.15 | 6.25 | 6.34 | 6.43 | 6.52 | 6.62 | 6.71 | 6.8 | 6.89
38.93 | 3.21 | 3.3 | 3.39 | 3.48 | 3.57 | 3.67 | 3.76 | 3.85 | 3.94 | 4.03 | 4.13 | 4.22 | 4.31 | 4.4 | 4.5 | 4.59 | 4.68 | 4.77 | 4.86 | 4.96 | 5.05 | 5.14 | 5.23 | 5.32 | 5.42 | 5.51 | 5.6 | 5.69 | 5.79 | 5.88 | 5.97 | 6.06 | 6.15 | 6.25 | 6.34 | 6.43 | 6.52 | 6.61 | 6.71 | 6.8 | 6.89
38.94 | 3.2 | 3.3 | 3.39 | 3.48 | 3.57 | 3.67 | 3.76 | 3.85 | 3.94 | 4.03 | 4.13 | 4.22 | 4.31 | 4.4 | 4.49 | 4.59 | 4.68 | 4.77 | 4.86 | 4.96 | 5.05 | 5.14 | 5.23 | 5.32 | 5.42 | 5.51 | 5.6 | 5.69 | 5.78 | 5.88 | 5.97 | 6.06 | 6.15 | 6.25 | 6.34 | 6.43 | 6.52 | 6.61 | 6.71 | 6.8 | 6.89
38.95 | 3.2 | 3.3 | 3.39 | 3.48 | 3.57 | 3.66 | 3.76 | 3.85 | 3.94 | 4.03 | 4.13 | 4.22 | 4.31 | 4.4 | 4.49 | 4.59 | 4.68 | 4.77 | 4.86 | 4.95 | 5.05 | 5.14 | 5.23 | 5.32 | 5.42 | 5.51 | 5.6 | 5.69 | 5.78 | 5.88 | 5.97 | 6.06 | 6.15 | 6.24 | 6.34 | 6.43 | 6.52 | 6.61 | 6.7 | 6.8 | 6.89
38.96 | 3.2 | 3.3 | 3.39 | 3.48 | 3.57 | 3.66 | 3.76 | 3.85 | 3.94 | 4.03 | 4.12 | 4.22 | 4.31 | 4.4 | 4.49 | 4.59 | 4.68 | 4.77 | 4.86 | 4.95 | 5.05 | 5.14 | 5.23 | 5.32 | 5.41 | 5.51 | 5.6 | 5.69 | 5.78 | 5.87 | 5.97 | 6.06 | 6.15 | 6.24 | 6.34 | 6.43 | 6.52 | 6.61 | 6.7 | 6.8 | 6.89
38.97 | 3.2 | 3.3 | 3.39 | 3.48 | 3.57 | 3.66 | 3.76 | 3.85 | 3.94 | 4.03 | 4.12 | 4.22 | 4.31 | 4.4 | 4.49 | 4.58 | 4.68 | 4.77 | 4.86 | 4.95 | 5.04 | 5.14 | 5.23 | 5.32 | 5.41 | 5.51 | 5.6 | 5.69 | 5.78 | 5.87 | 5.97 | 6.06 | 6.15 | 6.24 | 6.33 | 6.43 | 6.52 | 6.61 | 6.7 | 6.79 | 6.89
38.98 | 3.2 | 3.29 | 3.39 | 3.48 | 3.57 | 3.66 | 3.75 | 3.85 | 3.94 | 4.03 | 4.12 | 4.22 | 4.31 | 4.4 | 4.49 | 4.58 | 4.68 | 4.77 | 4.86 | 4.95 | 5.04 | 5.14 | 5.23 | 5.32 | 5.41 | 5.5 | 5.6 | 5.69 | 5.78 | 5.87 | 5.96 | 6.06 | 6.15 | 6.24 | 6.33 | 6.43 | 6.52 | 6.61 | 6.7 | 6.79 | 6.89
38.99 | 3.2 | 3.29 | 3.39 | 3.48 | 3.57 | 3.66 | 3.75 | 3.85 | 3.94 | 4.03 | 4.12 | 4.21 | 4.31 | 4.4 | 4.49 | 4.58 | 4.67 | 4.77 | 4.86 | 4.95 | 5.04 | 5.14 | 5.23 | 5.32 | 5.41 | 5.5 | 5.6 | 5.69 | 5.78 | 5.87 | 5.96 | 6.06 | 6.15 | 6.24 | 6.33 | 6.42 | 6.52 | 6.61 | 6.7 | 6.79 | 6.88
39 | 3.2 | 3.29 | 3.39 | 3.48 | 3.57 | 3.66 | 3.75 | 3.85 | 3.94 | 4.03 | 4.12 | 4.21 | 4.31 | 4.4 | 4.49 | 4.58 | 4.67 | 4.77 | 4.86 | 4.95 | 5.04 | 5.13 | 5.23 | 5.32 | 5.41 | 5.5 | 5.59 | 5.69 | 5.78 | 5.87 | 5.96 | 6.05 | 6.15 | 6.24 | 6.33 | 6.42 | 6.52 | 6.61 | 6.7 | 6.79 | 6.88
39.01 | 3.2 | 3.29 | 3.38 | 3.48 | 3.57 | 3.66 | 3.75 | 3.84 | 3.94 | 4.03 | 4.12 | 4.21 | 4.31 | 4.4 | 4.49 | 4.58 | 4.67 | 4.77 | 4.86 | 4.95 | 5.04 | 5.13 | 5.23 | 5.32 | 5.41 | 5.5 | 5.59 | 5.69 | 5.78 | 5.87 | 5.96 | 6.05 | 6.15 | 6.24 | 6.33 | 6.42 | 6.51 | 6.61 | 6.7 | 6.79 | 6.88
39.02 | 3.2 | 3.29 | 3.38 | 3.48 | 3.57 | 3.66 | 3.75 | 3.84 | 3.94 | 4.03 | 4.12 | 4.21 | 4.3 | 4.4 | 4.49 | 4.58 | 4.67 | 4.76 | 4.86 | 4.95 | 5.04 | 5.13 | 5.22 | 5.32 | 5.41 | 5.5 | 5.59 | 5.68 | 5.78 | 5.87 | 5.96 | 6.05 | 6.14 | 6.24 | 6.33 | 6.42 | 6.51 | 6.61 | 6.7 | 6.79 | 6.88
39.03 | 3.2 | 3.29 | 3.38 | 3.48 | 3.57 | 3.66 | 3.75 | 3.84 | 3.94 | 4.03 | 4.12 | 4.21 | 4.3 | 4.4 | 4.49 | 4.58 | 4.67 | 4.76 | 4.86 | 4.95 | 5.04 | 5.13 | 5.22 | 5.32 | 5.41 | 5.5 | 5.59 | 5.68 | 5.78 | 5.87 | 5.96 | 6.05 | 6.14 | 6.24 | 6.33 | 6.42 | 6.51 | 6.6 | 6.7 | 6.79 | 6.88
39.04 | 3.2 | 3.29 | 3.38 | 3.47 | 3.57 | 3.66 | 3.75 | 3.84 | 3.93 | 4.03 | 4.12 | 4.21 | 4.3 | 4.39 | 4.49 | 4.58 | 4.67 | 4.76 | 4.85 | 4.95 | 5.04 | 5.13 | 5.22 | 5.31 | 5.41 | 5.5 | 5.59 | 5.68 | 5.77 | 5.87 | 5.96 | 6.05 | 6.14 | 6.23 | 6.33 | 6.42 | 6.51 | 6.6 | 6.69 | 6.79 | 6.88
39.05 | 3.2 | 3.29 | 3.38 | 3.47 | 3.57 | 3.66 | 3.75 | 3.84 | 3.93 | 4.03 | 4.12 | 4.21 | 4.3 | 4.39 | 4.49 | 4.58 | 4.67 | 4.76 | 4.85 | 4.95 | 5.04 | 5.13 | 5.22 | 5.31 | 5.41 | 5.5 | 5.59 | 5.68 | 5.77 | 5.87 | 5.96 | 6.05 | 6.14 | 6.23 | 6.33 | 6.42 | 6.51 | 6.6 | 6.69 | 6.79 | 6.88
39.06 | 3.2 | 3.29 | 3.38 | 3.47 | 3.57 | 3.66 | 3.75 | 3.84 | 3.93 | 4.03 | 4.12 | 4.21 | 4.3 | 4.39 | 4.49 | 4.58 | 4.67 | 4.76 | 4.85 | 4.95 | 5.04 | 5.13 | 5.22 | 5.31 | 5.41 | 5.5 | 5.59 | 5.68 | 5.77 | 5.87 | 5.96 | 6.05 | 6.14 | 6.23 | 6.32 | 6.42 | 6.51 | 6.6 | 6.69 | 6.78 | 6.88
39.07 | 3.2 | 3.29 | 3.38 | 3.47 | 3.56 | 3.66 | 3.75 | 3.84 | 3.93 | 4.02 | 4.12 | 4.21 | 4.3 | 4.39 | 4.48 | 4.58 | 4.67 | 4.76 | 4.85 | 4.94 | 5.04 | 5.13 | 5.22 | 5.31 | 5.4 | 5.5 | 5.59 | 5.68 | 5.77 | 5.86 | 5.96 | 6.05 | 6.14 | 6.23 | 6.32 | 6.42 | 6.51 | 6.6 | 6.69 | 6.78 | 6.88
39.08 | 3.2 | 3.29 | 3.38 | 3.47 | 3.56 | 3.66 | 3.75 | 3.84 | 3.93 | 4.02 | 4.12 | 4.21 | 4.3 | 4.39 | 4.48 | 4.58 | 4.67 | 4.76 | 4.85 | 4.94 | 5.04 | 5.13 | 5.22 | 5.31 | 5.4 | 5.5 | 5.59 | 5.68 | 5.77 | 5.86 | 5.96 | 6.05 | 6.14 | 6.23 | 6.32 | 6.41 | 6.51 | 6.6 | 6.69 | 6.78 | 6.87
39.09 | 3.2 | 3.29 | 3.38 | 3.47 | 3.56 | 3.66 | 3.75 | 3.84 | 3.93 | 4.02 | 4.12 | 4.21 | 4.3 | 4.39 | 4.48 | 4.57 | 4.67 | 4.76 | 4.85 | 4.94 | 5.03 | 5.13 | 5.22 | 5.31 | 5.4 | 5.49 | 5.59 | 5.68 | 5.77 | 5.86 | 5.95 | 6.05 | 6.14 | 6.23 | 6.32 | 6.41 | 6.51 | 6.6 | 6.69 | 6.78 | 6.87
39.1 | 3.19 | 3.29 | 3.38 | 3.47 | 3.56 | 3.65 | 3.75 | 3.84 | 3.93 | 4.02 | 4.11 | 4.21 | 4.3 | 4.39 | 4.48 | 4.57 | 4.67 | 4.76 | 4.85 | 4.94 | 5.03 | 5.13 | 5.22 | 5.31 | 5.4 | 5.49 | 5.59 | 5.68 | 5.77 | 5.86 | 5.95 | 6.05 | 6.14 | 6.23 | 6.32 | 6.41 | 6.5 | 6.6 | 6.69 | 6.78 | 6.87
39.11 | 3.19 | 3.29 | 3.38 | 3.47 | 3.56 | 3.65 | 3.75 | 3.84 | 3.93 | 4.02 | 4.11 | 4.21 | 4.3 | 4.39 | 4.48 | 4.57 | 4.67 | 4.76 | 4.85 | 4.94 | 5.03 | 5.12 | 5.22 | 5.31 | 5.4 | 5.49 | 5.58 | 5.68 | 5.77 | 5.86 | 5.95 | 6.04 | 6.14 | 6.23 | 6.32 | 6.41 | 6.5 | 6.6 | 6.69 | 6.78 | 6.87
39.12 | 3.19 | 3.29 | 3.38 | 3.47 | 3.56 | 3.65 | 3.75 | 3.84 | 3.93 | 4.02 | 4.11 | 4.2 | 4.3 | 4.39 | 4.48 | 4.57 | 4.66 | 4.76 | 4.85 | 4.94 | 5.03 | 5.12 | 5.22 | 5.31 | 5.4 | 5.49 | 5.58 | 5.68 | 5.77 | 5.86 | 5.95 | 6.04 | 6.14 | 6.23 | 6.32 | 6.41 | 6.5 | 6.59 | 6.69 | 6.78 | 6.87
39.13 | 3.19 | 3.28 | 3.38 | 3.47 | 3.56 | 3.65 | 3.74 | 3.84 | 3.93 | 4.02 | 4.11 | 4.2 | 4.3 | 4.39 | 4.48 | 4.57 | 4.66 | 4.76 | 4.85 | 4.94 | 5.03 | 5.12 | 5.22 | 5.31 | 5.4 | 5.49 | 5.58 | 5.67 | 5.77 | 5.86 | 5.95 | 6.04 | 6.13 | 6.23 | 6.32 | 6.41 | 6.5 | 6.59 | 6.69 | 6.78 | 6.87
39.14 | 3.19 | 3.28 | 3.38 | 3.47 | 3.56 | 3.65 | 3.74 | 3.84 | 3.93 | 4.02 | 4.11 | 4.2 | 4.3 | 4.39 | 4.48 | 4.57 | 4.66 | 4.75 | 4.85 | 4.94 | 5.03 | 5.12 | 5.21 | 5.31 | 5.4 | 5.49 | 5.58 | 5.67 | 5.77 | 5.86 | 5.95 | 6.04 | 6.13 | 6.22 | 6.32 | 6.41 | 6.5 | 6.59 | 6.68 | 6.78 | 6.87
39.15 | 3.19 | 3.28 | 3.38 | 3.47 | 3.56 | 3.65 | 3.74 | 3.83 | 3.93 | 4.02 | 4.11 | 4.2 | 4.29 | 4.39 | 4.48 | 4.57 | 4.66 | 4.75 | 4.85 | 4.94 | 5.03 | 5.12 | 5.21 | 5.31 | 5.4 | 5.49 | 5.58 | 5.67 | 5.76 | 5.86 | 5.95 | 6.04 | 6.13 | 6.22 | 6.32 | 6.41 | 6.5 | 6.59 | 6.68 | 6.78 | 6.87
39.16 | 3.19 | 3.28 | 3.37 | 3.47 | 3.56 | 3.65 | 3.74 | 3.83 | 3.93 | 4.02 | 4.11 | 4.2 | 4.29 | 4.39 | 4.48 | 4.57 | 4.66 | 4.75 | 4.84 | 4.94 | 5.03 | 5.12 | 5.21 | 5.3 | 5.4 | 5.49 | 5.58 | 5.67 | 5.76 | 5.86 | 5.95 | 6.04 | 6.13 | 6.22 | 6.31 | 6.41 | 6.5 | 6.59 | 6.68 | 6.77 | 6.87
39.17 | 3.19 | 3.28 | 3.37 | 3.47 | 3.56 | 3.65 | 3.74 | 3.83 | 3.93 | 4.02 | 4.11 | 4.2 | 4.29 | 4.38 | 4.48 | 4.57 | 4.66 | 4.75 | 4.84 | 4.94 | 5.03 | 5.12 | 5.21 | 5.3 | 5.4 | 5.49 | 5.58 | 5.67 | 5.76 | 5.85 | 5.95 | 6.04 | 6.13 | 6.22 | 6.31 | 6.41 | 6.5 | 6.59 | 6.68 | 6.77 | 6.86
39.18 | 3.19 | 3.28 | 3.37 | 3.47 | 3.56 | 3.65 | 3.74 | 3.83 | 3.92 | 4.02 | 4.11 | 4.2 | 4.29 | 4.38 | 4.48 | 4.57 | 4.66 | 4.75 | 4.84 | 4.94 | 5.03 | 5.12 | 5.21 | 5.3 | 5.39 | 5.49 | 5.58 | 5.67 | 5.76 | 5.85 | 5.95 | 6.04 | 6.13 | 6.22 | 6.31 | 6.4 | 6.5 | 6.59 | 6.68 | 6.77 | 6.86
39.19 | 3.19 | 3.28 | 3.37 | 3.46 | 3.56 | 3.65 | 3.74 | 3.83 | 3.92 | 4.02 | 4.11 | 4.2 | 4.29 | 4.38 | 4.47 | 4.57 | 4.66 | 4.75 | 4.84 | 4.93 | 5.03 | 5.12 | 5.21 | 5.3 | 5.39 | 5.49 | 5.58 | 5.67 | 5.76 | 5.85 | 5.94 | 6.04 | 6.13 | 6.22 | 6.31 | 6.4 | 6.5 | 6.59 | 6.68 | 6.77 | 6.86
39.2 | 3.19 | 3.28 | 3.37 | 3.46 | 3.56 | 3.65 | 3.74 | 3.83 | 3.92 | 4.01 | 4.11 | 4.2 | 4.29 | 4.38 | 4.47 | 4.57 | 4.66 | 4.75 | 4.84 | 4.93 | 5.03 | 5.12 | 5.21 | 5.3 | 5.39 | 5.48 | 5.58 | 5.67 | 5.76 | 5.85 | 5.94 | 6.04 | 6.13 | 6.22 | 6.31 | 6.4 | 6.49 | 6.59 | 6.68 | 6.77 | 6.86
39.21 | 3.19 | 3.28 | 3.37 | 3.46 | 3.55 | 3.65 | 3.74 | 3.83 | 3.92 | 4.01 | 4.11 | 4.2 | 4.29 | 4.38 | 4.47 | 4.57 | 4.66 | 4.75 | 4.84 | 4.93 | 5.02 | 5.12 | 5.21 | 5.3 | 5.39 | 5.48 | 5.58 | 5.67 | 5.76 | 5.85 | 5.94 | 6.03 | 6.13 | 6.22 | 6.31 | 6.4 | 6.49 | 6.59 | 6.68 | 6.77 | 6.86
39.22 | 3.19 | 3.28 | 3.37 | 3.46 | 3.55 | 3.65 | 3.74 | 3.83 | 3.92 | 4.01 | 4.11 | 4.2 | 4.29 | 4.38 | 4.47 | 4.56 | 4.66 | 4.75 | 4.84 | 4.93 | 5.02 | 5.12 | 5.21 | 5.3 | 5.39 | 5.48 | 5.57 | 5.67 | 5.76 | 5.85 | 5.94 | 6.03 | 6.13 | 6.22 | 6.31 | 6.4 | 6.49 | 6.58 | 6.68 | 6.77 | 6.86
39.23 | 3.19 | 3.28 | 3.37 | 3.46 | 3.55 | 3.65 | 3.74 | 3.83 | 3.92 | 4.01 | 4.1 | 4.2 | 4.29 | 4.38 | 4.47 | 4.56 | 4.66 | 4.75 | 4.84 | 4.93 | 5.02 | 5.11 | 5.21 | 5.3 | 5.39 | 5.48 | 5.57 | 5.67 | 5.76 | 5.85 | 5.94 | 6.03 | 6.12 | 6.22 | 6.31 | 6.4 | 6.49 | 6.58 | 6.67 | 6.77 | 6.86
39.24 | 3.19 | 3.28 | 3.37 | 3.46 | 3.55 | 3.64 | 3.74 | 3.83 | 3.92 | 4.01 | 4.1 | 4.2 | 4.29 | 4.38 | 4.47 | 4.56 | 4.65 | 4.75 | 4.84 | 4.93 | 5.02 | 5.11 | 5.21 | 5.3 | 5.39 | 5.48 | 5.57 | 5.66 | 5.76 | 5.85 | 5.94 | 6.03 | 6.12 | 6.21 | 6.31 | 6.4 | 6.49 | 6.58 | 6.67 | 6.77 | 6.86
39.25 | 3.19 | 3.28 | 3.37 | 3.46 | 3.55 | 3.64 | 3.74 | 3.83 | 3.92 | 4.01 | 4.1 | 4.19 | 4.29 | 4.38 | 4.47 | 4.56 | 4.65 | 4.75 | 4.84 | 4.93 | 5.02 | 5.11 | 5.2 | 5.3 | 5.39 | 5.48 | 5.57 | 5.66 | 5.76 | 5.85 | 5.94 | 6.03 | 6.12 | 6.21 | 6.31 | 6.4 | 6.49 | 6.58 | 6.67 | 6.76 | 6.86
39.26 | 3.18 | 3.28 | 3.37 | 3.46 | 3.55 | 3.64 | 3.74 | 3.83 | 3.92 | 4.01 | 4.1 | 4.19 | 4.29 | 4.38 | 4.47 | 4.56 | 4.65 | 4.74 | 4.84 | 4.93 | 5.02 | 5.11 | 5.2 | 5.3 | 5.39 | 5.48 | 5.57 | 5.66 | 5.75 | 5.85 | 5.94 | 6.03 | 6.12 | 6.21 | 6.3 | 6.4 | 6.49 | 6.58 | 6.67 | 6.76 | 6.86
39.27 | 3.18 | 3.28 | 3.37 | 3.46 | 3.55 | 3.64 | 3.73 | 3.83 | 3.92 | 4.01 | 4.1 | 4.19 | 4.29 | 4.38 | 4.47 | 4.56 | 4.65 | 4.74 | 4.84 | 4.93 | 5.02 | 5.11 | 5.2 | 5.29 | 5.39 | 5.48 | 5.57 | 5.66 | 5.75 | 5.85 | 5.94 | 6.03 | 6.12 | 6.21 | 6.3 | 6.4 | 6.49 | 6.58 | 6.67 | 6.76 | 6.85
39.28 | 3.18 | 3.27 | 3.37 | 3.46 | 3.55 | 3.64 | 3.73 | 3.83 | 3.92 | 4.01 | 4.1 | 4.19 | 4.28 | 4.38 | 4.47 | 4.56 | 4.65 | 4.74 | 4.83 | 4.93 | 5.02 | 5.11 | 5.2 | 5.29 | 5.39 | 5.48 | 5.57 | 5.66 | 5.75 | 5.84 | 5.94 | 6.03 | 6.12 | 6.21 | 6.3 | 6.39 | 6.49 | 6.58 | 6.67 | 6.76 | 6.85
39.29 | 3.18 | 3.27 | 3.37 | 3.46 | 3.55 | 3.64 | 3.73 | 3.82 | 3.92 | 4.01 | 4.1 | 4.19 | 4.28 | 4.38 | 4.47 | 4.56 | 4.65 | 4.74 | 4.83 | 4.93 | 5.02 | 5.11 | 5.2 | 5.29 | 5.38 | 5.48 | 5.57 | 5.66 | 5.75 | 5.84 | 5.93 | 6.03 | 6.12 | 6.21 | 6.3 | 6.39 | 6.49 | 6.58 | 6.67 | 6.76 | 6.85
39.3 | 3.18 | 3.27 | 3.37 | 3.46 | 3.55 | 3.64 | 3.73 | 3.82 | 3.92 | 4.01 | 4.1 | 4.19 | 4.28 | 4.37 | 4.47 | 4.56 | 4.65 | 4.74 | 4.83 | 4.92 | 5.02 | 5.11 | 5.2 | 5.29 | 5.38 | 5.48 | 5.57 | 5.66 | 5.75 | 5.84 | 5.93 | 6.03 | 6.12 | 6.21 | 6.3 | 6.39 | 6.48 | 6.58 | 6.67 | 6.76 | 6.85
39.31 | 3.18 | 3.27 | 3.36 | 3.46 | 3.55 | 3.64 | 3.73 | 3.82 | 3.92 | 4.01 | 4.1 | 4.19 | 4.28 | 4.37 | 4.47 | 4.56 | 4.65 | 4.74 | 4.83 | 4.92 | 5.02 | 5.11 | 5.2 | 5.29 | 5.38 | 5.47 | 5.57 | 5.66 | 5.75 | 5.84 | 5.93 | 6.02 | 6.12 | 6.21 | 6.3 | 6.39 | 6.48 | 6.57 | 6.67 | 6.76 | 6.85
39.32 | 3.18 | 3.27 | 3.36 | 3.46 | 3.55 | 3.64 | 3.73 | 3.82 | 3.91 | 4.01 | 4.1 | 4.19 | 4.28 | 4.37 | 4.46 | 4.56 | 4.65 | 4.74 | 4.83 | 4.92 | 5.01 | 5.11 | 5.2 | 5.29 | 5.38 | 5.47 | 5.57 | 5.66 | 5.75 | 5.84 | 5.93 | 6.02 | 6.12 | 6.21 | 6.3 | 6.39 | 6.48 | 6.57 | 6.67 | 6.76 | 6.85
39.33 | 3.18 | 3.27 | 3.36 | 3.46 | 3.55 | 3.64 | 3.73 | 3.82 | 3.91 | 4.01 | 4.1 | 4.19 | 4.28 | 4.37 | 4.46 | 4.56 | 4.65 | 4.74 | 4.83 | 4.92 | 5.01 | 5.11 | 5.2 | 5.29 | 5.38 | 5.47 | 5.56 | 5.66 | 5.75 | 5.84 | 5.93 | 6.02 | 6.11 | 6.21 | 6.3 | 6.39 | 6.48 | 6.57 | 6.66 | 6.76 | 6.85
39.34 | 3.18 | 3.27 | 3.36 | 3.45 | 3.55 | 3.64 | 3.73 | 3.82 | 3.91 | 4 | 4.1 | 4.19 | 4.28 | 4.37 | 4.46 | 4.55 | 4.65 | 4.74 | 4.83 | 4.92 | 5.01 | 5.1 | 5.2 | 5.29 | 5.38 | 5.47 | 5.56 | 5.66 | 5.75 | 5.84 | 5.93 | 6.02 | 6.11 | 6.21 | 6.3 | 6.39 | 6.48 | 6.57 | 6.66 | 6.76 | 6.85
39.35 | 3.18 | 3.27 | 3.36 | 3.45 | 3.55 | 3.64 | 3.73 | 3.82 | 3.91 | 4 | 4.1 | 4.19 | 4.28 | 4.37 | 4.46 | 4.55 | 4.65 | 4.74 | 4.83 | 4.92 | 5.01 | 5.1 | 5.2 | 5.29 | 5.38 | 5.47 | 5.56 | 5.65 | 5.75 | 5.84 | 5.93 | 6.02 | 6.11 | 6.2 | 6.3 | 6.39 | 6.48 | 6.57 | 6.66 | 6.75 | 6.85
39.36 | 3.18 | 3.27 | 3.36 | 3.45 | 3.54 | 3.64 | 3.73 | 3.82 | 3.91 | 4 | 4.09 | 4.19 | 4.28 | 4.37 | 4.46 | 4.55 | 4.64 | 4.74 | 4.83 | 4.92 | 5.01 | 5.1 | 5.19 | 5.29 | 5.38 | 5.47 | 5.56 | 5.65 | 5.74 | 5.84 | 5.93 | 6.02 | 6.11 | 6.2 | 6.29 | 6.39 | 6.48 | 6.57 | 6.66 | 6.75 | 6.84
39.37 | 3.18 | 3.27 | 3.36 | 3.45 | 3.54 | 3.64 | 3.73 | 3.82 | 3.91 | 4 | 4.09 | 4.19 | 4.28 | 4.37 | 4.46 | 4.55 | 4.64 | 4.74 | 4.83 | 4.92 | 5.01 | 5.1 | 5.19 | 5.29 | 5.38 | 5.47 | 5.56 | 5.65 | 5.74 | 5.84 | 5.93 | 6.02 | 6.11 | 6.2 | 6.29 | 6.39 | 6.48 | 6.57 | 6.66 | 6.75 | 6.84
39.38 | 3.18 | 3.27 | 3.36 | 3.45 | 3.54 | 3.64 | 3.73 | 3.82 | 3.91 | 4 | 4.09 | 4.19 | 4.28 | 4.37 | 4.46 | 4.55 | 4.64 | 4.73 | 4.83 | 4.92 | 5.01 | 5.1 | 5.19 | 5.28 | 5.38 | 5.47 | 5.56 | 5.65 | 5.74 | 5.83 | 5.93 | 6.02 | 6.11 | 6.2 | 6.29 | 6.38 | 6.48 | 6.57 | 6.66 | 6.75 | 6.84
39.39 | 3.18 | 3.27 | 3.36 | 3.45 | 3.54 | 3.63 | 3.73 | 3.82 | 3.91 | 4 | 4.09 | 4.18 | 4.28 | 4.37 | 4.46 | 4.55 | 4.64 | 4.73 | 4.83 | 4.92 | 5.01 | 5.1 | 5.19 | 5.28 | 5.38 | 5.47 | 5.56 | 5.65 | 5.74 | 5.83 | 5.93 | 6.02 | 6.11 | 6.2 | 6.29 | 6.38 | 6.47 | 6.57 | 6.66 | 6.75 | 6.84
39.4 | 3.18 | 3.27 | 3.36 | 3.45 | 3.54 | 3.63 | 3.73 | 3.82 | 3.91 | 4 | 4.09 | 4.18 | 4.28 | 4.37 | 4.46 | 4.55 | 4.64 | 4.73 | 4.82 | 4.92 | 5.01 | 5.1 | 5.19 | 5.28 | 5.37 | 5.47 | 5.56 | 5.65 | 5.74 | 5.83 | 5.92 | 6.02 | 6.11 | 6.2 | 6.29 | 6.38 | 6.47 | 6.57 | 6.66 | 6.75 | 6.84
39.41 | 3.17 | 3.27 | 3.36 | 3.45 | 3.54 | 3.63 | 3.72 | 3.82 | 3.91 | 4 | 4.09 | 4.18 | 4.27 | 4.37 | 4.46 | 4.55 | 4.64 | 4.73 | 4.82 | 4.92 | 5.01 | 5.1 | 5.19 | 5.28 | 5.37 | 5.47 | 5.56 | 5.65 | 5.74 | 5.83 | 5.92 | 6.01 | 6.11 | 6.2 | 6.29 | 6.38 | 6.47 | 6.56 | 6.66 | 6.75 | 6.84
39.42 | 3.17 | 3.27 | 3.36 | 3.45 | 3.54 | 3.63 | 3.72 | 3.82 | 3.91 | 4 | 4.09 | 4.18 | 4.27 | 4.37 | 4.46 | 4.55 | 4.64 | 4.73 | 4.82 | 4.91 | 5.01 | 5.1 | 5.19 | 5.28 | 5.37 | 5.46 | 5.56 | 5.65 | 5.74 | 5.83 | 5.92 | 6.01 | 6.11 | 6.2 | 6.29 | 6.38 | 6.47 | 6.56 | 6.65 | 6.75 | 6.84
39.43 | 3.17 | 3.27 | 3.36 | 3.45 | 3.54 | 3.63 | 3.72 | 3.81 | 3.91 | 4 | 4.09 | 4.18 | 4.27 | 4.36 | 4.46 | 4.55 | 4.64 | 4.73 | 4.82 | 4.91 | 5.01 | 5.1 | 5.19 | 5.28 | 5.37 | 5.46 | 5.56 | 5.65 | 5.74 | 5.83 | 5.92 | 6.01 | 6.1 | 6.2 | 6.29 | 6.38 | 6.47 | 6.56 | 6.65 | 6.75 | 6.84
39.44 | 3.17 | 3.26 | 3.36 | 3.45 | 3.54 | 3.63 | 3.72 | 3.81 | 3.91 | 4 | 4.09 | 4.18 | 4.27 | 4.36 | 4.46 | 4.55 | 4.64 | 4.73 | 4.82 | 4.91 | 5 | 5.1 | 5.19 | 5.28 | 5.37 | 5.46 | 5.55 | 5.65 | 5.74 | 5.83 | 5.92 | 6.01 | 6.1 | 6.2 | 6.29 | 6.38 | 6.47 | 6.56 | 6.65 | 6.74 | 6.84
39.45 | 3.17 | 3.26 | 3.36 | 3.45 | 3.54 | 3.63 | 3.72 | 3.81 | 3.9 | 4 | 4.09 | 4.18 | 4.27 | 4.36 | 4.45 | 4.55 | 4.64 | 4.73 | 4.82 | 4.91 | 5 | 5.1 | 5.19 | 5.28 | 5.37 | 5.46 | 5.55 | 5.64 | 5.74 | 5.83 | 5.92 | 6.01 | 6.1 | 6.19 | 6.29 | 6.38 | 6.47 | 6.56 | 6.65 | 6.74 | 6.83
39.46 | 3.17 | 3.26 | 3.35 | 3.45 | 3.54 | 3.63 | 3.72 | 3.81 | 3.9 | 4 | 4.09 | 4.18 | 4.27 | 4.36 | 4.45 | 4.55 | 4.64 | 4.73 | 4.82 | 4.91 | 5 | 5.09 | 5.19 | 5.28 | 5.37 | 5.46 | 5.55 | 5.64 | 5.74 | 5.83 | 5.92 | 6.01 | 6.1 | 6.19 | 6.28 | 6.38 | 6.47 | 6.56 | 6.65 | 6.74 | 6.83
39.47 | 3.17 | 3.26 | 3.35 | 3.45 | 3.54 | 3.63 | 3.72 | 3.81 | 3.9 | 4 | 4.09 | 4.18 | 4.27 | 4.36 | 4.45 | 4.54 | 4.64 | 4.73 | 4.82 | 4.91 | 5 | 5.09 | 5.19 | 5.28 | 5.37 | 5.46 | 5.55 | 5.64 | 5.73 | 5.83 | 5.92 | 6.01 | 6.1 | 6.19 | 6.28 | 6.38 | 6.47 | 6.56 | 6.65 | 6.74 | 6.83
39.48 | 3.17 | 3.26 | 3.35 | 3.45 | 3.54 | 3.63 | 3.72 | 3.81 | 3.9 | 3.99 | 4.09 | 4.18 | 4.27 | 4.36 | 4.45 | 4.54 | 4.64 | 4.73 | 4.82 | 4.91 | 5 | 5.09 | 5.18 | 5.28 | 5.37 | 5.46 | 5.55 | 5.64 | 5.73 | 5.83 | 5.92 | 6.01 | 6.1 | 6.19 | 6.28 | 6.37 | 6.47 | 6.56 | 6.65 | 6.74 | 6.83
39.49 | 3.17 | 3.26 | 3.35 | 3.44 | 3.54 | 3.63 | 3.72 | 3.81 | 3.9 | 3.99 | 4.09 | 4.18 | 4.27 | 4.36 | 4.45 | 4.54 | 4.63 | 4.73 | 4.82 | 4.91 | 5 | 5.09 | 5.18 | 5.28 | 5.37 | 5.46 | 5.55 | 5.64 | 5.73 | 5.82 | 5.92 | 6.01 | 6.1 | 6.19 | 6.28 | 6.37 | 6.46 | 6.56 | 6.65 | 6.74 | 6.83
39.5 | 3.17 | 3.26 | 3.35 | 3.44 | 3.54 | 3.63 | 3.72 | 3.81 | 3.9 | 3.99 | 4.08 | 4.18 | 4.27 | 4.36 | 4.45 | 4.54 | 4.63 | 4.73 | 4.82 | 4.91 | 5 | 5.09 | 5.18 | 5.27 | 5.37 | 5.46 | 5.55 | 5.64 | 5.73 | 5.82 | 5.91 | 6.01 | 6.1 | 6.19 | 6.28 | 6.37 | 6.46 | 6.56 | 6.65 | 6.74 | 6.83
39.51 | 3.17 | 3.26 | 3.35 | 3.44 | 3.53 | 3.63 | 3.72 | 3.81 | 3.9 | 3.99 | 4.08 | 4.18 | 4.27 | 4.36 | 4.45 | 4.54 | 4.63 | 4.72 | 4.82 | 4.91 | 5 | 5.09 | 5.18 | 5.27 | 5.36 | 5.46 | 5.55 | 5.64 | 5.73 | 5.82 | 5.91 | 6.01 | 6.1 | 6.19 | 6.28 | 6.37 | 6.46 | 6.55 | 6.65 | 6.74 | 6.83
39.52 | 3.17 | 3.26 | 3.35 | 3.44 | 3.53 | 3.63 | 3.72 | 3.81 | 3.9 | 3.99 | 4.08 | 4.17 | 4.27 | 4.36 | 4.45 | 4.54 | 4.63 | 4.72 | 4.81 | 4.91 | 5 | 5.09 | 5.18 | 5.27 | 5.36 | 5.46 | 5.55 | 5.64 | 5.73 | 5.82 | 5.91 | 6 | 6.1 | 6.19 | 6.28 | 6.37 | 6.46 | 6.55 | 6.64 | 6.74 | 6.83
39.53 | 3.17 | 3.26 | 3.35 | 3.44 | 3.53 | 3.62 | 3.72 | 3.81 | 3.9 | 3.99 | 4.08 | 4.17 | 4.27 | 4.36 | 4.45 | 4.54 | 4.63 | 4.72 | 4.81 | 4.91 | 5 | 5.09 | 5.18 | 5.27 | 5.36 | 5.45 | 5.55 | 5.64 | 5.73 | 5.82 | 5.91 | 6 | 6.09 | 6.19 | 6.28 | 6.37 | 6.46 | 6.55 | 6.64 | 6.73 | 6.83
39.54 | 3.17 | 3.26 | 3.35 | 3.44 | 3.53 | 3.62 | 3.72 | 3.81 | 3.9 | 3.99 | 4.08 | 4.17 | 4.26 | 4.36 | 4.45 | 4.54 | 4.63 | 4.72 | 4.81 | 4.9 | 5 | 5.09 | 5.18 | 5.27 | 5.36 | 5.45 | 5.54 | 5.64 | 5.73 | 5.82 | 5.91 | 6 | 6.09 | 6.19 | 6.28 | 6.37 | 6.46 | 6.55 | 6.64 | 6.73 | 6.83
39.55 | 3.17 | 3.26 | 3.35 | 3.44 | 3.53 | 3.62 | 3.71 | 3.81 | 3.9 | 3.99 | 4.08 | 4.17 | 4.26 | 4.36 | 4.45 | 4.54 | 4.63 | 4.72 | 4.81 | 4.9 | 5 | 5.09 | 5.18 | 5.27 | 5.36 | 5.45 | 5.54 | 5.64 | 5.73 | 5.82 | 5.91 | 6 | 6.09 | 6.18 | 6.28 | 6.37 | 6.46 | 6.55 | 6.64 | 6.73 | 6.82
39.56 | 3.17 | 3.26 | 3.35 | 3.44 | 3.53 | 3.62 | 3.71 | 3.81 | 3.9 | 3.99 | 4.08 | 4.17 | 4.26 | 4.35 | 4.45 | 4.54 | 4.63 | 4.72 | 4.81 | 4.9 | 4.99 | 5.09 | 5.18 | 5.27 | 5.36 | 5.45 | 5.54 | 5.63 | 5.73 | 5.82 | 5.91 | 6 | 6.09 | 6.18 | 6.27 | 6.37 | 6.46 | 6.55 | 6.64 | 6.73 | 6.82
39.57 | 3.16 | 3.26 | 3.35 | 3.44 | 3.53 | 3.62 | 3.71 | 3.8 | 3.9 | 3.99 | 4.08 | 4.17 | 4.26 | 4.35 | 4.44 | 4.54 | 4.63 | 4.72 | 4.81 | 4.9 | 4.99 | 5.09 | 5.18 | 5.27 | 5.36 | 5.45 | 5.54 | 5.63 | 5.73 | 5.82 | 5.91 | 6 | 6.09 | 6.18 | 6.27 | 6.37 | 6.46 | 6.55 | 6.64 | 6.73 | 6.82
39.58 | 3.16 | 3.26 | 3.35 | 3.44 | 3.53 | 3.62 | 3.71 | 3.8 | 3.9 | 3.99 | 4.08 | 4.17 | 4.26 | 4.35 | 4.44 | 4.54 | 4.63 | 4.72 | 4.81 | 4.9 | 4.99 | 5.08 | 5.18 | 5.27 | 5.36 | 5.45 | 5.54 | 5.63 | 5.72 | 5.82 | 5.91 | 6 | 6.09 | 6.18 | 6.27 | 6.36 | 6.46 | 6.55 | 6.64 | 6.73 | 6.82
39.59 | 3.16 | 3.25 | 3.35 | 3.44 | 3.53 | 3.62 | 3.71 | 3.8 | 3.89 | 3.99 | 4.08 | 4.17 | 4.26 | 4.35 | 4.44 | 4.53 | 4.63 | 4.72 | 4.81 | 4.9 | 4.99 | 5.08 | 5.17 | 5.27 | 5.36 | 5.45 | 5.54 | 5.63 | 5.72 | 5.81 | 5.91 | 6 | 6.09 | 6.18 | 6.27 | 6.36 | 6.45 | 6.55 | 6.64 | 6.73 | 6.82
39.6 | 3.16 | 3.25 | 3.35 | 3.44 | 3.53 | 3.62 | 3.71 | 3.8 | 3.89 | 3.99 | 4.08 | 4.17 | 4.26 | 4.35 | 4.44 | 4.53 | 4.63 | 4.72 | 4.81 | 4.9 | 4.99 | 5.08 | 5.17 | 5.27 | 5.36 | 5.45 | 5.54 | 5.63 | 5.72 | 5.81 | 5.91 | 6 | 6.09 | 6.18 | 6.27 | 6.36 | 6.45 | 6.54 | 6.64 | 6.73 | 6.82
39.61 | 3.16 | 3.25 | 3.34 | 3.44 | 3.53 | 3.62 | 3.71 | 3.8 | 3.89 | 3.98 | 4.08 | 4.17 | 4.26 | 4.35 | 4.44 | 4.53 | 4.62 | 4.72 | 4.81 | 4.9 | 4.99 | 5.08 | 5.17 | 5.26 | 5.36 | 5.45 | 5.54 | 5.63 | 5.72 | 5.81 | 5.9 | 6 | 6.09 | 6.18 | 6.27 | 6.36 | 6.45 | 6.54 | 6.64 | 6.73 | 6.82
39.62 | 3.16 | 3.25 | 3.34 | 3.44 | 3.53 | 3.62 | 3.71 | 3.8 | 3.89 | 3.98 | 4.08 | 4.17 | 4.26 | 4.35 | 4.44 | 4.53 | 4.62 | 4.72 | 4.81 | 4.9 | 4.99 | 5.08 | 5.17 | 5.26 | 5.35 | 5.45 | 5.54 | 5.63 | 5.72 | 5.81 | 5.9 | 5.99 | 6.09 | 6.18 | 6.27 | 6.36 | 6.45 | 6.54 | 6.63 | 6.73 | 6.82
39.63 | 3.16 | 3.25 | 3.34 | 3.44 | 3.53 | 3.62 | 3.71 | 3.8 | 3.89 | 3.98 | 4.07 | 4.17 | 4.26 | 4.35 | 4.44 | 4.53 | 4.62 | 4.71 | 4.81 | 4.9 | 4.99 | 5.08 | 5.17 | 5.26 | 5.35 | 5.45 | 5.54 | 5.63 | 5.72 | 5.81 | 5.9 | 5.99 | 6.08 | 6.18 | 6.27 | 6.36 | 6.45 | 6.54 | 6.63 | 6.72 | 6.82
39.64 | 3.16 | 3.25 | 3.34 | 3.43 | 3.53 | 3.62 | 3.71 | 3.8 | 3.89 | 3.98 | 4.07 | 4.17 | 4.26 | 4.35 | 4.44 | 4.53 | 4.62 | 4.71 | 4.8 | 4.9 | 4.99 | 5.08 | 5.17 | 5.26 | 5.35 | 5.44 | 5.54 | 5.63 | 5.72 | 5.81 | 5.9 | 5.99 | 6.08 | 6.18 | 6.27 | 6.36 | 6.45 | 6.54 | 6.63 | 6.72 | 6.81
39.65 | 3.16 | 3.25 | 3.34 | 3.43 | 3.53 | 3.62 | 3.71 | 3.8 | 3.89 | 3.98 | 4.07 | 4.16 | 4.26 | 4.35 | 4.44 | 4.53 | 4.62 | 4.71 | 4.8 | 4.9 | 4.99 | 5.08 | 5.17 | 5.26 | 5.35 | 5.44 | 5.53 | 5.63 | 5.72 | 5.81 | 5.9 | 5.99 | 6.08 | 6.17 | 6.27 | 6.36 | 6.45 | 6.54 | 6.63 | 6.72 | 6.81
39.66 | 3.16 | 3.25 | 3.34 | 3.43 | 3.52 | 3.62 | 3.71 | 3.8 | 3.89 | 3.98 | 4.07 | 4.16 | 4.26 | 4.35 | 4.44 | 4.53 | 4.62 | 4.71 | 4.8 | 4.89 | 4.99 | 5.08 | 5.17 | 5.26 | 5.35 | 5.44 | 5.53 | 5.63 | 5.72 | 5.81 | 5.9 | 5.99 | 6.08 | 6.17 | 6.26 | 6.36 | 6.45 | 6.54 | 6.63 | 6.72 | 6.81
39.67 | 3.16 | 3.25 | 3.34 | 3.43 | 3.52 | 3.62 | 3.71 | 3.8 | 3.89 | 3.98 | 4.07 | 4.16 | 4.25 | 4.35 | 4.44 | 4.53 | 4.62 | 4.71 | 4.8 | 4.89 | 4.99 | 5.08 | 5.17 | 5.26 | 5.35 | 5.44 | 5.53 | 5.62 | 5.72 | 5.81 | 5.9 | 5.99 | 6.08 | 6.17 | 6.26 | 6.35 | 6.45 | 6.54 | 6.63 | 6.72 | 6.81
39.68 | 3.16 | 3.25 | 3.34 | 3.43 | 3.52 | 3.61 | 3.71 | 3.8 | 3.89 | 3.98 | 4.07 | 4.16 | 4.25 | 4.35 | 4.44 | 4.53 | 4.62 | 4.71 | 4.8 | 4.89 | 4.98 | 5.08 | 5.17 | 5.26 | 5.35 | 5.44 | 5.53 | 5.62 | 5.71 | 5.81 | 5.9 | 5.99 | 6.08 | 6.17 | 6.26 | 6.35 | 6.45 | 6.54 | 6.63 | 6.72 | 6.81
39.69 | 3.16 | 3.25 | 3.34 | 3.43 | 3.52 | 3.61 | 3.71 | 3.8 | 3.89 | 3.98 | 4.07 | 4.16 | 4.25 | 4.34 | 4.44 | 4.53 | 4.62 | 4.71 | 4.8 | 4.89 | 4.98 | 5.07 | 5.17 | 5.26 | 5.35 | 5.44 | 5.53 | 5.62 | 5.71 | 5.81 | 5.9 | 5.99 | 6.08 | 6.17 | 6.26 | 6.35 | 6.44 | 6.54 | 6.63 | 6.72 | 6.81
39.7 | 3.16 | 3.25 | 3.34 | 3.43 | 3.52 | 3.61 | 3.7 | 3.8 | 3.89 | 3.98 | 4.07 | 4.16 | 4.25 | 4.34 | 4.43 | 4.53 | 4.62 | 4.71 | 4.8 | 4.89 | 4.98 | 5.07 | 5.17 | 5.26 | 5.35 | 5.44 | 5.53 | 5.62 | 5.71 | 5.8 | 5.9 | 5.99 | 6.08 | 6.17 | 6.26 | 6.35 | 6.44 | 6.53 | 6.63 | 6.72 | 6.81
39.71 | 3.16 | 3.25 | 3.34 | 3.43 | 3.52 | 3.61 | 3.7 | 3.79 | 3.89 | 3.98 | 4.07 | 4.16 | 4.25 | 4.34 | 4.43 | 4.53 | 4.62 | 4.71 | 4.8 | 4.89 | 4.98 | 5.07 | 5.16 | 5.26 | 5.35 | 5.44 | 5.53 | 5.62 | 5.71 | 5.8 | 5.89 | 5.99 | 6.08 | 6.17 | 6.26 | 6.35 | 6.44 | 6.53 | 6.62 | 6.72 | 6.81
39.72 | 3.16 | 3.25 | 3.34 | 3.43 | 3.52 | 3.61 | 3.7 | 3.79 | 3.89 | 3.98 | 4.07 | 4.16 | 4.25 | 4.34 | 4.43 | 4.52 | 4.62 | 4.71 | 4.8 | 4.89 | 4.98 | 5.07 | 5.16 | 5.25 | 5.35 | 5.44 | 5.53 | 5.62 | 5.71 | 5.8 | 5.89 | 5.98 | 6.08 | 6.17 | 6.26 | 6.35 | 6.44 | 6.53 | 6.62 | 6.71 | 6.81
39.73 | 3.15 | 3.25 | 3.34 | 3.43 | 3.52 | 3.61 | 3.7 | 3.79 | 3.88 | 3.98 | 4.07 | 4.16 | 4.25 | 4.34 | 4.43 | 4.52 | 4.61 | 4.71 | 4.8 | 4.89 | 4.98 | 5.07 | 5.16 | 5.25 | 5.35 | 5.44 | 5.53 | 5.62 | 5.71 | 5.8 | 5.89 | 5.98 | 6.08 | 6.17 | 6.26 | 6.35 | 6.44 | 6.53 | 6.62 | 6.71 | 6.81
39.74 | 3.15 | 3.25 | 3.34 | 3.43 | 3.52 | 3.61 | 3.7 | 3.79 | 3.88 | 3.98 | 4.07 | 4.16 | 4.25 | 4.34 | 4.43 | 4.52 | 4.61 | 4.71 | 4.8 | 4.89 | 4.98 | 5.07 | 5.16 | 5.25 | 5.34 | 5.44 | 5.53 | 5.62 | 5.71 | 5.8 | 5.89 | 5.98 | 6.07 | 6.17 | 6.26 | 6.35 | 6.44 | 6.53 | 6.62 | 6.71 | 6.8
39.75 | 3.15 | 3.24 | 3.34 | 3.43 | 3.52 | 3.61 | 3.7 | 3.79 | 3.88 | 3.97 | 4.07 | 4.16 | 4.25 | 4.34 | 4.43 | 4.52 | 4.61 | 4.7 | 4.8 | 4.89 | 4.98 | 5.07 | 5.16 | 5.25 | 5.34 | 5.43 | 5.53 | 5.62 | 5.71 | 5.8 | 5.89 | 5.98 | 6.07 | 6.16 | 6.26 | 6.35 | 6.44 | 6.53 | 6.62 | 6.71 | 6.8
39.76 | 3.15 | 3.24 | 3.34 | 3.43 | 3.52 | 3.61 | 3.7 | 3.79 | 3.88 | 3.97 | 4.07 | 4.16 | 4.25 | 4.34 | 4.43 | 4.52 | 4.61 | 4.7 | 4.8 | 4.89 | 4.98 | 5.07 | 5.16 | 5.25 | 5.34 | 5.43 | 5.52 | 5.62 | 5.71 | 5.8 | 5.89 | 5.98 | 6.07 | 6.16 | 6.25 | 6.35 | 6.44 | 6.53 | 6.62 | 6.71 | 6.8
39.77 | 3.15 | 3.24 | 3.33 | 3.43 | 3.52 | 3.61 | 3.7 | 3.79 | 3.88 | 3.97 | 4.06 | 4.16 | 4.25 | 4.34 | 4.43 | 4.52 | 4.61 | 4.7 | 4.79 | 4.89 | 4.98 | 5.07 | 5.16 | 5.25 | 5.34 | 5.43 | 5.52 | 5.62 | 5.71 | 5.8 | 5.89 | 5.98 | 6.07 | 6.16 | 6.25 | 6.34 | 6.44 | 6.53 | 6.62 | 6.71 | 6.8
39.78 | 3.15 | 3.24 | 3.33 | 3.43 | 3.52 | 3.61 | 3.7 | 3.79 | 3.88 | 3.97 | 4.06 | 4.15 | 4.25 | 4.34 | 4.43 | 4.52 | 4.61 | 4.7 | 4.79 | 4.88 | 4.98 | 5.07 | 5.16 | 5.25 | 5.34 | 5.43 | 5.52 | 5.61 | 5.71 | 5.8 | 5.89 | 5.98 | 6.07 | 6.16 | 6.25 | 6.34 | 6.44 | 6.53 | 6.62 | 6.71 | 6.8
39.79 | 3.15 | 3.24 | 3.33 | 3.42 | 3.52 | 3.61 | 3.7 | 3.79 | 3.88 | 3.97 | 4.06 | 4.15 | 4.25 | 4.34 | 4.43 | 4.52 | 4.61 | 4.7 | 4.79 | 4.88 | 4.97 | 5.07 | 5.16 | 5.25 | 5.34 | 5.43 | 5.52 | 5.61 | 5.7 | 5.8 | 5.89 | 5.98 | 6.07 | 6.16 | 6.25 | 6.34 | 6.43 | 6.53 | 6.62 | 6.71 | 6.8
39.8 | 3.15 | 3.24 | 3.33 | 3.42 | 3.51 | 3.61 | 3.7 | 3.79 | 3.88 | 3.97 | 4.06 | 4.15 | 4.24 | 4.34 | 4.43 | 4.52 | 4.61 | 4.7 | 4.79 | 4.88 | 4.97 | 5.07 | 5.16 | 5.25 | 5.34 | 5.43 | 5.52 | 5.61 | 5.7 | 5.79 | 5.89 | 5.98 | 6.07 | 6.16 | 6.25 | 6.34 | 6.43 | 6.52 | 6.62 | 6.71 | 6.8
39.81 | 3.15 | 3.24 | 3.33 | 3.42 | 3.51 | 3.61 | 3.7 | 3.79 | 3.88 | 3.97 | 4.06 | 4.15 | 4.24 | 4.33 | 4.43 | 4.52 | 4.61 | 4.7 | 4.79 | 4.88 | 4.97 | 5.06 | 5.16 | 5.25 | 5.34 | 5.43 | 5.52 | 5.61 | 5.7 | 5.79 | 5.89 | 5.98 | 6.07 | 6.16 | 6.25 | 6.34 | 6.43 | 6.52 | 6.61 | 6.71 | 6.8
39.82 | 3.15 | 3.24 | 3.33 | 3.42 | 3.51 | 3.6 | 3.7 | 3.79 | 3.88 | 3.97 | 4.06 | 4.15 | 4.24 | 4.33 | 4.43 | 4.52 | 4.61 | 4.7 | 4.79 | 4.88 | 4.97 | 5.06 | 5.15 | 5.25 | 5.34 | 5.43 | 5.52 | 5.61 | 5.7 | 5.79 | 5.88 | 5.98 | 6.07 | 6.16 | 6.25 | 6.34 | 6.43 | 6.52 | 6.61 | 6.7 | 6.8
39.83 | 3.15 | 3.24 | 3.33 | 3.42 | 3.51 | 3.6 | 3.7 | 3.79 | 3.88 | 3.97 | 4.06 | 4.15 | 4.24 | 4.33 | 4.42 | 4.52 | 4.61 | 4.7 | 4.79 | 4.88 | 4.97 | 5.06 | 5.15 | 5.25 | 5.34 | 5.43 | 5.52 | 5.61 | 5.7 | 5.79 | 5.88 | 5.97 | 6.07 | 6.16 | 6.25 | 6.34 | 6.43 | 6.52 | 6.61 | 6.7 | 6.79
39.84 | 3.15 | 3.24 | 3.33 | 3.42 | 3.51 | 3.6 | 3.69 | 3.79 | 3.88 | 3.97 | 4.06 | 4.15 | 4.24 | 4.33 | 4.42 | 4.51 | 4.61 | 4.7 | 4.79 | 4.88 | 4.97 | 5.06 | 5.15 | 5.24 | 5.34 | 5.43 | 5.52 | 5.61 | 5.7 | 5.79 | 5.88 | 5.97 | 6.06 | 6.16 | 6.25 | 6.34 | 6.43 | 6.52 | 6.61 | 6.7 | 6.79
39.85 | 3.15 | 3.24 | 3.33 | 3.42 | 3.51 | 3.6 | 3.69 | 3.78 | 3.88 | 3.97 | 4.06 | 4.15 | 4.24 | 4.33 | 4.42 | 4.51 | 4.61 | 4.7 | 4.79 | 4.88 | 4.97 | 5.06 | 5.15 | 5.24 | 5.33 | 5.43 | 5.52 | 5.61 | 5.7 | 5.79 | 5.88 | 5.97 | 6.06 | 6.15 | 6.25 | 6.34 | 6.43 | 6.52 | 6.61 | 6.7 | 6.79
39.86 | 3.15 | 3.24 | 3.33 | 3.42 | 3.51 | 3.6 | 3.69 | 3.78 | 3.88 | 3.97 | 4.06 | 4.15 | 4.24 | 4.33 | 4.42 | 4.51 | 4.6 | 4.7 | 4.79 | 4.88 | 4.97 | 5.06 | 5.15 | 5.24 | 5.33 | 5.42 | 5.52 | 5.61 | 5.7 | 5.79 | 5.88 | 5.97 | 6.06 | 6.15 | 6.24 | 6.34 | 6.43 | 6.52 | 6.61 | 6.7 | 6.79
39.87 | 3.15 | 3.24 | 3.33 | 3.42 | 3.51 | 3.6 | 3.69 | 3.78 | 3.87 | 3.97 | 4.06 | 4.15 | 4.24 | 4.33 | 4.42 | 4.51 | 4.6 | 4.69 | 4.79 | 4.88 | 4.97 | 5.06 | 5.15 | 5.24 | 5.33 | 5.42 | 5.51 | 5.61 | 5.7 | 5.79 | 5.88 | 5.97 | 6.06 | 6.15 | 6.24 | 6.33 | 6.43 | 6.52 | 6.61 | 6.7 | 6.79
39.88 | 3.14 | 3.24 | 3.33 | 3.42 | 3.51 | 3.6 | 3.69 | 3.78 | 3.87 | 3.97 | 4.06 | 4.15 | 4.24 | 4.33 | 4.42 | 4.51 | 4.6 | 4.69 | 4.79 | 4.88 | 4.97 | 5.06 | 5.15 | 5.24 | 5.33 | 5.42 | 5.51 | 5.61 | 5.7 | 5.79 | 5.88 | 5.97 | 6.06 | 6.15 | 6.24 | 6.33 | 6.42 | 6.52 | 6.61 | 6.7 | 6.79
39.89 | 3.14 | 3.24 | 3.33 | 3.42 | 3.51 | 3.6 | 3.69 | 3.78 | 3.87 | 3.96 | 4.06 | 4.15 | 4.24 | 4.33 | 4.42 | 4.51 | 4.6 | 4.69 | 4.78 | 4.88 | 4.97 | 5.06 | 5.15 | 5.24 | 5.33 | 5.42 | 5.51 | 5.6 | 5.7 | 5.79 | 5.88 | 5.97 | 6.06 | 6.15 | 6.24 | 6.33 | 6.42 | 6.51 | 6.61 | 6.7 | 6.79
39.9 | 3.14 | 3.23 | 3.33 | 3.42 | 3.51 | 3.6 | 3.69 | 3.78 | 3.87 | 3.96 | 4.05 | 4.15 | 4.24 | 4.33 | 4.42 | 4.51 | 4.6 | 4.69 | 4.78 | 4.87 | 4.97 | 5.06 | 5.15 | 5.24 | 5.33 | 5.42 | 5.51 | 5.6 | 5.69 | 5.79 | 5.88 | 5.97 | 6.06 | 6.15 | 6.24 | 6.33 | 6.42 | 6.51 | 6.6 | 6.7 | 6.79
39.91 | 3.14 | 3.23 | 3.33 | 3.42 | 3.51 | 3.6 | 3.69 | 3.78 | 3.87 | 3.96 | 4.05 | 4.15 | 4.24 | 4.33 | 4.42 | 4.51 | 4.6 | 4.69 | 4.78 | 4.87 | 4.96 | 5.06 | 5.15 | 5.24 | 5.33 | 5.42 | 5.51 | 5.6 | 5.69 | 5.78 | 5.88 | 5.97 | 6.06 | 6.15 | 6.24 | 6.33 | 6.42 | 6.51 | 6.6 | 6.7 | 6.79
39.92 | 3.14 | 3.23 | 3.32 | 3.42 | 3.51 | 3.6 | 3.69 | 3.78 | 3.87 | 3.96 | 4.05 | 4.14 | 4.24 | 4.33 | 4.42 | 4.51 | 4.6 | 4.69 | 4.78 | 4.87 | 4.96 | 5.06 | 5.15 | 5.24 | 5.33 | 5.42 | 5.51 | 5.6 | 5.69 | 5.78 | 5.87 | 5.97 | 6.06 | 6.15 | 6.24 | 6.33 | 6.42 | 6.51 | 6.6 | 6.69 | 6.78
39.93 | 3.14 | 3.23 | 3.32 | 3.42 | 3.51 | 3.6 | 3.69 | 3.78 | 3.87 | 3.96 | 4.05 | 4.14 | 4.23 | 4.33 | 4.42 | 4.51 | 4.6 | 4.69 | 4.78 | 4.87 | 4.96 | 5.05 | 5.15 | 5.24 | 5.33 | 5.42 | 5.51 | 5.6 | 5.69 | 5.78 | 5.87 | 5.96 | 6.06 | 6.15 | 6.24 | 6.33 | 6.42 | 6.51 | 6.6 | 6.69 | 6.78
39.94 | 3.14 | 3.23 | 3.32 | 3.41 | 3.51 | 3.6 | 3.69 | 3.78 | 3.87 | 3.96 | 4.05 | 4.14 | 4.23 | 4.32 | 4.42 | 4.51 | 4.6 | 4.69 | 4.78 | 4.87 | 4.96 | 5.05 | 5.14 | 5.24 | 5.33 | 5.42 | 5.51 | 5.6 | 5.69 | 5.78 | 5.87 | 5.96 | 6.05 | 6.15 | 6.24 | 6.33 | 6.42 | 6.51 | 6.6 | 6.69 | 6.78
39.95 | 3.14 | 3.23 | 3.32 | 3.41 | 3.5 | 3.6 | 3.69 | 3.78 | 3.87 | 3.96 | 4.05 | 4.14 | 4.23 | 4.32 | 4.42 | 4.51 | 4.6 | 4.69 | 4.78 | 4.87 | 4.96 | 5.05 | 5.14 | 5.23 | 5.33 | 5.42 | 5.51 | 5.6 | 5.69 | 5.78 | 5.87 | 5.96 | 6.05 | 6.14 | 6.24 | 6.33 | 6.42 | 6.51 | 6.6 | 6.69 | 6.78
39.96 | 3.14 | 3.23 | 3.32 | 3.41 | 3.5 | 3.6 | 3.69 | 3.78 | 3.87 | 3.96 | 4.05 | 4.14 | 4.23 | 4.32 | 4.41 | 4.51 | 4.6 | 4.69 | 4.78 | 4.87 | 4.96 | 5.05 | 5.14 | 5.23 | 5.32 | 5.42 | 5.51 | 5.6 | 5.69 | 5.78 | 5.87 | 5.96 | 6.05 | 6.14 | 6.23 | 6.33 | 6.42 | 6.51 | 6.6 | 6.69 | 6.78
39.97 | 3.14 | 3.23 | 3.32 | 3.41 | 3.5 | 3.59 | 3.69 | 3.78 | 3.87 | 3.96 | 4.05 | 4.14 | 4.23 | 4.32 | 4.41 | 4.5 | 4.6 | 4.69 | 4.78 | 4.87 | 4.96 | 5.05 | 5.14 | 5.23 | 5.32 | 5.41 | 5.51 | 5.6 | 5.69 | 5.78 | 5.87 | 5.96 | 6.05 | 6.14 | 6.23 | 6.32 | 6.42 | 6.51 | 6.6 | 6.69 | 6.78
39.98 | 3.14 | 3.23 | 3.32 | 3.41 | 3.5 | 3.59 | 3.68 | 3.78 | 3.87 | 3.96 | 4.05 | 4.14 | 4.23 | 4.32 | 4.41 | 4.5 | 4.59 | 4.69 | 4.78 | 4.87 | 4.96 | 5.05 | 5.14 | 5.23 | 5.32 | 5.41 | 5.5 | 5.6 | 5.69 | 5.78 | 5.87 | 5.96 | 6.05 | 6.14 | 6.23 | 6.32 | 6.41 | 6.51 | 6.6 | 6.69 | 6.78
39.99 | 3.14 | 3.23 | 3.32 | 3.41 | 3.5 | 3.59 | 3.68 | 3.78 | 3.87 | 3.96 | 4.05 | 4.14 | 4.23 | 4.32 | 4.41 | 4.5 | 4.59 | 4.69 | 4.78 | 4.87 | 4.96 | 5.05 | 5.14 | 5.23 | 5.32 | 5.41 | 5.5 | 5.59 | 5.69 | 5.78 | 5.87 | 5.96 | 6.05 | 6.14 | 6.23 | 6.32 | 6.41 | 6.5 | 6.6 | 6.69 | 6.78
40 | 3.14 | 3.23 | 3.32 | 3.41 | 3.5 | 3.59 | 3.68 | 3.77 | 3.87 | 3.96 | 4.05 | 4.14 | 4.23 | 4.32 | 4.41 | 4.5 | 4.59 | 4.68 | 4.78 | 4.87 | 4.96 | 5.05 | 5.14 | 5.23 | 5.32 | 5.41 | 5.5 | 5.59 | 5.69 | 5.78 | 5.87 | 5.96 | 6.05 | 6.14 | 6.23 | 6.32 | 6.41 | 6.5 | 6.59 | 6.69 | 6.78

## DO-Cond_Table
DOTABLES Result
Salinity Correction Factor Table
Salinity correction factors for dissolved oxygen in water based on specific conductance
[Factors are dimensionless. Values based on published equations by Benson and Krause (1984).
°C, degrees Celsius; µS/cm, microsiemens/cm]
Temp. | Specific conductance (µS/cm)
(°C) | 0 | 100 | 200 | 300 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | 1100 | 1200 | 1300 | 1400 | 1500 | 1600 | 1700 | 1800 | 1900 | 2000 | 2100 | 2200 | 2300 | 2400 | 2500 | 2600 | 2700 | 2800 | 2900 | 3000 | 3100 | 3200 | 3300 | 3400 | 3500 | 3600 | 3700 | 3800 | 3900 | 4000 | 4100 | 4200 | 4300 | 4400 | 4500 | 4600 | 4700 | 4800 | 4900 | 5000 | 5100 | 5200 | 5300 | 5400 | 5500 | 5600 | 5700 | 5800 | 5900 | 6000 | 6100 | 6200 | 6300 | 6400 | 6500 | 6600 | 6700 | 6800 | 6900 | 7000 | 7100 | 7200 | 7300 | 7400 | 7500 | 7600 | 7700 | 7800 | 7900 | 8000 | 8100 | 8200 | 8300 | 8400 | 8500 | 8600 | 8700 | 8800 | 8900 | 9000 | 9100 | 9200 | 9300 | 9400 | 9500 | 9600 | 9700 | 9800 | 9900 | 10000 | 10100 | 10200 | 10300 | 10400 | 10500 | 10600 | 10700 | 10800 | 10900 | 11000 | 11100 | 11200 | 11300 | 11400 | 11500 | 11600 | 11700 | 11800 | 11900 | 12000 | 12100 | 12200 | 12300 | 12400 | 12500 | 12600 | 12700 | 12800 | 12900 | 13000 | 13100 | 13200 | 13300 | 13400 | 13500 | 13600 | 13700 | 13800 | 13900 | 14000 | 14100 | 14200 | 14300 | 14400 | 14500 | 14600 | 14700 | 14800 | 14900 | 15000 | 15100 | 15200 | 15300 | 15400 | 15500 | 15600 | 15700 | 15800 | 15900 | 16000 | 16100 | 16200 | 16300 | 16400 | 16500 | 16600 | 16700 | 16800 | 16900 | 17000 | 17100 | 17200 | 17300 | 17400 | 17500 | 17600 | 17700 | 17800 | 17900 | 18000 | 18100 | 18200 | 18300 | 18400 | 18500 | 18600 | 18700 | 18800 | 18900 | 19000 | 19100 | 19200 | 19300 | 19400 | 19500 | 19600 | 19700 | 19800 | 19900 | 20000 | 20100 | 20200 | 20300 | 20400 | 20500 | 20600 | 20700 | 20800 | 20900 | 21000 | 21100 | 21200 | 21300 | 21400 | 21500 | 21600 | 21700 | 21800 | 21900 | 22000 | 22100 | 22200 | 22300 | 22400 | 22500 | 22600 | 22700 | 22800 | 22900 | 23000 | 23100 | 23200 | 23300 | 23400 | 23500 | 23600 | 23700 | 23800 | 23900 | 24000 | 24100 | 24200 | 24300 | 24400 | 24500 | 24600 | 24700 | 24800 | 24900 | 25000 | 25100 | 25200 | 25300 | 25400 | 25500 | 25600 | 25700 | 25800 | 25900 | 26000 | 26100 | 26200 | 26300 | 26400 | 26500 | 26600 | 26700 | 26800 | 26900 | 27000 | 27100 | 27200 | 27300 | 27400 | 27500 | 27600 | 27700 | 27800 | 27900 | 28000 | 28100 | 28200 | 28300 | 28400 | 28500 | 28600 | 28700 | 28800 | 28900 | 29000 | 29100 | 29200 | 29300 | 29400 | 29500 | 29600 | 29700 | 29800 | 29900 | 30000 | 30100 | 30200 | 30300 | 30400 | 30500 | 30600 | 30700 | 30800 | 30900 | 31000 | 31100 | 31200 | 31300 | 31400 | 31500 | 31600 | 31700 | 31800 | 31900 | 32000 | 32100 | 32200 | 32300 | 32400 | 32500 | 32600 | 32700 | 32800 | 32900 | 33000 | 33100 | 33200 | 33300 | 33400 | 33500 | 33600 | 33700 | 33800 | 33900 | 34000 | 34100 | 34200 | 34300 | 34400 | 34500 | 34600 | 34700 | 34800 | 34900 | 35000 | 35100 | 35200 | 35300 | 35400 | 35500 | 35600 | 35700 | 35800 | 35900 | 36000 | 36100 | 36200 | 36300 | 36400 | 36500 | 36600 | 36700 | 36800 | 36900 | 37000 | 37100 | 37200 | 37300 | 37400 | 37500 | 37600 | 37700 | 37800 | 37900 | 38000 | 38100 | 38200 | 38300 | 38400 | 38500 | 38600 | 38700 | 38800 | 38900 | 39000 | 39100 | 39200 | 39300 | 39400 | 39500 | 39600 | 39700 | 39800 | 39900 | 40000
| 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80 | 81 | 82 | 83 | 84 | 85 | 86 | 87 | 88 | 89 | 90 | 91 | 92 | 93 | 94 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 117 | 118 | 119 | 120 | 121 | 122 | 123 | 124 | 125 | 126 | 127 | 128 | 129 | 130 | 131 | 132 | 133 | 134 | 135 | 136 | 137 | 138 | 139 | 140 | 141 | 142 | 143 | 144 | 145 | 146 | 147 | 148 | 149 | 150 | 151 | 152 | 153 | 154 | 155 | 156 | 157 | 158 | 159 | 160 | 161 | 162 | 163 | 164 | 165 | 166 | 167 | 168 | 169 | 170 | 171 | 172 | 173 | 174 | 175 | 176 | 177 | 178 | 179 | 180 | 181 | 182 | 183 | 184 | 185 | 186 | 187 | 188 | 189 | 190 | 191 | 192 | 193 | 194 | 195 | 196 | 197 | 198 | 199 | 200 | 201 | 202 | 203 | 204 | 205 | 206 | 207 | 208 | 209 | 210 | 211 | 212 | 213 | 214 | 215 | 216 | 217 | 218 | 219 | 220 | 221 | 222 | 223 | 224 | 225 | 226 | 227 | 228 | 229 | 230 | 231 | 232 | 233 | 234 | 235 | 236 | 237 | 238 | 239 | 240 | 241 | 242 | 243 | 244 | 245 | 246 | 247 | 248 | 249 | 250 | 251 | 252 | 253 | 254 | 255 | 256 | 257 | 258 | 259 | 260 | 261 | 262 | 263 | 264 | 265 | 266 | 267 | 268 | 269 | 270 | 271 | 272 | 273 | 274 | 275 | 276 | 277 | 278 | 279 | 280 | 281 | 282 | 283 | 284 | 285 | 286 | 287 | 288 | 289 | 290 | 291 | 292 | 293 | 294 | 295 | 296 | 297 | 298 | 299 | 300 | 301 | 302 | 303 | 304 | 305 | 306 | 307 | 308 | 309 | 310 | 311 | 312 | 313 | 314 | 315 | 316 | 317 | 318 | 319 | 320 | 321 | 322 | 323 | 324 | 325 | 326 | 327 | 328 | 329 | 330 | 331 | 332 | 333 | 334 | 335 | 336 | 337 | 338 | 339 | 340 | 341 | 342 | 343 | 344 | 345 | 346 | 347 | 348 | 349 | 350 | 351 | 352 | 353 | 354 | 355 | 356 | 357 | 358 | 359 | 360 | 361 | 362 | 363 | 364 | 365 | 366 | 367 | 368 | 369 | 370 | 371 | 372 | 373 | 374 | 375 | 376 | 377 | 378 | 379 | 380 | 381 | 382 | 383 | 384 | 385 | 386 | 387 | 388 | 389 | 390 | 391 | 392 | 393 | 394 | 395 | 396 | 397 | 398 | 399 | 400 | 401 | 402
0 | 1 | 1 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.866 | 0.866 | 0.865 | 0.865 | 0.865 | 0.864 | 0.864 | 0.863 | 0.863 | 0.863 | 0.862 | 0.862 | 0.861 | 0.861 | 0.86 | 0.86 | 0.86 | 0.859 | 0.859 | 0.858 | 0.858 | 0.858 | 0.857 | 0.857 | 0.856 | 0.856 | 0.855 | 0.855 | 0.855 | 0.854 | 0.854 | 0.853 | 0.853 | 0.852 | 0.852 | 0.852 | 0.851 | 0.851 | 0.85 | 0.85 | 0.85 | 0.849 | 0.849 | 0.848 | 0.848 | 0.847 | 0.847 | 0.847 | 0.846 | 0.846 | 0.845 | 0.845 | 0.844 | 0.844 | 0.844 | 0.843 | 0.843 | 0.842 | 0.842 | 0.842 | 0.841 | 0.841 | 0.84 | 0.84 | 0.839 | 0.839 | 0.839 | 0.838 | 0.838 | 0.837 | 0.837 | 0.837
1 | 1 | 1 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867 | 0.866 | 0.866 | 0.865 | 0.865 | 0.865 | 0.864 | 0.864 | 0.863 | 0.863 | 0.862 | 0.862 | 0.862 | 0.861 | 0.861 | 0.86 | 0.86 | 0.86 | 0.859 | 0.859 | 0.858 | 0.858 | 0.857 | 0.857 | 0.857 | 0.856 | 0.856 | 0.855 | 0.855 | 0.855 | 0.854 | 0.854 | 0.853 | 0.853 | 0.852 | 0.852 | 0.852 | 0.851 | 0.851 | 0.85 | 0.85 | 0.85 | 0.849 | 0.849 | 0.848 | 0.848 | 0.847 | 0.847 | 0.847 | 0.846 | 0.846 | 0.845 | 0.845 | 0.845 | 0.844 | 0.844 | 0.843 | 0.843 | 0.842 | 0.842 | 0.842 | 0.841 | 0.841 | 0.84 | 0.84 | 0.84 | 0.839 | 0.839 | 0.838 | 0.838
2 | 1 | 1 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867 | 0.866 | 0.866 | 0.865 | 0.865 | 0.864 | 0.864 | 0.864 | 0.863 | 0.863 | 0.862 | 0.862 | 0.862 | 0.861 | 0.861 | 0.86 | 0.86 | 0.86 | 0.859 | 0.859 | 0.858 | 0.858 | 0.857 | 0.857 | 0.857 | 0.856 | 0.856 | 0.855 | 0.855 | 0.855 | 0.854 | 0.854 | 0.853 | 0.853 | 0.853 | 0.852 | 0.852 | 0.851 | 0.851 | 0.85 | 0.85 | 0.85 | 0.849 | 0.849 | 0.848 | 0.848 | 0.848 | 0.847 | 0.847 | 0.846 | 0.846 | 0.845 | 0.845 | 0.845 | 0.844 | 0.844 | 0.843 | 0.843 | 0.843 | 0.842 | 0.842 | 0.841 | 0.841 | 0.841 | 0.84 | 0.84 | 0.839
3 | 1 | 1 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.866 | 0.866 | 0.866 | 0.865 | 0.865 | 0.864 | 0.864 | 0.864 | 0.863 | 0.863 | 0.862 | 0.862 | 0.862 | 0.861 | 0.861 | 0.86 | 0.86 | 0.859 | 0.859 | 0.859 | 0.858 | 0.858 | 0.857 | 0.857 | 0.857 | 0.856 | 0.856 | 0.855 | 0.855 | 0.855 | 0.854 | 0.854 | 0.853 | 0.853 | 0.853 | 0.852 | 0.852 | 0.851 | 0.851 | 0.85 | 0.85 | 0.85 | 0.849 | 0.849 | 0.848 | 0.848 | 0.848 | 0.847 | 0.847 | 0.846 | 0.846 | 0.846 | 0.845 | 0.845 | 0.844 | 0.844 | 0.844 | 0.843 | 0.843 | 0.842 | 0.842 | 0.841 | 0.841 | 0.841
4 | 1 | 1 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.866 | 0.866 | 0.866 | 0.865 | 0.865 | 0.864 | 0.864 | 0.864 | 0.863 | 0.863 | 0.862 | 0.862 | 0.861 | 0.861 | 0.861 | 0.86 | 0.86 | 0.859 | 0.859 | 0.859 | 0.858 | 0.858 | 0.857 | 0.857 | 0.857 | 0.856 | 0.856 | 0.855 | 0.855 | 0.855 | 0.854 | 0.854 | 0.853 | 0.853 | 0.853 | 0.852 | 0.852 | 0.851 | 0.851 | 0.851 | 0.85 | 0.85 | 0.849 | 0.849 | 0.848 | 0.848 | 0.848 | 0.847 | 0.847 | 0.846 | 0.846 | 0.846 | 0.845 | 0.845 | 0.844 | 0.844 | 0.844 | 0.843 | 0.843 | 0.842 | 0.842
5 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867 | 0.866 | 0.866 | 0.865 | 0.865 | 0.865 | 0.864 | 0.864 | 0.863 | 0.863 | 0.863 | 0.862 | 0.862 | 0.861 | 0.861 | 0.861 | 0.86 | 0.86 | 0.859 | 0.859 | 0.859 | 0.858 | 0.858 | 0.857 | 0.857 | 0.857 | 0.856 | 0.856 | 0.855 | 0.855 | 0.855 | 0.854 | 0.854 | 0.853 | 0.853 | 0.853 | 0.852 | 0.852 | 0.851 | 0.851 | 0.851 | 0.85 | 0.85 | 0.849 | 0.849 | 0.849 | 0.848 | 0.848 | 0.847 | 0.847 | 0.846 | 0.846 | 0.846 | 0.845 | 0.845 | 0.844 | 0.844 | 0.844 | 0.843
6 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867 | 0.866 | 0.866 | 0.865 | 0.865 | 0.865 | 0.864 | 0.864 | 0.863 | 0.863 | 0.863 | 0.862 | 0.862 | 0.861 | 0.861 | 0.861 | 0.86 | 0.86 | 0.859 | 0.859 | 0.859 | 0.858 | 0.858 | 0.857 | 0.857 | 0.857 | 0.856 | 0.856 | 0.855 | 0.855 | 0.855 | 0.854 | 0.854 | 0.853 | 0.853 | 0.853 | 0.852 | 0.852 | 0.851 | 0.851 | 0.851 | 0.85 | 0.85 | 0.849 | 0.849 | 0.849 | 0.848 | 0.848 | 0.847 | 0.847 | 0.847 | 0.846 | 0.846 | 0.845 | 0.845 | 0.845
7 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.866 | 0.866 | 0.866 | 0.865 | 0.865 | 0.864 | 0.864 | 0.864 | 0.863 | 0.863 | 0.863 | 0.862 | 0.862 | 0.861 | 0.861 | 0.861 | 0.86 | 0.86 | 0.859 | 0.859 | 0.859 | 0.858 | 0.858 | 0.857 | 0.857 | 0.857 | 0.856 | 0.856 | 0.855 | 0.855 | 0.855 | 0.854 | 0.854 | 0.853 | 0.853 | 0.853 | 0.852 | 0.852 | 0.851 | 0.851 | 0.851 | 0.85 | 0.85 | 0.849 | 0.849 | 0.849 | 0.848 | 0.848 | 0.847 | 0.847 | 0.847 | 0.846 | 0.846
8 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.866 | 0.866 | 0.866 | 0.865 | 0.865 | 0.864 | 0.864 | 0.864 | 0.863 | 0.863 | 0.862 | 0.862 | 0.862 | 0.861 | 0.861 | 0.86 | 0.86 | 0.86 | 0.859 | 0.859 | 0.858 | 0.858 | 0.858 | 0.857 | 0.857 | 0.857 | 0.856 | 0.856 | 0.855 | 0.855 | 0.855 | 0.854 | 0.854 | 0.853 | 0.853 | 0.853 | 0.852 | 0.852 | 0.851 | 0.851 | 0.851 | 0.85 | 0.85 | 0.849 | 0.849 | 0.849 | 0.848 | 0.848 | 0.847 | 0.847
9 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867 | 0.866 | 0.866 | 0.865 | 0.865 | 0.865 | 0.864 | 0.864 | 0.864 | 0.863 | 0.863 | 0.862 | 0.862 | 0.862 | 0.861 | 0.861 | 0.86 | 0.86 | 0.86 | 0.859 | 0.859 | 0.858 | 0.858 | 0.858 | 0.857 | 0.857 | 0.856 | 0.856 | 0.856 | 0.855 | 0.855 | 0.855 | 0.854 | 0.854 | 0.853 | 0.853 | 0.853 | 0.852 | 0.852 | 0.851 | 0.851 | 0.851 | 0.85 | 0.85 | 0.849 | 0.849 | 0.849 | 0.848
10 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867 | 0.866 | 0.866 | 0.865 | 0.865 | 0.865 | 0.864 | 0.864 | 0.863 | 0.863 | 0.863 | 0.862 | 0.862 | 0.862 | 0.861 | 0.861 | 0.86 | 0.86 | 0.86 | 0.859 | 0.859 | 0.858 | 0.858 | 0.858 | 0.857 | 0.857 | 0.856 | 0.856 | 0.856 | 0.855 | 0.855 | 0.854 | 0.854 | 0.854 | 0.853 | 0.853 | 0.853 | 0.852 | 0.852 | 0.851 | 0.851 | 0.851 | 0.85 | 0.85 | 0.849
11 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.866 | 0.866 | 0.866 | 0.865 | 0.865 | 0.865 | 0.864 | 0.864 | 0.863 | 0.863 | 0.863 | 0.862 | 0.862 | 0.861 | 0.861 | 0.861 | 0.86 | 0.86 | 0.86 | 0.859 | 0.859 | 0.858 | 0.858 | 0.858 | 0.857 | 0.857 | 0.856 | 0.856 | 0.856 | 0.855 | 0.855 | 0.854 | 0.854 | 0.854 | 0.853 | 0.853 | 0.853 | 0.852 | 0.852 | 0.851 | 0.851 | 0.851
12 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867 | 0.866 | 0.866 | 0.866 | 0.865 | 0.865 | 0.864 | 0.864 | 0.864 | 0.863 | 0.863 | 0.863 | 0.862 | 0.862 | 0.861 | 0.861 | 0.861 | 0.86 | 0.86 | 0.859 | 0.859 | 0.859 | 0.858 | 0.858 | 0.858 | 0.857 | 0.857 | 0.856 | 0.856 | 0.856 | 0.855 | 0.855 | 0.854 | 0.854 | 0.854 | 0.853 | 0.853 | 0.853 | 0.852 | 0.852
13 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867 | 0.866 | 0.866 | 0.865 | 0.865 | 0.865 | 0.864 | 0.864 | 0.864 | 0.863 | 0.863 | 0.862 | 0.862 | 0.862 | 0.861 | 0.861 | 0.861 | 0.86 | 0.86 | 0.859 | 0.859 | 0.859 | 0.858 | 0.858 | 0.857 | 0.857 | 0.857 | 0.856 | 0.856 | 0.856 | 0.855 | 0.855 | 0.854 | 0.854 | 0.854 | 0.853 | 0.853
14 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.866 | 0.866 | 0.866 | 0.865 | 0.865 | 0.865 | 0.864 | 0.864 | 0.863 | 0.863 | 0.863 | 0.862 | 0.862 | 0.862 | 0.861 | 0.861 | 0.86 | 0.86 | 0.86 | 0.859 | 0.859 | 0.859 | 0.858 | 0.858 | 0.857 | 0.857 | 0.857 | 0.856 | 0.856 | 0.855 | 0.855 | 0.855 | 0.854 | 0.854
15 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867 | 0.866 | 0.866 | 0.866 | 0.865 | 0.865 | 0.864 | 0.864 | 0.864 | 0.863 | 0.863 | 0.863 | 0.862 | 0.862 | 0.861 | 0.861 | 0.861 | 0.86 | 0.86 | 0.86 | 0.859 | 0.859 | 0.858 | 0.858 | 0.858 | 0.857 | 0.857 | 0.857 | 0.856 | 0.856 | 0.855 | 0.855
16 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867 | 0.866 | 0.866 | 0.865 | 0.865 | 0.865 | 0.864 | 0.864 | 0.864 | 0.863 | 0.863 | 0.862 | 0.862 | 0.862 | 0.861 | 0.861 | 0.861 | 0.86 | 0.86 | 0.86 | 0.859 | 0.859 | 0.858 | 0.858 | 0.858 | 0.857 | 0.857 | 0.857 | 0.856
17 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.866 | 0.866 | 0.866 | 0.865 | 0.865 | 0.865 | 0.864 | 0.864 | 0.864 | 0.863 | 0.863 | 0.862 | 0.862 | 0.862 | 0.861 | 0.861 | 0.861 | 0.86 | 0.86 | 0.859 | 0.859 | 0.859 | 0.858 | 0.858 | 0.858 | 0.857
18 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867 | 0.866 | 0.866 | 0.866 | 0.865 | 0.865 | 0.864 | 0.864 | 0.864 | 0.863 | 0.863 | 0.863 | 0.862 | 0.862 | 0.862 | 0.861 | 0.861 | 0.86 | 0.86 | 0.86 | 0.859 | 0.859 | 0.859 | 0.858
19 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867 | 0.866 | 0.866 | 0.865 | 0.865 | 0.865 | 0.864 | 0.864 | 0.864 | 0.863 | 0.863 | 0.863 | 0.862 | 0.862 | 0.861 | 0.861 | 0.861 | 0.86 | 0.86 | 0.86 | 0.859
20 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.866 | 0.866 | 0.866 | 0.865 | 0.865 | 0.865 | 0.864 | 0.864 | 0.864 | 0.863 | 0.863 | 0.862 | 0.862 | 0.862 | 0.861 | 0.861 | 0.861 | 0.86
21 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867 | 0.866 | 0.866 | 0.866 | 0.865 | 0.865 | 0.864 | 0.864 | 0.864 | 0.863 | 0.863 | 0.863 | 0.862 | 0.862 | 0.862 | 0.861
22 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.866 | 0.866 | 0.866 | 0.865 | 0.865 | 0.865 | 0.864 | 0.864 | 0.864 | 0.863 | 0.863 | 0.863 | 0.862
23 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867 | 0.866 | 0.866 | 0.866 | 0.865 | 0.865 | 0.865 | 0.864 | 0.864 | 0.863 | 0.863
24 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867 | 0.866 | 0.866 | 0.865 | 0.865 | 0.865 | 0.864 | 0.864
25 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867 | 0.866 | 0.866 | 0.866 | 0.865 | 0.865
26 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867 | 0.866 | 0.866
27 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868 | 0.867 | 0.867 | 0.867
28 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868 | 0.868 | 0.868
29 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869 | 0.869 | 0.869 | 0.868
30 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87 | 0.87 | 0.869
31 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871 | 0.87 | 0.87
32 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872 | 0.871 | 0.871 | 0.871
33 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872 | 0.872 | 0.872
34 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873 | 0.873 | 0.872
35 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874 | 0.873 | 0.873
36 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.874 | 0.874 | 0.874
37 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875 | 0.875 | 0.875 | 0.875
38 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876 | 0.876 | 0.875
39 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877 | 0.876 | 0.876
40 | 1 | 1 | 0.999 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.997 | 0.997 | 0.997 | 0.997 | 0.996 | 0.996 | 0.996 | 0.995 | 0.995 | 0.995 | 0.995 | 0.994 | 0.994 | 0.994 | 0.993 | 0.993 | 0.993 | 0.992 | 0.992 | 0.992 | 0.992 | 0.991 | 0.991 | 0.991 | 0.99 | 0.99 | 0.99 | 0.99 | 0.989 | 0.989 | 0.989 | 0.988 | 0.988 | 0.988 | 0.988 | 0.987 | 0.987 | 0.987 | 0.986 | 0.986 | 0.986 | 0.985 | 0.985 | 0.985 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.983 | 0.983 | 0.983 | 0.982 | 0.982 | 0.982 | 0.981 | 0.981 | 0.981 | 0.98 | 0.98 | 0.98 | 0.98 | 0.979 | 0.979 | 0.979 | 0.978 | 0.978 | 0.978 | 0.977 | 0.977 | 0.977 | 0.977 | 0.976 | 0.976 | 0.976 | 0.975 | 0.975 | 0.975 | 0.975 | 0.974 | 0.974 | 0.974 | 0.973 | 0.973 | 0.973 | 0.972 | 0.972 | 0.972 | 0.972 | 0.971 | 0.971 | 0.971 | 0.97 | 0.97 | 0.97 | 0.969 | 0.969 | 0.969 | 0.969 | 0.968 | 0.968 | 0.968 | 0.967 | 0.967 | 0.967 | 0.966 | 0.966 | 0.966 | 0.966 | 0.965 | 0.965 | 0.965 | 0.964 | 0.964 | 0.964 | 0.963 | 0.963 | 0.963 | 0.963 | 0.962 | 0.962 | 0.962 | 0.961 | 0.961 | 0.961 | 0.96 | 0.96 | 0.96 | 0.959 | 0.959 | 0.959 | 0.959 | 0.958 | 0.958 | 0.958 | 0.957 | 0.957 | 0.957 | 0.956 | 0.956 | 0.956 | 0.956 | 0.955 | 0.955 | 0.955 | 0.954 | 0.954 | 0.954 | 0.953 | 0.953 | 0.953 | 0.952 | 0.952 | 0.952 | 0.952 | 0.951 | 0.951 | 0.951 | 0.95 | 0.95 | 0.95 | 0.949 | 0.949 | 0.949 | 0.948 | 0.948 | 0.948 | 0.948 | 0.947 | 0.947 | 0.947 | 0.946 | 0.946 | 0.946 | 0.945 | 0.945 | 0.945 | 0.944 | 0.944 | 0.944 | 0.944 | 0.943 | 0.943 | 0.943 | 0.942 | 0.942 | 0.942 | 0.941 | 0.941 | 0.941 | 0.94 | 0.94 | 0.94 | 0.94 | 0.939 | 0.939 | 0.939 | 0.938 | 0.938 | 0.938 | 0.937 | 0.937 | 0.937 | 0.936 | 0.936 | 0.936 | 0.936 | 0.935 | 0.935 | 0.935 | 0.934 | 0.934 | 0.934 | 0.933 | 0.933 | 0.933 | 0.932 | 0.932 | 0.932 | 0.931 | 0.931 | 0.931 | 0.931 | 0.93 | 0.93 | 0.93 | 0.929 | 0.929 | 0.929 | 0.928 | 0.928 | 0.928 | 0.927 | 0.927 | 0.927 | 0.926 | 0.926 | 0.926 | 0.926 | 0.925 | 0.925 | 0.925 | 0.924 | 0.924 | 0.924 | 0.923 | 0.923 | 0.923 | 0.922 | 0.922 | 0.922 | 0.921 | 0.921 | 0.921 | 0.92 | 0.92 | 0.92 | 0.92 | 0.919 | 0.919 | 0.919 | 0.918 | 0.918 | 0.918 | 0.917 | 0.917 | 0.917 | 0.916 | 0.916 | 0.916 | 0.915 | 0.915 | 0.915 | 0.914 | 0.914 | 0.914 | 0.914 | 0.913 | 0.913 | 0.913 | 0.912 | 0.912 | 0.912 | 0.911 | 0.911 | 0.911 | 0.91 | 0.91 | 0.91 | 0.909 | 0.909 | 0.909 | 0.908 | 0.908 | 0.908 | 0.907 | 0.907 | 0.907 | 0.907 | 0.906 | 0.906 | 0.906 | 0.905 | 0.905 | 0.905 | 0.904 | 0.904 | 0.904 | 0.903 | 0.903 | 0.903 | 0.902 | 0.902 | 0.902 | 0.901 | 0.901 | 0.901 | 0.9 | 0.9 | 0.9 | 0.899 | 0.899 | 0.899 | 0.899 | 0.898 | 0.898 | 0.898 | 0.897 | 0.897 | 0.897 | 0.896 | 0.896 | 0.896 | 0.895 | 0.895 | 0.895 | 0.894 | 0.894 | 0.894 | 0.893 | 0.893 | 0.893 | 0.892 | 0.892 | 0.892 | 0.891 | 0.891 | 0.891 | 0.89 | 0.89 | 0.89 | 0.89 | 0.889 | 0.889 | 0.889 | 0.888 | 0.888 | 0.888 | 0.887 | 0.887 | 0.887 | 0.886 | 0.886 | 0.886 | 0.885 | 0.885 | 0.885 | 0.884 | 0.884 | 0.884 | 0.883 | 0.883 | 0.883 | 0.882 | 0.882 | 0.882 | 0.881 | 0.881 | 0.881 | 0.88 | 0.88 | 0.88 | 0.879 | 0.879 | 0.879 | 0.879 | 0.878 | 0.878 | 0.878 | 0.877 | 0.877 | 0.877

## Documentation
pH Correction:
pH probes are composed a reference electrode and an H+ ion sensitive electrode. The voltage potential between the two electrodes (E) is related to the decimal logarithm of the reciprocal
of Hydrogen Ion activity (pH) by a linear equation. Therefore calibration corrections applied to pH data are linear in nature.
The point of zero potential between the electrodes (E=0) occurs in a perfectly neutral solution, a pH of 7 (roughly, depending on temperature). Deviation from
a pH of 7 is corrected by applying on offset to the y-intercept in the equation. This, in effect, shifts the entire line up or down (the green arrow in the figure below).
Deviation from a pH of 4 or 10 (or any other combination of standards used) involves adjusting the slope of the equation (the red arrow).
In the correction routine in this spreadsheet, the y-intercept correction is applied first followed by the slope correction for both pH 10 and 4 (if applicable).